## Raw Intensities
| Sample Id | R | Acquisition Time | QC Status | Dataset File | Method File | Na KED 23/23 Helium KED_Low (cps) | Mg KED 24/24 Helium KED_Low (cps) | Al 27/27 (cps) | Al KED 27/27 Helium KED_Low (cps) | P Oshift 31/47 Oxygen DRC (cps) | S (34) -1 34/34 Helium KED_Low (cps) | K KED 39/39 Helium KED_Low (cps) | Ca KED 43/43 Helium KED_Low (cps) | V KED 51/51 Helium KED_Low (cps) | Cr (52) 52/52 Helium KED_Low (cps) | Cr (53) NH3 53/53 Ammonia DRC (cps) | Fe (54) NH3 54/54 Ammonia DRC (cps) | Mn Oshift 55/71 Oxygen DRC (cps) | Fe (56) KED 56/56 Helium KED_Low (cps) | Co KED 59/59 Helium KED_Low (cps) | Ni KED 60/60 Helium KED_Low (cps) | Cu (63) KED 63/63 Helium KED_Low (cps) | Cu (65) NH3 65/65 Ammonia DRC (cps) | Zn NH3 66/66 Ammonia DRC (cps) | As 75/75 (cps) | As KED 75/75 Helium KED_Low (cps) | As RH Oshift 75/91 Oxygen DRC (cps) | As Oshift 75/91 Oxygen DRC (cps) | Se (77) 77/77 (cps) | Se (77) Oshift 77/93 Oxygen DRC (cps) | Se (78) Oshift 78/94 Oxygen DRC (cps) | Se (80) Oshift 80/96 Oxygen DRC (cps) | Se (82) 82/82 (cps) | Se (82) Oshift 82/98 Oxygen DRC (cps) | Mo (95) NH3 95/95 Ammonia DRC (cps) | Mo (95) OShift 95/127 Oxygen DRC (cps) | Rh (IS1) 103/103 (IS) (cps) | Rh (IS2) 103/103 (IS) Oxygen DRC (cps) | Rh (IS3) 103/103 (IS) Helium KED_Low (cps) | Rh (IS4) 103/103 (IS) Ammonia DRC (cps) | Rh (IS5) 103/103 (IS) Ammonia DRC (cps) | Rh (IS6) 103/103 (IS) Oxygen DRC (cps) | Rh OShift (IS1) 103/119 (IS) Oxygen DRC (cps) | Cd (111) NH3 111/111 Ammonia DRC (cps) | Cd (114) NH3 114/114 Ammonia DRC (cps) | Pb 208/208 (cps) | Pb NH3 208/208 Ammonia DRC (cps)
1 | Sac Blank |  | 06/08/2023 13:23:24 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\Sac Blank.356 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0 | 0 | 5155.5973 | 2.6667 | 2627.5751 | 644.6812 | 15.3333 | 0 | 1.3333 | 14.6667 | 40.0001 | 454.2131 | 16.6667 | 1408.0694 | 2 | 6 | 10.6667 | 108.0004 | 384.0052 | -3039.8034 | 0 | 36.6667 | 36.6667 | 1099.3757 | 1.3333 | 4 | 6.6667 | 53.4977 | 2 | 15.3333 | 10 | 182769.2655 | 265085.5927 | 14337.1948 | 272750.8925 | 224928.416 | 238625.9044 | 5129.5876 | 2 | 11.3333 | 18.6667 | 50
2 | sac blank cal |  | 06/08/2023 13:29:13 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\sac blank cal.357 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.3333 | 0 | 5670.52 | 58.0003 | 2129.4921 | 642.0145 | 14 | 0 | 0.66667 | 9.3333 | 22 | 509.4667 | 9.3333 | 1426.7379 | 0.66667 | 6.6667 | 4 | 80.6669 | 306.0033 | -2967.631 | 0 | 32 | 32 | 1045.3719 | 0 | 1.3333 | 4 | 57.395 | 0.66667 | 10.6667 | 6.6667 | 196830.8838 | 258233.7298 | 14098.9543 | 280289.0073 | 226489.4263 | 235016.4348 | 5122.2517 | 0.66667 | 6 | 36 | 62.0001
3 | CAL BLANK |  | 06/08/2023 13:35:02 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CAL BLANK.358 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.66667 | 0 | 3125.6753 | 0.66667 | 1866.122 | 636.0142 | 17.3333 | 0.66667 | 0.66667 | 6.6667 | 24.6667 | 819.0943 | 13.3333 | 1537.4161 | 3.3333 | 10.6667 | 12 | 74.0002 | 358.6712 | -2829.9787 | 0.66667 | 27.3334 | 27.3334 | 1025.3702 | 4.6667 | 1.3333 | 2.6667 | 37.4705 | 0 | 6 | 11.3333 | 180102.215 | 255065.2033 | 13990.1811 | 276357.489 | 220254.432 | 230485.1249 | 5076.2351 | 1.3333 | 8 | 35.3334 | 69.3334
4 | LOW 0.01 |  | 06/08/2023 13:40:50 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LOW 0.01.359 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 12 | 6.6667 | 44668.3995 | 24 | 5359.0051 | 641.3477 | 30 | 4 | 9.3333 | 24.6667 | 120.6672 | 1406.8037 | 611.3464 | 1585.4214 | 32 | 20 | 58.6668 | 461.3408 | 1529.4152 | -2708.5655 | 2 | 129.3339 | 129.3339 | 1044.0382 | 2.6667 | 9.3333 | 16.6667 | 60.9124 | 3.3333 | 413.3393 | 392.0054 | 179124.7892 | 253794.7497 | 13848.7099 | 273488.6462 | 218102.3684 | 224211.1157 | 5017.5478 | 89.3336 | 236.002 | 405.3355 | 841.3429
5 | LOW 0.05 |  | 06/08/2023 13:46:39 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LOW 0.05.360 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 102.0004 | 29.3334 | 50507.1304 | 43.3334 | 7830.8125 | 653.3483 | 55.3334 | 20.6667 | 58.0001 | 106.6671 | 536.0102 | 1457.3375 | 872.6934 | 1641.4279 | 157.3342 | 60.6668 | 157.3342 | 1308.06 | 1286.058 | -2304.6464 | 8.6667 | 446.6737 | 446.6737 | 1040.0379 | 12.6667 | 29.3334 | 66.0002 | 89.5254 | 10.6667 | 875.3602 | 843.3583 | 174172.7239 | 248800.1366 | 13522.3981 | 265464.5214 | 219779.0539 | 224388.5039 | 4909.5102 | 464.0076 | 1091.3751 | 1634.7039 | 3518.8382
6 | LOW 0.1 |  | 06/08/2023 13:52:27 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LOW 0.1.361 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 164.0009 | 78.6669 | 66705.413 | 46.0001 | 8693.3136 | 659.3486 | 105.3337 | 41.3334 | 110.0004 | 171.3344 | 1029.3704 | 3335.8806 | 1810.1153 | 1940.1317 | 368.6715 | 102.0004 | 278.6694 | 2316.1885 | 2824.9461 | -2248.9273 | 16.6667 | 869.3598 | 869.3598 | 1109.3764 | 24.6667 | 60.6668 | 119.3338 | 76.8314 | 24 | 1879.4571 | 1617.4251 | 176741.3101 | 236645.9464 | 13715.924 | 272424.4973 | 218327.9206 | 221955.0675 | 4684.1037 | 929.3636 | 2114.1565 | 3362.8216 | 7140.0382
7 | LOW 0.5 |  | 06/08/2023 13:58:15 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LOW 0.5.362 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 882.0273 | 383.3385 | 99432.183 | 56.0001 | 9936.1218 | 761.3536 | 436.0067 | 172.6677 | 634.0141 | 872.0266 | 4945.5232 | 7274.1848 | 8289.071 | 2952.305 | 1807.4478 | 490.6751 | 1282.7243 | 10720.6888 | 6597.5232 | 250.5916 | 60.6668 | 4205.9525 | 4205.9525 | 1398.7352 | 86.6669 | 302.6699 | 620.0135 | 360.3421 | 112.6671 | 8826.7264 | 8056.2715 | 177339.2711 | 244575.8035 | 13070.6542 | 269232.914 | 215426.8278 | 216603.764 | 4858.8262 | 4433.3546 | 10474.5056 | 16156.2598 | 34825.3697
8 | STD 1 |  | 06/08/2023 14:04:03 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 1.363 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1802.1139 | 814.0233 | 146490.8033 | 86.0003 | 9510.4983 | 866.693 | 848.0252 | 302.6699 | 1285.3913 | 1750.1073 | 9583.8808 | 12181.861 | 16607.6506 | 4208.6201 | 3553.7754 | 1021.3699 | 2548.2273 | 20225.6413 | 10842.78 | 3726.9536 | 126.6672 | 8275.063 | 8275.063 | 1667.4309 | 179.3345 | 567.3446 | 1161.3806 | 688.5412 | 194.668 | 17572.1355 | 16083.7197 | 174759.7721 | 242942.2923 | 13358.9102 | 263003.7451 | 210478.1675 | 215783.7034 | 4674.7656 | 8557.8962 | 20443.9538 | 31153.2334 | 69305.3098
9 | STD 5 |  | 06/08/2023 14:09:52 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 5.364 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 8684.6417 | 3717.1506 | 416833.8581 | 268.6692 | 22680.6568 | 1910.1279 | 4039.9051 | 1458.0745 | 5701.8053 | 8207.0243 | 45248.2132 | 31960.5664 | 77622.3288 | 8896.7705 | 16875.2982 | 4775.4651 | 11366.5217 | 89657.8063 | 44189.5734 | 28171.4802 | 551.344 | 39065.3411 | 39065.3411 | 4110.5916 | 866.0263 | 2676.2507 | 5650.4508 | 3025.7673 | 1029.3705 | 81877.31 | 75792.5342 | 157415.1079 | 233196.5948 | 13165.3972 | 250165.4314 | 206727.7047 | 210675.8643 | 4400.0108 | 40227.2255 | 95509.5602 | 146996.4675 | 316176.9324
10 | STD 10 |  | 06/08/2023 14:15:40 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 10.365 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 16670.388 | 7285.8579 | 768832.5419 | 450.0071 | 36950.3947 | 3082.3325 | 8015.5814 | 2986.3121 | 11623.3938 | 15884.1595 | 91523.6015 | 62454.0389 | 153868.9827 | 15586.4987 | 34047.8635 | 9207.6337 | 21743.87 | 175978.1335 | 85257.6687 | 58258.0374 | 1166.7143 | 77316.6683 | 77316.6683 | 7156.459 | 1690.7667 | 5036.2218 | 10966.8759 | 5665.4931 | 1916.7953 | 165122.1669 | 148778.7622 | 165063.4829 | 232818.3318 | 12664.9452 | 247069.2922 | 201124.1935 | 205356.1414 | 4470.6996 | 80195.8244 | 190070.9239 | 282433.5407 | 619828.2861
11 | STD 25 |  | 06/08/2023 14:21:28 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 25.366 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 40476.6085 | 16942.0405 | 1722994.4893 | 1027.3703 | 80624.2061 | 7007.0523 | 19600.771 | 7468.6189 | 28698.8109 | 38975.7668 | 220107.5529 | 146615.7512 | 373271.0684 | 35845.5823 | 81750.0452 | 22103.7641 | 52573.2461 | 411631.9279 | 201730.3702 | 150241.6423 | 2682.2518 | 190239.0081 | 190239.0081 | 15815.4166 | 4153.9372 | 12916.505 | 27286.0349 | 14373.9624 | 4842.1539 | 399860.9347 | 367015.4368 | 157929.5161 | 215280.0777 | 12048.412 | 230930.2945 | 204800.4543 | 189602.7374 | 4067.9133 | 192229.3648 | 449953.9835 | 661124.4668 | 1399423.5267
12 | STD 50 |  | 06/08/2023 14:27:16 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 50.367 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 76425.9306 | 32020.5346 | 3627090.622 | 1948.7999 | 148818.5178 | 12662.2772 | 38116.7916 | 14488.0104 | 56698.9666 | 76336.7772 | 425000.1488 | 279365.2806 | 717760.789 | 67581.5237 | 157590.6039 | 42068.5372 | 99105.9439 | 788237.7808 | 384654.607 | 293529.6568 | 5179.6056 | 368342.2298 | 368342.2298 | 29672.1209 | 7878.8393 | 24732.0579 | 52548.4739 | 28434.904 | 9605.2286 | 791676.1484 | 723447.2399 | 153121.6643 | 203870.2062 | 11699.4572 | 222303.2713 | 177857.0724 | 183748.8782 | 3909.202 | 363418.7811 | 852160.9888 | 1231633.6038 | 2609125.1138
13 | STD 100 |  | 06/08/2023 14:33:03 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 100.368 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 145329.508 | 57372.979 | 6926229.4415 | 3519.1001 | 264978.7222 | 23618.8444 | 72761.513 | 28018.7864 | 109655.943 | 148295.2136 | 815748.1678 | 521167.7008 | 1356369.9948 | 126055.0416 | 305295.8694 | 79005.2236 | 185640.4577 | 1444648.995 | 696665.7816 | 567668.2431 | 9648.5907 | 676574.9678 | 676574.9678 | 55784.0457 | 15083.2926 | 47003.1996 | 98836.0613 | 54956.7234 | 17941.9272 | 1543840.7207 | 1380690.1981 | 146596.3508 | 191260.51 | 10986.8901 | 210352.8534 | 168709.7639 | 170137.9821 | 3577.7814 | 677576.0939 | 1600530.5892 | 2276795.2079 | 5070073.1318
14 | Calibration Curves |  | 06/08/2023 14:33:03
15 | LBLANK |  | 06/08/2023 14:38:52 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.369 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4 | 1.3333 | 4680.0999 | 2.6667 | 2530.891 | 706.6842 | 16 | 0 | 5.3333 | 18 | 48.0001 | 664.6813 | 47.3334 | 1545.417 | 4.6667 | 34 | 23.3334 | 139.334 | 408.6725 | -3155.6423 | 2 | 116.6671 | 116.6671 | 1190.7164 | 0.66667 | 3.3333 | 12 | 70.8692 | 2 | 715.3513 | 225.3351 | 185301.4246 | 255627.5175 | 14510.0361 | 278697.7125 | 203255.3257 | 226092.5268 | 5073.5683 | 14.6667 | 40.6667 | 38 | 77.3334
16 | LBLANK |  | 06/08/2023 14:57:55 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.370 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 6.6667 | 0 | 6419.684 | 1.3333 | 1396.735 | 574.6783 | 13.3333 | 0 | 3.3333 | 16.6667 | 49.3334 | 639.8196 | 18 | 1375.3996 | 2 | 12.6667 | 12.6667 | 111.3338 | 456.674 | -2810.4308 | 3.3333 | 36.6667 | 36.6667 | 1055.3724 | 2 | 4.6667 | 5.3333 | 37.4488 | 1.3333 | 119.3338 | 85.3336 | 172111.9419 | 241266.9573 | 12453.4771 | 265282.5642 | 216199.9653 | 212227.489 | 4894.1716 | 2.6667 | 15.3333 | 20.6667 | 46.6667
17 | ICV PE |  | 06/08/2023 15:03:43 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\ICV PE.371 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 16684.4046 | 7089.0923 | 774275.9702 | 432.6732 | 65057.1864 | 3088.334 | 8154.9947 | 3102.3369 | 11644.0769 | 16534.2303 | 88760.8962 | 62935.3046 | 155037.4566 | 16325.3231 | 34176.1695 | 9509.1648 | 22148.494 | 172587.6022 | 81713.702 | 59129.0904 | 1013.3693 | 77360.9168 | 77360.9168 | 5653.7853 | 1698.1009 | 5268.305 | 11053.6082 | 6013.3929 | 2012.1419 | 158310.3841 | 145349.0789 | 158845.0618 | 223452.7512 | 12724.3322 | 244373.328 | 210139.3361 | 200004.7924 | 4343.9939 | 78627.8015 | 187234.3403 | 278696.3399 | 604263.0145
18 | ICV PE |  | 06/08/2023 15:09:59 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\ICV PE.372 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 16339.3393 | 7195.145 | 768096.4265 | 440.6735 | 65737.5786 | 3127.6757 | 7932.202 | 3110.3388 | 11648.7519 | 16482.8394 | 89764.4734 | 63815.3573 | 155726.8649 | 16529.5607 | 34124.7095 | 9471.8061 | 22082.3886 | 175014.9735 | 82815.358 | 59287.1121 | 1018.0364 | 76929.9153 | 76929.9153 | 5740.4868 | 1660.0966 | 5080.9034 | 10971.5455 | 5774.6919 | 1913.4616 | 157487.327 | 145864.2649 | 165620.5593 | 223078.3379 | 12597.5533 | 241561.9807 | 198734.2322 | 198046.47 | 4524.7169 | 79200.3011 | 189253.4366 | 279583.6379 | 601583.9981
19 | CCV 1ppb TRACE |  | 06/08/2023 15:15:47 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.373 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1714.1028 | 746.6862 | 140455.8294 | 78.6669 | 6071.2901 | 954.0319 | 842.0248 | 315.3368 | 1258.7221 | 1740.7729 | 9858.7342 | 7748.4111 | 16453.4746 | 3128.3425 | 3627.7941 | 1030.0372 | 2505.5532 | 20387.8715 | 9605.2287 | 3589.9854 | 103.3337 | 8221.0318 | 8221.0318 | 1757.4415 | 176.6678 | 563.3445 | 1224.7192 | 692.5037 | 182.0012 | 17953.9461 | 16141.1143 | 174225.4615 | 240593.6612 | 13270.1609 | 264432.4115 | 211804.6168 | 208274.8251 | 4784.1345 | 8505.199 | 20543.4366 | 30327.9551 | 66958.9512
20 | CCV 1ppb TRACE |  | 06/08/2023 15:22:03 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.374 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1802.1137 | 806.0227 | 137176.157 | 76.0002 | 4646.7559 | 876.6936 | 848.0252 | 296.6698 | 1198.0504 | 1715.4363 | 9396.4256 | 7827.767 | 16091.0615 | 3191.0233 | 3630.4614 | 1036.7044 | 2588.9014 | 19697.5712 | 9232.3157 | 4098.2103 | 117.3338 | 8002.9078 | 8002.9078 | 1708.769 | 173.3344 | 514.0093 | 1116.0436 | 646.4709 | 220.0017 | 17223.0457 | 15326.8864 | 169822.8241 | 229717.0459 | 13235.4618 | 253338.6182 | 208333.467 | 202986.7037 | 4569.3978 | 8403.1381 | 19825.7521 | 29447.9015 | 64214.5447
21 | LFB |  | 06/08/2023 15:27:51 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.375 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 15701.2906 | 6699.5722 | 699916.6104 | 410.0059 | 36228.5471 | 3135.6776 | 7560.667 | 2782.9379 | 11168.3648 | 15408.3067 | 85961.1919 | 58875.9919 | 145765.2441 | 15238.7916 | 33108.9967 | 9173.6139 | 21369.9721 | 168140.4574 | 82206.5437 | 56238.9346 | 1071.3735 | 74813.4222 | 74813.4222 | 6910.3382 | 1592.7555 | 4925.5158 | 10541.8895 | 5543.6127 | 1853.4536 | 157657.9688 | 142850.0816 | 160831.1039 | 214239.1821 | 12368.6856 | 229659.7313 | 192086.1967 | 193641.9106 | 4284.6425 | 75249.0359 | 179747.7703 | 258277.9906 | 563814.5687
22 | LLOQ 0.01 |  | 06/08/2023 15:33:39 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LLOQ 0.01.376 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 21.3334 | 9.3333 | 42070.5309 | 26 | 4616.079 | 678.0161 | 27.3334 | 4.6667 | 13.3333 | 27.3334 | 124.0006 | 1532.5806 | 534.01 | 1700.7679 | 32.6667 | 21.3333 | 51.3334 | 467.341 | 1472.7426 | -2545.8866 | 1.3333 | 102.0004 | 102.0004 | 1046.0384 | 4.6667 | 9.3333 | 14 | -8.4849 | 4 | 455.3406 | 392.0054 | 178985.5082 | 236009.8031 | 13201.4379 | 259771.9315 | 214332.7099 | 209699.8883 | 4760.1264 | 89.3336 | 226.6685 | 407.3356 | 767.3415
23 | CCV 10 ppb MAJORS |  | 06/08/2023 15:39:28 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.377 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2 | 1.3333 | 2142.1607 | 0.66667 | 428.6731 | 72.6669 | 10 | 0.66667 | 0.66667 | 12 | 32 | 587.392 | 10 | 624.6803 | 10.6667 | 10 | 12.6667 | 85.3336 | 240.6687 | 988.6525 | 0 | 15.3333 | 15.3333 | 218.0017 | 1.3333 | 2 | 5.3333 | 23.8433 | 1.3333 | 37.3334 | 8.6667 | 192941.5958 | 264644.2699 | 15222.1099 | 311491.2091 | 245395.0214 | 226129.2022 | 5131.5911 | 4 | 8 | 11.3333 | 20
24 | CCV 10 ppb MAJORS |  | 06/08/2023 15:45:42 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.378 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 16255.2471 | 7031.0639 | 670148.785 | 416.0061 | 29895.2707 | 3093.6683 | 8150.9918 | 2978.9777 | 11574.6897 | 16025.6506 | 86914.2721 | 59552.607 | 147251.0391 | 15342.2371 | 33245.3133 | 9179.6157 | 21974.8893 | 167900.9678 | 79828.4223 | 58322.8656 | 1084.0413 | 73270.7554 | 73270.7554 | 6782.2765 | 1582.0876 | 4866.1625 | 10208.9805 | 5546.1515 | 1817.4493 | 158677.7667 | 144315.3431 | 159516.3546 | 213695.1862 | 11627.3988 | 235110.433 | 189586.4368 | 190975.2933 | 4217.956 | 75522.4615 | 178894.0373 | 261604.1156 | 573255.6808
25 | LLOQ 0.1 |  | 06/08/2023 15:51:31 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LLOQ 0.1.379 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 202.0014 | 88.6669 | 63666.2519 | 34.6667 | 5945.2369 | 670.0158 | 101.3337 | 43.3334 | 116.6671 | 202.6681 | 1006.0355 | 3198.5032 | 1764.109 | 2067.4833 | 342.6708 | 126.6672 | 306.67 | 2319.5216 | 2679.5846 | -1651.7455 | 8 | 790.0219 | 790.0219 | 1084.7079 | 10 | 60.6668 | 123.3339 | 111.0469 | 26.6667 | 1818.7825 | 1616.7582 | 170695.8654 | 230021.1196 | 13451.6644 | 256347.5928 | 207106.0947 | 204106.3219 | 4639.4205 | 801.3559 | 1976.8036 | 2928.1175 | 6478.5737
26 | STD 0.5 |  | 06/08/2023 15:57:19 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 0.5.380 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 876.6936 | 380.0051 | 92176.4801 | 68.0002 | 8530.5482 | 760.0203 | 400.6723 | 158.0009 | 608.6797 | 922.6965 | 4763.4609 | 7177.6199 | 8064.9433 | 2952.305 | 1755.4413 | 502.6756 | 1299.3926 | 10442.4852 | 6354.0796 | 818.7054 | 54.0001 | 3960.5495 | 3960.5495 | 1354.0643 | 78.6669 | 284.0028 | 585.3454 | 306.4177 | 80.6669 | 8615.2639 | 7571.3409 | 165637.4035 | 224791.7929 | 12877.8031 | 227709.2961 | 204238.282 | 201097.0991 | 4578.7339 | 4186.6135 | 9836.7203 | 14437.5105 | 32369.8017
27 | 8_2_21_UF_1 |  | 06/08/2023 16:03:07 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_UF_1.381 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 8599.255 | 3641.131 | 154805.6044 | 92.667 | 22174.5387 | 3533.7706 | 931.3637 | 8032.2575 | 429.3398 | 101.3337 | 507.3424 | 5081.484 | 205191.9229 | 2828.9476 | 60.0001 | 214.6683 | 1993.4726 | 15309.5336 | 12689.6348 | 22615.1074 | 422.0062 | 30655.5312 | 30655.5312 | 1067.3734 | 8.6667 | 24.6667 | 52.0001 | 201.4016 | 6.6667 | 75399.1401 | 69069.2386 | 150700.1075 | 203902.2025 | 11794.2009 | 224297.3137 | 184519.3858 | 184880.9711 | 3865.1895 | 51.3334 | 127.3339 | 441.3359 | 1043.3478
28 | 8_2_21_UF_6 |  | 06/08/2023 16:08:55 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_UF_6.382 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 16200.5148 | 4483.3702 | 63849.1721 | 36 | 15690.6131 | 3396.4045 | 1550.7512 | 9617.905 | 1068.04 | 154.6675 | 728.6853 | 5731.6805 | 181316.2638 | 2849.6177 | 63.3335 | 183.3345 | 4439.3572 | 34774.9483 | 18202.2559 | 30421.9112 | 531.3433 | 41827.8302 | 41827.8302 | 1062.0395 | 16.6667 | 52.6668 | 119.3338 | 246.0267 | 20.6667 | 47785.7969 | 44243.0753 | 147345.3276 | 198637.4546 | 11471.2707 | 213822.6619 | 180092.8703 | 183839.3902 | 3903.867 | 129.3339 | 278.0027 | 914.6782 | 2027.3911
29 | 8_2_21_UF_6_SP |  | 06/08/2023 16:14:43 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_UF_6_SP.383 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 24239.2188 | 8067.6125 | 418416.1362 | 242.0021 | 31817.3934 | 4508.0447 | 5611.7693 | 11016.9144 | 6732.2528 | 7910.8571 | 44646.9978 | 36789.0412 | 258945.8603 | 10253.6861 | 16592.2973 | 4676.7655 | 14771.6342 | 116224.2112 | 55959.3895 | 59244.1744 | 1112.0433 | 77410.5154 | 77410.5154 | 3906.5344 | 805.3561 | 2478.215 | 5239.628 | 3062.6258 | 951.365 | 124982.4078 | 116466.9447 | 146550.5408 | 195912.2096 | 11266.4411 | 213861.1387 | 175224.3329 | 159294.8273 | 3925.8728 | 37440.3337 | 87523.9755 | 133557.7456 | 298539.4689
30 | STD 1 |  | 06/08/2023 16:20:31 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 1.384 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1744.7732 | 726.0185 | 139982.4974 | 78.6669 | 6923.6782 | 918.0295 | 875.3602 | 323.337 | 1244.721 | 1689.4335 | 8776.6953 | 12315.7118 | 15342.2349 | 3830.5135 | 3423.744 | 929.3636 | 2332.1907 | 18319.7403 | 9651.2595 | 3675.9652 | 111.3338 | 7946.8768 | 7946.8768 | 1643.4279 | 165.3343 | 489.3418 | 1060.0394 | 631.832 | 173.3344 | 16038.3326 | 14849.7155 | 159581.9135 | 211702.4534 | 12275.2732 | 235254.7311 | 196503.6768 | 188082.9369 | 4170.6101 | 7902.1861 | 18586.7512 | 28543.194 | 63068.2561
31 | 8_2_21_UF_6_SP_LD |  | 06/08/2023 16:26:19 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_UF_6_SP_LD.385 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 24432.8933 | 8218.3646 | 480109.6514 | 260.6691 | 35023.555 | 4638.753 | 5511.7307 | 11337.1647 | 6742.9247 | 8074.2812 | 44404.9159 | 33587.8873 | 248253.0978 | 9545.1879 | 16193.8438 | 4684.1013 | 15033.9083 | 115435.959 | 56245.1754 | 60026.5935 | 1092.0418 | 76283.1326 | 76283.1326 | 3885.8618 | 746.0195 | 2402.8689 | 5251.6318 | 3236.7117 | 938.0308 | 125358.3067 | 112893.71 | 146648.2037 | 193026.7149 | 11320.4842 | 211799.357 | 175886.384 | 174758.1496 | 3808.5077 | 37435.6527 | 87623.2473 | 132818.9744 | 295851.1191
32 | 8_2_21_UF_6_SP_LSP |  | 06/08/2023 16:32:06 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_UF_6_SP_LSP.386 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 95307.535 | 36170.4002 | 1292131.5168 | 714.6846 | 69869.8048 | 15128.6737 | 40196.4725 | 24034.2018 | 19401.1658 | 25154.1292 | 136618.9371 | 98867.0746 | 398716.5047 | 25190.1933 | 50990.1763 | 13799.3284 | 34847.7857 | 273792.1192 | 132745.9161 | 122922.7607 | 2102.8215 | 150769.4302 | 150769.4302 | 10337.0729 | 2420.2051 | 7654.0501 | 16476.165 | 9508.1137 | 2928.3004 | 290702.406 | 263237.83 | 136689.0046 | 178097.9886 | 9751.3272 | 191979.4359 | 161286.7315 | 158762.707 | 3511.0981 | 110171.2088 | 261077.4516 | 386092.2611 | 853123.0062
33 | 8_2_21_UF_8 |  | 06/08/2023 16:37:54 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_UF_8.387 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 17885.8571 | 4630.0836 | 60751.5793 | 30.6667 | 32110.0611 | 3103.0039 | 1790.1122 | 9161.6064 | 1480.0767 | 265.3358 | 1270.0565 | 9228.9502 | 183998.0771 | 3775.1655 | 62.6668 | 166.001 | 4488.0384 | 34712.1241 | 8544.5549 | 34934.1157 | 655.3485 | 47302.8553 | 47302.8553 | 1116.7103 | 12.6667 | 56.6668 | 128.6673 | 319.4903 | 18.6667 | 40507.3598 | 37491.1362 | 149884.2088 | 201549.8253 | 11741.4903 | 219251.2124 | 180413.526 | 182606.8062 | 4045.2398 | 95.3337 | 212.0016 | 1120.0175 | 2476.0846
34 | 8_2_21_UF_9 |  | 06/08/2023 16:43:42 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_UF_9.388 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 12884.4807 | 4107.2571 | 59972.3008 | 38.6667 | 27057.6013 | 2557.5623 | 1366.0653 | 7740.7642 | 1022.0367 | 138.0007 | 611.3465 | 3952.718 | 88258.4691 | 2446.8765 | 58.6668 | 150.0008 | 3047.6584 | 24067.5933 | 7229.8291 | 23965.5324 | 430.0065 | 32748.1608 | 32748.1608 | 1062.7062 | 13.3333 | 42.6667 | 86.0003 | 209.4432 | 18.6667 | 27829.7526 | 25416.5908 | 148400.844 | 198288.0317 | 11520.6455 | 217387.5642 | 181235.0806 | 182653.6809 | 3924.5391 | 73.3335 | 171.3344 | 1247.3548 | 2768.1054
35 | LBLANK |  | 06/08/2023 16:49:31 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.389 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 6 | 0 | 5151.5964 | 2 | 880.694 | 634.0141 | 13.3333 | 0.66667 | 2 | 8.6667 | 42.0001 | 593.1509 | 18 | 1304.7262 | 2 | 20.6667 | 12.6667 | 95.3337 | 302.67 | -2317.0864 | 3.3333 | 264.6691 | 264.6691 | 961.3658 | 3.3333 | 8 | 11.3333 | 51.0214 | 2.6667 | 46.0001 | 29.3334 | 155367.6581 | 209031.676 | 12534.8534 | 229849.8427 | 188542.031 | 184944.8598 | 4180.6116 | 2 | 14.6667 | 16 | 50.6667
36 | CCV 1ppb TRACE |  | 06/08/2023 16:55:19 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.390 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1633.4267 | 703.3507 | 129439.2518 | 78.6669 | 3666.4707 | 873.3601 | 794.6888 | 322.0037 | 1165.3809 | 1588.7551 | 8775.3632 | 7189.6556 | 15195.4112 | 2808.2761 | 3357.728 | 934.6973 | 2307.5198 | 18010.6797 | 8748.6785 | 3413.3622 | 106.6671 | 7452.6104 | 7452.6104 | 1543.4167 | 153.3342 | 457.3407 | 1034.7042 | 562.575 | 195.3347 | 15898.1756 | 14725.5857 | 154180.2794 | 204695.4891 | 12720.9987 | 226998.6517 | 186328.8333 | 180473.103 | 4047.2401 | 7914.1926 | 18789.6868 | 27403.0059 | 61683.8841
37 | CCV 1ppb TRACE |  | 06/08/2023 17:01:37 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.391 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1636.7604 | 749.353 | 129482.4587 | 69.3335 | 3689.1441 | 924.6966 | 811.3565 | 326.0037 | 1218.052 | 1576.0871 | 8947.4681 | 7192.151 | 15153.3668 | 2858.9527 | 3381.0669 | 1032.704 | 2326.8562 | 18311.0624 | 8756.6831 | 3328.1543 | 112.6671 | 7561.3344 | 7561.3344 | 1514.0802 | 161.3342 | 516.676 | 1135.3785 | 572.2159 | 211.3349 | 16036.3302 | 14572.0951 | 154669.5814 | 208221.206 | 12073.7672 | 228916.9389 | 186871.6718 | 182822.6778 | 4047.907 | 7833.4806 | 18506.6483 | 27265.5626 | 61618.9678
38 | CCV 10 ppb MAJORS |  | 06/08/2023 17:07:25 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.392 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 15533.7757 | 6784.2783 | 644161.4149 | 383.3385 | 28262.6071 | 3002.3157 | 7511.3085 | 2884.9584 | 11066.285 | 15033.2394 | 83078.8913 | 55640.2805 | 138231.6537 | 14918.4538 | 32035.2151 | 8808.0481 | 20456.6408 | 160489.8573 | 76183.9329 | 54091.1023 | 1048.7053 | 70311.9453 | 70311.9453 | 6545.5001 | 1459.408 | 4536.0534 | 9868.0745 | 5479.261 | 1740.7727 | 150145.0617 | 134400.6042 | 147179.8263 | 193137.6825 | 11233.7491 | 213718.8203 | 177136.8634 | 174505.559 | 3791.1697 | 71375.8787 | 169504.3991 | 248410.2938 | 560144.4148
39 | CCV 10 ppb MAJORS |  | 06/08/2023 17:13:42 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.393 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 15689.9454 | 6712.2446 | 655558.2497 | 406.6725 | 28400.8995 | 3089.6675 | 7506.6387 | 2929.6338 | 11102.3127 | 15459.696 | 83190.2466 | 56578.2296 | 141542.4344 | 14993.8649 | 32822.9994 | 9019.5132 | 20857.2151 | 161246.9549 | 76595.4605 | 55096.2515 | 1038.0378 | 70600.7085 | 70600.7085 | 6791.6143 | 1471.4092 | 4743.4542 | 9978.82 | 5601.1302 | 1788.1119 | 151500.4339 | 137248.8188 | 150429.3186 | 197867.6336 | 11463.9321 | 218072.5314 | 180179.1539 | 175562.3409 | 3999.2267 | 72000.9941 | 170410.3798 | 253585.4152 | 556721.2458
40 | 8_2_21_UF_10 |  | 06/08/2023 17:19:31 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_UF_10.394 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 13398.2802 | 5821.8531 | 61108.432 | 36 | 32404.7109 | 4095.9205 | 1593.4222 | 10116.2496 | 1017.3696 | 102.667 | 438.6734 | 4088.6344 | 74859.6487 | 2521.556 | 68.0002 | 170.6677 | 2755.5992 | 20428.5966 | 8195.6838 | 27717.8092 | 532.6766 | 36960.4187 | 36960.4187 | 1072.707 | 12.6667 | 55.3335 | 108.6671 | 258.5869 | 20.6667 | 27235.9437 | 24818.2084 | 141714.8952 | 191441.6223 | 10568.5826 | 206090.6642 | 172985.7128 | 166552.644 | 3851.8527 | 54.6668 | 147.3341 | 1621.3686 | 3573.512
41 | 8_2_21_UF_11 |  | 06/08/2023 17:25:19 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_UF_11.395 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 12193.2018 | 5746.4904 | 54553.9794 | 31.3334 | 29199.1438 | 4026.5675 | 1573.4201 | 10140.9375 | 1088.7082 | 1198.0503 | 6544.8326 | 38592.9152 | 81928.2657 | 10920.8409 | 86.6669 | 165.3343 | 2320.1885 | 18794.355 | 6860.9808 | 26827.426 | 546.6771 | 37076.7188 | 37076.7188 | 1048.0385 | 11.3333 | 43.3334 | 90.667 | 234.5322 | 13.3333 | 25778.5744 | 23799.8095 | 138918.181 | 188629.9576 | 10415.2098 | 204933.4859 | 169878.2416 | 167941.6015 | 3778.4999 | 42.6667 | 102.0004 | 2638.7626 | 5887.1479
42 | 8_2_21_UF_11_DUP |  | 06/08/2023 17:31:09 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_UF_11_DUP.396 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 12109.131 | 5783.1706 | 53506.0238 | 28 | 28731.5362 | 3825.1791 | 1518.7476 | 9665.2706 | 1002.7019 | 88.6669 | 423.3396 | 2816.3639 | 80964.8507 | 2132.8259 | 58.6668 | 130.0006 | 2313.5208 | 18854.437 | 7092.427 | 26451.5199 | 479.3414 | 35830.8801 | 35830.8801 | 1036.7043 | 15.3333 | 45.3334 | 94.667 | 243.8063 | 24 | 25240.9461 | 22946.4162 | 141184.2324 | 187988.3203 | 11100.3119 | 204619.8945 | 172105.852 | 168189.8171 | 3686.4759 | 35.3334 | 98.0003 | 2612.7596 | 5787.7881
43 | 8_2_21_F_1 |  | 06/08/2023 17:36:57 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_F_1.397 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 7149.79 | 3836.5152 | 229717.6727 | 137.334 | 16171.8217 | 3833.8482 | 902.0285 | 8307.082 | 395.3388 | 81.3336 | 388.672 | 36285.7296 | 214394.8662 | 10437.1465 | 57.3334 | 250.0022 | 2538.2256 | 19718.9336 | 15198.7475 | 22818.9304 | 428.0065 | 32451.488 | 32451.488 | 990.0344 | 8 | 30 | 56.6668 | 201.9727 | 12.6667 | 81108.2681 | 74232.4065 | 143002.9033 | 191407.137 | 11253.7654 | 209945.4742 | 173891.2714 | 171254.2122 | 3697.1454 | 52.6668 | 145.3341 | 1171.352 | 2583.4224
44 | 8_2_21_F_6 |  | 06/08/2023 17:44:00 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_F_6.398 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 14483.339 | 4632.0844 | 178465.2212 | 100.667 | 11231.7475 | 3631.795 | 1594.089 | 9984.8229 | 1010.7024 | 604.0128 | 3325.0536 | 53774.3285 | 183798.4129 | 14556.0803 | 84.0003 | 226.0018 | 8003.5749 | 61954.0508 | 25823.3196 | 30874.5094 | 588.6788 | 41272.2705 | 41272.2705 | 988.0342 | 20 | 58.0001 | 104.0004 | 288.5176 | 25.3334 | 48629.2979 | 44126.7528 | 138960.7794 | 206382.0338 | 11087.6348 | 204228.8262 | 170392.8388 | 165946.9844 | 4103.9333 | 133.334 | 299.3365 | 2396.0792 | 5123.027
45 | 8_2_21_F_8 |  | 06/08/2023 17:49:48 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_F_8.399 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 16435.4503 | 4523.3827 | 194236.2505 | 132.6673 | 28225.1987 | 3112.3393 | 1836.7849 | 9323.0414 | 1410.0698 | 184.0012 | 1012.7026 | 33466.4294 | 192812.223 | 9779.349 | 70.0002 | 175.3344 | 8391.7981 | 64333.8731 | 16939.3709 | 35652.4813 | 660.682 | 47044.0111 | 47044.0111 | 1060.0394 | 12.6667 | 60.6668 | 96.667 | 333.8201 | 17.3333 | 42589.3977 | 39033.2733 | 141115.7477 | 196680.3371 | 10802.0825 | 204996.4034 | 173271.152 | 167617.0527 | 3868.524 | 80.0002 | 228.6685 | 2512.7516 | 5565.0871
46 | 8_2_21_F_8_DUP |  | 06/08/2023 17:55:36 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_F_8_DUP.400 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 17038.1608 | 4655.4259 | 197087.4531 | 111.3338 | 28105.6229 | 3199.6927 | 1811.4483 | 9869.4133 | 1419.4039 | 83.3336 | 423.3396 | 31247.417 | 191153.1213 | 9436.4519 | 85.3336 | 178.0011 | 8736.0043 | 64948.6462 | 18052.7328 | 35818.264 | 628.0139 | 47218.5805 | 47218.5805 | 1090.0417 | 20.6667 | 48.0001 | 132.6673 | 337.3048 | 24 | 42180.8663 | 38132.1748 | 143629.8085 | 190364.7214 | 9710.0266 | 202265.8802 | 173683.4356 | 172098.6823 | 3740.4897 | 116.0005 | 257.3357 | 2620.7604 | 5825.1275
47 | 8_2_21_F_8_DUP_LD |  | 06/08/2023 18:01:24 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_F_8_DUP_LD.401 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 16313.3108 | 4729.4495 | 306348.6836 | 196.668 | 34246.3366 | 3111.6723 | 1787.4452 | 9531.8468 | 1450.0737 | 126.6672 | 629.3472 | 70903.3153 | 195283.7466 | 18733.6091 | 79.3336 | 198.6681 | 8645.9511 | 65230.6012 | 20238.3276 | 35362.2732 | 616.68 | 46623.3196 | 46623.3196 | 1057.3725 | 18.6667 | 52.0001 | 117.3338 | 332.4975 | 21.3334 | 42217.6345 | 39324.0593 | 139953.0035 | 191332.3436 | 11197.7203 | 214642.7606 | 173848.8998 | 147442.4127 | 3777.1661 | 98.0003 | 232.6686 | 2762.1073 | 6069.8412
48 | 8_2_21_F_8_DUP_LD |  | 06/08/2023 18:07:40 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_F_8_DUP_LD.402 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 16742.4726 | 4598.0734 | 319023.7935 | 202.0014 | 32986.0419 | 3167.0179 | 1810.1149 | 9432.4482 | 1438.0724 | 134.0006 | 620.6802 | 71478.5463 | 195804.3189 | 19361.7936 | 72.6669 | 190.6679 | 8717.3258 | 66353.7469 | 20525.4043 | 35732.4986 | 592.0123 | 47218.5739 | 47218.5739 | 1042.705 | 20.6667 | 50.0001 | 114.0005 | 291.2654 | 13.3333 | 41960.2003 | 38816.6719 | 142691.7677 | 189445.2161 | 11046.2701 | 210031.9512 | 173362.9562 | 167332.3386 | 3637.1296 | 98.0003 | 252.6689 | 2750.103 | 6216.528
49 | 8_2_21_F_8_DUP_LSP |  | 06/08/2023 18:13:28 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_F_8_DUP_LSP.403 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 91406.1867 | 33971.6799 | 1114874.8423 | 648.6814 | 66882.2154 | 14520.0435 | 38097.3992 | 22149.1593 | 14462.6532 | 17929.2443 | 97483.51 | 97984.6373 | 340993.9804 | 25202.2289 | 36531.9861 | 9944.127 | 30037.5529 | 228912.1351 | 98190.9787 | 100004.7064 | 1768.1095 | 121970.4842 | 121970.4842 | 7506.6386 | 1753.441 | 5389.0164 | 11343.837 | 6712.7788 | 2018.8093 | 210792.4796 | 190858.5097 | 130290.8385 | 167895.5146 | 9833.3833 | 182819.0457 | 152266.5472 | 149065.8359 | 3362.3958 | 77632.4312 | 181758.399 | 269282.1865 | 585744.7678
50 | LBLANK |  | 06/08/2023 18:19:16 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.404 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 11.3333 | 2.6667 | 4116.593 | 2 | 1714.7697 | 680.6829 | 16 | 0 | 0.66667 | 11.3333 | 51.3334 | 577.7711 | 26 | 1346.7302 | 2 | 24.6667 | 15.3333 | 121.3339 | 312.0034 | -2575.3967 | 1.3333 | 71.3335 | 71.3335 | 1023.3701 | 0 | 1.3333 | 3.3333 | 52.1536 | 2 | 100.667 | 50.6668 | 157424.6455 | 209045.8127 | 12277.9406 | 237920.4186 | 191342.6014 | 182190.4784 | 4319.3198 | 2 | 19.3333 | 30.6667 | 58.6667
51 | CCV 1ppb TRACE |  | 06/08/2023 18:25:05 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.405 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1771.4432 | 751.3532 | 134348.1932 | 69.3335 | 3846.5179 | 973.3666 | 888.0276 | 338.004 | 1220.0523 | 1710.1024 | 9188.9555 | 7488.7224 | 15906.187 | 3094.3355 | 3516.4329 | 1037.3712 | 2522.8895 | 18626.1396 | 8861.4144 | 3963.7819 | 115.3338 | 7635.3741 | 7635.3741 | 1574.0867 | 177.3344 | 510.6758 | 1150.0463 | 672.9254 | 183.3345 | 16691.7466 | 14860.3924 | 155603.599 | 205001.1585 | 11812.885 | 230954.4828 | 189019.91 | 177227.3535 | 4156.6055 | 7852.8249 | 18680.8739 | 27344.2102 | 62242.2311
52 | CCV 1ppb TRACE |  | 06/08/2023 18:31:23 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.406 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1752.7743 | 730.0187 | 139737.4679 | 76.6669 | 3796.5048 | 947.3649 | 834.0244 | 324.6704 | 1297.3923 | 1696.7675 | 9130.9197 | 7309.2152 | 15587.1657 | 3127.0089 | 3551.1081 | 995.368 | 2479.5487 | 18794.3563 | 8738.0057 | 3926.2022 | 98.667 | 7758.1071 | 7758.1071 | 1555.4182 | 164.0009 | 500.6754 | 1068.7068 | 614.5869 | 177.3344 | 16018.9768 | 14654.1804 | 155818.6164 | 205852.185 | 12191.2006 | 229669.5428 | 189213.0606 | 183852.91 | 4202.6181 | 7800.7962 | 18641.49 | 27583.085 | 61410.6967
53 | LFB |  | 06/08/2023 18:37:11 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.407 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 15647.2318 | 6777.6074 | 679655.9113 | 393.3388 | 35017.535 | 3206.36 | 7493.9662 | 2994.314 | 11143.0119 | 15552.4659 | 82447.9427 | 56969.8441 | 140962.834 | 15109.989 | 33138.3914 | 9066.8764 | 21231.774 | 162404.6333 | 78584.8999 | 55484.1329 | 1049.3719 | 71081.0777 | 71081.0777 | 6806.9548 | 1494.0782 | 4602.7422 | 10102.907 | 5751.4792 | 1838.785 | 147711.145 | 136229.1656 | 148419.0478 | 192719.3654 | 11637.405 | 213285.1017 | 177100.4344 | 171743.0406 | 3847.8517 | 70822.4549 | 168267.8764 | 243277.2813 | 540246.1006
54 | CCV 10 ppb MAJORS |  | 06/08/2023 18:42:59 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.408 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 16266.5929 | 6852.3101 | 667008.6946 | 388.672 | 28730.8674 | 3191.0236 | 7691.4037 | 3012.9844 | 11327.8234 | 15795.3945 | 86267.0323 | 58996.4603 | 144982.0252 | 15387.6179 | 33251.9885 | 9307.6996 | 21857.3759 | 168050.7301 | 79251.8888 | 56243.4149 | 1042.0381 | 71169.517 | 71169.517 | 6824.9638 | 1559.4186 | 4744.7878 | 10114.9135 | 5532.0583 | 1866.1219 | 154897.2995 | 138504.3164 | 150614.591 | 194488.9655 | 11758.1732 | 214766.9989 | 178196.5142 | 168840.6463 | 3908.5348 | 73156.1809 | 172454.6909 | 249332.5576 | 553352.3173
55 | CCV 10 ppb MAJORS |  | 06/08/2023 18:49:21 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.409 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 16368.7057 | 6868.318 | 651651.3131 | 404.6724 | 28564.5312 | 3201.6921 | 7894.1811 | 3050.9927 | 11589.3677 | 16333.9994 | 86827.2188 | 58668.9264 | 145450.737 | 15556.4657 | 33701.7079 | 9576.5422 | 21883.4148 | 168504.6826 | 79066.209 | 55953.8052 | 1091.375 | 71355.7662 | 71355.7662 | 6636.2087 | 1550.7509 | 4680.1002 | 10078.8878 | 5580.8796 | 1840.1186 | 155080.6221 | 138104.5251 | 164932.4617 | 194418.8386 | 11044.2771 | 216782.4593 | 176700.1186 | 170397.5888 | 3857.8543 | 73233.9127 | 172665.8728 | 245357.1361 | 551985.896
56 | LBLANK |  | 06/08/2023 19:26:31 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.410 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2 | 0 | 4284.6427 | 2 | 798.689 | 529.3432 | 14.6667 | 0 | 0.66667 | 13.3333 | 49.3334 | 604.8692 | 22.6667 | 1314.7272 | 1.3333 | 12.6667 | 15.3333 | 99.3337 | 440.6736 | -1898.5167 | 0.66667 | 38.0001 | 38.0001 | 1000.7019 | 0.66667 | 3.3333 | 6 | 79.192 | 0 | 54.0001 | 45.3334 | 153146.8073 | 198249.6115 | 11550.6679 | 226570.886 | 187776.0628 | 177993.3655 | 4279.3082 | 2 | 15.3333 | 20 | 45.3334
57 | PE ICV |  | 06/08/2023 19:32:19 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\PE ICV.411 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 15868.8107 | 6792.9488 | 689777.3171 | 423.3396 | 56955.3145 | 3143.6792 | 7595.352 | 2946.9711 | 11189.0476 | 15571.8164 | 83633.4293 | 58481.7035 | 143556.4151 | 15490.3948 | 32625.212 | 9043.5304 | 21351.2782 | 163343.2568 | 76404.4465 | 55053.7404 | 950.6983 | 69525.4602 | 69525.4602 | 5082.9046 | 1444.0731 | 4719.4461 | 9982.8211 | 5610.9977 | 1788.112 | 145906.6792 | 130261.9687 | 150832.7773 | 189196.6808 | 11243.7564 | 210836.4536 | 178062.3108 | 166278.8323 | 3805.1743 | 71787.9662 | 169828.8631 | 248177.4692 | 550444.6597
58 | 8_2_21_F_9 |  | 06/08/2023 19:38:07 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_F_9.412 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 11354.5108 | 4181.9454 | 176972.9342 | 116.6671 | 23769.0913 | 2736.2621 | 1448.7401 | 8287.0698 | 1062.7063 | 202.0014 | 933.3638 | 29858.1504 | 98673.6348 | 8964.1454 | 60.6668 | 166.001 | 6700.9049 | 49232.7126 | 12101.7906 | 25874.5783 | 480.0081 | 34396.0252 | 34396.0252 | 1089.3749 | 8.6667 | 36 | 83.3336 | 219.0942 | 14.6667 | 30224.6068 | 27630.0297 | 146125.5808 | 189581.738 | 11090.9724 | 207808.5665 | 173119.4761 | 167317.4741 | 3868.5244 | 84.0003 | 222.0017 | 5518.4157 | 11987.9539
59 | 8_2_21_F_10 |  | 06/08/2023 19:43:55 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_F_10.413 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 11183.0446 | 5719.1447 | 280248.7553 | 162.6676 | 30333.511 | 4173.2765 | 1604.0901 | 10373.7666 | 1031.3706 | 94.0003 | 436.0067 | 24493.6582 | 89167.4292 | 7755.4393 | 92.0003 | 183.3345 | 4980.8682 | 37789.9248 | 18276.3502 | 28799.6313 | 554.6775 | 38246.4634 | 38246.4634 | 1141.3789 | 10 | 38.0001 | 105.3337 | 284.4087 | 20.6667 | 28889.1812 | 26346.9435 | 143542.9568 | 186466.3898 | 10458.5104 | 203880.5971 | 174174.7306 | 165691.4187 | 3763.1623 | 66.0002 | 188.6679 | 3504.1695 | 7864.8563
60 | 8_2_21_F_10_DUP |  | 06/08/2023 19:49:43 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_F_10_DUP.414 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 11539.9932 | 5672.4595 | 276622.1942 | 154.6675 | 29646.7316 | 4208.6201 | 1589.4218 | 10219.6554 | 992.7012 | 384.6719 | 1954.8004 | 30375.6933 | 89839.6702 | 8925.4543 | 74.6669 | 185.3345 | 4934.1854 | 38441.658 | 17592.8269 | 28975.0913 | 540.6769 | 37784.5694 | 37784.5694 | 1076.0405 | 13.3333 | 44.6667 | 102.0004 | 283.1455 | 22.6667 | 29144.3695 | 26477.8492 | 141721.5329 | 186065.4945 | 10902.1588 | 204107.7743 | 173744.2505 | 167025.4711 | 3844.5197 | 78.6669 | 174.6677 | 3478.8341 | 7621.4598
61 | 8_2_21_F_11 |  | 06/08/2023 19:55:32 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_F_11.415 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10912.1728 | 5872.5411 | 300675.6393 | 175.3344 | 26917.3421 | 4204.6187 | 1583.4216 | 10365.7597 | 1079.3742 | 119.3338 | 494.6753 | 26165.3813 | 98942.7931 | 7856.8267 | 76.0002 | 190.6679 | 5378.3457 | 42162.1292 | 11859.5889 | 28753.5326 | 495.3419 | 37570.6757 | 37570.6757 | 1088.0415 | 20.6667 | 52.0001 | 99.3337 | 266.9396 | 15.3333 | 26889.9508 | 24481.6271 | 142883.0785 | 185509.3924 | 10924.8426 | 202828.5259 | 157711.7236 | 164075.1074 | 3740.4897 | 46.6667 | 117.3338 | 9957.3576 | 21848.5559
62 | 8_2_21_F_11_SP |  | 06/08/2023 20:01:20 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_F_11_SP.416 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 19396.4934 | 9286.3507 | 636444.491 | 396.0055 | 42238.361 | 5429.6984 | 5641.7807 | 11816.8867 | 6919.6758 | 8322.4235 | 45646.1473 | 61887.1821 | 172520.8946 | 15940.2237 | 16359.3619 | 4606.0759 | 15350.9119 | 117633.7507 | 48849.3851 | 57442.6324 | 1023.37 | 72557.7983 | 72557.7983 | 3805.1734 | 780.0213 | 2419.5383 | 5016.881 | 3032.1585 | 886.0275 | 103903.8912 | 94861.9274 | 136851.2987 | 182981.896 | 10730.0293 | 199557.4819 | 167023.4015 | 164584.0651 | 3646.4654 | 36471.1842 | 85404.55 | 139402.5615 | 306948.6343
63 | 8_2_21_W_1 |  | 06/08/2023 20:07:08 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_W_1.417 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1479.4099 | 791.3553 | 515320.1121 | 348.0042 | 9356.3972 | 1304.7263 | 215.335 | 1682.7659 | 88.0003 | 60.0001 | 277.336 | 45304.1864 | 66444.8343 | 12233.9031 | 33.3334 | 90.667 | 946.6981 | 7793.4586 | 2557.5623 | 3662.0627 | 85.3336 | 7007.7189 | 7007.7189 | 992.0346 | 4.6667 | 8 | 19.3333 | 63.1753 | 2.6667 | 16621.6671 | 14965.8381 | 150415.0533 | 193127.6652 | 11473.2722 | 218350.7842 | 184506.5458 | 173453.4059 | 3837.8488 | 18.6667 | 62.6668 | 1022.0145 | 2132.7284
64 | 8_2_21_W_6 |  | 06/08/2023 20:12:56 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_W_6.418 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2944.9718 | 932.0304 | 910491.0518 | 576.0116 | 10279.032 | 1235.3868 | 372.6715 | 2058.1485 | 244.0021 | 86.0003 | 422.6729 | 85101.4398 | 81812.9353 | 22066.3654 | 42.6667 | 48.0001 | 3237.7003 | 24486.3046 | 8995.4981 | 5404.4333 | 136.0007 | 9353.0628 | 9353.0628 | 1032.704 | 5.3333 | 12 | 27.3334 | 97.4035 | 6.6667 | 10426.4711 | 9267.0058 | 149818.8704 | 196984.4239 | 11312.4788 | 217093.8882 | 180223.854 | 172807.0388 | 3889.8629 | 54.0001 | 126.6672 | 3468.165 | 7392.0844
65 | 8_2_21_W_6_DUP |  | 06/08/2023 20:18:44 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_W_6_DUP.419 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2910.2964 | 946.0314 | 806455.4457 | 490.6751 | 10005.5038 | 1213.3849 | 278.0027 | 1970.1359 | 219.335 | 70.0002 | 342.6708 | 73641.7963 | 77097.4953 | 19588.7606 | 28.6667 | 74.0002 | 2952.972 | 22996.4955 | 7786.1227 | 5152.901 | 112.6671 | 9089.5576 | 9089.5576 | 984.0343 | 1.3333 | 17.3333 | 25.3334 | 84.585 | 1.3333 | 10200.976 | 9140.2573 | 149390.6352 | 192939.5012 | 11236.4177 | 216034.316 | 179134.816 | 171321.8523 | 3919.8712 | 53.3334 | 114.6671 | 3186.8068 | 6638.6065
66 | LBLANK |  | 06/08/2023 20:24:33 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.420 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.66667 | 0 | 4568.7354 | 2.6667 | 1269.3898 | 615.3466 | 17.3333 | 0 | 1.3333 | 12.6667 | 39.3334 | 552.9718 | 15.3333 | 1310.7269 | 1.3333 | 6.6667 | 10 | 103.3337 | 307.3366 | -2068.4102 | 0.66667 | 26.6667 | 26.6667 | 956.032 | 0.66667 | 1.3333 | 5.3333 | 52.6328 | 0 | 23.3334 | 8.6667 | 149603.9947 | 199196.0122 | 11633.4032 | 220990.5323 | 183355.1755 | 176524.0803 | 3990.5574 | 2 | 14.6667 | 18.6667 | 39.3334
67 | CCV 1ppb TRACE |  | 06/08/2023 20:30:21 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.421 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1710.7693 | 684.6831 | 132468.7534 | 81.3336 | 3627.794 | 864.0262 | 820.0236 | 319.3369 | 1266.7229 | 1631.4265 | 8748.6781 | 7221.7023 | 14869.7366 | 2916.9646 | 3409.0734 | 954.6986 | 2309.5202 | 17929.2444 | 8518.5403 | 3935.3608 | 112.0004 | 7242.5021 | 7242.5021 | 1586.7549 | 137.334 | 485.3416 | 1069.3735 | 619.2754 | 178.6678 | 16077.7107 | 13982.8401 | 145024.7405 | 197237.8309 | 11592.704 | 218079.963 | 183122.8984 | 174560.7156 | 3929.207 | 7744.1007 | 18193.5781 | 27678.4585 | 60941.5627
68 | CCV 1ppb TRACE |  | 06/08/2023 20:36:39 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.422 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1662.0967 | 701.3506 | 130098.4478 | 78.0002 | 3474.4225 | 902.0286 | 768.0207 | 308.67 | 1222.719 | 1622.0922 | 8925.457 | 7222.9592 | 15003.2078 | 3056.3269 | 3588.451 | 971.3664 | 2409.5366 | 18520.6656 | 8531.2134 | 3797.4108 | 118.0005 | 7372.5689 | 7372.5689 | 1586.755 | 140.6674 | 464.6742 | 1063.3729 | 599.2954 | 206.0015 | 15922.8708 | 14221.7425 | 152999.6571 | 197114.6703 | 11760.8397 | 222417.4376 | 181983.2261 | 173412.0472 | 4079.9163 | 7786.1219 | 18449.2405 | 27327.4895 | 60792.9873
69 | CCV 10 ppb MAJORS |  | 06/08/2023 20:42:27 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.423 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 15941.5587 | 6542.165 | 656825.2517 | 434.6733 | 27654.7489 | 3066.3291 | 7403.9185 | 2902.2949 | 11055.6098 | 15215.4344 | 83271.3428 | 57145.922 | 140502.9318 | 14613.4733 | 31707.8476 | 8843.4033 | 20757.0721 | 162323.6805 | 77469.4948 | 54615.3472 | 982.7005 | 68844.8432 | 68844.8432 | 6558.8386 | 1472.0759 | 4506.7111 | 9715.9701 | 5428.1953 | 1786.1117 | 150253.474 | 135032.5695 | 148737.1186 | 186170.5961 | 11652.7747 | 207116.6258 | 175274.824 | 165644.8926 | 3819.1771 | 71292.8226 | 168637.5499 | 254411.3333 | 550363.1497
70 | CCV 10 ppb MAJORS |  | 06/08/2023 20:48:44 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.424 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 16101.7363 | 6582.1829 | 654651.646 | 398.0055 | 28113.6397 | 3109.6718 | 7597.3531 | 2966.308 | 11248.4294 | 15547.1241 | 84239.6469 | 57145.5181 | 141691.9274 | 15070.6165 | 33251.9916 | 9101.5663 | 21299.201 | 162612.9561 | 77326.734 | 55835.2583 | 1072.0403 | 69859.7344 | 69859.7344 | 6800.2848 | 1470.0757 | 4694.1045 | 10026.1849 | 5504.1473 | 1791.4457 | 150577.457 | 137376.7684 | 148366.8781 | 188017.1987 | 11362.518 | 209434.1675 | 176612.7415 | 165476.9444 | 3781.8339 | 70960.4692 | 169558.3371 | 252875.6706 | 542510.2258
71 | 8_2_21_W_8 |  | 06/08/2023 20:54:33 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_W_8.425 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3397.0705 | 956.032 | 535408.8806 | 324.6704 | 14446.6348 | 1164.0474 | 350.671 | 1982.8044 | 304.6699 | 121.3339 | 610.013 | 58583.1928 | 101585.2809 | 15659.9131 | 32.6667 | 57.3335 | 2744.2636 | 20650.25 | 5574.4208 | 6451.1089 | 141.334 | 10393.1126 | 10393.1126 | 1030.0372 | 3.3333 | 18 | 33.3334 | 95.0626 | 5.3333 | 8767.3562 | 7846.8216 | 151008.6145 | 196944.9727 | 11736.8207 | 209014.0186 | 179169.2427 | 178018.3857 | 3943.878 | 34 | 103.3337 | 1970.0546 | 4286.9219
72 | 8_2_21_W_9 |  | 06/08/2023 21:00:21 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_W_9.426 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2323.5223 | 856.0257 | 482351.2412 | 294.6697 | 12947.2 | 1028.7038 | 276.6694 | 1726.7711 | 202.0014 | 46.0001 | 250.6689 | 44438.0462 | 63444.6019 | 12309.9685 | 25.3334 | 53.3334 | 1769.443 | 14107.6295 | 3688.4762 | 3812.8475 | 98.667 | 7143.7864 | 7143.7864 | 976.7002 | 4 | 11.3333 | 18.6667 | 47.1044 | 2 | 6164.6632 | 5603.0987 | 153587.9282 | 195787.9059 | 11634.7376 | 214507.6995 | 181533.241 | 171918.3455 | 4014.5642 | 29.3334 | 72.0002 | 1173.3519 | 2620.7608
73 | 8_2_21_W_9_DUP |  | 06/08/2023 21:06:09 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_W_9_DUP.427 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2304.8527 | 850.0253 | 450362.4136 | 319.3369 | 11654.7524 | 1020.0366 | 282.6695 | 1603.4234 | 210.0016 | 38.0001 | 142.0007 | 40685.6421 | 50069.5965 | 11157.689 | 18.6667 | 56.6668 | 1784.1114 | 13587.1262 | 3789.1692 | 3996.5387 | 98.667 | 7147.1216 | 7147.1216 | 1096.7088 | 1.3333 | 12.6667 | 20.6667 | 60.883 | 6 | 6250.0336 | 5535.0723 | 151475.7488 | 199085.0722 | 11688.1129 | 218782.5761 | 181715.0784 | 172982.3857 | 3957.2158 | 23.3334 | 70.0002 | 1132.0181 | 2475.4161
74 | 8_2_21_W_10 |  | 06/08/2023 21:11:57 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_W_10.428 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2299.5188 | 1158.7137 | 1584179.5799 | 975.3667 | 15459.7023 | 1305.3931 | 348.6709 | 2069.4832 | 225.3351 | 63.3335 | 324.6704 | 108826.4961 | 90091.8775 | 28324.7347 | 53.3334 | 58.0001 | 2222.84 | 17018.1322 | 9334.3823 | 4588.569 | 110.0004 | 8205.0234 | 8205.0234 | 1060.7061 | 4.6667 | 8.6667 | 31.3334 | 99.7997 | 2.6667 | 5791.1736 | 5329.6609 | 153897.9236 | 192903.149 | 11392.5414 | 214494.3991 | 172239.4765 | 171600.1991 | 3902.5331 | 48.6668 | 114.6671 | 3278.1482 | 7248.0517
75 | 8_2_21_W_11 |  | 06/08/2023 21:17:45 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_W_11.429 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2138.827 | 1170.7147 | 2615601.8029 | 1446.74 | 17152.968 | 1326.0619 | 300.0032 | 2026.1438 | 274.6693 | 82.0002 | 450.6738 | 167446.8141 | 148854.9294 | 41554.3526 | 64.6668 | 59.3335 | 2782.9377 | 22225.9473 | 16541.5718 | 5126.3813 | 104.0004 | 8406.4731 | 8406.4731 | 1075.3738 | 2 | 9.3333 | 29.3334 | 87.0731 | 7.3333 | 5593.0947 | 4999.5418 | 151160.2085 | 191623.62 | 11433.9086 | 212452.1725 | 179374.0728 | 170905.0172 | 3917.2039 | 81.3336 | 162.6676 | 5119.6952 | 11071.0169
76 | 5_10_23_DB |  | 06/08/2023 21:23:33 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_10_23_DB.430 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4 | 0 | 32563.0753 | 19.3333 | 13008.5882 | 612.0133 | 11.3333 | 0 | 0 | 39.3334 | 127.3339 | 1447.0289 | 38.6667 | 1466.742 | 1.3333 | 14 | 101.3337 | 716.018 | 294.6697 | -2076.2931 | 0.66667 | 32 | 32 | 996.0347 | 0.66667 | 1.3333 | 4 | 46.5028 | 2 | 10.6667 | 6 | 148073.5955 | 193943.6244 | 11582.027 | 218987.3978 | 181805.5748 | 170013.9311 | 4001.8943 | 0.66667 | 7.3333 | 92.6668 | 211.334
77 | 5_10_23_MFB |  | 06/08/2023 21:29:21 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_10_23_MFB.431 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 7708.079 | 3291.7126 | 364465.6808 | 216.6683 | 33012.7876 | 1746.1068 | 3642.4643 | 1410.7364 | 5525.0687 | 7808.1333 | 41934.1287 | 30639.0735 | 70150.4837 | 8593.2519 | 16003.6297 | 4337.3251 | 10640.6291 | 81800.204 | 38498.4733 | 25995.0315 | 480.0081 | 33634.8877 | 33634.8877 | 3863.8561 | 712.6845 | 2315.5211 | 4857.493 | 2733.4175 | 870.0265 | 72634.2245 | 65694.7255 | 151965.338 | 187628.8113 | 10844.7824 | 211356.8588 | 177157.7026 | 165793.4124 | 3833.181 | 35681.8397 | 83755.4782 | 129724.2364 | 285146.7978
78 | 5_10_23_T-239 |  | 06/08/2023 21:35:10 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_10_23_T-239.432 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 71140.0475 | 7985.5651 | 7675754.473 | 4317.9859 | 38773.2137 | 744.6861 | 2748.9312 | 13721.2535 | 2020.8097 | 2046.1466 | 10881.4771 | 221546.8421 | 47051.3605 | 55533.1075 | 4426.0192 | 1400.0686 | 1550.7509 | 12758.363 | 43539.5825 | 11156.9605 | 217.335 | 15304.1946 | 15304.1946 | 3540.4389 | 266.0025 | 827.3573 | 1789.4454 | 1182.049 | 330.0038 | 15087.2961 | 13382.9335 | 142701.2124 | 175083.7789 | 10433.1448 | 200398.5356 | 165380.551 | 153588.5626 | 3561.1106 | 23037.8948 | 52648.8744 | 30254.5107 | 64509.1673
79 | 5_10_23_DB |  | 06/08/2023 21:40:58 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_10_23_DB.433 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 6 | 0 | 33753.8316 | 22.6667 | 14835.6998 | 629.3473 | 12.6667 | 2 | 2 | 13.3333 | 32.6667 | 943.5782 | 21.3333 | 1376.733 | 1.3333 | 13.3333 | 65.3335 | 500.0088 | 261.3357 | -1979.6352 | 1.3333 | 28 | 28 | 1042.7047 | 0 | 0.66667 | 6.6667 | 39.7114 | 0.66667 | 27.3334 | 9.3333 | 173509.0536 | 206492.0689 | 11761.5104 | 229050.3853 | 190236.6092 | 177987.5134 | 4372.0026 | 2.6667 | 7.3333 | 62.6667 | 154.0003
80 | 8_16_21_UF_1 |  | 06/08/2023 21:46:46 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_UF_1.434 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 15151.3747 | 3761.162 | 135782.4811 | 88.6669 | 67466.2839 | 6525.4913 | 1598.7562 | 10589.2578 | 498.0087 | 120.6672 | 633.3475 | 3893.261 | 177215.7845 | 2590.9017 | 92.667 | 295.3364 | 1466.7421 | 11471.9377 | 7059.0778 | 33436.9043 | 598.0126 | 42353.3606 | 42353.3606 | 1186.716 | 7.3333 | 27.3334 | 66.0002 | 252.4451 | 11.3333 | 262360.1484 | 237557.7379 | 146394.5047 | 183870.3976 | 10848.1206 | 204865.8915 | 170801.193 | 164289.3307 | 3816.5098 | 57.3334 | 144.6674 | 433.336 | 840.0094
81 | LBLANK |  | 06/08/2023 21:52:35 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.435 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.6667 | 0 | 4590.0726 | 2.6667 | 2623.5743 | 596.0125 | 13.3333 | 1.3333 | 1.3333 | 12 | 46.6667 | 574.4646 | 14 | 1300.0592 | 0.66667 | 14.6667 | 11.3333 | 103.3337 | 316.6702 | -2053.107 | 0.66667 | 105.3337 | 105.3337 | 994.7013 | 1.3333 | 3.3333 | 3.3333 | 34.1939 | 2 | 47.3334 | 44.0001 | 154284.8747 | 199157.9982 | 11917.6367 | 222626.8242 | 185305.4694 | 176873.9887 | 4075.9151 | 0 | 9.3333 | 14.6667 | 38.6667
82 | CCV 1ppb TRACE |  | 06/08/2023 21:58:24 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.436 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1686.7665 | 721.3516 | 132492.3111 | 94.667 | 4159.9403 | 852.6921 | 809.3563 | 314.6701 | 1210.0514 | 1675.4317 | 8928.1225 | 7368.5077 | 15453.6875 | 2959.6399 | 3488.426 | 997.3682 | 2352.194 | 18351.7828 | 8511.8691 | 3901.5131 | 108.0004 | 7382.5751 | 7382.5751 | 1609.424 | 156.6675 | 485.3416 | 1017.3696 | 637.7874 | 203.3348 | 15963.5811 | 14313.8396 | 153867.6887 | 195447.924 | 11630.7343 | 218924.7673 | 183883.909 | 174643.8236 | 4020.5657 | 7757.4406 | 18082.105 | 27795.3284 | 61397.2007
83 | CCV 1ppb TRACE |  | 06/08/2023 22:04:41 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.437 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1687.433 | 709.351 | 135050.2088 | 76.0002 | 3825.8457 | 872.6934 | 786.0219 | 308.0033 | 1224.0525 | 1677.432 | 8946.1337 | 7205.6019 | 15092.638 | 3023.6534 | 3526.4353 | 930.6971 | 2512.221 | 18457.9222 | 8721.3287 | 3713.3572 | 109.3338 | 7369.2338 | 7369.2338 | 1661.4301 | 164.6676 | 472.0079 | 1058.7059 | 629.3354 | 180.6678 | 15812.7471 | 14507.363 | 157807.5934 | 196153.6249 | 11787.5288 | 220287.0751 | 175559.214 | 186761.9388 | 3989.8908 | 7664.7224 | 17894.5345 | 28605.1991 | 62101.9641
84 | LFB |  | 06/08/2023 22:10:29 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.438 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 15487.7269 | 6432.7814 | 668381.7526 | 424.673 | 34357.947 | 3040.9905 | 6975.0371 | 2839.6156 | 10786.7376 | 14834.3648 | 82782.5004 | 57177.129 | 137820.9812 | 14345.8671 | 31697.1296 | 8646.6163 | 20316.4363 | 161815.9673 | 78025.8392 | 55481.997 | 1028.037 | 68582.9085 | 68582.9085 | 6728.2513 | 1459.4079 | 4670.7637 | 9907.4345 | 5682.7128 | 1808.1145 | 149971.1359 | 132301.8475 | 148815.4723 | 186452.3577 | 11102.3126 | 210449.5088 | 173192.189 | 165984.0832 | 3772.4981 | 70617.4451 | 166945.2549 | 248588.4594 | 535271.8238
85 | CCV 10 ppb MAJORS |  | 06/08/2023 22:16:17 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.439 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 15971.5992 | 6722.9149 | 654144.4322 | 409.3392 | 29358.8083 | 3078.3319 | 7580.011 | 3030.9886 | 11219.0714 | 15916.8624 | 85704.3309 | 58191.0843 | 143442.8289 | 15025.8984 | 33272.0379 | 9112.9062 | 21124.2789 | 164294.0974 | 77446.0373 | 55763.5251 | 1023.3701 | 69429.0021 | 69429.0021 | 6868.3179 | 1454.7407 | 4520.7152 | 9927.4488 | 5427.1428 | 1802.7805 | 151980.2909 | 137132.5579 | 150395.0798 | 188900.064 | 11344.5042 | 208679.1408 | 174541.9607 | 167381.5579 | 3835.8485 | 70961.8089 | 168068.2014 | 253644.47 | 550438.7248
86 | CCV 10 ppb MAJORS |  | 06/08/2023 22:22:34 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.440 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 15913.5255 | 6741.5906 | 653341.5897 | 394.6721 | 28343.4344 | 3085.6666 | 7596.0194 | 3001.6488 | 11130.3346 | 15697.2871 | 86847.2166 | 58525.017 | 142683.7073 | 14997.8704 | 33213.2349 | 8915.4494 | 21079.5457 | 166357.051 | 78766.5523 | 56426.9532 | 1056.7058 | 69990.3756 | 69990.3756 | 6953.0255 | 1507.413 | 4748.7899 | 9876.7474 | 5615.0517 | 1746.7735 | 153736.9029 | 137982.5081 | 149959.6381 | 187757.9117 | 11565.3474 | 211660.6552 | 174267.939 | 166663.2808 | 3884.5283 | 71946.0468 | 169337.0609 | 253874.3085 | 550032.3795
87 | 8_16_21_UF_6 |  | 06/08/2023 22:28:23 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_UF_6.441 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 16269.927 | 4123.2617 | 77924.6465 | 58.0001 | 39245.1668 | 4720.1136 | 1456.741 | 9890.0895 | 881.3605 | 93.3336 | 430.0065 | 2975.9912 | 155410.7991 | 2270.1806 | 68.6668 | 614.6799 | 3259.0389 | 25416.5916 | 16283.943 | 29800.7599 | 536.6768 | 38061.3083 | 38061.3083 | 1221.3856 | 15.3333 | 36 | 100.667 | 315.9487 | 12.6667 | 129624.1231 | 117785.611 | 145853.4517 | 183913.6891 | 11098.3094 | 203017.8234 | 168273.9935 | 166532.4473 | 3751.8262 | 103.3337 | 235.3353 | 2726.1019 | 5918.4766
88 | 8_16_21_UF_8 |  | 06/08/2023 22:34:11 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_UF_8.442 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 15984.2728 | 4332.6573 | 76675.8963 | 41.3334 | 51304.642 | 3911.8695 | 1540.7499 | 9136.256 | 1068.7067 | 144.6674 | 650.6815 | 4183.8062 | 158127.7566 | 2435.5411 | 59.3335 | 348.0042 | 3494.4277 | 27978.7156 | 10194.303 | 29712.3609 | 562.6778 | 37635.5131 | 37635.5131 | 1158.7137 | 11.3333 | 31.3334 | 78.0002 | 287.7314 | 17.3333 | 102553.4823 | 93089.6829 | 144637.8543 | 183308.6429 | 10733.3644 | 201475.4643 | 166790.8451 | 166001.7068 | 3704.4804 | 82.0002 | 181.3345 | 1227.3539 | 2694.7647
89 | 8_16_21_UF_9 |  | 06/08/2023 22:39:59 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_UF_9.443 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 13618.4883 | 4126.5964 | 104265.1169 | 64.0001 | 50693.1253 | 3466.4205 | 1228.7197 | 8220.3656 | 905.3621 | 94.667 | 484.0082 | 2591.081 | 86829.0811 | 1998.8066 | 49.3334 | 642.6811 | 2597.5696 | 20825.838 | 5540.4081 | 22517.421 | 346.6709 | 29815.0805 | 29815.0805 | 1154.7136 | 10 | 37.3334 | 65.3335 | 215.0129 | 8 | 67466.2726 | 62707.3333 | 146112.1135 | 183269.5479 | 10950.2069 | 202531.2955 | 166667.9997 | 164515.2169 | 3737.8223 | 46.0001 | 122.6672 | 760.0078 | 1554.0335
90 | 8_16_21_UF_10 |  | 06/08/2023 22:45:47 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_UF_10.444 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 13156.7301 | 5587.7594 | 138689.6431 | 77.3335 | 55656.2095 | 3956.5478 | 1394.0681 | 9983.4899 | 917.3629 | 140.0007 | 752.6865 | 5348.7127 | 50321.8147 | 2720.9258 | 46.0001 | 248.0022 | 2339.525 | 18701.5701 | 8146.9908 | 24222.5541 | 404.0058 | 31638.3294 | 31638.3294 | 1166.7144 | 18 | 44.6667 | 85.3336 | 238.1366 | 18.6667 | 32825.6725 | 30065.6062 | 145006.2357 | 184453.1832 | 10659.9811 | 200308.5813 | 164348.2471 | 164848.3349 | 3745.8245 | 54.6668 | 138.0007 | 1330.0242 | 2911.4493
91 | 8_16_21_UF_11 |  | 06/08/2023 22:51:35 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_UF_11.445 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 11870.2633 | 5499.0588 | 64359.3888 | 45.3334 | 49996.0268 | 3970.5527 | 1412.7366 | 9142.2632 | 888.0276 | 74.0002 | 328.6705 | 1882.8345 | 34027.8161 | 1910.1278 | 33.3334 | 122.0005 | 1882.7907 | 15047.255 | 7331.2142 | 23242.1965 | 412.006 | 30188.5303 | 30188.5303 | 1147.3794 | 13.3333 | 40.6667 | 79.3336 | 199.479 | 16 | 26457.812 | 24269.2646 | 141989.4967 | 182475.4969 | 10968.2289 | 199627.873 | 164480.1972 | 162984.1069 | 3730.4874 | 20.6667 | 66.0002 | 1473.3631 | 3086.799
92 | 8_16_21_F_1 |  | 06/08/2023 22:57:23 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_F_1.446 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 13473.6867 | 3845.851 | 1086811.6599 | 618.6801 | 39963.1551 | 6296.054 | 1564.0857 | 10308.3863 | 442.6736 | 103.3337 | 496.0086 | 30011.4019 | 196217.8169 | 8577.2419 | 104.6671 | 294.6697 | 1986.1382 | 16329.3275 | 8408.4744 | 34525.0852 | 616.0133 | 44194.9247 | 44194.9247 | 1185.3826 | 8.6667 | 29.3334 | 70.0002 | 263.5899 | 12.6667 | 273214.0746 | 249334.5618 | 143605.6724 | 174878.3571 | 10888.1484 | 198571.1648 | 166624.8557 | 154279.6997 | 3571.781 | 54.0001 | 118.0005 | 700.6735 | 1558.6999
93 | 8_16_21_F_1_DUP |  | 06/08/2023 23:03:11 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_F_1_DUP.447 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 13260.8188 | 3856.5206 | 159753.7113 | 98.0003 | 38882.1768 | 6344.0758 | 1537.4162 | 10404.4545 | 431.3399 | 106.0004 | 538.0102 | 29531.1219 | 194600.6735 | 8365.1154 | 96.667 | 276.6694 | 2202.8365 | 17755.0275 | 8310.4168 | 33724.3552 | 590.0122 | 43472.0491 | 43472.0491 | 1089.375 | 8.6667 | 32 | 72.0002 | 226.1522 | 10.6667 | 268986.7815 | 244506.0172 | 143505.8625 | 183033.7676 | 10559.2349 | 198436.695 | 166562.1039 | 165975.1206 | 3639.1303 | 48.6667 | 122.6672 | 863.3432 | 1916.7185
94 | 8_16_21_F_1_DUP_LD |  | 06/08/2023 23:08:58 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_F_1_DUP_LD.448 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 13006.5863 | 3679.8076 | 243381.3667 | 125.3339 | 42907.0149 | 6322.0653 | 1520.7476 | 10243.0042 | 414.006 | 112.6671 | 588.0121 | 30332.8466 | 191902.4776 | 8358.4447 | 115.3338 | 300.0032 | 2223.5065 | 18339.0988 | 8904.1082 | 34778.9188 | 586.6788 | 43563.0041 | 43563.0041 | 1086.0413 | 7.3333 | 23.3334 | 58.0001 | 254.2129 | 13.3333 | 266134.8024 | 246061.2205 | 142068.9374 | 180872.5584 | 11028.2656 | 199426.4782 | 165730.5151 | 164186.149 | 3695.1445 | 51.3334 | 138.6673 | 866.6768 | 1982.7193
95 | 8_16_21_F_1_DUP_LD |  | 06/08/2023 23:15:14 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_F_1_DUP_LD.449 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 13220.785 | 3813.8425 | 245000.604 | 133.334 | 42951.8068 | 6264.7092 | 1558.7517 | 10393.1128 | 451.3405 | 120.6672 | 589.3455 | 30351.224 | 194102.6236 | 8900.1058 | 105.3337 | 331.3372 | 2278.8487 | 18153.5284 | 8739.3409 | 34033.0837 | 606.0129 | 43321.5876 | 43321.5876 | 1153.38 | 7.3333 | 30.6667 | 58.6668 | 283.8338 | 16.6667 | 266316.8323 | 246606.496 | 142759.0929 | 183040.6045 | 10669.3186 | 198842.1051 | 168291.5005 | 162047.9195 | 3812.5087 | 64.6668 | 138.6673 | 956.6789 | 2012.7219
96 | 8_16_21_F_1_DUP_LSP |  | 06/08/2023 23:21:03 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_F_1_DUP_LSP.450 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 85880.0467 | 32668.6544 | 1088715.5452 | 598.0125 | 77344.1426 | 16935.3678 | 36360.8843 | 23133.3838 | 13660.5292 | 17728.3274 | 99352.9691 | 98482.9783 | 340946.3004 | 23991.4628 | 36805.3531 | 9960.805 | 24400.8258 | 188827.1241 | 89765.8535 | 100028.8409 | 1696.1007 | 119472.1749 | 119472.1749 | 7754.1046 | 1720.7705 | 5311.6542 | 11359.8485 | 6871.1786 | 2049.4808 | 425607.7307 | 380393.5122 | 130991.1669 | 162453.1648 | 9491.1523 | 177873.9432 | 163572.6962 | 143774.5962 | 3245.0352 | 75985.5718 | 180166.5995 | 268086.4363 | 571891.2333
97 | LBLANK |  | 06/08/2023 23:26:52 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.451 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 6 | 0.66667 | 4163.2736 | 1.3333 | 1729.4385 | 683.3497 | 16 | 0.66667 | 0.66667 | 8.6667 | 43.3334 | 612.993 | 19.3333 | 1380.0667 | 2.6667 | 18.6667 | 18 | 92.667 | 318.0035 | -2363.2314 | 0.66667 | 76.6669 | 76.6669 | 1170.0482 | 1.3333 | 2.6667 | 8 | 97.9025 | 2 | 184.0012 | 93.3336 | 151805.4569 | 206019.8779 | 11674.1019 | 233794.9761 | 190370.0657 | 182678.618 | 4385.3399 | 1.3333 | 20.6667 | 26 | 38.6667
98 | CCV 1ppb TRACE |  | 06/08/2023 23:32:40 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.452 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1708.7689 | 730.6854 | 139758.4262 | 84.6669 | 3847.8517 | 903.3619 | 832.0243 | 328.0038 | 1226.7193 | 1667.4309 | 9241.6553 | 7755.3646 | 15829.4319 | 3004.984 | 3472.4225 | 996.7016 | 2456.2112 | 19128.1333 | 8946.1343 | 4046.9855 | 109.3338 | 7642.7107 | 7642.7107 | 1776.7772 | 162.6676 | 512.6759 | 1037.3711 | 645.7988 | 206.0015 | 16876.6297 | 14829.028 | 158698.786 | 202687.2153 | 11868.9288 | 225014.6365 | 187276.2877 | 177635.8258 | 4200.6177 | 7641.3766 | 18661.5194 | 28237.5184 | 61257.837
99 | CCV 1ppb TRACE |  | 06/08/2023 23:38:58 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.453 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1723.4373 | 730.6854 | 135281.1758 | 81.3336 | 3658.4688 | 953.3652 | 820.6902 | 327.3371 | 1241.3873 | 1653.4291 | 9145.5956 | 7420.7489 | 15571.151 | 3119.6741 | 3533.7704 | 994.7013 | 2469.5468 | 18719.5908 | 8786.0351 | 3857.4705 | 118.0005 | 7551.329 | 7551.329 | 1751.4409 | 142.6674 | 495.3419 | 1140.7124 | 644.3307 | 202.6681 | 16452.1351 | 14618.8094 | 158892.8906 | 203710.2316 | 11976.3523 | 225889.6732 | 188049.6282 | 178297.8512 | 4220.6235 | 7816.805 | 18373.8093 | 28132.2103 | 61397.0754
100 | CCV 10 ppb MAJORS |  | 06/08/2023 23:44:46 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.454 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 15774.0395 | 6482.8051 | 664725.3698 | 400.0056 | 28885.8468 | 3081.6658 | 7162.462 | 2926.9666 | 11176.3704 | 15915.5284 | 86142.9705 | 58448.8578 | 145947.8516 | 14944.48 | 33296.7604 | 9125.5805 | 21361.9598 | 165987.4691 | 78680.08 | 56579.8064 | 1019.3698 | 70196.0482 | 70196.0482 | 7053.7417 | 1486.0774 | 4697.4389 | 10115.584 | 5722.755 | 1778.7774 | 153211.2071 | 138404.6303 | 152904.0275 | 192373.9894 | 11333.1613 | 211060.7211 | 174236.1495 | 167534.7402 | 3937.2095 | 71381.9068 | 168513.4237 | 257411.7331 | 553617.1693
101 | CCV 10 ppb MAJORS |  | 06/08/2023 23:51:04 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.455 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 15585.1682 | 6646.8794 | 661391.5262 | 394.0054 | 29307.366 | 3052.993 | 7402.5849 | 2998.9816 | 11339.1665 | 15555.1322 | 87185.9199 | 59355.8365 | 145884.4448 | 14999.2039 | 33692.3548 | 9102.9006 | 21205.0617 | 166677.3918 | 78935.4859 | 57606.2199 | 1088.0415 | 69910.6675 | 69910.6675 | 7116.4391 | 1495.4117 | 4646.7563 | 9890.0892 | 5739.7605 | 1850.1199 | 153672.1529 | 137837.8565 | 152966.1041 | 190688.2939 | 11556.6785 | 212473.2897 | 176740.6762 | 170801.7246 | 3889.196 | 71377.8837 | 170807.2128 | 260154.497 | 552113.7255
102 | 8_16_21_F_6 |  | 06/08/2023 23:56:52 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_F_6.456 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 14557.417 | 4259.3016 | 448311.8177 | 227.3351 | 34711.4533 | 5087.5726 | 1556.7517 | 10084.8929 | 833.3577 | 125.3339 | 577.345 | 41675.2996 | 172150.4022 | 11377.1984 | 92.0003 | 784.6882 | 6785.6114 | 52281.5265 | 38662.9258 | 33091.3078 | 552.6774 | 41260.1663 | 41260.1663 | 1243.3875 | 14 | 48.0001 | 92.0003 | 351.5995 | 12 | 143499.1801 | 130323.1064 | 140643.0302 | 183690.8585 | 10948.1937 | 201553.521 | 168587.6045 | 166661.9183 | 3816.5101 | 109.3338 | 302.6699 | 182420.3799 | 386176.3536
103 | 8_16_21_F_6_SP |  | 06/09/2023 00:02:41 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_F_6_SP.457 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 22304.0661 | 7441.2712 | 743593.3216 | 411.3393 | 53162.0765 | 6009.9309 | 5043.5569 | 11329.1575 | 6522.1559 | 7778.784 | 43203.9003 | 66889.0015 | 236151.2626 | 17215.0359 | 16368.7057 | 5075.5687 | 16368.712 | 128669.5353 | 75509.0634 | 61460.0967 | 1041.3713 | 73447.6615 | 73447.6615 | 4092.5862 | 762.0204 | 2327.5232 | 4857.4925 | 3069.2098 | 931.3638 | 211183.6036 | 191990.1986 | 146933.8133 | 181844.2791 | 10369.7656 | 196539.514 | 165737.2397 | 160337.5133 | 3776.4992 | 35070.3344 | 82917.9775 | 141010.3898 | 299744.2402
104 | 8_16_21_F_8 |  | 06/09/2023 00:08:29 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_F_8.458 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 15198.7479 | 4598.0736 | 223307.1701 | 141.334 | 50776.1283 | 4468.6989 | 1565.4192 | 9940.1239 | 1205.3842 | 107.3337 | 564.6778 | 38074.3406 | 178423.7286 | 10435.811 | 76.0002 | 447.3404 | 7737.4286 | 59662.3326 | 17350.5301 | 33034.4707 | 580.0118 | 41285.5732 | 41285.5732 | 1176.0485 | 8 | 53.3334 | 98.667 | 322.4267 | 12.6667 | 114032.6534 | 103204.7931 | 148334.8289 | 183929.9477 | 10610.6121 | 206282.1662 | 169214.4402 | 166853.4334 | 3859.1879 | 112.0004 | 208.6682 | 3893.5411 | 8220.2718
105 | 8_16_21_F_8_DUP |  | 06/09/2023 00:14:18 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_F_8_DUP.459 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 14917.1177 | 4269.972 | 223961.6345 | 135.334 | 48237.3069 | 4263.3028 | 1512.0802 | 9156.9342 | 1037.371 | 156.6675 | 752.0198 | 38315.6298 | 170067.6534 | 10457.1599 | 82.6669 | 441.3402 | 7285.1907 | 56777.9411 | 18208.9312 | 32151.0405 | 528.0098 | 40354.9269 | 40354.9269 | 1177.3819 | 15.3333 | 43.3334 | 85.3336 | 263.6925 | 19.3333 | 111399.3277 | 100782.2823 | 146150.4905 | 181524.745 | 10626.6181 | 200667.6459 | 171527.392 | 162111.3451 | 3776.4992 | 93.3336 | 213.3349 | 3969.5487 | 8332.2782
106 | 8_16_21_F_9 |  | 06/09/2023 00:20:06 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_F_9.460 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 12650.2653 | 4035.2365 | 371371.3882 | 242.6687 | 50456.9567 | 3938.5433 | 1265.3894 | 8621.9364 | 940.0309 | 362.0046 | 1738.7725 | 34128.1869 | 97777.4758 | 9409.7652 | 72.0002 | 1020.0364 | 5785.8383 | 44921.1839 | 10908.1642 | 25859.3597 | 447.3404 | 32732.7978 | 32732.7978 | 1185.3825 | 16.6667 | 38.0001 | 91.3336 | 226.981 | 16.6667 | 75776.4673 | 69536.8408 | 146340.505 | 184519.367 | 10719.3537 | 201528.1908 | 169107.6903 | 158704.1271 | 3874.5257 | 56.6668 | 161.3343 | 2682.7652 | 5766.451
107 | 8_16_21_F_10 |  | 06/09/2023 00:25:54 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_F_10.461 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 11373.8602 | 5651.7847 | 252004.1447 | 140.6674 | 48321.6047 | 4352.6632 | 1506.7462 | 10379.7703 | 956.6987 | 76.6669 | 474.6746 | 22359.9228 | 58010.2326 | 6864.3161 | 60.6668 | 145.3341 | 4450.0264 | 34742.868 | 11908.9618 | 26471.0496 | 438.0067 | 33954.9719 | 33954.9719 | 1193.3834 | 9.3333 | 38.6667 | 114.6672 | 274.2192 | 15.3333 | 35550.8587 | 32638.5769 | 146472.4056 | 179483.427 | 10625.9509 | 201351.284 | 166230.9905 | 161828.1226 | 3701.1465 | 79.3336 | 164.6676 | 2668.0998 | 5615.1002
108 | 8_16_21_F_10_DUP |  | 06/09/2023 00:31:42 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_F_10_DUP.462 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10988.8943 | 5474.3823 | 254690.4895 | 147.3341 | 47725.5881 | 4206.6194 | 1441.4061 | 9945.4629 | 901.3619 | 105.3337 | 570.0114 | 21195.0567 | 55913.88 | 6586.852 | 54.6668 | 402.0057 | 4561.3948 | 34738.8536 | 10893.4857 | 26104.8541 | 428.6731 | 32627.2226 | 32627.2226 | 1200.7172 | 17.3333 | 52.0001 | 79.3336 | 268.2459 | 13.3333 | 34480.2394 | 31343.0195 | 145883.1163 | 188496.433 | 10777.3976 | 206909.6969 | 171108.0873 | 167377.3039 | 3895.8652 | 66.6668 | 169.3343 | 2662.7619 | 5574.4216
109 | 8_16_21_F_11 |  | 06/09/2023 00:37:30 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_F_11.463 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10970.2131 | 5636.4456 | 311158.5171 | 204.6681 | 45389.3502 | 4428.6868 | 1446.0732 | 9877.4143 | 965.366 | 78.0002 | 472.6745 | 22589.5653 | 42832.7922 | 6899.666 | 65.3335 | 166.6676 | 4317.986 | 33525.9648 | 12155.1691 | 26290.5847 | 432.0066 | 33277.3818 | 33277.3818 | 1204.0508 | 17.3333 | 48.0001 | 106.6671 | 277.0048 | 18.6667 | 29888.5686 | 26960.086 | 144666.9391 | 180493.8691 | 10726.6984 | 198145.6165 | 167471.2832 | 165066.9248 | 3801.8392 | 30.6667 | 90.0003 | 3166.8049 | 6417.8977
110 | 8_16_21_W_1 |  | 06/09/2023 00:43:18 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_W_1.464 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2730.928 | 773.3543 | 280705.714 | 171.3344 | 21302.5378 | 1953.4669 | 307.3366 | 2196.8356 | 88.0003 | 30 | 149.3341 | 29780.0708 | 62480.3372 | 8218.366 | 35.3334 | 75.3335 | 643.3478 | 5228.2899 | 1936.7982 | 5735.5422 | 132.6673 | 9508.4974 | 9508.4974 | 1077.374 | 1.3333 | 10.6667 | 11.3333 | 97.1276 | 4 | 57955.3407 | 51920.8494 | 151712.7536 | 192632.0577 | 11662.1026 | 195194.3093 | 176202.049 | 164380.6292 | 3897.1982 | 20 | 43.3334 | 602.0049 | 1294.6895
111 | 8_16_21_W_6 |  | 06/09/2023 00:49:05 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_W_6.465 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2988.3126 | 883.3607 | 662971.3088 | 394.0054 | 10984.8884 | 1492.0779 | 310.6701 | 2146.1613 | 187.3346 | 48.0001 | 214.0016 | 61302.9438 | 64124.2626 | 15453.0237 | 24 | 60.6668 | 2436.2079 | 19261.6439 | 5809.1809 | 5405.2963 | 126.6672 | 9084.221 | 9084.221 | 1096.7088 | 5.3333 | 15.3333 | 21.3333 | 140.695 | 4.6667 | 30152.4657 | 26623.4552 | 159036.6443 | 193000.3809 | 11276.5336 | 214954.5927 | 178160.0998 | 171570.2899 | 4012.5635 | 35.3334 | 88.0003 | 2873.4446 | 5812.4675
112 | LBLANK |  | 06/09/2023 00:54:53 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.466 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2 | 0 | 3695.8114 | 2.6667 | 1796.7797 | 631.3473 | 11.3333 | 0 | 0 | 8 | 38.0001 | 571.7063 | 16.6667 | 1306.7264 | 0.66667 | 8.6667 | 12.6667 | 101.3337 | 295.3364 | -2176.787 | 0 | 16 | 16 | 1015.3696 | 0 | 1.3333 | 3.3333 | 13.689 | 2 | 26.6667 | 26 | 154609.6876 | 197141.0332 | 12136.4873 | 219060.2398 | 177084.2991 | 175147.5835 | 4040.5715 | 0.66667 | 12.6667 | 24 | 42.6667
113 | CCV 1ppb TRACE |  | 06/09/2023 01:00:42 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.467 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1621.4254 | 676.6828 | 133851.4883 | 78.6669 | 4029.235 | 886.0276 | 783.3548 | 296.0031 | 1165.3809 | 1581.4209 | 8801.3774 | 7265.1037 | 14681.5415 | 2938.3022 | 3386.4014 | 948.6982 | 2326.8562 | 18098.1258 | 8461.172 | 3618.0496 | 105.3337 | 6976.3698 | 6976.3698 | 1694.1006 | 166.001 | 479.3414 | 1044.7049 | 593.9132 | 172.6677 | 15731.9928 | 13929.4545 | 159935.097 | 190152.5741 | 11445.2525 | 218982.0206 | 181401.1956 | 167418.0196 | 3811.8422 | 7470.6204 | 17904.5499 | 28804.0421 | 61105.7456
114 | CCV 1ppb TRACE |  | 06/09/2023 01:06:59 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.468 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1628.093 | 694.6836 | 127250.9618 | 79.3336 | 3681.8078 | 860.6926 | 740.686 | 305.3366 | 1134.045 | 1616.0915 | 8840.7361 | 7414.7388 | 14650.1751 | 2792.273 | 3372.3984 | 943.3645 | 2302.8526 | 17915.8953 | 8375.7884 | 3584.1624 | 104.6671 | 7028.396 | 7028.396 | 1636.0937 | 149.3341 | 460.0074 | 1023.37 | 622.7866 | 199.3347 | 15649.2335 | 14137.66 | 155123.7462 | 193065.505 | 11474.609 | 220523.6063 | 173170.0192 | 169710.2093 | 3948.546 | 7499.9686 | 17997.9979 | 28319.5941 | 61414.1951
115 | LFB |  | 06/09/2023 01:12:47 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.469 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 14732.2618 | 6067.9553 | 646387.9506 | 406.0058 | 33787.9196 | 2965.6414 | 6882.9911 | 2808.943 | 10366.4273 | 14610.1377 | 80455.935 | 54278.5926 | 134757.3362 | 14032.2258 | 30935.4687 | 8287.0701 | 19820.4099 | 156375.8899 | 74290.0086 | 55014.6285 | 1000.0351 | 65531.3106 | 65531.3106 | 6716.2453 | 1420.0706 | 4530.052 | 9402.4266 | 5562.2133 | 1742.1064 | 143164.5195 | 127694.182 | 152840.9525 | 180707.3417 | 10693.3349 | 203713.53 | 167346.5458 | 161909.3781 | 3777.833 | 68688.741 | 161584.0532 | 252820.6269 | 528559.5672
116 | CCV 10 ppb MAJORS |  | 06/09/2023 01:18:35 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.470 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 15074.6165 | 6121.9792 | 638520.185 | 382.0052 | 28816.3852 | 2915.631 | 7097.0966 | 2840.9492 | 10951.5302 | 14899.7661 | 83339.0814 | 56597.9037 | 140226.9282 | 14413.9349 | 31792.6717 | 8665.9614 | 20536.0842 | 161002.8589 | 75676.5835 | 56172.8821 | 1050.0386 | 67751.6274 | 67751.6274 | 7007.7195 | 1472.0759 | 4516.714 | 9531.1794 | 5514.9099 | 1738.7726 | 149623.4702 | 133185.9682 | 152522.6461 | 184431.6044 | 10843.4475 | 206672.4951 | 170578.3672 | 166636.2405 | 3867.857 | 70108.9548 | 164869.4376 | 261215.093 | 551678.7898
117 | CCV 10 ppb MAJORS |  | 06/09/2023 01:24:53 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.471 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 15619.2015 | 6371.4208 | 642714.0924 | 366.6714 | 28883.1742 | 3028.9882 | 7194.4786 | 2828.9467 | 10982.2203 | 15344.9048 | 84497.8261 | 58217.6222 | 140173.717 | 14359.8819 | 32524.3219 | 8943.4655 | 20891.9355 | 163453.1462 | 76671.8768 | 56640.346 | 974.0332 | 67485.024 | 67485.024 | 6973.035 | 1490.0777 | 4574.0656 | 9808.0343 | 5661.5436 | 1775.4437 | 151159.535 | 134749.2108 | 151829.2381 | 182936.5618 | 11046.9374 | 207304.6763 | 170545.4041 | 161884.7959 | 3917.8705 | 70819.8087 | 167310.783 | 261183.7521 | 545952.9556
118 | 8_16_21_W_6_DUP |  | 06/09/2023 01:30:41 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_W_6_DUP.472 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2899.6277 | 864.0262 | 662466.1501 | 438.0067 | 9152.9322 | 1471.4091 | 278.6694 | 2081.4851 | 184.0012 | 112.0004 | 574.0115 | 63519.0315 | 62844.6027 | 16914.6746 | 35.3334 | 55.3334 | 2450.2103 | 18784.3438 | 6267.3752 | 4971.2613 | 126.0006 | 8771.3598 | 8771.3598 | 1104.7095 | 0.66667 | 12.6667 | 25.3334 | 81.0622 | 4 | 28700.8028 | 26016.3355 | 157186.0673 | 191150.6317 | 11651.4181 | 215234.9897 | 178146.2796 | 168181.2275 | 3904.534 | 34 | 102.667 | 2539.4224 | 5203.0362
119 | 8_16_21_W_8 |  | 06/09/2023 01:36:29 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_W_8.473 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2938.969 | 853.3589 | 661758.177 | 405.3391 | 14443.9659 | 1275.3903 | 298.6698 | 1903.4604 | 235.3353 | 36 | 227.3351 | 60124.0409 | 81617.168 | 15697.2888 | 35.3334 | 49.3334 | 2356.8613 | 17476.6842 | 5627.1081 | 5171.0328 | 96.667 | 8641.9466 | 8641.9466 | 1084.0411 | 6.6667 | 16 | 25.3334 | 130.7595 | 3.3333 | 23032.557 | 20518.7251 | 153664.0594 | 186870.3252 | 11763.5155 | 210412.3552 | 174385.4129 | 167083.5198 | 3907.2013 | 36 | 92.0003 | 2240.0687 | 4574.2863
120 | 8_16_21_W_9 |  | 06/09/2023 01:42:18 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_W_9.474 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2475.548 | 828.024 | 707151.8601 | 415.3394 | 12068.4289 | 1258.0554 | 258.0023 | 1804.1139 | 208.6682 | 61.3335 | 298.6698 | 63902.2089 | 67659.86 | 16323.3219 | 27.3334 | 50.6668 | 1816.1157 | 14823.0237 | 4282.6419 | 3835.2002 | 101.3337 | 7013.0555 | 7013.0555 | 1048.7052 | 6 | 9.3333 | 23.3334 | 107.9732 | 3.3333 | 16212.5282 | 14191.0484 | 152249.6904 | 189506.9578 | 11139.6759 | 185701.4303 | 174261.0732 | 166642.354 | 3981.2218 | 33.3334 | 62.6668 | 1887.3828 | 3966.2147
121 | 8_16_21_W_9_DUP |  | 06/09/2023 01:48:06 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_W_9_DUP.475 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2349.5266 | 789.3552 | 770468.1612 | 486.675 | 11588.6995 | 1202.0508 | 251.3356 | 1672.0979 | 196.668 | 52.6668 | 224.0018 | 68527.9166 | 72943.089 | 17707.6353 | 38.6667 | 52.0001 | 1895.4591 | 14605.4625 | 4512.0462 | 3710.9561 | 90.0003 | 6796.9505 | 6796.9505 | 1028.037 | 0 | 14.6667 | 21.3333 | 42.2969 | 4.6667 | 15400.298 | 13534.4087 | 151239.4751 | 186157.7555 | 11336.4975 | 211220.3146 | 171844.9131 | 165230.0716 | 3779.1669 | 22.6667 | 69.3335 | 2069.3913 | 4294.2542
122 | 8_16_21_W_10 |  | 06/09/2023 01:53:55 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_W_10.476 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2199.5027 | 1119.3773 | 852180.8955 | 526.6764 | 11127.667 | 1276.0571 | 292.003 | 2023.4768 | 199.3347 | 155.3342 | 856.0257 | 63081.2277 | 47214.5613 | 16597.6371 | 20.6667 | 46.6667 | 1372.0659 | 11360.5163 | 4289.9777 | 3929.2327 | 82.6669 | 6968.3659 | 6968.3659 | 1124.7111 | 5.3333 | 6 | 33.3334 | 107.8163 | 4 | 7243.1697 | 6481.47 | 151226.2238 | 185767.4561 | 10973.5464 | 209357.7602 | 172670.002 | 165164.7253 | 4021.2327 | 27.3334 | 54.6668 | 1992.7204 | 4166.2354
123 | 8_16_21_W_11 |  | 06/09/2023 01:59:43 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_W_11.477 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2098.1541 | 1132.7116 | 435997.2379 | 271.3359 | 8895.437 | 1321.3947 | 273.336 | 2025.477 | 179.3345 | 25.3334 | 147.3341 | 29923.2875 | 23200.8272 | 8235.7067 | 12 | 39.3334 | 1176.0485 | 9405.0953 | 3075.6647 | 3861.5286 | 80.0002 | 6802.9533 | 6802.9533 | 1044.7049 | 3.3333 | 7.3333 | 29.3334 | 121.8491 | 4.6667 | 6297.3878 | 5517.0661 | 151125.9096 | 183538.3254 | 10980.2183 | 209996.9283 | 164320.9463 | 165821.031 | 3845.1846 | 14 | 39.3334 | 1176.6858 | 2484.0846
124 | 5_10_23_DB |  | 06/09/2023 02:05:32 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_10_23_DB.478 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0 | 0 | 27927.2862 | 16.6667 | 12816.4147 | 571.3448 | 10 | 0 | 0 | 15.3333 | 52.6668 | 1029.459 | 25.3334 | 1324.7282 | 1.3333 | 13.3333 | 55.3334 | 480.0081 | 244.0021 | -2246.2342 | 0.66667 | 8 | 8 | 1026.7036 | 0.66667 | 0.66667 | 2.6667 | -47.3312 | 1.3333 | 12 | 8 | 153193.6997 | 189671.5077 | 11423.9047 | 216446.2047 | 173479.6869 | 165689.4295 | 4045.9064 | 0 | 6.6667 | 68.6667 | 126.6669
125 | 5_10_23_MFB |  | 06/09/2023 02:11:19 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_10_23_MFB.479 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 7311.8719 | 3070.9969 | 345980.7044 | 211.3349 | 29561.8888 | 1742.1062 | 3365.7301 | 1386.7341 | 5266.9709 | 7407.2538 | 41109.7335 | 30266.7788 | 68611.7164 | 8036.9286 | 15516.4284 | 4211.9546 | 10125.589 | 80446.5575 | 37791.256 | 26005.1051 | 472.0078 | 32603.8304 | 32603.8304 | 3961.216 | 634.6808 | 2208.8375 | 4676.7655 | 2701.3727 | 846.0251 | 72126.9583 | 63588.555 | 151540.2265 | 182216.8451 | 11169.0312 | 201993.8553 | 169098.5126 | 163205.0453 | 3781.8342 | 35071.6699 | 83152.6363 | 133183.2989 | 279807.4337
126 | 5_10_23_T-239 |  | 06/09/2023 02:17:07 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_10_23_T-239.480 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 65227.9602 | 7111.7707 | 7623910.5738 | 3775.8324 | 34480.2313 | 692.0168 | 2439.5416 | 13139.3734 | 1891.4586 | 1914.795 | 10678.6565 | 207763.5493 | 45174.6518 | 52386.5473 | 4109.2578 | 1345.3968 | 1393.4013 | 11900.9551 | 41052.2339 | 11678.6016 | 198.6681 | 14374.5626 | 14374.5626 | 3655.1342 | 267.3358 | 748.6863 | 1736.1056 | 1228.9192 | 325.337 | 14463.9904 | 12728.3351 | 134950.3434 | 166261.9233 | 10125.5876 | 192095.5733 | 157767.1281 | 148297.0806 | 3601.7874 | 21818.6495 | 51548.8424 | 32146.0721 | 62959.5833
127 | 5_10_23_DB |  | 06/09/2023 02:22:55 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_10_23_DB.481 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4.6667 | 0.66667 | 53518.7579 | 27.3334 | 15408.3054 | 625.347 | 16 | 1.3333 | 2.6667 | 32.6667 | 165.3343 | 1905.9426 | 32 | 1644.7613 | 0.66667 | 6 | 82.6669 | 731.3521 | 264.0025 | -2399.374 | 0 | 22.6667 | 22.6667 | 1118.0438 | 0.66667 | 0 | 2.6667 | 16.3394 | 1.3333 | 44.0001 | 9.3333 | 165083.9809 | 195350.1009 | 11618.7269 | 227238.3974 | 180133.9764 | 169696.6886 | 4140.6005 | 1.3333 | 6.6667 | 91.3334 | 211.3339
128 | LBLANK |  | 06/09/2023 02:28:44 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.482 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.3333 | 0 | 3701.813 | 0.66667 | 3517.0996 | 614.0133 | 12.6667 | 1.3333 | 0.66667 | 13.3333 | 35.3334 | 554.4099 | 13.3333 | 1295.3921 | 0 | 8 | 10.6667 | 90.0003 | 310.0034 | -2260.465 | 0 | 13.3333 | 13.3333 | 1098.0423 | 1.3333 | 1.3333 | 6 | 25.7434 | 0 | 30 | 17.3333 | 159508.4482 | 191692.2178 | 11479.9447 | 233592.4566 | 181720.6605 | 170107.4668 | 4047.2406 | 0 | 8 | 22.6667 | 37.3334
129 | CCV 1ppb TRACE |  | 06/09/2023 02:34:32 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.483 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1624.0925 | 661.3487 | 127163.5412 | 86.0003 | 3927.8733 | 866.0263 | 682.683 | 268.0025 | 1184.7159 | 1672.7646 | 8767.3565 | 7041.29 | 14335.1891 | 2746.2641 | 3352.3934 | 950.6983 | 2284.1826 | 18072.0922 | 8305.7475 | 3785.5674 | 106.6671 | 7101.0984 | 7101.0984 | 1658.0963 | 134.0006 | 469.341 | 1003.3686 | 646.5518 | 192.0013 | 15637.8879 | 13521.7301 | 151904.7743 | 192980.2487 | 11530.6529 | 216206.6737 | 177311.6285 | 166390.0633 | 4249.9657 | 7449.2751 | 17855.8217 | 28815.3892 | 60076.3601
130 | CCV 1ppb TRACE |  | 06/09/2023 02:40:49 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.484 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1621.4254 | 646.0146 | 124381.8923 | 75.3335 | 3587.1171 | 841.3582 | 733.3522 | 285.3362 | 1185.3827 | 1592.7555 | 8795.3753 | 7202.7625 | 14496.0191 | 2840.9492 | 3415.0749 | 950.0317 | 2329.5234 | 18100.127 | 8529.2121 | 3690.1832 | 92.0003 | 7041.0682 | 7041.0682 | 1745.44 | 161.3343 | 450.6738 | 1040.7047 | 613.9631 | 184.6679 | 15766.0325 | 14013.5367 | 157277.7462 | 188288.4277 | 11631.4006 | 218625.1381 | 176456.4382 | 168787.9704 | 4091.2535 | 7575.3413 | 17946.6009 | 28294.92 | 59079.2629
131 | CCV 10 ppb MAJORS |  | 06/09/2023 02:46:38 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.485 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 14844.3768 | 6029.9391 | 611375.7811 | 372.6715 | 27560.5631 | 2945.6373 | 6905.0018 | 2777.6038 | 10865.4643 | 14895.0957 | 81311.4624 | 55175.9584 | 135279.5873 | 14013.5395 | 31561.4939 | 8689.3092 | 20373.1855 | 156573.4145 | 74719.5772 | 55545.4551 | 1016.7029 | 65467.6679 | 65467.6679 | 7033.0644 | 1423.4043 | 4361.9993 | 9475.8084 | 5476.3926 | 1699.4344 | 147407.4647 | 129873.0149 | 152080.6684 | 176782.485 | 10836.7791 | 204328.9365 | 165990.1326 | 160080.0665 | 3778.5001 | 68781.1855 | 164553.8029 | 261673.2928 | 540303.0108
132 | CCV 10 ppb MAJORS |  | 06/09/2023 02:52:55 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.486 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 14719.5803 | 6001.9281 | 626936.0771 | 368.6714 | 27680.7937 | 2926.3 | 6811.6238 | 2870.9553 | 10657.3081 | 14724.2525 | 82971.6294 | 55852.7723 | 136635.6946 | 14228.4157 | 31415.8459 | 8503.1967 | 20339.14 | 160503.3031 | 74987.6376 | 55881.7943 | 1040.7046 | 66481.0112 | 66481.0112 | 7083.089 | 1436.0722 | 4361.9997 | 9278.3455 | 5694.1506 | 1670.0978 | 147196.4929 | 132500.3087 | 151154.3482 | 182573.949 | 10828.7695 | 204679.2284 | 169119.1861 | 154172.1581 | 3773.8319 | 69127.5252 | 162844.8864 | 260385.7669 | 540546.0397
133 | MT 34 D |  | 06/09/2023 02:58:43 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\MT 34 D.487 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4.6667 | 0 | 72955.1697 | 50.0001 | 19307.7063 | 654.6817 | 12 | 4.6667 | 3.3333 | 26 | 142.6674 | 1809.1754 | 45.3334 | 1716.7701 | 5.3333 | 38.0001 | 46.0001 | 396.0055 | 816.0234 | -2419.4953 | 0 | 26 | 26 | 1121.3775 | 0 | 1.3333 | 2 | 53.7875 | 0 | 404.6724 | 322.6703 | 161493.0446 | 199449.4328 | 11861.591 | 230691.2597 | 187731.6967 | 174468.5825 | 4384.6729 | 0.66667 | 25.3334 | 804.0095 | 1702.7068
134 | MT 35 D |  | 06/09/2023 03:04:31 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\MT 35 D.488 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.3333 | 0 | 87094.7212 | 50.0001 | 21102.9098 | 597.3458 | 9.3333 | 2.6667 | 0.66667 | 19.3333 | 96.0003 | 1123.3173 | 42.0001 | 1477.4098 | 0.66667 | 22.6667 | 46.0001 | 386.6719 | 638.681 | -2496.9968 | 0.66667 | 19.3333 | 19.3333 | 1051.3721 | 0.66667 | 0 | 2.6667 | 15.57 | 2.6667 | 77.3335 | 18.6667 | 161450.6199 | 195089.748 | 12105.127 | 224045.4971 | 182896.0398 | 169988.8256 | 4179.9452 | 0.66667 | 22 | 823.3427 | 1673.3721
135 | MT 36 D |  | 06/09/2023 03:10:18 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\MT 36 D.489 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.3333 | 0 | 69691.5888 | 30 | 21113.5924 | 643.3479 | 14 | 2.6667 | 0 | 35.3334 | 140.6674 | 1058.3784 | 48.0001 | 1415.4035 | 0 | 12.6667 | 49.3334 | 416.0061 | 813.3565 | -2280.5898 | 0 | 22 | 22 | 1024.0367 | 0.66667 | 0.66667 | 4 | 117.2364 | 0.66667 | 82.0002 | 43.3334 | 161483.6097 | 195860.1693 | 11742.1582 | 222677.5172 | 183453.1857 | 169181.3664 | 4063.245 | 2.6667 | 16 | 780.675 | 1654.0373
136 | MT 37 D |  | 06/09/2023 03:16:06 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\MT 37 D.490 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 7.3333 | 0.66667 | 81755.9995 | 42.6667 | 21917.4707 | 598.6792 | 10.6667 | 2.6667 | 0 | 20.6667 | 90.667 | 856.5131 | 40.0001 | 1385.4005 | 2.6667 | 13.3333 | 40.0001 | 346.0042 | 770.0208 | -2237.9359 | 0 | 13.3333 | 13.3333 | 1066.0398 | 0 | 1.3333 | 5.3333 | 19.1254 | 0 | 73.3335 | 32.6667 | 159608.0099 | 193151.7987 | 11854.9192 | 222976.4641 | 186045.171 | 168698.5418 | 4221.9573 | 2.6667 | 15.3333 | 786.0082 | 1689.3723
137 | MT 38 D |  | 06/09/2023 03:21:55 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\MT 38 D.491 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 25.3334 | 0.66667 | 107985.2601 | 77.3335 | 22620.566 | 567.3446 | 18.6667 | 4 | 1.3333 | 24 | 125.3339 | 22520.3818 | 332.0039 | 6748.2604 | 4.6667 | 24.6667 | 220.0017 | 1866.1219 | 3670.4715 | -2323.9591 | 2 | 19.3333 | 19.3333 | 1083.3746 | 0 | 4 | 4 | 80.6768 | 0 | 72.0002 | 46.0001 | 159955.2136 | 192834.127 | 11434.5754 | 220888.7015 | 180703.1438 | 161098.7126 | 4053.9085 | 2.6667 | 16.6667 | 805.3421 | 1688.0392
138 | MT 39 D |  | 06/09/2023 03:27:43 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\MT 39 D.492 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.3333 | 0 | 74046.7835 | 43.3334 | 21888.7566 | 610.0131 | 12.6667 | 2 | 2.6667 | 29.3334 | 130.6673 | 6441.6673 | 108.0004 | 2523.5568 | 0.66667 | 11.3333 | 63.3335 | 494.6753 | 714.6846 | -2432.3691 | 0 | 12.6667 | 12.6667 | 1100.7091 | 0 | 1.3333 | 1.3333 | 26.9686 | 0 | 223.3351 | 180.0012 | 157944.3707 | 190602.0175 | 11961.0059 | 211904.2564 | 181300.54 | 166649.7679 | 4106.5902 | 5.3333 | 24.6667 | 789.3415 | 1684.7051
139 | MT 40 D |  | 06/09/2023 03:33:31 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\MT 40 D.493 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2 | 0 | 79111.7849 | 47.3334 | 23402.4873 | 632.0141 | 10.6667 | 2 | 0 | 19.3333 | 93.3336 | 1151.6631 | 66.6668 | 1492.7447 | 6 | 16 | 56.0001 | 431.34 | 747.3529 | -2529.3283 | 0 | 16 | 16 | 1064.7064 | 0 | 0.66667 | 3.3333 | 29.0228 | 0 | 46.6667 | 43.3334 | 161051.3949 | 194074.1994 | 11839.5726 | 222378.1882 | 174125.4025 | 169613.7934 | 4106.5905 | 4 | 21.3333 | 858.6773 | 1686.7059
140 | MT 32 D |  | 06/09/2023 03:39:20 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\MT 32 D.494 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.3333 | 0 | 103643.3032 | 62.0001 | 21839.3497 | 584.012 | 10.6667 | 3.3333 | 1.3333 | 22.6667 | 100.667 | 1296.389 | 130.6673 | 1439.4063 | 0 | 12 | 116.6671 | 937.3641 | 810.6897 | -2213.0927 | 0.66667 | 24.6667 | 24.6667 | 1030.7039 | 1.3333 | 1.3333 | 3.3333 | 27.0768 | 0 | 46.6667 | 42.0001 | 160923.418 | 192090.4047 | 11185.7198 | 223452.8091 | 183556.4562 | 170866.1154 | 4131.931 | 1.3333 | 18.6667 | 1124.017 | 2437.4136
141 | MT 32 D LD |  | 06/09/2023 03:45:07 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\MT 32 D LD.495 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.3333 | 0 | 74649.2977 | 42.0001 | 21762.5674 | 602.0127 | 14.6667 | 1.3333 | 0 | 30.6667 | 126.6672 | 1042.7309 | 46.0001 | 1369.3991 | 0.66667 | 12.6667 | 98.0003 | 781.3547 | 994.0346 | -2449.4071 | 0 | 17.3333 | 17.3333 | 1042.038 | 0.66667 | 0.66667 | 2.6667 | 24.9414 | 0 | 56.6668 | 45.3334 | 157008.2093 | 188898.9592 | 11499.9602 | 226528.6723 | 170615.265 | 192084.2645 | 4091.2525 | 0 | 26 | 1059.3493 | 2255.4047
142 | MT 32 D LD |  | 06/09/2023 03:51:24 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\MT 32 D LD.496 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.6667 | 0.66667 | 72695.8515 | 45.3334 | 21731.1847 | 638.0142 | 14.6667 | 1.3333 | 0.66667 | 26.6667 | 119.3338 | 974.6614 | 41.3334 | 1344.73 | 0 | 13.3333 | 90.0003 | 745.3528 | 962.6991 | -2204.9919 | 0 | 18.6667 | 18.6667 | 1004.702 | 1.3333 | 0.66667 | 4.6667 | 50.4974 | 2.6667 | 51.3334 | 43.3334 | 156307.1775 | 185459.4917 | 11422.5649 | 219152.9643 | 176112.24 | 162841.207 | 3962.5518 | 4 | 18.6667 | 1070.0158 | 2247.4029
143 | LBLANK |  | 06/09/2023 03:57:12 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.497 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2 | 0 | 3456.4183 | 2 | 3479.7572 | 598.6792 | 16 | 0.66667 | 0.66667 | 9.3333 | 40.0001 | 569.0959 | 13.3333 | 1296.7258 | 0 | 10.6667 | 23.3334 | 99.3337 | 295.3364 | -2125.3133 | 0 | 12 | 12 | 1023.3701 | 1.3333 | 2 | 2.6667 | 25.7976 | 0 | 18.6667 | 7.3333 | 157064.7519 | 173960.1467 | 11847.5779 | 227799.8033 | 180086.8477 | 168707.9526 | 3641.1344 | 2 | 7.3333 | 22 | 39.3334
144 | CCV 1ppb TRACE |  | 06/09/2023 04:03:01 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.498 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1570.0863 | 634.0141 | 119828.4668 | 64.0001 | 4093.2533 | 836.6912 | 690.0167 | 277.336 | 1160.7139 | 1541.4166 | 8459.8377 | 6930.8037 | 13930.1232 | 2726.2605 | 3301.715 | 964.6992 | 2192.1683 | 17747.6843 | 8201.6874 | 3655.7736 | 98.0003 | 6825.6311 | 6825.6311 | 1669.4309 | 169.3343 | 458.674 | 970.0329 | 578.4872 | 189.3346 | 15529.1033 | 13400.2822 | 154572.681 | 185131.2128 | 11311.8107 | 212380.4019 | 172943.2455 | 164252.2592 | 4095.2537 | 7265.1803 | 17608.8464 | 28281.6478 | 59309.7571
145 | CCV 1ppb TRACE |  | 06/09/2023 04:09:18 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.499 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1492.0779 | 625.347 | 119971.0858 | 86.6669 | 3537.1046 | 798.0224 | 680.0162 | 270.0026 | 1148.0462 | 1549.4174 | 8546.5579 | 7052.8479 | 13974.1653 | 2746.9309 | 3261.0389 | 907.3622 | 2252.1776 | 17455.9964 | 8114.3052 | 3628.2287 | 95.3337 | 6782.9433 | 6782.9433 | 1658.0962 | 157.3342 | 475.3413 | 1010.036 | 643.2614 | 154.6675 | 15144.0239 | 13613.8184 | 159567.9868 | 184953.5998 | 11684.111 | 211545.504 | 174184.1932 | 162126.9152 | 4091.2528 | 7443.2721 | 17297.1328 | 28050.17 | 59745.4682
146 | LFB |  | 06/09/2023 04:15:06 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.500 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 13635.1708 | 5610.4349 | 592059.9383 | 366.6714 | 31695.1216 | 2700.2553 | 6238.6955 | 2630.2423 | 9840.0549 | 13920.7826 | 78294.641 | 52991.8662 | 127220.046 | 13285.5084 | 29280.6466 | 8141.6531 | 19501.9697 | 150673.1067 | 72106.8565 | 53844.4726 | 934.6973 | 63208.2286 | 63208.2286 | 6567.5099 | 1385.4007 | 4285.3097 | 9211.6357 | 5367.4274 | 1598.7564 | 139308.6708 | 123250.7336 | 149287.4518 | 171337.0217 | 10671.3178 | 198549.629 | 162675.7784 | 153381.075 | 3623.793 | 66761.6649 | 157076.1706 | 250284.4514 | 514459.6243
147 | LFB |  | 06/09/2023 04:21:22 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.501 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 13764.6281 | 5573.0869 | 617085.1084 | 384.6719 | 32696.7108 | 2708.2569 | 6206.6814 | 2642.2445 | 10134.9276 | 13910.7701 | 79001.8636 | 53430.7536 | 130697.8118 | 13549.7591 | 29968.0738 | 8187.0126 | 19285.0092 | 152837.2759 | 73822.9226 | 55187.5998 | 988.0342 | 64092.1184 | 64092.1184 | 6862.9819 | 1414.7368 | 4387.3405 | 9371.0747 | 5563.1395 | 1693.4337 | 141585.6035 | 124927.878 | 149086.14 | 176177.9428 | 10772.0605 | 201436.2583 | 163549.5724 | 159836.5468 | 3783.1679 | 67323.6095 | 160225.0039 | 251636.3206 | 515270.3913
148 | CCV 10 ppb MAJORS |  | 06/09/2023 04:27:10 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.502 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 14392.581 | 5837.8601 | 605681.0544 | 368.0048 | 27003.4987 | 2845.6169 | 6622.2013 | 2734.9285 | 10632.6227 | 15061.9403 | 83172.101 | 56769.0002 | 134187.9971 | 13958.8226 | 31003.607 | 8609.2602 | 19870.477 | 159496.9494 | 74576.1638 | 55617.9797 | 993.3679 | 64667.3832 | 64667.3832 | 7055.7423 | 1383.4003 | 4418.6833 | 9260.3341 | 5373.1629 | 1680.7656 | 148897.3659 | 129520.5139 | 154462.7777 | 176586.1935 | 10881.4764 | 200509.839 | 164486.2834 | 158628.5185 | 3757.8278 | 68534.6793 | 162542.898 | 261297.5714 | 530476.3829
149 | CCV 10 ppb MAJORS |  | 06/09/2023 04:33:27 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.503 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 14518.0402 | 6009.2637 | 618715.0141 | 377.3384 | 27166.4797 | 2867.6216 | 6676.2266 | 2826.9464 | 10963.5388 | 14841.7059 | 81840.6228 | 55397.0218 | 135509.6763 | 14317.8383 | 31345.6877 | 8531.214 | 20064.0815 | 158774.1831 | 74270.0526 | 56101.8774 | 974.0335 | 65139.5089 | 65139.5089 | 7181.1379 | 1371.3992 | 4385.3397 | 9263.6696 | 5479.1931 | 1694.7672 | 148335.0673 | 130035.2921 | 152288.8117 | 178980.9075 | 10780.7346 | 217471.477 | 169664.2779 | 158755.9251 | 3749.8256 | 68138.9002 | 163625.134 | 261640.3987 | 538983.1999
150 | MT 33 D |  | 06/09/2023 04:39:15 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\MT 33 D.504 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 6.6667 | 0 | 76524.559 | 48.6668 | 19062.0431 | 599.3459 | 15.3333 | 0.66667 | 2 | 22 | 98.667 | 1015.2633 | 61.3335 | 1404.069 | 2.6667 | 12.6667 | 110.0004 | 816.0233 | 630.014 | -2235.3272 | 0 | 31.3334 | 31.3334 | 1188.7161 | 0 | 0.66667 | 2.6667 | 75.6854 | 1.3333 | 294.003 | 63.3335 | 157169.3217 | 193333.4825 | 11542.6627 | 223903.7031 | 179278.6031 | 169033.6076 | 4152.6036 | 6 | 27.3334 | 895.3441 | 1734.0416
151 | MT 33 D LSP |  | 06/09/2023 04:45:03 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\MT 33 D LSP.505 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 35959.2059 | 14143.6659 | 1592841.4496 | 979.367 | 83186.1963 | 6052.6154 | 16802.5445 | 6843.639 | 26487.2001 | 36560.7245 | 202587.6862 | 137421.0277 | 331751.2276 | 32228.3133 | 75649.1395 | 20014.6808 | 47560.3754 | 374448.3275 | 171330.6828 | 136270.5583 | 2184.8339 | 151997.0101 | 151997.0101 | 14604.1294 | 3218.3627 | 10172.9552 | 21719.1645 | 12921.9194 | 3843.1836 | 359158.0305 | 318575.8091 | 140417.4447 | 169745.2168 | 10445.1523 | 193889.6775 | 160903.1263 | 160461.6438 | 3572.447 | 159713.2183 | 375747.6143 | 618578.1898 | 1247314.527
152 | cal blank |  | 06/09/2023 04:50:51 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\cal blank.506 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.3333 | 0 | 2274.8479 | 0 | 2742.9305 | 594.0124 | 10 | 0 | 1.3333 | 5.3333 | 16.6667 | 500.9208 | 14.6667 | 1311.3935 | 1.3333 | 13.3333 | 4.6667 | 82.6669 | 242.0021 | -2421.5372 | 0.66667 | 41.3334 | 41.3334 | 1100.7092 | 1.3333 | 4.6667 | 1.3333 | 57.2563 | 1.3333 | 546.0105 | 60.0001 | 156570.9012 | 194908.0099 | 11419.8971 | 218492.9724 | 175194.5306 | 167602.1307 | 4056.576 | 3.3333 | 10 | 31.3333 | 60.6667
153 | STD 0.01 |  | 06/09/2023 04:56:40 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 0.01.507 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 19.3333 | 11.3333 | 38896.235 | 26 | 4041.9052 | 562.6778 | 16.6667 | 3.3333 | 8.6667 | 33.3334 | 108.0004 | 1335.7857 | 536.6768 | 1487.4109 | 43.3334 | 16.6667 | 53.3334 | 390.0053 | 1377.3997 | -2269.1378 | 0 | 109.3338 | 109.3338 | 1150.0464 | 1.3333 | 6 | 10 | 61.2889 | 2 | 632.014 | 348.6709 | 157612.2063 | 190424.8422 | 11286.4569 | 217771.4967 | 177936.0379 | 163157.2009 | 4127.93 | 60.6668 | 173.3344 | 342.6683 | 718.6737
154 | std 0.05 |  | 06/09/2023 05:02:29 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\std 0.05.508 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 84.0002 | 42.6667 | 42601.4737 | 32.6667 | 5960.577 | 572.6782 | 50.6668 | 20.6667 | 58.0001 | 100.0004 | 484.0082 | 1440.5918 | 753.3532 | 1439.406 | 196.0014 | 62.6668 | 154.0008 | 1131.3782 | 1086.0413 | -1867.5249 | 7.3333 | 375.3383 | 375.3383 | 1088.7084 | 7.3333 | 28.6667 | 64.6668 | 91.675 | 8.6667 | 953.3652 | 706.0175 | 156044.4225 | 190232.9325 | 11310.4761 | 215163.7124 | 173056.9367 | 167774.3412 | 3991.2249 | 394.0054 | 880.6938 | 1436.6945 | 3070.7968
155 | std 0.1 |  | 06/09/2023 05:08:17 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\std 0.1.509 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 160.0009 | 62.6668 | 22314.0823 | 10 | 3517.0997 | 638.681 | 78.0002 | 32.6667 | 124.0005 | 162.0009 | 884.0274 | 2382.4062 | 1521.4144 | 1796.1129 | 352.0043 | 106.0004 | 288.0029 | 2024.1434 | 1768.1096 | -1621.0235 | 12.6667 | 707.3509 | 707.3509 | 1216.0518 | 18 | 44.6667 | 108.0004 | 100.3421 | 18 | 1688.7666 | 1403.4023 | 157956.4651 | 191635.5126 | 11592.7024 | 219999.8112 | 178343.3103 | 168154.1191 | 4199.284 | 742.0193 | 1764.7757 | 2845.4476 | 5924.4824
156 | std 0.5 |  | 06/09/2023 05:14:05 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\std 0.5.510 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 778.6879 | 335.3373 | 84480.4534 | 46.6667 | 7021.7257 | 729.352 | 376.005 | 183.3345 | 550.6773 | 834.6911 | 4501.3762 | 6677.2426 | 7375.9037 | 2715.5915 | 1724.7708 | 517.3427 | 1248.0545 | 9533.847 | 5693.1348 | 707.4434 | 46.6667 | 3540.4387 | 3540.4387 | 1370.0658 | 80.6669 | 244.6688 | 498.6754 | 326.3917 | 88.6669 | 8008.9112 | 6976.3699 | 157815.5821 | 191228.2808 | 11349.1761 | 220630.1822 | 176173.6669 | 167704.3018 | 4109.9245 | 3769.8308 | 8739.3392 | 14362.8163 | 29196.3154
157 | std 1 |  | 06/09/2023 05:19:54 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\std 1.511 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1514.747 | 626.0138 | 126141.1574 | 74.6669 | 6306.0594 | 787.3551 | 682.683 | 304.6699 | 1155.3801 | 1570.0863 | 8754.6823 | 8997.7784 | 14368.5565 | 3309.7167 | 3348.3925 | 982.7005 | 2222.8396 | 17749.6885 | 9094.2273 | 3358.729 | 104.6671 | 6780.2766 | 6780.2766 | 1609.424 | 144.0007 | 448.6737 | 1021.37 | 545.0111 | 172.6677 | 15611.8595 | 13552.4302 | 156218.8549 | 187083.8969 | 11381.2006 | 206843.3169 | 181069.7969 | 165431.8387 | 4018.5652 | 7245.8372 | 17235.7252 | 27736.5601 | 56717.362
158 | std 5 |  | 06/09/2023 05:25:42 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\std 5.512 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 7153.1239 | 2962.9743 | 346825.7769 | 205.3348 | 18650.1668 | 1722.1038 | 3405.0725 | 1349.3972 | 5338.9975 | 7454.6113 | 41278.8913 | 29323.4532 | 67799.8627 | 7913.5246 | 15497.0712 | 4323.9885 | 10040.1945 | 80992.9905 | 38821.3565 | 26364.9242 | 476.008 | 32907.8613 | 32907.8613 | 4024.5672 | 694.0169 | 2232.8412 | 4731.4505 | 2746.6285 | 842.0249 | 73540.1757 | 64052.6099 | 157039.8582 | 181181.7057 | 10974.2139 | 192757.2274 | 167886.087 | 148275.1615 | 3783.8345 | 34883.2139 | 82879.0978 | 130172.3365 | 268673.6002
159 | std 10 |  | 06/09/2023 05:31:30 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\std 10.513 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 14409.2651 | 5855.8674 | 621807.1397 | 359.3379 | 29615.3347 | 2794.9404 | 6414.1066 | 2686.9195 | 10462.4974 | 14554.0797 | 81298.6931 | 55388.7765 | 133391.2929 | 13993.5184 | 30950.1584 | 8505.8651 | 19797.0433 | 158160.7411 | 73792.7748 | 54844.7184 | 970.6997 | 63988.9943 | 63988.9943 | 6999.0485 | 1378.0665 | 4215.9555 | 9109.5702 | 5426.4874 | 1629.4263 | 145877.7312 | 126686.6086 | 148900.0915 | 177732.8159 | 10631.2883 | 201299.9614 | 165117.4928 | 154503.0031 | 3825.8458 | 67924.438 | 160709.6024 | 254717.2795 | 521454.6908
160 | LBLANK |  | 06/09/2023 05:37:19 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.514 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.3333 | 0.66667 | 3630.4613 | 2 | 1577.4205 | 606.6796 | 9.3333 | 0 | 0.66667 | 10.6667 | 48.0001 | 595.0217 | 16 | 1310.7268 | 3.3333 | 10.6667 | 12 | 96.0003 | 276.6694 | -2148.5436 | 1.3333 | 22 | 22 | 1111.3767 | 0 | 0.66667 | 2 | 45.8523 | 0 | 266.6692 | 42.0001 | 154786.778 | 190070.8361 | 11415.2266 | 219988.3727 | 169769.3334 | 165354.9203 | 4104.5897 | 1.3333 | 6 | 24.6667 | 39.3334
161 | CCV 1ppb TRACE |  | 06/09/2023 05:43:07 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.515 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1572.7534 | 652.6816 | 126166.0389 | 82.0003 | 3721.1522 | 826.024 | 731.3521 | 290.6697 | 1194.0499 | 1616.7582 | 8870.7539 | 7240.0203 | 14203.7244 | 2980.311 | 3468.4211 | 987.3675 | 2306.8534 | 18114.1491 | 8259.0534 | 3861.8861 | 124.0005 | 6986.3749 | 6986.3749 | 1717.4367 | 140.0007 | 486.675 | 972.0331 | 611.7207 | 176.0011 | 15794.061 | 13686.5578 | 159196.888 | 187135.0609 | 10774.0736 | 216810.1785 | 174450.1546 | 167448.3915 | 4079.9159 | 7466.6183 | 17686.9429 | 28058.1752 | 58122.1941
162 | CCV 1ppb TRACE |  | 06/09/2023 05:49:25 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.516 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1558.085 | 614.6799 | 123295.8122 | 80.6669 | 3400.4051 | 828.6907 | 693.3502 | 265.3358 | 1190.0498 | 1603.4234 | 8792.0384 | 7084.0046 | 14635.494 | 2894.2932 | 3308.3838 | 881.3605 | 2216.8387 | 17876.5114 | 8165 | 3721.7692 | 99.3337 | 6971.7008 | 6971.7008 | 1670.7646 | 152.0008 | 476.0079 | 994.7013 | 642.5512 | 170.6677 | 15490.3942 | 14164.352 | 155607.6932 | 180676.6122 | 11362.5172 | 208313.3726 | 172798.3227 | 165632.8745 | 3799.1723 | 7157.1258 | 17295.7993 | 27920.0032 | 58150.4275
163 | CCV 10 ppb MAJORS |  | 06/09/2023 05:55:13 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.517 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 14311.1648 | 5820.5192 | 591833.0647 | 354.0044 | 26897.9676 | 2840.2824 | 6632.873 | 2680.2514 | 10664.6466 | 14689.5487 | 80774.384 | 54270.3535 | 132872.4453 | 13978.1697 | 30691.6016 | 8195.0169 | 19739.6285 | 156104.2987 | 72261.6442 | 54522.1321 | 1015.3694 | 64126.2727 | 64126.2727 | 6921.01 | 1280.7241 | 4279.3076 | 9054.8688 | 5351.3427 | 1612.758 | 143284.368 | 127334.9909 | 147337.2497 | 175272.6682 | 10731.3629 | 200944.846 | 134163.6594 | 157002.2312 | 3690.4773 | 67003.4331 | 159121.3533 | 254198.4219 | 541632.7512
164 | CCV 10 ppb MAJORS |  | 06/09/2023 06:01:30 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.518 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 14339.8608 | 5765.1632 | 609052.3262 | 352.671 | 26823.1692 | 2767.6016 | 6530.159 | 2692.9206 | 10334.4041 | 14395.9174 | 81009.7107 | 54120.5538 | 133920.163 | 13499.709 | 30292.7576 | 8301.7455 | 19389.1509 | 156043.1133 | 73061.6896 | 55120.9357 | 1016.7029 | 64286.3257 | 64286.3257 | 6882.9918 | 1370.7325 | 4367.3344 | 9166.9406 | 5351.5857 | 1680.7656 | 144293.0677 | 128304.3755 | 147490.2418 | 174108.6037 | 10716.6913 | 199590.6324 | 164111.9679 | 155449.3013 | 3698.4788 | 66831.2944 | 157845.9252 | 252658.3812 | 519659.315
165 | std 25 |  | 06/09/2023 06:07:19 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\std 25.519 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 34382.6596 | 13463.0086 | 1472271.4546 | 842.6915 | 68867.601 | 6005.9312 | 16140.4469 | 6540.832 | 25538.8093 | 34908.5997 | 197801.306 | 133612.2784 | 320159.9256 | 30670.2296 | 73956.945 | 19582.748 | 45981.2316 | 371387.7457 | 177283.2872 | 140812.4125 | 2346.8596 | 158406.0804 | 158406.0804 | 15268.1559 | 3519.7669 | 10964.8734 | 22698.0187 | 13639.5409 | 4079.2492 | 353619.1358 | 312187.1919 | 141982.8347 | 165340.0478 | 10245.0083 | 188267.1813 | 151611.6471 | 149413.9832 | 3492.4268 | 159072.1039 | 379038.2608 | 594657.6583 | 1180147.4148
166 | std 50 |  | 06/09/2023 06:13:06 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\std 50.520 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 66656.5047 | 26308.2065 | 3486435.2483 | 1686.0995 | 131128.5607 | 11273.7801 | 31332.9923 | 13104.0077 | 50641.6028 | 68754.393 | 381635.2471 | 235606.5579 | 623615.98 | 58878.4277 | 142479.0033 | 37632.8401 | 87980.7786 | 685697.0019 | 329391.8199 | 272715.8992 | 4507.378 | 305626.1153 | 305626.1153 | 28504.4098 | 6581.5161 | 20915.9683 | 43930.7817 | 26016.4877 | 7944.2086 | 675336.8265 | 603651.9695 | 135239.8081 | 148310.2448 | 9762.6687 | 176054.3368 | 144712.6205 | 139015.1891 | 3093.6685 | 297102.7598 | 699009.7241 | 1100152.6998 | 2181219.0592
167 | std 50 |  | 06/09/2023 06:19:23 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\std 50.521 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 68477.0608 | 27090.9976 | 3569258.0947 | 1734.1053 | 135179.9751 | 11734.821 | 31903.5923 | 13246.8067 | 51394.2971 | 70046.0172 | 382351.0768 | 234987.2865 | 622611.3775 | 59647.6226 | 142629.8542 | 37912.9101 | 89369.3423 | 686787.3784 | 329457.3554 | 273775.3728 | 4542.7223 | 301535.2505 | 301535.2505 | 28397.5302 | 6501.4792 | 20805.1447 | 44021.0546 | 26445.7632 | 7933.5362 | 670346.9568 | 604174.7518 | 134595.8068 | 156061.2002 | 9015.5131 | 176618.8651 | 143217.7041 | 138680.5929 | 3281.7104 | 292409.7436 | 694985.1807 | 1091139.4032 | 2151540.1234
168 | std 100 |  | 06/09/2023 06:25:11 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\std 100.522 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 127618.1718 | 49649.4636 | 6619440.2364 | 3119.0071 | 236991.2667 | 21653.7332 | 61558.3564 | 25043.942 | 98016.4745 | 131128.4385 | 716554.1374 | 426842.8666 | 1169745.4725 | 107651.4443 | 266402.8193 | 69469.1812 | 162259.6827 | 1210347.888 | 582678.7035 | 502404.587 | 8413.1439 | 547667.108 | 547667.108 | 50510.4754 | 12378.6956 | 38060.6399 | 80857.5277 | 47977.798 | 14564.7554 | 1250986.0802 | 1122811.0601 | 120715.3029 | 140686.0347 | 8557.8968 | 157305.5139 | 130202.6722 | 125507.0919 | 2918.2981 | 525505.7743 | 1237627.179 | 1879544.852 | 3793194.5676
169 | std 100 |  | 06/09/2023 06:31:28 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\std 100.523 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 127776.274 | 50603.4907 | 6750773.3519 | 3117.6735 | 239713.8562 | 22053.678 | 61604.5481 | 25535.4704 | 100495.6326 | 133717.6487 | 721373.0405 | 436770.0517 | 1177942.136 | 110058.3545 | 268054.9636 | 69171.0642 | 165607.7302 | 1230231.0264 | 583392.4602 | 507412.256 | 8513.8699 | 556542.745 | 556542.745 | 50980.8245 | 12374.0247 | 38535.239 | 81775.386 | 48870.3332 | 14743.6047 | 1266287.4285 | 1136217.029 | 123629.387 | 142489.1334 | 8438.4922 | 158408.0955 | 130628.5452 | 125956.8827 | 3022.3197 | 524381.1921 | 1236483.4982 | 1912634.4476 | 4294956.9683
170 | LBLANK |  | 06/09/2023 06:37:16 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.524 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 8 | 1.3333 | 3943.8777 | 3.3333 | 1928.797 | 637.3476 | 16 | 0 | 4 | 12 | 46.0001 | 595.3906 | 30 | 1378.7332 | 14.6667 | 20 | 19.3333 | 122.0005 | 274.6693 | -2419.9325 | 6 | 219.335 | 219.335 | 1117.3771 | 2 | 5.3333 | 11.3333 | 36.9905 | 0 | 1592.7555 | 262.6691 | 157106.9357 | 187116.3508 | 11189.7223 | 214600.6064 | 174616.9113 | 166481.1169 | 4063.2446 | 10.6667 | 28 | 25.3333 | 74.0001
171 | PE ICV |  | 06/09/2023 06:43:04 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\PE ICV.525 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 14558.7526 | 6145.3219 | 637771.8355 | 376.6716 | 55971.4425 | 2848.951 | 6750.2618 | 2772.9359 | 10457.1614 | 14415.27 | 82511.6194 | 56179.6503 | 136665.3636 | 13844.7064 | 30218.594 | 8425.1506 | 19715.5961 | 154459.2008 | 71048.2486 | 55286.8288 | 856.0256 | 64098.8278 | 64098.8278 | 5381.0136 | 1372.7326 | 4396.6767 | 9285.6844 | 5528.3998 | 1668.0976 | 138590.3457 | 121847.5004 | 149682.131 | 175567.142 | 10213.6574 | 198430.7182 | 164337.1699 | 156597.1076 | 3779.1665 | 65797.8486 | 156276.8155 | 248120.8712 | 508069.2205
172 | 5ppb CCV |  | 06/09/2023 06:48:52 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5ppb CCV.526 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 7405.9192 | 3106.3378 | 355452.3442 | 228.6685 | 18684.2151 | 1790.1122 | 3378.3995 | 1389.4011 | 5413.0255 | 7530.6511 | 41743.5665 | 30158.295 | 68316.9726 | 7912.8575 | 15784.0493 | 4288.6436 | 10348.4142 | 81267.8475 | 38487.1139 | 27506.637 | 490.6751 | 32961.99 | 32961.99 | 4149.9361 | 728.6853 | 2298.8517 | 4823.4808 | 2834.4948 | 828.024 | 73336.4231 | 65012.2988 | 146905.2432 | 179876.3598 | 10960.203 | 203162.4332 | 164918.5365 | 159012.9743 | 3958.5487 | 33545.3408 | 79981.2934 | 127694.7421 | 257477.8812
173 | 25ppb CCV |  | 06/09/2023 06:54:40 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\25ppb CCV.527 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 35288.2236 | 14130.3212 | 1485726.9142 | 908.0289 | 70695.8455 | 6218.0203 | 16300.6285 | 6776.2753 | 26390.3545 | 35925.1201 | 198495.538 | 132630.5694 | 320607.5894 | 31145.9377 | 74578.8493 | 20070.0901 | 47784.4527 | 367022.8972 | 175452.368 | 143415.4781 | 2448.2101 | 158254.3991 | 158254.3991 | 15618.5332 | 3397.7375 | 10730.6967 | 22774.8073 | 14086.2252 | 4171.2756 | 350346.87 | 307953.9428 | 141419.2986 | 163972.4334 | 9870.0757 | 184495.8479 | 150914.9618 | 145452.1598 | 3470.4215 | 157765.0837 | 372539.1899 | 588072.341 | 1169803.3478
174 | 8_30_21_W_1 |  | 06/09/2023 07:00:27 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_W_1.528 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2093.4873 | 693.3502 | 1488724.2382 | 849.3587 | 16742.6434 | 1553.4178 | 258.669 | 1842.1189 | 78.0002 | 56.0001 | 280.6695 | 36862.5095 | 42180.1826 | 9722.6418 | 30 | 74.0002 | 688.6833 | 5346.3336 | 3446.4157 | 3495.0386 | 94.667 | 6481.4704 | 6481.4704 | 1144.0458 | 2 | 6.6667 | 10.6667 | 58.2542 | 2.6667 | 43251.3833 | 37991.116 | 154875.1973 | 177708.5803 | 10856.1252 | 207322.7759 | 168202.6371 | 155482.2986 | 3757.8276 | 15.3333 | 42.0001 | 1130.6846 | 2254.0686
175 | LBLANK |  | 06/09/2023 07:06:16 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.529 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.3333 | 0.66667 | 3645.1325 | 2.6667 | 2393.534 | 604.6795 | 17.3333 | 0 | 1.3333 | 11.3333 | 30.6667 | 564.4103 | 17.3333 | 1309.3933 | 0.66667 | 16 | 11.3333 | 88.6669 | 294.003 | -2408.5508 | 0.66667 | 37.3334 | 37.3334 | 1080.0409 | 0.66667 | 0.66667 | 6 | 5.7104 | 0.66667 | 613.3465 | 56.6668 | 145640.501 | 180032.0656 | 11136.3409 | 208867.196 | 169715.6274 | 159602.0507 | 3883.1945 | 2.6667 | 10.6667 | 28.6667 | 42.6667
176 | CCV 1ppb TRACE |  | 06/09/2023 07:12:04 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.530 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1548.0839 | 649.3482 | 122515.1674 | 76.6669 | 3607.1222 | 782.6881 | 713.3512 | 294.6697 | 1172.7149 | 1544.0836 | 8519.207 | 6988.7813 | 13702.5736 | 2776.9371 | 3208.3604 | 917.3628 | 2213.5049 | 17315.8264 | 7928.1994 | 3360.9636 | 88.0003 | 6598.8572 | 6598.8572 | 1698.7677 | 154.6675 | 438.0067 | 972.0331 | 632.2015 | 154.0008 | 15315.5413 | 12957.8754 | 155464.839 | 187777.7061 | 10998.8992 | 199842.7774 | 165632.8403 | 152632.8781 | 4063.9127 | 6902.3342 | 16827.9076 | 26726.434 | 55554.9815
177 | CCV 1ppb TRACE |  | 06/09/2023 07:18:22 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.531 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.6667 | 0 | 1794.7798 | 0 | 240.6687 | 74.0002 | 6 | 0 | 0 | 7.3333 | 27.3334 | 462.4059 | 6.6667 | 498.0087 | 6 | 5.3333 | 7.3333 | 64.0001 | 188.6679 | 938.6123 | 0.66667 | 2.6667 | 2.6667 | 174.6677 | 0.66667 | 1.3333 | 0.66667 | -9.5131 | 0 | 557.3442 | 3.3333 | 165357.9153 | 190031.871 | 12569.5375 | 232457.9999 | 188271.8174 | 156045.3393 | 4049.9081 | 2.6667 | 4.6667 | 12.6667 | 19.3333
178 | LFB |  | 06/09/2023 07:24:08 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.532 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 13965.4897 | 5612.4364 | 594203.3692 | 372.0048 | 31781.3126 | 2672.2502 | 6250.7009 | 2504.8865 | 10092.231 | 14100.956 | 77404.4852 | 53374.7501 | 127229.3787 | 13037.2814 | 29076.2272 | 7870.1674 | 18830.4063 | 148498.5278 | 70802.3572 | 53600.3394 | 909.3623 | 61899.1603 | 61899.1603 | 6627.5379 | 1336.7292 | 4229.2926 | 9016.8452 | 5408.8379 | 1524.748 | 135793.2269 | 119899.7382 | 151337.2326 | 167137.4519 | 10294.3758 | 191588.2543 | 156344.8833 | 148907.4403 | 3592.4522 | 63678.2777 | 150553.94 | 237623.4428 | 482576.9156
179 | LFB |  | 06/09/2023 07:30:24 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.533 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 13556.4314 | 5635.7781 | 599999.2993 | 364.6713 | 31026.9908 | 2659.5813 | 6197.3444 | 2578.2326 | 10012.1753 | 13781.3158 | 78129.7554 | 52749.5917 | 127801.8083 | 13234.7966 | 29055.5182 | 7848.1555 | 18844.4226 | 147515.1299 | 70510.9264 | 54053.6681 | 914.0293 | 62419.4114 | 62419.4114 | 6810.2908 | 1425.4045 | 4187.9471 | 8853.4097 | 5262.321 | 1637.4272 | 136587.2064 | 121042.0705 | 145004.9207 | 167536.209 | 10230.9967 | 190355.9021 | 157376.7803 | 144830.4943 | 3612.4567 | 62883.4534 | 148660.8492 | 238504.1459 | 482131.8033
180 | CCV 10 ppb MAJORS |  | 06/09/2023 07:36:12 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.534 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 14175.03 | 5729.1486 | 592129.021 | 361.3379 | 26641.4856 | 2870.9553 | 6502.8132 | 2763.6008 | 10639.9622 | 14624.1485 | 80093.9001 | 54195.78 | 132276.2757 | 13727.2596 | 30660.2014 | 8341.7686 | 19636.8222 | 152543.4424 | 71507.871 | 53891.1901 | 971.3664 | 62949.7274 | 62949.7274 | 6768.2699 | 1352.7308 | 4239.9625 | 8849.407 | 5371.3104 | 1658.7632 | 140887.9912 | 125539.8799 | 149186.4271 | 168394.7296 | 9946.1314 | 193726.1363 | 157914.8244 | 148606.2723 | 3613.1238 | 65519.2386 | 155548.9242 | 246796.7897 | 497069.7372
181 | CCV 10 ppb MAJORS |  | 06/09/2023 07:42:30 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.535 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 14492.6817 | 5860.5353 | 593480.7726 | 372.6715 | 26391.0232 | 2794.9401 | 6546.8337 | 2586.2341 | 10341.0756 | 14397.9189 | 80027.5693 | 54497.1535 | 131972.8955 | 13797.3295 | 30379.6455 | 8222.3657 | 19568.0661 | 153322.5229 | 71630.4752 | 54001.8495 | 928.0302 | 62606.8927 | 62606.8927 | 6913.0061 | 1291.3918 | 4126.596 | 8886.7633 | 5275.5626 | 1562.0855 | 142518.8216 | 124876.7727 | 146961.4027 | 168933.8475 | 10128.9273 | 192408.6597 | 149013.4214 | 150194.8965 | 3563.1111 | 65621.7077 | 153809.605 | 247792.5127 | 502914.0493
182 | 8_30_21_W_6 |  | 06/09/2023 07:48:18 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_W_6.536 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2367.5302 | 824.6906 | 633967.0785 | 386.6719 | 10294.375 | 1333.3956 | 258.669 | 1858.7877 | 158.0009 | 50.6668 | 269.3359 | 46332.8682 | 42209.6109 | 12061.7596 | 23.3334 | 85.3336 | 1828.7838 | 14132.9891 | 5132.9227 | 3221.9572 | 73.3335 | 6554.8369 | 6554.8369 | 1092.7085 | 6 | 12.6667 | 18 | 55.3473 | 8.6667 | 23548.7291 | 20613.5287 | 149406.6056 | 173925.0338 | 10468.5261 | 198941.1726 | 155244.7533 | 153147.2831 | 3750.4925 | 28 | 77.3335 | 2090.726 | 4313.5875
183 | 8_30_21_W_6_LD |  | 06/09/2023 07:54:05 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_W_6_LD.537 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2279.5153 | 776.0211 | 835471.3929 | 491.3418 | 12628.2458 | 1314.7273 | 252.0023 | 1795.4462 | 151.3341 | 38.6667 | 208.0015 | 45585.1238 | 42237.6871 | 11758.1705 | 21.3333 | 93.3336 | 1746.1068 | 13886.747 | 5101.5779 | 3350.5885 | 68.6668 | 6494.1425 | 6494.1425 | 1152.0466 | 2.6667 | 12.6667 | 20.6667 | 146.4271 | 4 | 23089.9802 | 20672.2833 | 150621.2853 | 158629.5763 | 10493.8533 | 198241.8724 | 167095.3 | 152701.2267 | 3326.3886 | 49.3334 | 106.6671 | 2232.0681 | 4410.9328
184 | 8_30_21_W_6_LD |  | 06/09/2023 08:00:20 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_W_6_LD.538 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2273.5144 | 792.022 | 778270.5709 | 485.3416 | 13138.0389 | 1322.0614 | 255.3356 | 1887.4581 | 152.0008 | 40.0001 | 199.3347 | 44597.2794 | 41480.1374 | 12099.1223 | 22 | 85.3336 | 1796.7797 | 13608.4788 | 5040.8893 | 3163.3831 | 80.0002 | 6432.1149 | 6432.1149 | 1048.7053 | 4 | 10.6667 | 22 | 95.5916 | 6 | 22834.9082 | 20234.3213 | 145182.7198 | 169908.4411 | 9872.0843 | 192877.3723 | 140220.5532 | 149360.8168 | 3691.8103 | 36 | 76.6669 | 2218.7364 | 4432.9418
185 | 8_30_21_W_6_DUP |  | 06/09/2023 08:06:08 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_W_6_DUP.539 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2295.5178 | 776.6878 | 769112.6047 | 478.6747 | 13378.2634 | 1327.395 | 252.6689 | 1795.4464 | 128.0006 | 36.6667 | 199.3347 | 48350.6142 | 41772.9868 | 12781.0517 | 22 | 81.3336 | 1930.7971 | 15128.6769 | 5113.5818 | 3104.924 | 75.3335 | 6212.0171 | 6212.0171 | 1048.0385 | 2 | 4.6667 | 22.6667 | 110.9794 | 2.6667 | 22935.0694 | 20328.4551 | 147359.4521 | 169787.1242 | 10476.5069 | 193953.283 | 163512.6069 | 150696.0416 | 3660.4694 | 28.6667 | 80.6669 | 2557.4234 | 4835.6544
186 | 8_30_21_W_6 |  | 06/09/2023 08:11:56 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_W_6.540 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2172.1653 | 773.3543 | 610797.3386 | 357.3378 | 12646.266 | 1295.3921 | 243.3355 | 1814.782 | 130.6673 | 44.0001 | 254.669 | 43916.1245 | 40924.5378 | 11493.2899 | 18.6667 | 88.667 | 1693.4338 | 13276.8344 | 4862.1606 | 3053.0904 | 95.3337 | 6385.427 | 6385.427 | 1030.7039 | 0.66667 | 5.3333 | 16 | 62.382 | 4 | 22601.8655 | 20228.9799 | 150790.3772 | 169304.0782 | 9877.415 | 190506.5364 | 157168.124 | 157471.6131 | 3622.4596 | 34.6667 | 91.3336 | 2160.0645 | 4150.239
187 | 8_30_21_W_6_DUP_LSP |  | 06/09/2023 08:17:43 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_W_6_DUP_LSP.541 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 28195.7999 | 11089.6366 | 1809409.8133 | 1081.3744 | 59793.5832 | 5547.7441 | 11951.6662 | 6814.292 | 19162.1779 | 25979.6155 | 145776.6916 | 144311.1226 | 274492.7772 | 34457.5077 | 53683.3678 | 14906.4412 | 35858.9536 | 282965.8107 | 131550.2506 | 109190.1642 | 1764.1091 | 119309.4786 | 119309.4786 | 11663.436 | 2566.2305 | 7742.7646 | 16282.61 | 9998.6482 | 3002.9824 | 280603.7595 | 246603.8146 | 138130.4884 | 152982.1015 | 8740.6751 | 174316.4669 | 142274.3473 | 135158.6776 | 3334.3891 | 118634.6136 | 279202.0386 | 455652.6737 | 897332.1715
188 | 8_30_21_W_6_DUP_LSP |  | 06/09/2023 08:24:01 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_W_6_DUP_LSP.542 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 28535.1386 | 11383.2002 | 1859362.2233 | 1104.7095 | 61701.6398 | 5666.4576 | 12309.3019 | 6931.0149 | 19732.9528 | 26922.6788 | 148207.5265 | 145950.9596 | 282431.4076 | 35211.3423 | 55571.2306 | 15257.4774 | 36838.7768 | 289989.8664 | 135176.5372 | 111991.3326 | 1882.1241 | 122474.113 | 122474.113 | 11967.0104 | 2521.5559 | 7949.545 | 16775.1776 | 10227.2427 | 3056.3271 | 284916.1252 | 252932.8404 | 136944.5899 | 154600.0637 | 9270.3431 | 177227.2315 | 144773.2118 | 138068.7103 | 3279.043 | 119110.5861 | 282180.4715 | 453707.4664 | 901559.5623
189 | 8_30_21_W_8 |  | 06/09/2023 08:29:49 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_W_8.543 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2470.2139 | 804.0227 | 633982.8003 | 400.0056 | 14100.292 | 1310.0601 | 253.3356 | 1775.4437 | 157.3342 | 66.0002 | 287.3362 | 38180.4455 | 45291.6897 | 10608.6043 | 26 | 54.0001 | 1816.1156 | 13686.5535 | 3565.7783 | 3691.4003 | 88.6669 | 6593.5213 | 6593.5213 | 1043.3714 | 1.3333 | 10 | 22.6667 | 90.3826 | 4 | 19903.1889 | 17488.6994 | 146493.2843 | 167701.2629 | 9810.7183 | 185225.8458 | 157585.0562 | 121553.3444 | 3689.8102 | 28.6667 | 70.6668 | 1665.3713 | 3312.1529
190 | 8_30_21_W_8 |  | 06/09/2023 08:36:07 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_W_8.544 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2360.8618 | 726.6853 | 624955.2809 | 382.0052 | 12823.7544 | 1270.7234 | 244.6688 | 1782.7779 | 187.3346 | 58.0001 | 280.0028 | 37218.3039 | 44342.7165 | 10076.8868 | 17.3333 | 42.0001 | 1791.4459 | 13370.9228 | 3659.8021 | 3458.3672 | 88.0003 | 6575.5134 | 6575.5134 | 1029.3705 | 4.6667 | 9.3333 | 16.6667 | 85.0383 | 1.3333 | 19420.5264 | 17308.4797 | 142407.6454 | 162685.7771 | 10230.3291 | 187488.5767 | 154796.3362 | 143734.2788 | 3515.0993 | 24.6667 | 80.6669 | 1608.0359 | 3324.8178
191 | 8_30_21_W_8_SP |  | 06/09/2023 08:41:55 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_W_8_SP.545 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3711.8157 | 1308.06 | 671680.8237 | 431.3399 | 15713.9736 | 1418.7372 | 896.6948 | 2029.4778 | 1208.0515 | 1497.4119 | 8192.3489 | 40097.2065 | 55126.8266 | 10574.5807 | 3035.6559 | 910.6957 | 3569.7804 | 27702.1721 | 10548.5598 | 9157.9754 | 188.6679 | 12879.8042 | 12879.8042 | 1560.0853 | 137.334 | 444.0069 | 934.6973 | 560.9667 | 177.3344 | 33266.0218 | 29283.9842 | 139602.8258 | 159451.6435 | 9715.3099 | 188522.4853 | 150926.4931 | 141145.9553 | 3401.7384 | 6734.254 | 15856.7966 | 27696.5351 | 57301.1368
192 | LBLANK |  | 06/09/2023 08:47:44 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.546 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.66667 | 0 | 3423.7436 | 1.3333 | 2371.5307 | 540.0102 | 9.3333 | 0 | 0 | 9.3333 | 39.3334 | 555.5012 | 18 | 1094.7087 | 2 | 22.6667 | 8 | 81.3336 | 280.0028 | -2446.1901 | 0.66667 | 21.3334 | 21.3334 | 1006.0354 | 0.66667 | 1.3333 | 4.6667 | 54.5659 | 0 | 65.3335 | 27.3334 | 146077.7297 | 170535.4398 | 10438.4848 | 206665.3545 | 160056.34 | 140850.6715 | 3745.8264 | 2 | 8 | 21.3333 | 32
193 | CCV 1ppb TRACE |  | 06/09/2023 08:53:33 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.547 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1417.4036 | 564.6779 | 115148.2306 | 73.3335 | 3451.7514 | 796.6889 | 688.0166 | 266.0025 | 1058.0392 | 1504.0792 | 7936.2047 | 6347.232 | 13015.2616 | 2468.8801 | 3004.9829 | 828.024 | 1995.4728 | 16221.2064 | 7363.8977 | 3172.1645 | 96.0003 | 6224.0225 | 6224.0225 | 1531.4154 | 135.334 | 432.6733 | 934.6972 | 565.3225 | 151.3341 | 14047.5706 | 12329.9855 | 139102.7437 | 165912.5776 | 10287.7137 | 190525.4522 | 154163.044 | 144623.2522 | 3619.1251 | 6662.2202 | 15485.7254 | 26262.7658 | 52502.4665
194 | CCV 1ppb TRACE |  | 06/09/2023 08:59:50 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.548 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1470.7425 | 580.0118 | 113406.5143 | 64.6668 | 3166.351 | 824.0239 | 653.3483 | 272.0026 | 1063.3729 | 1430.7383 | 7828.1451 | 6429.1292 | 12796.3954 | 2578.8995 | 2931.6341 | 836.0245 | 2097.4873 | 15935.5506 | 7399.2495 | 3331.9338 | 88.6669 | 5991.2568 | 5991.2568 | 1529.4156 | 126.6672 | 452.6738 | 900.0284 | 586.7974 | 173.3344 | 13975.5007 | 12370.0223 | 143992.0534 | 163308.1811 | 10189.6379 | 187778.1753 | 154068.3852 | 144204.8541 | 3506.4305 | 6325.4001 | 15475.0463 | 25807.0459 | 52225.3583
195 | CCV 10 ppb MAJORS |  | 06/09/2023 09:05:39 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.549 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 13398.2813 | 5418.3611 | 545738.1709 | 364.6713 | 24249.2302 | 2700.2555 | 6067.9553 | 2546.2269 | 9736.6503 | 13488.3692 | 74604.3087 | 50787.4396 | 121082.3148 | 12806.4045 | 28097.617 | 7666.7254 | 18143.5176 | 142649.3544 | 67114.6183 | 51127.9015 | 892.6946 | 57589.1789 | 57589.1789 | 6570.8448 | 1286.0579 | 3871.8581 | 8185.6783 | 5054.5879 | 1495.4117 | 133243.9163 | 115366.7066 | 137774.8852 | 155351.8355 | 9238.9936 | 180607.9567 | 146073.6886 | 136056.8645 | 3413.0745 | 61428.457 | 146809.2022 | 234087.025 | 468923.3769
196 | CCV 10 ppb MAJORS |  | 06/09/2023 09:11:56 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.550 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 13501.7114 | 5462.3787 | 550753.2786 | 353.3377 | 24369.436 | 2775.6031 | 6167.9984 | 2572.8984 | 9870.0753 | 13623.1595 | 75096.203 | 51297.9674 | 123193.6693 | 13089.9952 | 28376.8245 | 7798.1285 | 18502.642 | 144380.615 | 67460.9234 | 52508.8606 | 906.6955 | 59316.899 | 59316.899 | 6540.1636 | 1236.7203 | 3923.206 | 8377.7896 | 5248.697 | 1499.4124 | 134593.9655 | 117906.643 | 140225.5466 | 157235.3302 | 9821.3751 | 180583.1551 | 148733.0281 | 138385.754 | 3339.0574 | 61860.3263 | 146871.8368 | 234568.1492 | 473519.0965
197 | 8_30_21_W_9 |  | 06/09/2023 09:17:44 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_W_9.551 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2065.4827 | 740.6859 | 950203.4849 | 608.6797 | 12200.5412 | 1109.3765 | 233.3353 | 1712.1028 | 172.6677 | 51.3334 | 255.3357 | 73538.92 | 60039.909 | 19257.6381 | 35.3334 | 40.0001 | 1921.4628 | 14400.5885 | 5257.6344 | 2478.7831 | 80.6669 | 5484.3866 | 5484.3866 | 992.7013 | 2.6667 | 6.6667 | 24.6667 | 75.8509 | 3.3333 | 14155.6767 | 12258.5914 | 142005.8827 | 162761.3844 | 10196.3048 | 188945.3119 | 157727.3767 | 146899.4537 | 3488.4259 | 21.3334 | 74.6669 | 2343.4102 | 4666.2993
198 | 8_30_21_W_10 |  | 06/09/2023 09:23:32 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_W_10.552 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1954.8006 | 954.6986 | 860899.2233 | 509.3424 | 14180.3684 | 1237.3871 | 250.0022 | 1924.7964 | 173.3344 | 38.6667 | 228.6685 | 55595.4456 | 38027.2287 | 14688.2151 | 28.6667 | 39.3334 | 1426.0714 | 10652.6384 | 4883.5021 | 2760.2161 | 80.0002 | 5493.0562 | 5493.0562 | 938.0308 | 4 | 12 | 21.3334 | 92.48 | 3.3333 | 8844.7372 | 7872.8365 | 143255.8303 | 159111.9429 | 9711.3019 | 185788.0793 | 150884.6424 | 143142.0652 | 3350.3929 | 28 | 76.0002 | 2066.0585 | 3968.8805
199 | 8_30_21_W_11 |  | 06/09/2023 09:29:20 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_W_11.553 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1980.1373 | 904.0286 | 994988.0665 | 596.6791 | 14186.3739 | 1248.7213 | 240.002 | 1744.1066 | 198.6681 | 40.0001 | 201.3348 | 62968.8684 | 43075.5311 | 16190.5067 | 23.3334 | 48.0001 | 1419.4038 | 11241.7553 | 5259.6358 | 2883.2358 | 69.3335 | 5517.7322 | 5517.7322 | 940.0309 | 1.3333 | 7.3333 | 22 | 85.9431 | 3.3333 | 8277.7321 | 7183.1392 | 141494.6974 | 157955.833 | 10029.5196 | 182388.4181 | 145814.266 | 138908.8901 | 3433.0792 | 27.3334 | 100.0004 | 6183.1868 | 12534.8168
200 | 8_30_21_F_1 |  | 06/09/2023 09:35:08 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_F_1.554 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9186.9536 | 3030.3214 | 300434.2762 | 192.0013 | 22378.851 | 4839.4863 | 1138.0454 | 8589.9156 | 309.3367 | 160.6676 | 923.3632 | 18825.5837 | 128623.1664 | 5837.1946 | 82.6669 | 321.337 | 1528.0818 | 12139.823 | 35028.8934 | 22433.3257 | 364.0046 | 26275.4752 | 26275.4752 | 1095.3754 | 7.3333 | 16.6667 | 40.6667 | 227.7086 | 4 | 186852.2628 | 165923.4956 | 135488.9286 | 150937.8691 | 9196.9612 | 174879.9116 | 143267.5115 | 138261.8851 | 3436.4135 | 44.6667 | 108.0004 | 2322.7405 | 4666.2945
201 | 8_30_21_F_1 |  | 06/09/2023 09:41:25 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_F_1.555 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9323.7086 | 2912.297 | 309207.3203 | 189.3346 | 21018.7847 | 4967.5307 | 1126.0444 | 8414.4799 | 322.0036 | 174.6677 | 958.6989 | 19199.3815 | 130999.9136 | 5807.1802 | 84.6669 | 313.3368 | 1525.4148 | 12287.9494 | 34993.4745 | 22652.7603 | 386.0052 | 26825.8323 | 26825.8323 | 1043.3715 | 8.6667 | 12 | 36.6667 | 193.6752 | 13.3333 | 188038.8627 | 168883.7699 | 134448.4145 | 154494.249 | 9762.6797 | 173234.7122 | 142973.9136 | 138336.2719 | 3337.7233 | 33.3334 | 116.0005 | 2370.7423 | 4644.9649
202 | 8_30_21_F_1_DUP |  | 06/09/2023 09:47:13 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_F_1_DUP.556 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9628.5793 | 3026.3207 | 318258.482 | 194.668 | 20899.9441 | 4948.8571 | 1088.0415 | 8470.5105 | 329.3371 | 165.3343 | 814.6899 | 19012.6756 | 128727.3799 | 5905.2213 | 81.3336 | 330.0038 | 1557.4182 | 12418.0623 | 35054.9573 | 22424.0649 | 367.3381 | 26565.3458 | 26565.3458 | 1069.3735 | 6.6667 | 21.3334 | 35.3334 | 231.8492 | 7.3333 | 187352.6074 | 168414.9545 | 125248.1746 | 154788.079 | 9396.4262 | 174900.8757 | 145856.2936 | 136968.2361 | 3309.0499 | 44.0001 | 106.6671 | 2350.7405 | 4808.9848
203 | 8_30_21_F_1_DUP |  | 06/09/2023 09:53:30 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_F_1_DUP.557 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9559.1991 | 3000.9819 | 311664.9992 | 208.6682 | 20467.988 | 4893.5051 | 1122.7108 | 8313.7535 | 325.337 | 156.0009 | 836.6913 | 18744.1823 | 130099.1242 | 5694.4682 | 76.0002 | 317.3369 | 1504.7459 | 12338.6598 | 35036.2449 | 21804.1503 | 368.6714 | 26636.1425 | 26636.1425 | 1036.7043 | 4.6667 | 16 | 38.6667 | 219.2038 | 9.3333 | 185483.8838 | 167260.1736 | 134013.1678 | 153482.832 | 9447.1242 | 173912.8774 | 144006.8963 | 136154.5398 | 3305.049 | 46.6667 | 96.667 | 2321.4059 | 4817.6521
204 | 8_30_21_F_6 |  | 06/09/2023 09:59:18 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_F_6.558 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10717.3522 | 3569.113 | 200945.4741 | 115.3338 | 7013.0545 | 4208.6199 | 1134.045 | 8330.4285 | 636.6809 | 66.0002 | 417.3394 | 28997.3652 | 126687.2763 | 7992.2357 | 70.0002 | 184.0012 | 4732.784 | 37751.1482 | 21175.0152 | 22756.8446 | 405.3391 | 27724.8781 | 27724.8781 | 1038.0377 | 8.6667 | 37.3334 | 66.0002 | 203.6862 | 18 | 105083.0777 | 94047.2299 | 132673.2622 | 154082.5815 | 9600.5602 | 172166.0171 | 121183.5689 | 136890.1629 | 3291.0464 | 116.0005 | 233.3353 | 3010.121 | 6216.5329
205 | 8_30_21_F_6 |  | 06/09/2023 10:05:38 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_F_6.559 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10500.5259 | 3631.795 | 200089.5697 | 116.0005 | 6892.3291 | 4291.978 | 1090.0416 | 8643.2808 | 622.0136 | 82.0002 | 418.6728 | 28735.6622 | 129788.9624 | 8077.6176 | 66.0002 | 178.0011 | 4943.522 | 37640.1917 | 21132.9535 | 22832.2256 | 382.6718 | 27833.0879 | 27833.0879 | 1093.3752 | 11.3333 | 39.3334 | 72.6669 | 213.1225 | 14 | 104884.3026 | 94898.8258 | 134181.2189 | 153182.2632 | 9563.2318 | 169663.5949 | 143883.7791 | 139349.6789 | 3313.7184 | 117.3338 | 240.6687 | 2962.1208 | 6069.162
206 | 8_30_21_F_8 |  | 06/09/2023 10:11:26 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_F_8.560 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 11445.253 | 3541.1055 | 136923.1204 | 86.0003 | 10176.2903 | 4031.2359 | 1193.3835 | 8423.1497 | 800.0225 | 90.0003 | 410.6726 | 30766.3288 | 141736.4548 | 8816.7206 | 64.6668 | 180.0012 | 5445.0375 | 42327.2864 | 11033.5953 | 24766.1619 | 407.3392 | 30080.313 | 30080.313 | 1055.3725 | 6.6667 | 33.3334 | 65.3335 | 231.1015 | 11.3333 | 93763.3717 | 85737.2773 | 132809.8683 | 141995.4328 | 9601.2285 | 170832.7465 | 141558.0105 | 136273.5505 | 3096.3362 | 99.3337 | 199.3347 | 2470.7483 | 4970.3385
207 | 8_30_21_F_8 |  | 06/09/2023 10:17:43 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_F_8.561 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 11118.9931 | 3560.4442 | 137073.9355 | 64.0001 | 9754.6631 | 3973.8861 | 1156.0468 | 8502.5297 | 770.0208 | 85.3336 | 418.6728 | 30890.5809 | 138029.6337 | 8761.3537 | 67.3335 | 160.6676 | 5433.033 | 41615.207 | 11083.6322 | 24926.8702 | 436.6734 | 29077.5633 | 29077.5633 | 1083.3744 | 12.6667 | 36 | 74.0002 | 245.7578 | 12 | 94092.8766 | 83296.8178 | 132822.6086 | 155467.2235 | 9254.9971 | 172077.5975 | 144428.5365 | 134589.083 | 3303.7156 | 88.0003 | 212.6683 | 2558.09 | 5035.0129
208 | LBLANK |  | 06/09/2023 10:23:31 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.562 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.3333 | 0 | 3475.7571 | 3.3333 | 462.6742 | 562.6778 | 10 | 0 | 0.66667 | 6.6667 | 35.3334 | 490.7004 | 17.3333 | 1158.7138 | 0 | 13.3333 | 6.6667 | 86.0003 | 250.0022 | -2365.9369 | 0 | 31.3334 | 31.3334 | 953.3651 | 0 | 2 | 6 | 35.8719 | 1.3333 | 53.3334 | 45.3334 | 140983.1136 | 163242.9673 | 9959.4706 | 177274.6418 | 150627.5382 | 141895.9901 | 3597.1195 | 0.66667 | 8.6667 | 12.6667 | 36.6667
209 | CCV 1ppb TRACE |  | 06/09/2023 10:29:20 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.563 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1426.0714 | 566.0113 | 112185.502 | 66.6668 | 2935.6351 | 819.3568 | 624.0137 | 296.6698 | 1050.0386 | 1422.0708 | 7888.1778 | 6426.1498 | 12642.2585 | 2558.8959 | 3000.3154 | 848.6919 | 2062.1489 | 15689.2777 | 7488.6295 | 3156.3636 | 94.0003 | 6083.9621 | 6083.9621 | 1478.7434 | 142.0007 | 405.3391 | 890.0277 | 569.4253 | 145.3341 | 13871.3981 | 12213.2187 | 141737.7045 | 161751.9551 | 9828.7139 | 185708.0486 | 152713.9277 | 139299.2505 | 3493.0937 | 6548.8345 | 15463.0315 | 24765.0555 | 51096.5209
210 | CCV 1ppb TRACE |  | 06/09/2023 10:35:37 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.564 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1439.4059 | 612.6798 | 113268.766 | 72.6669 | 2834.2814 | 815.3567 | 618.0134 | 281.3361 | 1074.7071 | 1468.7422 | 7959.5503 | 6442.7371 | 12579.5369 | 2615.5729 | 3017.6522 | 850.692 | 2072.1503 | 15763.3589 | 7381.9071 | 3000.4208 | 93.3336 | 6101.9697 | 6101.9697 | 1538.7495 | 129.3339 | 428.0064 | 886.0275 | 593.6853 | 144.0007 | 13860.0538 | 12395.3761 | 143472.211 | 163173.4062 | 9865.413 | 185901.1078 | 153517.8255 | 142429.9098 | 3587.1172 | 6554.8369 | 15384.2812 | 24986.4574 | 51677.1735
211 | LFB |  | 06/09/2023 10:41:25 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.565 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 12430.0721 | 5070.233 | 534067.7989 | 356.0044 | 29395.5473 | 2588.9013 | 5687.1323 | 2454.2109 | 9190.2893 | 12745.6837 | 69932.0973 | 46912.4011 | 113265.2539 | 12009.0455 | 26376.9962 | 7308.537 | 17269.7678 | 131682.8643 | 64068.0312 | 49285.0669 | 830.6908 | 55940.651 | 55940.651 | 6015.9333 | 1210.0512 | 3835.8484 | 8009.5793 | 5052.2595 | 1472.7426 | 123326.0784 | 109919.9635 | 135856.9594 | 149746.7804 | 9551.8589 | 172889.2251 | 141707.3804 | 133988.7977 | 3361.0621 | 57489.4544 | 137542.3128 | 216923.5782 | 437154.1594
212 | LFB |  | 06/09/2023 10:47:41 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.566 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 12781.7156 | 5018.8815 | 547022.4537 | 350.0043 | 29898.5933 | 2628.9086 | 5831.8573 | 2426.8729 | 9302.3619 | 12923.845 | 70591.9913 | 47520.2888 | 115943.3553 | 12149.1655 | 27195.864 | 7428.598 | 17620.1935 | 134340.1451 | 64922.5315 | 50844.8338 | 884.6941 | 56642.7448 | 56642.7448 | 6236.0275 | 1233.3866 | 3873.8586 | 8306.4146 | 5107.5755 | 1490.0778 | 126244.7455 | 111592.1597 | 135624.1461 | 152638.5361 | 9719.3057 | 175064.9123 | 143577.9427 | 131328.9584 | 3269.041 | 58579.2021 | 139437.8832 | 217473.0336 | 444915.6242
213 | CCV 10 ppb MAJORS |  | 06/09/2023 10:53:35 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.567 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 13669.2055 | 5503.7267 | 558367.6665 | 345.3375 | 24753.4351 | 2801.6081 | 6216.6858 | 2634.9097 | 10028.1859 | 13782.6457 | 75442.0178 | 50727.372 | 122258.3526 | 13280.1716 | 28764.931 | 7774.1173 | 18395.839 | 141660.2539 | 66632.3709 | 52629.8529 | 907.3622 | 58408.4972 | 58408.4972 | 6644.8792 | 1276.057 | 3945.2116 | 8405.8058 | 5296.6075 | 1464.7418 | 133681.2661 | 116515.2258 | 140834.169 | 152930.3529 | 9772.6763 | 180007.8255 | 147427.5373 | 137080.5732 | 3391.0691 | 60743.5423 | 144359.0595 | 231438.3505 | 465101.4152
214 | LBLANK |  | 06/09/2023 11:28:06 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.568 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.66667 | 0 | 18437.9234 | 9.3333 | 836.6912 | 628.0139 | 7.3333 | 0 | 0.66667 | 14 | 40.0001 | 608.3556 | 11.3333 | 1318.7276 | 0.66667 | 11.3333 | 11.3333 | 92.667 | 1100.0428 | -2796.9924 | 0 | 9.3333 | 9.3333 | 1104.0427 | 0 | 2 | 0 | 54.1644 | 0 | 20 | 16 | 176517.9465 | 197865.5394 | 12640.9241 | 234349.9989 | 192913.21 | 172238.9985 | 4402.012 | 0.66667 | 17.3333 | 18 | 39.3334
215 | Cal blank |  | 06/09/2023 11:33:55 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\Cal blank.569 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0 | 0 | 15222.1422 | 6.6667 | 624.6803 | 742.0193 | 8.6667 | 0 | 0.66667 | 10 | 22 | 607.0108 | 10.6667 | 1318.0609 | 3.3333 | 18.6667 | 11.3333 | 60.6668 | 798.0224 | -2769.9418 | 0.66667 | 8.6667 | 8.6667 | 1038.7046 | 0 | 0.66667 | 2 | 64.0671 | 1.3333 | 14.6667 | 6.6667 | 173145.1592 | 198833.7946 | 12686.2991 | 235708.7554 | 194173.4048 | 174881.419 | 4366.6677 | 2.6667 | 7.3333 | 30.6667 | 62.0001
216 | CAL BLANK |  | 06/09/2023 11:39:43 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CAL BLANK.570 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.66667 | 0.66667 | 12476.8066 | 8.6667 | 547.3438 | 706.0176 | 10 | 0 | 0 | 7.3333 | 21.3333 | 648.0744 | 15.3333 | 1244.7209 | 2 | 11.3333 | 6.6667 | 87.3336 | 2218.8526 | -3006.808 | 0 | 3.3333 | 3.3333 | 1070.7069 | 0 | 1.3333 | 2.6667 | 59.5411 | 0 | 8.6667 | 4 | 175923.9588 | 197782.4116 | 12444.7517 | 230195.4537 | 188175.9642 | 172199.7869 | 4290.6443 | 0 | 7.3333 | 20.6667 | 56.6667
217 | LOW 0.01 |  | 06/09/2023 11:45:32 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LOW 0.01.571 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 14 | 8 | 48037.325 | 29.3334 | 3570.4464 | 707.3509 | 18 | 2 | 9.3333 | 28.6667 | 124.6672 | 1432.2223 | 488.675 | 1448.7405 | 31.3334 | 18.6667 | 51.3334 | 374.6716 | 3159.6832 | -3016.16 | 0 | 65.3335 | 65.3335 | 1074.0404 | 3.3333 | 11.3333 | 12 | 42.8201 | 2 | 330.6705 | 330.6705 | 175443.415 | 197349.4216 | 12621.5735 | 232419.4697 | 188365.2744 | 173930.4519 | 4250.6324 | 63.3335 | 192.0013 | 368.0019 | 680.673
218 | LOW 0.05 |  | 06/09/2023 11:51:20 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LOW 0.05.572 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 80.0002 | 28.6667 | 45682.3478 | 30.6667 | 5643.1146 | 709.351 | 43.3334 | 18.6667 | 57.3334 | 82.6669 | 462.6742 | 1537.2696 | 672.6826 | 1469.4089 | 162.0009 | 59.3335 | 132.6673 | 1032.0373 | 2461.5467 | -2587.502 | 4 | 318.6703 | 318.6703 | 1125.3777 | 5.3333 | 22.6667 | 45.3334 | 60.8041 | 16 | 754.6867 | 626.6804 | 176694.8327 | 195555.4608 | 12197.2049 | 230936.3626 | 188325.8823 | 173952.1167 | 4305.9824 | 339.3374 | 862.0261 | 1408.0268 | 2982.1235
219 | LOW 0.1 |  | 06/09/2023 11:57:09 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LOW 0.1.573 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 146.6674 | 54.0001 | 25102.7055 | 10.6667 | 3324.3868 | 701.3506 | 92.0003 | 32 | 88.0003 | 176.6678 | 859.3592 | 2350.9294 | 1283.391 | 1611.4243 | 320.0036 | 93.3336 | 229.3352 | 1784.7783 | 2160.1634 | -2259.9576 | 6 | 576.6783 | 576.6783 | 1161.3806 | 20.6667 | 42.6667 | 88.0003 | 94.0352 | 18.6667 | 1426.738 | 1284.0577 | 169485.0049 | 192628.1385 | 12231.9015 | 226869.9285 | 185109.8056 | 166298.4401 | 4314.6516 | 695.3503 | 1634.0935 | 2751.4376 | 5419.7366
220 | LOW 0.5 |  | 06/09/2023 12:02:57 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LOW 0.5.574 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 716.6847 | 290.003 | 80381.0378 | 52.0001 | 6644.8786 | 782.6881 | 358.0045 | 135.334 | 507.3424 | 728.0186 | 3883.1949 | 5957.3154 | 6425.4451 | 2576.2323 | 1546.7507 | 437.34 | 1090.7083 | 8543.2224 | 5291.6466 | -230.1296 | 50.6668 | 2958.9734 | 2958.9734 | 1337.396 | 70.0002 | 224.0018 | 470.6744 | 279.5645 | 85.3336 | 7143.1208 | 6226.6911 | 171573.4124 | 194178.1122 | 12417.3945 | 226097.4724 | 158207.1445 | 167798.4167 | 4323.3211 | 3379.7335 | 7940.8733 | 13277.7465 | 27624.4502
221 | STD 1 |  | 06/09/2023 12:08:45 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 1.575 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1518.0807 | 594.0124 | 117266.6786 | 76.0002 | 5569.086 | 903.3619 | 648.0147 | 274.6693 | 1053.3722 | 1498.7453 | 7885.5099 | 8808.0566 | 12699.6424 | 3168.3517 | 3129.6761 | 808.0229 | 2060.8153 | 15970.2568 | 8493.8582 | 2707.9366 | 96.667 | 5981.919 | 5981.919 | 1577.4204 | 127.3339 | 423.3396 | 892.0279 | 557.5862 | 168.6677 | 14218.4069 | 12596.8849 | 174104.5338 | 189099.3622 | 12252.5882 | 223026.1748 | 182543.7305 | 164892.5193 | 4141.2671 | 6811.6245 | 15639.8899 | 26091.3245 | 54420.4449
222 | STD 5 |  | 06/09/2023 12:14:33 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 5.576 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 6794.9492 | 2750.2648 | 319962.3271 | 230.6685 | 17101.5673 | 1749.4406 | 3151.6811 | 1384.7338 | 4974.8661 | 6891.662 | 37671.6049 | 26530.9118 | 61210.8768 | 7537.9892 | 14422.6102 | 3893.1971 | 9629.2442 | 72246.2313 | 34714.1289 | 24770.1551 | 438.6735 | 29020.7815 | 29020.7815 | 3897.8652 | 616.0133 | 1990.1389 | 4315.9863 | 2734.528 | 760.0202 | 66994.7282 | 59038.4111 | 168752.8865 | 182406.8136 | 11901.6228 | 213341.9135 | 173902.2226 | 162384.5163 | 3950.5464 | 31811.3857 | 75348.2025 | 123189.6485 | 249220.4478
223 | STD 10 |  | 06/09/2023 12:20:21 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 10.577 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 13605.1424 | 5426.3638 | 567125.8798 | 385.3385 | 27215.2389 | 2880.2909 | 6413.4396 | 2772.9358 | 9918.1084 | 13696.5631 | 75149.1619 | 51374.8533 | 121557.6803 | 12959.8764 | 28939.285 | 7974.8924 | 18487.9557 | 142788.1255 | 67243.9066 | 51586.1815 | 868.6931 | 57417.1663 | 57417.1663 | 6439.4519 | 1283.391 | 3847.8515 | 8269.7268 | 5181.1191 | 1494.745 | 135504.9914 | 118323.341 | 163515.8753 | 178238.4195 | 11359.8517 | 208310.5196 | 169219.0075 | 158155.3068 | 3888.5293 | 61723.7261 | 147459.8193 | 235704.4309 | 481560.9128
224 | STD 25 |  | 06/09/2023 12:26:09 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 25.578 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 33228.606 | 13201.431 | 1398943.1789 | 874.0268 | 63793.4861 | 6402.101 | 15651.9055 | 6678.2278 | 24886.9979 | 33907.5304 | 183765.1454 | 124493.0342 | 294481.0714 | 30405.6573 | 71497.1473 | 19036.0089 | 45023.5216 | 341540.985 | 163989.2008 | 138801.798 | 2302.1856 | 144546.274 | 144546.274 | 15120.0058 | 3091.0016 | 9938.1227 | 20869.8999 | 13449.7796 | 3857.8542 | 331597.3134 | 291905.5055 | 145258.2501 | 170786.9197 | 11163.6982 | 198765.9467 | 162498.379 | 151561.2151 | 3779.1666 | 150651.5374 | 356186.6311 | 568104.9727 | 1110472.1244
225 | STD 50 |  | 06/09/2023 12:31:57 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 50.579 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 65120.7607 | 25129.4193 | 3380265.0863 | 1678.0986 | 121216.8007 | 11666.7626 | 30572.692 | 12995.2416 | 48874.129 | 65708.7714 | 359474.0008 | 218210.8223 | 579893.0985 | 57983.4613 | 137441.3262 | 36753.2209 | 86069.1675 | 639948.9469 | 307136.3601 | 262490.5655 | 4351.3294 | 280248.3825 | 280248.3825 | 27097.0228 | 6203.3467 | 19413.1823 | 40982.7115 | 25402.4393 | 7359.2286 | 638719.1486 | 573932.7021 | 153680.8801 | 165955.1023 | 10521.207 | 190482.8915 | 143988.2514 | 144207.632 | 3458.4188 | 279705.4323 | 659310.2653 | 1018031.163 | 2020095.8795
226 | STD 100 |  | 06/09/2023 12:37:44 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 100.580 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 128398.721 | 48812.6076 | 6618946.1448 | 3199.0249 | 224181.9888 | 22497.0355 | 61449.2823 | 25505.4206 | 97360.104 | 131074.812 | 687417.9313 | 404951.8916 | 1110672.4748 | 109345.0507 | 267896.6308 | 69792.1322 | 162672.5262 | 1156015.7497 | 548596.2301 | 493956.1394 | 8316.4202 | 518837.044 | 518837.044 | 50216.1746 | 11639.4078 | 36590.1331 | 77168.5584 | 48220.4396 | 13814.0094 | 1204271.2206 | 1085349.0932 | 140896.1143 | 151051.0732 | 9675.2863 | 168924.2849 | 142293.1523 | 135144.2426 | 3227.0312 | 498866.4182 | 1180631.3299 | 1780843.7527 | 3553622.6489
227 | Calibration Curves |  | 06/09/2023 12:37:44
228 | LBLANK |  | 06/09/2023 12:43:32 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.581 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4.6667 | 0.66667 | 9923.4641 | 4.6667 | 1662.7635 | 780.6881 | 12.6667 | 0 | 5.3333 | 13.3333 | 40.0001 | 646.0915 | 28 | 1449.4071 | 11.3333 | 17.3333 | 20.6667 | 108.6671 | 495.342 | -3169.6714 | 1.3333 | 42.0001 | 42.0001 | 1188.7164 | 0 | 4 | 6 | 55.5194 | 2 | 1478.0767 | 126.0006 | 164451.7702 | 198035.9042 | 12542.1702 | 236784.2628 | 192278.7257 | 181885.2717 | 4476.0347 | 4.6667 | 24 | 29.3333 | 58
229 | LBLANK |  | 06/09/2023 13:13:17 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.583 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 12 | 0 | 8876.4116 | 9.3333 | 1207.3844 | 556.0114 | 13.3333 | 0 | 2.6667 | 11.3333 | 34.6667 | 644.4513 | 17.3333 | 1213.3853 | 5.3333 | 13.3333 | 16 | 100.667 | 1614.1027 | -2839 | 1.3333 | 13.3333 | 13.3333 | 1083.3745 | 1.3333 | 4 | 0.66667 | 46.1535 | 0 | 694.6836 | 66.6668 | 167029.6465 | 189368.1968 | 12176.526 | 218481.4505 | 183016.9291 | 163111.2583 | 4249.966 | 1.3333 | 11.3333 | 20.6667 | 43.3334
230 | LBLANK |  | 06/09/2023 13:19:06 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.584 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.3333 | 0 | 4188.6143 | 2.6667 | 689.35 | 742.0193 | 13.3333 | 0 | 0.66667 | 4 | 26.6667 | 576.8435 | 14.6667 | 1250.0548 | 4 | 13.3333 | 11.3333 | 87.3336 | 398.6723 | -2892.0726 | 0 | 10 | 10 | 1072.0403 | 0 | 4 | 3.3333 | 29.5895 | 0 | 642.6812 | 20.6667 | 175238.3405 | 189243.8936 | 12108.4649 | 223811.9604 | 186084.1879 | 145761.3934 | 4190.6146 | 0.66667 | 10 | 22.6667 | 35.3334
231 | ICV PE |  | 06/09/2023 13:24:53 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\ICV PE.585 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 13932.7933 | 5669.1256 | 582739.9929 | 368.6715 | 49158.4418 | 2933.6346 | 6379.4248 | 2752.9319 | 10106.9077 | 13940.7988 | 75065.4087 | 52612.4537 | 123118.2874 | 13603.1431 | 29702.198 | 8102.2973 | 18884.4741 | 141289.2606 | 65408.0787 | 51460.7969 | 810.023 | 58421.8834 | 58421.8834 | 5012.8794 | 1275.3903 | 4034.5697 | 8407.8072 | 5321.7724 | 1535.4159 | 128922.4793 | 114393.55 | 169053.537 | 180471.6125 | 10930.1849 | 209168.8781 | 171884.5702 | 157104.4866 | 3958.5486 | 60573.4822 | 144039.1881 | 230649.8154 | 465933.6908
232 | ICV PE |  | 06/09/2023 13:31:10 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\ICV PE.586 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 13984.8421 | 5879.8766 | 592613.4422 | 372.6715 | 49439.4014 | 2866.9543 | 6419.4422 | 2852.2848 | 10112.2449 | 14112.9678 | 77609.071 | 54087.9338 | 122728.9339 | 13785.9822 | 29597.3018 | 8286.4029 | 19389.8165 | 146637.6889 | 68694.1122 | 51679.8019 | 817.3568 | 58379.0507 | 58379.0507 | 5065.5646 | 1262.7225 | 3922.5385 | 8435.1567 | 5268.4394 | 1520.0809 | 132645.3112 | 113892.8786 | 167031.7105 | 179179.4538 | 11273.1152 | 190352.0582 | 178124.3157 | 158277.3523 | 3931.8745 | 62621.7368 | 148900.0504 | 232438.6956 | 473768.8344
233 | CCV 1ppb TRACE |  | 06/09/2023 13:36:58 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.587 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1531.4154 | 617.3468 | 121206.7742 | 80.6669 | 3634.4626 | 952.6986 | 743.3527 | 294.003 | 1120.7106 | 1524.0813 | 8229.0362 | 6660.6544 | 13110.6812 | 2768.9355 | 3161.6833 | 869.3598 | 2200.1698 | 16560.2619 | 7794.7938 | 2555.3169 | 102.667 | 6235.3605 | 6235.3605 | 1729.4383 | 137.334 | 399.3389 | 908.6956 | 546.7302 | 180.6678 | 14755.6178 | 12705.648 | 174867.7828 | 191773.9823 | 12717.0281 | 221887.507 | 182124.9223 | 165839.0968 | 4267.9709 | 6733.5867 | 15902.1805 | 25657.0097 | 52929.6721
234 | CCV 1ppb TRACE |  | 06/09/2023 13:43:15 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.588 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1532.0822 | 620.6802 | 115755.1037 | 80.0002 | 3205.0263 | 954.6986 | 737.3524 | 299.3365 | 1138.7121 | 1548.084 | 8240.3763 | 6685.5258 | 13165.3984 | 2893.6265 | 3187.689 | 868.6932 | 2223.5065 | 16058.3551 | 7604.0242 | 3606.5367 | 104.6671 | 6051.2816 | 6051.2816 | 1673.4317 | 133.334 | 436.6734 | 930.697 | 627.2384 | 156.6675 | 14665.5246 | 12743.0144 | 172274.1229 | 187383.6332 | 12250.5853 | 216997.7681 | 179712.6665 | 164046.5937 | 4152.6035 | 6729.585 | 15772.7029 | 24699.6686 | 51534.4766
235 | LFB |  | 06/09/2023 13:49:03 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.589 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 13515.0614 | 5409.024 | 558679.5549 | 354.0044 | 30193.878 | 2948.9712 | 6188.0067 | 2640.9111 | 9669.273 | 13664.534 | 73039.581 | 49612.3433 | 117105.4268 | 12989.9039 | 28328.0602 | 7804.1314 | 18766.3203 | 140904.8233 | 66892.2835 | 51688.4996 | 942.6978 | 56796.0133 | 56796.0133 | 6593.5217 | 1297.3923 | 3895.1977 | 8351.107 | 5331.5342 | 1472.7427 | 131254.8915 | 113284.7646 | 165891.7449 | 176375.5944 | 11571.3517 | 207108.501 | 170242.4573 | 159151.6726 | 3954.5475 | 60173.8023 | 142983.3298 | 219913.206 | 444081.3422
236 | LLOQ 0.01 |  | 06/09/2023 13:54:51 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LLOQ 0.01.590 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 14 | 7.3333 | 38029.8928 | 28.6667 | 3935.2092 | 660.682 | 24.6667 | 3.3333 | 22 | 23.3334 | 104.6671 | 1453.386 | 476.008 | 1441.4061 | 30 | 20.6667 | 47.3334 | 366.6714 | 1600.757 | -2470.1727 | 0.66667 | 73.3335 | 73.3335 | 1098.0424 | 2 | 6.6667 | 6.6667 | 34.9039 | 2 | 620.6802 | 272.6693 | 172294.9636 | 188598.3873 | 12837.0989 | 224144.3613 | 182511.9433 | 165423.6253 | 4112.5926 | 76.0002 | 154.6675 | 320.0014 | 646.0059
237 | CCV 10 ppb MAJORS |  | 06/09/2023 14:00:40 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.591 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 14038.2285 | 5527.7371 | 547087.5867 | 324.0037 | 25147.4498 | 3010.984 | 6452.7914 | 2690.9202 | 10114.9137 | 14063.5878 | 74790.6199 | 50663.5975 | 120587.5289 | 13425.6395 | 29662.098 | 8106.9671 | 19078.0642 | 141085.9685 | 65764.3654 | 53024.0735 | 974.6999 | 57169.5047 | 57169.5047 | 6591.5209 | 1224.0527 | 3781.8342 | 8185.011 | 5288.5929 | 1430.7385 | 132978.7697 | 116759.8858 | 164650.1675 | 174070.1064 | 11066.2875 | 216230.2278 | 167364.7994 | 155332.7551 | 3756.494 | 60573.5107 | 144658.762 | 229257.516 | 457113.3232
238 | CCV 10 ppb MAJORS |  | 06/09/2023 14:06:57 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.592 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 13939.4681 | 5493.7241 | 556541.7628 | 345.3375 | 24704.683 | 2926.2998 | 6518.1543 | 2692.2537 | 10125.5876 | 13914.1066 | 76405.7804 | 52215.0325 | 122026.9665 | 13499.71 | 29956.0443 | 8120.9746 | 19126.1297 | 143362.4365 | 66110.64 | 53179.2192 | 940.0309 | 57318.0922 | 57318.0922 | 6821.6285 | 1224.0526 | 3960.5489 | 8190.348 | 5243.4281 | 1464.7418 | 135727.0558 | 118190.2356 | 164944.1004 | 174527.6692 | 11487.2838 | 204894.916 | 169757.3701 | 156259.9837 | 3838.5157 | 61534.2477 | 145167.1976 | 230550.5838 | 459838.1133
239 | LLOQ 0.1 |  | 06/09/2023 14:12:45 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LLOQ 0.1.593 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 160.6676 | 69.3335 | 23005.8456 | 18.6667 | 2060.1488 | 690.6834 | 86.0003 | 24 | 108.0004 | 165.3343 | 850.692 | 2432.4059 | 1394.7348 | 1828.1173 | 360.0045 | 96.0003 | 266.6692 | 1838.7851 | 4655.4745 | -1658.4366 | 13.3333 | 634.6808 | 634.6808 | 1151.3797 | 19.3333 | 39.3334 | 94.667 | 92.8353 | 18 | 1704.1016 | 1257.3887 | 172099.3916 | 189418.2249 | 12191.2001 | 220946.8479 | 183963.0479 | 167602.804 | 4159.939 | 652.0149 | 1552.0843 | 2546.0884 | 5169.0327
240 | 8_30_21_F_9 |  | 06/09/2023 14:19:27 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_F_9.594 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10276.3624 | 3677.1401 | 109174.5107 | 87.3336 | 11387.8707 | 3738.4898 | 1061.3728 | 8428.4876 | 711.3511 | 91.3336 | 454.6739 | 21645.2167 | 75375.0137 | 6836.9692 | 48.6668 | 136.0006 | 4976.8669 | 36947.7224 | 6707.5749 | 20344.2457 | 342.0041 | 23581.4488 | 23581.4488 | 1151.3799 | 6 | 28.6667 | 66.6668 | 254.2993 | 14 | 66363.8119 | 58836.249 | 157654.5672 | 170918.3974 | 11379.8642 | 198497.5764 | 165564.5667 | 153600.8079 | 3666.4706 | 66.6668 | 166.001 | 2182.7316 | 4452.2702
241 | 8_30_21_F_10 |  | 06/09/2023 14:25:14 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_F_10.595 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10239.0029 | 4521.3822 | 184638.4236 | 126.6672 | 17004.1144 | 4155.271 | 1266.0562 | 9565.2023 | 778.0212 | 106.6671 | 550.0106 | 21002.5336 | 55405.9032 | 6601.5264 | 46.6667 | 137.334 | 4027.2344 | 29506.4535 | 15710.6432 | 21426.0743 | 358.0045 | 25210.2256 | 25210.2256 | 1096.7088 | 12 | 33.3334 | 74.6669 | 283.6016 | 10.6667 | 42694.391 | 38780.5674 | 154888.495 | 170596.5753 | 10997.5684 | 199749.3069 | 159227.9494 | 154807.6712 | 3727.8197 | 59.3335 | 148.0008 | 2379.4098 | 4936.3325
242 | 8_30_21_F_11 |  | 06/09/2023 14:31:03 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_F_11.596 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10118.916 | 4848.8228 | 190528.9722 | 119.3338 | 11897.6201 | 4122.5948 | 1228.7196 | 9184.9525 | 855.359 | 108.6671 | 525.343 | 18092.3186 | 44029.7513 | 6067.2882 | 50.6668 | 146.0008 | 3972.5524 | 29105.6285 | 10218.3199 | 22253.1716 | 376.6716 | 25734.4942 | 25734.4942 | 1228.7195 | 15.3333 | 33.3334 | 93.3336 | 290.6304 | 16 | 40645.0809 | 36762.5771 | 156392.0727 | 172010.0595 | 10974.2145 | 196433.4078 | 174833.348 | 150443.4887 | 3832.5142 | 50.0001 | 124.0005 | 2483.4172 | 4928.3299
243 | 8_30_21_UF_1 |  | 06/09/2023 14:36:51 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_UF_1.597 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 11049.6054 | 3031.6553 | 153442.3959 | 99.3337 | 52899.7659 | 4938.8541 | 1187.3828 | 8824.0586 | 361.3379 | 87.3336 | 476.0079 | 2604.1417 | 121703.6126 | 2050.814 | 70.6668 | 290.6696 | 1166.7143 | 8968.1477 | 16685.7392 | 21645.771 | 374.0049 | 24854.2702 | 24854.2702 | 1240.7205 | 4.6667 | 20 | 30.6667 | 207.0775 | 5.3333 | 175159.9719 | 157334.3227 | 154904.7603 | 171594.4283 | 10861.4699 | 198745.7775 | 169308.8116 | 156555.9593 | 3670.472 | 43.3334 | 100.667 | 1856.7138 | 3629.5122
244 | 8_30_21_UF_6 |  | 06/09/2023 14:42:39 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_UF_6.598 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 12722.3293 | 3631.795 | 95046.7345 | 50.0001 | 39377.5363 | 4335.3244 | 1184.0491 | 9132.2518 | 612.6798 | 88.0003 | 447.3404 | 4083.6768 | 123039.0029 | 2427.54 | 66.0002 | 204.0015 | 3216.3621 | 23377.1185 | 10306.3861 | 22923.8437 | 367.3381 | 26645.4971 | 26645.4971 | 1250.0547 | 10.6667 | 22.6667 | 83.3336 | 252.9605 | 9.3333 | 102443.5768 | 92490.4697 | 157758.9775 | 169231.0863 | 10898.1556 | 204934.1458 | 162283.2728 | 147752.8932 | 3660.4744 | 84.6669 | 226.6685 | 1989.3859 | 4181.5735
245 | 8_30_21_UF_8 |  | 06/09/2023 14:48:27 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_UF_8.599 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 12861.1237 | 3754.4934 | 126691.3625 | 85.3336 | 44914.5059 | 4155.9387 | 1260.0557 | 8830.0631 | 790.6886 | 108.0004 | 409.3392 | 2372.3336 | 130662.2016 | 1974.8032 | 56.0001 | 158.6676 | 3431.7455 | 25190.8585 | 6652.2157 | 23765.6932 | 413.3393 | 27539.1854 | 27539.1854 | 1256.0553 | 16 | 38.6667 | 62.6668 | 271.172 | 10.6667 | 88239.0164 | 78995.1394 | 157932.9208 | 171141.7007 | 11002.236 | 196669.8639 | 165470.8352 | 153854.8571 | 3707.1476 | 84.0003 | 182.6679 | 1514.6981 | 2966.7877
246 | 8_30_21_UF_9 |  | 06/09/2023 14:54:15 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_UF_9.600 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 11062.2823 | 3367.7305 | 59955.6188 | 36.6667 | 43032.7202 | 3361.0621 | 1005.3688 | 7772.7818 | 657.3485 | 86.6669 | 400.0056 | 2153.5709 | 70783.5961 | 1937.4649 | 52.6668 | 149.3341 | 3044.9915 | 22625.2363 | 21223.403 | 18698.9121 | 368.6715 | 22343.462 | 22343.462 | 1120.044 | 8 | 27.3334 | 57.3334 | 243.2032 | 8.6667 | 60523.272 | 54406.7459 | 159191.3271 | 169927.184 | 10757.3901 | 195032.3293 | 162931.215 | 159564.9216 | 3755.8271 | 57.3334 | 126.6672 | 1501.3641 | 3086.1307
247 | 8_30_21_UF_9_LD |  | 06/09/2023 15:00:03 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_UF_9_LD.601 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10738.0343 | 3417.0753 | 104925.2686 | 76.0002 | 47241.324 | 3426.4112 | 1006.0354 | 7851.4921 | 661.3487 | 75.3335 | 391.3387 | 2504.7412 | 69546.22 | 1912.128 | 47.3334 | 132.0006 | 2894.96 | 22452.2993 | 10365.83 | 18369.8984 | 300.0032 | 22051.0059 | 22051.0059 | 1110.7099 | 9.3333 | 26.6667 | 75.3336 | 194.3509 | 8.6667 | 60340.5099 | 53524.0837 | 158190.0522 | 171767.1448 | 10917.5051 | 185790.9352 | 161741.1304 | 144922.1089 | 3685.142 | 58.0001 | 127.3339 | 1629.3689 | 3244.8124
248 | 8_30_21_UF_9_LD |  | 06/09/2023 15:06:19 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_UF_9_LD.602 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10724.691 | 3454.4178 | 105221.5896 | 72.6669 | 45325.8072 | 3398.4045 | 1023.3701 | 7744.0985 | 688.6833 | 89.3336 | 430.0065 | 2442.5309 | 69416.9273 | 1914.795 | 32 | 150.6675 | 2975.6436 | 22283.3659 | 7017.7249 | 18021.9604 | 336.6707 | 22197.9012 | 22197.9012 | 1167.3812 | 11.3333 | 30.6667 | 61.3335 | 179.0061 | 16.6667 | 59520.4043 | 53664.6094 | 157726.0398 | 169691.8891 | 10998.236 | 194799.3252 | 160909.8404 | 150391.517 | 3794.5039 | 51.3334 | 147.3341 | 1576.7005 | 3144.1353
249 | 8_30_21_UF_10 |  | 06/09/2023 15:12:07 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_UF_10.603 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 11328.49 | 4450.0264 | 157643.0551 | 105.3337 | 46632.6558 | 3889.1961 | 1208.7179 | 9332.3808 | 764.0205 | 101.3337 | 481.3415 | 6649.4107 | 49005.2515 | 2988.3126 | 58.0001 | 128.0006 | 2525.5567 | 19052.0294 | 9615.9032 | 20776.9892 | 332.6706 | 23827.1902 | 23827.1902 | 1173.3816 | 9.3333 | 25.3334 | 82.0002 | 275.6174 | 15.3333 | 40393.6988 | 36429.0553 | 159257.4515 | 168409.5991 | 10912.8338 | 193455.1561 | 162175.4354 | 150978.279 | 3672.4722 | 60.0001 | 136.6673 | 1982.7202 | 3890.205
250 | LBLANK |  | 06/09/2023 15:17:56 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.604 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.66667 | 0.66667 | 5022.8838 | 2.6667 | 1950.1331 | 684.6831 | 8 | 1.3333 | 0 | 6 | 32.6667 | 592.1586 | 14.6667 | 1226.0527 | 1.3333 | 12.6667 | 8.6667 | 76.6669 | 315.3368 | -2103.2914 | 2.6667 | 100.667 | 100.667 | 1074.0404 | 1.3333 | 0.66667 | 3.3333 | 43.7331 | 0.66667 | 33.3334 | 24 | 159256.6936 | 182698.1831 | 12436.0856 | 213509.0207 | 176272.9041 | 158008.4244 | 4083.9173 | 0.66667 | 6 | 21.3333 | 44
251 | CCV 1ppb TRACE |  | 06/09/2023 15:23:44 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.605 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1422.7376 | 607.3463 | 115728.3841 | 74.0002 | 3397.7375 | 904.6955 | 676.016 | 291.3364 | 1014.7028 | 1500.7458 | 7761.4413 | 6483.2801 | 12424.0684 | 2577.566 | 2979.6443 | 868.6931 | 2070.1501 | 15310.8681 | 7126.4441 | 3361.1645 | 104.6671 | 5954.5743 | 5954.5743 | 1660.0965 | 128.0006 | 377.3383 | 840.6914 | 637.6807 | 152.0008 | 13635.1708 | 12224.5616 | 167872.7012 | 181138.514 | 11538.6623 | 207255.5188 | 160040.6706 | 157844.5907 | 4085.9176 | 6314.063 | 15061.2698 | 25410.1589 | 51850.7959
252 | CCV 1ppb TRACE |  | 06/09/2023 15:30:02 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.606 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1448.0735 | 629.3472 | 113707.4323 | 57.3334 | 3087.0003 | 883.3607 | 681.3496 | 294.6697 | 1045.3717 | 1416.737 | 7640.7104 | 6562.9124 | 12297.959 | 2561.563 | 3004.9827 | 835.3578 | 2038.1456 | 15378.2737 | 7171.1334 | 3230.496 | 94.667 | 5836.5255 | 5836.5255 | 1583.4212 | 144.0007 | 417.3395 | 873.36 | 533.9032 | 128.0006 | 13537.746 | 12131.149 | 165207.267 | 176467.9117 | 11519.3092 | 202915.2345 | 169456.4926 | 155932.5772 | 3795.8378 | 6402.7679 | 15301.5264 | 25135.2539 | 50127.4616
253 | CCV 10 ppb MAJORS |  | 06/09/2023 15:35:51 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.607 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 13442.9893 | 5385.6818 | 546753.2019 | 367.3381 | 23126.0407 | 2842.2828 | 6202.0132 | 2639.5773 | 9836.7194 | 13288.1793 | 72270.3466 | 49079.7554 | 119200.7081 | 12543.5052 | 27619.3405 | 7580.6775 | 18066.7552 | 136120.8169 | 63246.3638 | 51663.758 | 891.3612 | 56705.667 | 56705.667 | 6497.4789 | 1208.0511 | 3810.5083 | 8202.3541 | 5163.7401 | 1492.7447 | 129489.5874 | 116560.9764 | 160752.0881 | 151558.9081 | 10878.8077 | 195473.6895 | 163351.9898 | 148964.2488 | 3369.7335 | 59564.5867 | 141914.099 | 232932.6287 | 464008.7237
254 | CCV 10 ppb MAJORS |  | 06/09/2023 15:42:08 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.608 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 13625.8285 | 5433.7 | 546015.0516 | 370.6715 | 22028.3249 | 2952.972 | 6229.3581 | 2752.932 | 9876.7469 | 13683.2169 | 73960.299 | 50929.3364 | 120467.1819 | 13114.6854 | 28463.6624 | 7941.5406 | 18534.0158 | 139214.3896 | 64452.4002 | 51130.7527 | 904.0286 | 56329.5058 | 56329.5058 | 6615.5327 | 1232.0532 | 3832.5143 | 8205.0223 | 5075.3508 | 1434.0721 | 132428.9951 | 116518.6303 | 166570.5814 | 171940.575 | 11047.6038 | 203038.0168 | 165543.1178 | 195893.0789 | 3723.1518 | 60774.353 | 144254.6775 | 233793.1267 | 467732.5562
255 | 8_30_21_UF_10_LSP |  | 06/09/2023 15:47:56 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_UF_10_LSP.609 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 22979.8032 | 9118.2424 | 697650.21 | 426.6731 | 59051.8111 | 5783.1705 | 7271.8504 | 10866.7983 | 10147.6037 | 12999.9125 | 70100.9397 | 52561.1758 | 155310.4169 | 14364.5517 | 27023.5355 | 7431.9329 | 19214.9143 | 143358.4148 | 67844.7454 | 70900.1607 | 1200.0504 | 73999.847 | 73999.847 | 6535.4951 | 1153.3799 | 3707.8145 | 7821.4744 | 5251.106 | 1422.0709 | 159680.2006 | 141854.1895 | 159485.9182 | 174468.6768 | 10783.4027 | 190077.5517 | 162659.7979 | 147925.2571 | 3797.1717 | 56107.3154 | 131925.0597 | 220124.1676 | 435966.6772
256 | 8_30_21_UF_10_DUP_SP |  | 06/09/2023 15:53:45 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_UF_10_DUP_SP.610 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 18456.5822 | 7222.4923 | 343856.8079 | 214.6683 | 61391.6597 | 5174.9378 | 4616.0791 | 11079.6283 | 5924.5619 | 7157.7935 | 38626.8185 | 27415.787 | 110742.2515 | 8015.5815 | 14769.6313 | 4180.6119 | 11812.2148 | 87450.1916 | 40978.7028 | 48286.6638 | 852.0254 | 52380.5235 | 52380.5235 | 4062.5776 | 635.3475 | 1996.1395 | 4224.6246 | 2827.3495 | 786.6884 | 107227.0431 | 94966.616 | 154990.2497 | 164720.891 | 10662.6449 | 183944.5001 | 155834.0925 | 147640.3538 | 3605.7884 | 30080.976 | 72222.7853 | 118238.0305 | 237880.9813
257 | 8_30_21_UF_11 |  | 06/09/2023 15:59:33 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_UF_11.611 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10768.724 | 4455.3618 | 79282.2143 | 52.0001 | 43738.8642 | 3960.549 | 1231.3864 | 8799.3759 | 782.0216 | 100.0004 | 428.6731 | 2390.1108 | 38014.5129 | 1990.1386 | 42.0001 | 154.0008 | 2608.2382 | 19050.7003 | 6326.0674 | 20968.8832 | 383.3385 | 24141.7182 | 24141.7182 | 1157.3802 | 11.3333 | 32 | 79.3336 | 191.4949 | 14 | 37795.2647 | 34175.5017 | 154873.2797 | 165714.27 | 10896.8211 | 190776.7468 | 159561.5185 | 147525.8433 | 3605.7885 | 52.0001 | 90.0003 | 1454.029 | 2987.4551
258 | 8_30_21_BB |  | 06/09/2023 16:05:21 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_BB.612 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4.6667 | 1.3333 | 231617.4562 | 139.334 | 2912.9637 | 680.6829 | 20.6667 | 2 | 4 | 46.6667 | 204.0015 | 6982.7345 | 152.6675 | 2910.9632 | 2.6667 | 14.6667 | 226.0018 | 1619.4252 | 1414.7368 | -2296.178 | 2 | 69.3335 | 69.3335 | 1033.3707 | 0 | 6.6667 | 3.3333 | 39.0717 | 0 | 54.6668 | 35.3334 | 164925.9305 | 180325.0025 | 11478.6099 | 218546.5256 | 172215.8312 | 158356.8941 | 3891.8634 | 1.3333 | 13.3333 | 928.6779 | 1896.7165
259 | 8_30_21_FB |  | 06/09/2023 16:11:09 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_FB.613 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 41.3334 | 6.6667 | 131936.5645 | 101.3337 | 2089.4865 | 668.6823 | 14.6667 | 10.6667 | 2.6667 | 232.0019 | 1174.0484 | 3993.4977 | 1632.76 | 1970.1359 | 20 | 7486.6279 | 264.6691 | 1891.4585 | 38301.9482 | -2283.3552 | 0 | 60.0001 | 60.0001 | 1052.7055 | 1.3333 | 3.3333 | 5.3333 | 25.8517 | 0.66667 | 56.0001 | 37.3334 | 162856.5311 | 176856.062 | 11381.1985 | 207952.0766 | 168893.8634 | 151726.0966 | 3905.8674 | 8.6667 | 23.3334 | 1632.0364 | 3300.8158
260 | 8_30_21_LAB_A |  | 06/09/2023 16:16:56 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_LAB_A.614 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2 | 0 | 30766.4487 | 26 | 1161.3806 | 658.0152 | 11.3333 | 0 | 3.3333 | 109.3338 | 634.6808 | 3558.5359 | 105.3337 | 1888.7916 | 2 | 234.6686 | 214.6683 | 1576.0869 | 1398.7352 | -2232.3373 | 0 | 37.3334 | 37.3334 | 1063.3729 | 0 | 1.3333 | 0.66667 | 31.1635 | 0 | 318.0035 | 271.3359 | 159584.5806 | 174301.6427 | 11404.551 | 207289.8331 | 169534.3961 | 152347.6691 | 3800.5055 | 1.3333 | 8 | 878.6774 | 1727.3748
261 | 8_30_21_LAB_B |  | 06/09/2023 16:22:44 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_LAB_B.615 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.3333 | 0 | 33291.4321 | 18.6667 | 1210.0513 | 636.6809 | 14 | 2 | 0 | 42.0001 | 234.6686 | 2485.4398 | 62.6668 | 1577.4205 | 1.3333 | 221.3351 | 170.6677 | 1418.7372 | 1094.042 | -2353.6648 | 0.66667 | 28.6667 | 28.6667 | 1053.3722 | 0 | 2 | 1.3333 | 25.9005 | 0.66667 | 18 | 12.6667 | 162564.5501 | 176420.8721 | 11461.9305 | 193529.9419 | 168388.6056 | 152089.3451 | 3783.1681 | 0 | 6 | 887.3438 | 1664.7049
262 | 9_13_21_W_1 |  | 06/09/2023 16:28:32 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_W_1.616 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1947.4661 | 632.6807 | 435476.8549 | 323.337 | 8980.1553 | 1572.7533 | 282.0028 | 1866.7886 | 56.0001 | 51.3334 | 285.3362 | 36455.8199 | 50907.2119 | 10126.2553 | 26.6667 | 62.6668 | 649.3481 | 4710.777 | 1755.4413 | 2128.3468 | 64.6668 | 4582.0681 | 4582.0681 | 1055.3724 | 3.3333 | 4.6667 | 9.3333 | 62.5032 | 2 | 40673.1576 | 35639.7354 | 163432.384 | 171261.7893 | 11281.7916 | 197333.7035 | 169715.5045 | 147401.2447 | 3789.8364 | 14 | 36.6667 | 941.3453 | 1792.0435
263 | 9_13_21_W_6 |  | 06/09/2023 16:34:20 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_W_6.617 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2022.1431 | 712.6845 | 386943.0241 | 259.3357 | 5687.799 | 1315.3939 | 206.6682 | 1709.4357 | 118.6672 | 56.0001 | 246.0021 | 38588.1078 | 36342.1689 | 10847.4501 | 16.6667 | 54.0001 | 1611.4243 | 11460.5971 | 4358.665 | 2204.1495 | 78.0002 | 4771.4634 | 4771.4634 | 1071.3735 | 4 | 10 | 12.6667 | 108.7805 | 4.6667 | 14766.2944 | 13057.2987 | 159159.7309 | 169279.1047 | 11295.7974 | 196473.1867 | 162303.4709 | 148564.7522 | 3664.4702 | 30 | 64.0001 | 1520.031 | 3039.4596
264 | 9_13_21_W_8 |  | 06/09/2023 16:40:07 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_W_8.618 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2125.4915 | 708.0176 | 393180.3839 | 277.336 | 6596.8563 | 1179.3823 | 238.002 | 1704.1017 | 158.0009 | 90.667 | 488.675 | 40094.2779 | 32576.4348 | 10760.0509 | 20 | 92.667 | 1258.7222 | 9899.4373 | 2871.6221 | 2827.8166 | 65.3335 | 4942.1888 | 4942.1888 | 1020.0364 | 2.6667 | 10 | 18 | 82.3685 | 6 | 12721.6739 | 11130.3343 | 158989.513 | 169779.954 | 10800.7485 | 176350.876 | 165772.9917 | 145308.7216 | 3776.4996 | 22 | 54.0001 | 1249.3545 | 2529.4222
265 | LBLANK |  | 06/09/2023 16:45:55 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.619 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.6667 | 0 | 2746.9311 | 2.6667 | 926.03 | 624.6804 | 10.6667 | 0 | 0.66667 | 8 | 30 | 427.0828 | 10.6667 | 1114.0435 | 0.66667 | 8 | 8 | 66.0002 | 251.3355 | -2083.1974 | 0 | 12 | 12 | 1020.0365 | 0 | 0.66667 | 2 | 11.7269 | 0 | 22.6667 | 16 | 159532.534 | 171701.7426 | 11383.8694 | 205271.1165 | 161143.1292 | 150040.478 | 3829.8468 | 2 | 8 | 12 | 22.6667
266 | CCV 1ppb TRACE |  | 06/09/2023 16:51:44 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.620 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1333.3957 | 556.0109 | 74368.4154 | 49.3334 | 2674.9172 | 812.0231 | 632.014 | 276.6694 | 939.3643 | 1322.728 | 7189.8089 | 6596.686 | 11024.2536 | 2652.2466 | 2782.9378 | 757.3535 | 1832.1175 | 13872.0656 | 6690.9005 | 2679.2258 | 92.667 | 5431.0323 | 5431.0323 | 1492.078 | 126.0006 | 353.3377 | 820.6904 | 460.9079 | 139.334 | 12657.6054 | 11135.6723 | 159453.0097 | 169633.8955 | 11212.4 | 199390.5577 | 162733.0951 | 148207.5225 | 3809.1745 | 5869.8728 | 14007.5308 | 23532.9436 | 47789.1562
267 | CCV 1ppb TRACE |  | 06/09/2023 16:58:02 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.621 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1317.3941 | 526.0097 | 73969.6855 | 54.0001 | 2542.2262 | 848.0252 | 609.3463 | 292.6697 | 998.0349 | 1371.3992 | 7147.7885 | 6657.6293 | 10972.2127 | 2650.2459 | 2795.6069 | 761.3536 | 1959.4677 | 13773.9714 | 6715.5781 | 2992.6795 | 77.3335 | 5369.0089 | 5369.0089 | 1535.4162 | 112.6671 | 349.3376 | 840.6914 | 522.6068 | 145.3341 | 12602.8909 | 11255.7659 | 161537.5976 | 171529.695 | 11388.538 | 198984.9287 | 165344.7391 | 146943.2919 | 3797.1713 | 6009.2637 | 14089.6115 | 23769.7307 | 48460.8713
268 | LFB |  | 06/09/2023 17:03:50 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.622 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 12219.2252 | 4901.5075 | 522143.7899 | 342.6708 | 27141.0986 | 2643.5779 | 5684.4658 | 2517.5555 | 8878.7587 | 12279.9426 | 66468.9546 | 45486.07 | 106215.3994 | 11874.2673 | 25708.4521 | 6967.0319 | 16730.4579 | 126778.0602 | 60232.7129 | 48638.8049 | 869.3599 | 51485.9454 | 51485.9454 | 6318.7306 | 1156.7135 | 3595.119 | 7614.6958 | 5059.1069 | 1391.4011 | 120514.9359 | 105730.496 | 160063.3364 | 162170.6476 | 10765.3885 | 192455.2681 | 160074.6067 | 141135.7754 | 3567.112 | 56618.6544 | 133807.1012 | 223752.0255 | 440689.3817
269 | CCV 10 ppb MAJORS |  | 06/09/2023 17:09:38 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.623 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 12963.214 | 5237.6273 | 524688.8101 | 333.3372 | 22149.1583 | 2944.3034 | 6137.3183 | 2655.5802 | 9705.2959 | 13023.2694 | 71905.8971 | 49098.2022 | 113011.3749 | 12622.908 | 27621.3587 | 7638.0421 | 17900.5434 | 134821.959 | 65389.9973 | 50487.4352 | 911.3625 | 52529.7506 | 52529.7506 | 6655.5511 | 1185.3827 | 3705.8143 | 7680.7313 | 5018.2486 | 1391.4014 | 127604.1533 | 110245.83 | 160314.4304 | 167684.7637 | 10935.5191 | 185986.6369 | 161683.2303 | 145919.4565 | 3688.4763 | 59263.3387 | 140484.1116 | 230838.6915 | 455530.7245
270 | CCV 10 ppb MAJORS |  | 06/09/2023 17:15:55 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.624 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 13303.5253 | 5242.9632 | 516337.3577 | 350.0043 | 22689.3397 | 2778.2703 | 6226.0238 | 2790.2726 | 9511.1663 | 13000.5803 | 70786.9453 | 48542.2322 | 115173.0752 | 12844.4391 | 27901.8903 | 7732.7592 | 18184.9023 | 133429.5823 | 65014.9411 | 49815.8553 | 910.6957 | 53754.9489 | 53754.9489 | 6356.0805 | 1200.7172 | 3639.1302 | 7752.7704 | 4971.142 | 1408.0694 | 127014.1719 | 111510.8763 | 161575.319 | 166616.7156 | 11254.4316 | 196405.5933 | 162412.0151 | 148356.5541 | 3726.486 | 58043.0189 | 138254.4645 | 226661.4056 | 447171.9372
271 | 9_13_21_W_9 |  | 06/09/2023 17:21:43 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_W_9.625 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2009.4747 | 741.3526 | 286954.2349 | 202.6681 | 5797.1761 | 1233.3866 | 207.3348 | 1639.4276 | 138.0007 | 38.0001 | 178.0011 | 24218.1165 | 19995.9859 | 7197.1461 | 16 | 37.3334 | 1138.7121 | 8243.0448 | 1777.4439 | 2012.4209 | 70.0002 | 4319.3199 | 4319.3199 | 1076.0406 | 2.6667 | 6.6667 | 20 | 62.4328 | 4.6667 | 9846.7259 | 8618.5999 | 149125.9397 | 170271.4113 | 11461.9301 | 199969.2262 | 165669.8014 | 150075.5308 | 3761.8287 | 13.3333 | 40.6667 | 832.6763 | 1719.3732
272 | 9_13_21_W_10 |  | 06/09/2023 17:27:31 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_W_10.626 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1810.7814 | 896.6948 | 700208.7451 | 464.0076 | 7471.2868 | 1208.0511 | 220.0017 | 1751.4408 | 159.3342 | 29.3334 | 161.3343 | 49734.8697 | 31855.4787 | 13186.0836 | 22 | 43.3334 | 1136.7119 | 8684.6397 | 3685.1419 | 2327.4675 | 64.0001 | 4861.4938 | 4861.4938 | 1066.7066 | 0.66667 | 11.3333 | 18.6667 | 83.2408 | 1.3333 | 6264.7067 | 5330.9948 | 159383.5299 | 164357.4043 | 11094.3063 | 195769.6515 | 162580.6617 | 148448.6508 | 3639.7971 | 17.3333 | 52.6668 | 1702.7062 | 3172.804
273 | 9_13_21_W_10_DUP |  | 06/09/2023 17:33:20 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_W_10_DUP.627 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1742.7731 | 912.6958 | 604225.5362 | 378.0051 | 7461.9492 | 1276.057 | 238.002 | 1725.4377 | 150.6675 | 39.3334 | 230.0019 | 41910.148 | 28066.8779 | 11732.1505 | 23.3334 | 36 | 954.6986 | 7691.4034 | 3350.3933 | 2286.1362 | 64.0002 | 4726.1151 | 4726.1151 | 1066.0398 | 3.3333 | 8.6667 | 16.6667 | 105.5336 | 3.3333 | 6111.9741 | 5371.0096 | 159025.5691 | 165309.0619 | 11086.9676 | 195164.7613 | 161241.1285 | 145331.5677 | 3673.1389 | 19.3333 | 48.6668 | 1970.7204 | 3875.5408
274 | 9_13_21_W_11 |  | 06/09/2023 17:39:08 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_W_11.628 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1652.0955 | 856.0258 | 1367940.6643 | 941.3644 | 7429.2647 | 1216.0519 | 208.6682 | 1677.4318 | 151.3341 | 57.3335 | 314.0035 | 84152.3617 | 39741.8737 | 22089.732 | 25.3334 | 45.3334 | 1538.7496 | 11166.3633 | 6190.6744 | 1986.2968 | 64.6668 | 4511.3794 | 4511.3794 | 1076.0405 | 2.6667 | 9.3333 | 18 | 58.4706 | 2.6667 | 5533.0715 | 4816.1454 | 159151.4813 | 163437.0557 | 10870.1342 | 191954.3842 | 160543.2201 | 143286.3335 | 3571.113 | 27.3334 | 68.6668 | 2261.4034 | 4309.5889
275 | 9_13_21_F_1 |  | 06/09/2023 17:44:56 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_F_1.629 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9525.8417 | 2952.3055 | 219879.6294 | 146.6674 | 21028.1317 | 5128.2538 | 1151.3798 | 9075.5493 | 293.3364 | 234.0019 | 1269.3898 | 23184.3199 | 145132.8423 | 6977.7056 | 92.0003 | 369.3381 | 1474.0761 | 11974.3499 | 10926.8452 | 18179.4039 | 302.0032 | 20739.7116 | 20739.7116 | 1180.7155 | 5.3333 | 20 | 34.6667 | 184.172 | 8 | 196284.1358 | 173172.642 | 145172.7875 | 161248.3629 | 10704.0089 | 190337.6361 | 158473.5079 | 144698.5206 | 3533.7706 | 50.6668 | 111.3338 | 2678.765 | 5353.0579
276 | 9_13_21_F_6 |  | 06/09/2023 17:50:44 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_F_6.630 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10117.5818 | 3649.7995 | 188943.3568 | 109.3338 | 6105.3044 | 3885.195 | 1094.7086 | 9068.2121 | 610.013 | 140.0007 | 648.6814 | 22314.1773 | 133032.669 | 6947.6892 | 74.0002 | 295.3364 | 4400.6784 | 33306.7942 | 17780.3927 | 19001.874 | 310.6701 | 21425.3882 | 21425.3882 | 1205.3842 | 10 | 30 | 74.6669 | 270.5539 | 12 | 74678.0173 | 65779.1134 | 157033.0306 | 162034.4927 | 10235.6745 | 191922.9625 | 157589.9569 | 145931.6365 | 3693.8113 | 87.3336 | 238.002 | 3278.8135 | 6620.5962
277 | 9_13_21_F_8 |  | 06/09/2023 17:56:32 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_F_8.631 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10894.8201 | 3465.7543 | 204406.1061 | 126.6672 | 9042.1943 | 3694.4779 | 1090.7084 | 8663.9605 | 782.6881 | 88.6669 | 496.6753 | 25811.764 | 119689.3326 | 7839.484 | 60.6668 | 199.3347 | 4723.4475 | 35393.1231 | 13539.7469 | 20822.6431 | 367.3381 | 23443.8882 | 23443.8882 | 1224.0525 | 7.3333 | 27.3334 | 61.3335 | 281.2967 | 10 | 62517.1629 | 55270.7204 | 154232.8321 | 162778.1336 | 10241.009 | 188700.1472 | 156249.8939 | 145106.0021 | 3617.7915 | 70.0002 | 173.3344 | 4448.272 | 8749.0588
278 | 9_13_21_F_9 |  | 06/09/2023 18:02:20 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_F_9.632 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9700.6263 | 3497.7615 | 430948.5932 | 267.3359 | 8390.4634 | 3261.0398 | 1014.7027 | 7838.8174 | 638.0142 | 94.667 | 490.0084 | 26677.4085 | 43609.1278 | 7840.1514 | 48.0001 | 144.6674 | 4025.235 | 29689.4909 | 32370.6399 | 17085.1793 | 268.6692 | 19344.424 | 19344.424 | 1189.3829 | 9.3333 | 27.3334 | 69.3335 | 222.3087 | 10.6667 | 45574.5984 | 40886.439 | 154150.5613 | 163531.4867 | 10322.4035 | 187335.2108 | 154546.8433 | 147849.22 | 3673.1401 | 44.0001 | 100.667 | 2612.7603 | 5318.3858
279 | 9_13_21_F_10 |  | 06/09/2023 18:08:08 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_F_10.633 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 8603.2565 | 4255.3004 | 212897.3776 | 146.6674 | 10192.3026 | 3593.1186 | 1094.042 | 8647.2834 | 689.35 | 90.667 | 454.0073 | 18410.9073 | 49613.3448 | 5803.8466 | 45.3334 | 141.334 | 3096.3358 | 23591.47 | 11480.613 | 18913.4563 | 332.0039 | 21675.098 | 21675.098 | 1134.7118 | 8.6667 | 33.3334 | 61.3335 | 212.5407 | 10.6667 | 29535.8474 | 26045.0562 | 150921.6929 | 158165.4578 | 10461.8318 | 179040.4429 | 152831.855 | 140991.68 | 3490.4264 | 47.3334 | 116.6672 | 2610.7584 | 5323.7238
280 | 9_13_21_F_11 |  | 06/09/2023 18:13:56 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_F_11.634 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 8275.73 | 4211.2873 | 589749.1413 | 386.6719 | 7772.7808 | 3669.8048 | 1067.3733 | 8503.8637 | 690.6834 | 91.3336 | 442.0069 | 20425.2796 | 52985.4174 | 6177.3353 | 39.3334 | 135.334 | 3651.8005 | 27762.2866 | 22649.9526 | 18420.7993 | 306.67 | 21144.3035 | 21144.3035 | 1150.7131 | 8 | 30 | 61.3335 | 251.1549 | 13.3333 | 27384.2246 | 24137.0408 | 152608.998 | 158049.4801 | 10329.0684 | 175760.739 | 151378.2202 | 137645.4125 | 3483.7584 | 53.3334 | 120.6672 | 2501.4175 | 4916.3317
281 | 9_13_21_F_11_LD |  | 06/09/2023 18:19:44 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_F_11_LD.635 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 8327.0936 | 4281.9751 | 514220.3783 | 324.0037 | 10148.9373 | 3626.4606 | 1086.0413 | 8607.2595 | 640.681 | 86.0003 | 420.0062 | 19753.8267 | 53032.2558 | 6077.2929 | 48.6668 | 132.6673 | 3700.4793 | 27873.8375 | 22004.2697 | 18490.041 | 293.3364 | 20971.3866 | 20971.3866 | 1139.3788 | 5.3333 | 31.3334 | 65.3335 | 190.9363 | 9.3333 | 27098.3522 | 23937.3733 | 148609.2518 | 155475.038 | 10319.0652 | 181123.636 | 153391.8415 | 138382.4885 | 3439.7474 | 50.0001 | 136.6673 | 2471.416 | 4922.9944
282 | 9_13_21_F_11_LD |  | 06/09/2023 18:26:00 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_F_11_LD.636 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 8337.7658 | 4287.3099 | 513985.9068 | 346.6709 | 10512.5382 | 3583.783 | 1103.376 | 8529.2124 | 655.3484 | 94.0003 | 396.6722 | 19518.1116 | 53440.4494 | 6093.2993 | 46.6667 | 151.3341 | 3699.1455 | 27439.0015 | 21798.623 | 18370.6871 | 312.0034 | 20891.9328 | 20891.9328 | 1147.3794 | 11.3333 | 32.6667 | 66.0002 | 239.7563 | 10 | 26670.8745 | 24093.637 | 151939.8729 | 157798.0554 | 10608.6044 | 182867.7624 | 152827.19 | 141087.3334 | 3419.7432 | 58.0001 | 109.3338 | 2406.0793 | 4832.3213
283 | LBLANK |  | 06/09/2023 18:31:48 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.637 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.3333 | 0 | 2914.9646 | 3.3333 | 820.0237 | 588.0121 | 15.3333 | 0 | 0.66667 | 4 | 26.6667 | 453.5057 | 10 | 1080.7075 | 0.66667 | 12.6667 | 8.6667 | 76.0002 | 218.0017 | -2356.3357 | 0 | 6.6667 | 6.6667 | 1038.0378 | 2.6667 | 2 | 1.3333 | 57.8905 | 0.66667 | 16.6667 | 16 | 158194.0697 | 170708.6534 | 11012.2433 | 200547.8689 | 162999.4166 | 146209.1395 | 3677.807 | 0.66667 | 7.3333 | 16 | 22.6667
284 | CCV 1ppb TRACE |  | 06/09/2023 18:37:37 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.638 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1289.3915 | 518.6762 | 76141.055 | 50.0001 | 2644.9115 | 818.0236 | 628.6805 | 250.0022 | 949.3649 | 1326.7283 | 7113.104 | 6544.9391 | 11110.3195 | 2554.2283 | 2726.927 | 762.687 | 1761.4421 | 13700.568 | 6457.4598 | 2659.2721 | 82.6669 | 5300.3165 | 5300.3165 | 1590.0885 | 119.3338 | 376.6716 | 769.3541 | 526.9196 | 145.3341 | 12516.8146 | 11098.3106 | 159893.1791 | 167710.0397 | 10877.4739 | 197601.9376 | 143975.8015 | 147160.7663 | 3587.7841 | 5778.5019 | 13647.8502 | 23995.1937 | 48599.0276
285 | CCV 1ppb TRACE |  | 06/09/2023 18:43:54 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.639 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1248.0547 | 498.6754 | 75698.1075 | 42.6667 | 2512.2213 | 858.6925 | 642.0145 | 257.3357 | 947.3648 | 1320.061 | 6963.0303 | 6677.7686 | 11046.9363 | 2618.9069 | 2674.2503 | 740.0192 | 1923.4628 | 13816.6796 | 6654.2167 | 2890.6592 | 86.0003 | 5276.3078 | 5276.3078 | 1584.7546 | 114.0005 | 387.3386 | 797.3556 | 518.6336 | 152.0008 | 12436.078 | 10855.4564 | 159704.5068 | 166951.9846 | 10670.6531 | 198876.1248 | 146275.5679 | 146509.4461 | 3786.502 | 5871.2064 | 13903.4323 | 24030.5129 | 49083.7001
286 | CCV 10 ppb MAJORS |  | 06/09/2023 18:49:43 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.640 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 12071.7669 | 5026.8856 | 499811.7106 | 284.0028 | 21474.1292 | 2672.2499 | 5778.5019 | 2513.5545 | 9010.8411 | 12297.9581 | 66386.5793 | 45794.2749 | 107336.4068 | 11754.835 | 25844.0284 | 6885.6593 | 16748.4802 | 125293.7422 | 60449.638 | 47941.6539 | 827.3573 | 51321.3563 | 51321.3563 | 6163.9963 | 1079.3742 | 3495.0943 | 7505.972 | 4845.8329 | 1322.0612 | 118844.312 | 105757.3518 | 156223.1342 | 146720.6129 | 10525.2093 | 187645.0088 | 154943.0842 | 142979.2812 | 3202.3608 | 54977.5899 | 131602.0659 | 224477.5481 | 444165.9544
287 | CCV 10 ppb MAJORS |  | 06/09/2023 18:56:01 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.641 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 12468.1064 | 5100.2441 | 510463.8964 | 332.6706 | 21876.0712 | 2760.2666 | 5801.8451 | 2594.2356 | 9151.5973 | 12626.9115 | 68204.5237 | 46774.9731 | 109766.1338 | 12197.8721 | 26949.4013 | 7108.4351 | 17203.6861 | 128863.3669 | 62282.1694 | 49560.9112 | 816.0233 | 51755.5832 | 51755.5832 | 6258.7041 | 1090.0416 | 3447.7494 | 7608.6941 | 4884.4086 | 1323.3947 | 122609.3349 | 106453.8448 | 155119.0767 | 159108.515 | 10870.1345 | 186237.0452 | 168712.2642 | 147164.8231 | 3492.4268 | 55895.8387 | 134115.3899 | 225133.8273 | 433154.1717
288 | 9_13_21_F_11_DUP |  | 06/09/2023 19:01:49 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_F_11_DUP.642 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 8009.5781 | 4080.5832 | 435557.463 | 290.6696 | 7854.1589 | 3571.7799 | 1033.3708 | 8490.5236 | 625.3471 | 128.6673 | 752.6865 | 20890.831 | 51923.5305 | 6571.5111 | 50.6668 | 132.0006 | 3701.8135 | 27744.2483 | 20726.3589 | 18026.8689 | 297.3365 | 20495.359 | 20495.359 | 1126.0444 | 7.3333 | 27.3334 | 66.0002 | 262.5644 | 15.3333 | 26387.6897 | 23532.0371 | 153665.3086 | 155265.8744 | 10353.7508 | 180977.1401 | 150526.2305 | 138162.5291 | 3398.4044 | 56.0001 | 130.6673 | 2614.7595 | 5345.0575
289 | 9_13_21_F_11_DUP_LSP |  | 06/09/2023 19:07:37 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_F_11_DUP_LSP.643 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 18785.0094 | 8269.0592 | 920322.9478 | 594.6791 | 29109.6287 | 5098.9103 | 6376.0897 | 9828.7139 | 9271.0078 | 11978.3536 | 64848.2066 | 62732.4046 | 151329.328 | 16394.7356 | 24941.7544 | 6762.2673 | 18865.7826 | 140117.1398 | 73036.2317 | 63773.1813 | 1076.7072 | 67995.4428 | 67995.4428 | 6014.5992 | 1051.3721 | 3358.3948 | 7119.7745 | 4705.3852 | 1219.3854 | 138329.8553 | 124397.9391 | 150813.302 | 143898.8762 | 10018.846 | 178887.201 | 145983.0033 | 133959.1477 | 3247.0358 | 51478.6031 | 123114.9227 | 211622.4111 | 413430.2325
290 | 9_13_21_UF_1 |  | 06/09/2023 19:13:24 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_UF_1.644 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10540.5549 | 2838.2821 | 103091.5158 | 62.6668 | 18453.245 | 4950.1915 | 1112.7101 | 8363.1141 | 272.6693 | 76.0002 | 424.673 | 4614.111 | 136270.8596 | 2374.8641 | 80.0002 | 355.3378 | 974.0334 | 7992.2351 | 8117.6404 | 17335.4879 | 285.3362 | 19744.9698 | 19744.9698 | 1133.3783 | 4.6667 | 14.6667 | 34 | 193.7166 | 6 | 181520.6991 | 161030.5628 | 136520.8102 | 155629.9396 | 10236.3341 | 180972.4119 | 150809.5251 | 148241.1838 | 3395.0702 | 46.0001 | 110.0004 | 1116.0168 | 2274.7369
291 | 9_13_21_UF_6 |  | 06/09/2023 19:19:12 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_UF_6.645 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10870.135 | 3313.0512 | 46389.2184 | 25.3334 | 7584.6796 | 3677.8069 | 1002.7019 | 8409.8085 | 583.3453 | 94.667 | 418.6728 | 3225.2893 | 124840.4725 | 2077.4844 | 50.0001 | 169.3343 | 2852.2847 | 22328.1036 | 11000.9007 | 17830.6322 | 308.67 | 20471.3274 | 20471.3274 | 1159.3805 | 9.3333 | 31.3334 | 55.3334 | 202.8395 | 10 | 68957.371 | 60712.0769 | 149410.6322 | 161346.5643 | 10325.0634 | 182078.9603 | 151179.1856 | 140666.5734 | 3509.0983 | 86.0003 | 208.0015 | 1504.0308 | 3066.129
292 | 9_13_21_UF_6_DUP |  | 06/09/2023 19:25:00 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_UF_6_DUP.646 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 11046.2692 | 3391.0691 | 44863.7071 | 26 | 7622.0328 | 3697.8121 | 1042.7047 | 8266.3911 | 544.0104 | 121.3339 | 570.0114 | 3684.4769 | 126402.1201 | 2222.8402 | 56.6668 | 177.3344 | 2932.9679 | 22048.3388 | 11420.5646 | 17561.5489 | 284.0028 | 20866.5653 | 20866.5653 | 1188.7163 | 10.6667 | 34 | 64.0001 | 214.4706 | 9.3333 | 68184.3684 | 62264.7468 | 149244.2293 | 138862.6069 | 10328.3993 | 183446.3607 | 150067.5837 | 140708.2695 | 3029.6549 | 86.6669 | 188.6679 | 1520.6982 | 3097.4645
293 | 9_13_21_UF_8 |  | 06/09/2023 19:30:49 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_UF_8.647 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 11153.6859 | 3405.7393 | 44917.8767 | 40.0001 | 8593.2508 | 3310.3836 | 1057.3725 | 7960.2183 | 702.0173 | 74.6669 | 355.3378 | 2780.8794 | 110108.0461 | 2030.1443 | 52.6668 | 158.6676 | 2619.5735 | 19847.1116 | 6706.2411 | 19201.7493 | 327.3371 | 22166.5195 | 22166.5195 | 1156.0469 | 9.3333 | 26 | 66.0002 | 205.7549 | 6.6667 | 57702.3174 | 51552.1905 | 154403.409 | 157092.5138 | 9906.7677 | 180468.1254 | 151445.203 | 139541.7489 | 3407.7398 | 57.3335 | 123.3339 | 1238.6881 | 2576.0905

## Raw Intensities RSDs
| Sample Id | R | Acquisition Time | QC Status | Dataset File | Method File | Na KED 23/23 Helium KED_Low | Mg KED 24/24 Helium KED_Low | Al 27/27 | Al KED 27/27 Helium KED_Low | P Oshift 31/47 Oxygen DRC | S (34) -1 34/34 Helium KED_Low | K KED 39/39 Helium KED_Low | Ca KED 43/43 Helium KED_Low | V KED 51/51 Helium KED_Low | Cr (52) 52/52 Helium KED_Low | Cr (53) NH3 53/53 Ammonia DRC | Fe (54) NH3 54/54 Ammonia DRC | Mn Oshift 55/71 Oxygen DRC | Fe (56) KED 56/56 Helium KED_Low | Co KED 59/59 Helium KED_Low | Ni KED 60/60 Helium KED_Low | Cu (63) KED 63/63 Helium KED_Low | Cu (65) NH3 65/65 Ammonia DRC | Zn NH3 66/66 Ammonia DRC | As 75/75 | As KED 75/75 Helium KED_Low | As RH Oshift 75/91 Oxygen DRC | As Oshift 75/91 Oxygen DRC | Se (77) 77/77 | Se (77) Oshift 77/93 Oxygen DRC | Se (78) Oshift 78/94 Oxygen DRC | Se (80) Oshift 80/96 Oxygen DRC | Se (82) 82/82 | Se (82) Oshift 82/98 Oxygen DRC | Mo (95) NH3 95/95 Ammonia DRC | Mo (95) OShift 95/127 Oxygen DRC | Rh (IS1) 103/103 (IS) | Rh (IS2) 103/103 (IS) Oxygen DRC | Rh (IS3) 103/103 (IS) Helium KED_Low | Rh (IS4) 103/103 (IS) Ammonia DRC | Rh (IS5) 103/103 (IS) Ammonia DRC | Rh (IS6) 103/103 (IS) Oxygen DRC | Rh OShift (IS1) 103/119 (IS) Oxygen DRC | Cd (111) NH3 111/111 Ammonia DRC | Cd (114) NH3 114/114 Ammonia DRC | Pb 208/208 | Pb NH3 208/208 Ammonia DRC
1 | Sac Blank |  | 6/8/2023 1:23:24 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\Sac Blank.356 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth |  |  | 2.812% | 114.6% | 3.03% | 3.829% | 39.85% |  | 86.6% | 51.63% | 18.03% | 5.185% | 6.928% | 2.584% | 100% | 33.33% | 28.64% | 21.83% | 5.512% | 3.66% |  | 8.332% | 8.332% | 3.835% | 86.6% | 0% | 62.45% | 34.19% | 0% | 41.93% | 52.92% | 1.627% | 0.6725% | 2.877% | 1.585% | 1.201% | 0.8656% | 1.379% | 100% | 26.96% | 16.37% | 16%
2 | sac blank cal |  | 6/8/2023 1:29:13 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\sac blank cal.357 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 86.6% |  | 28.63% | 170.2% | 2.996% | 10.48% | 75.59% |  | 173.2% | 32.73% | 31.49% | 4.339% | 12.37% | 1.696% | 173.2% | 45.83% | 50% | 12.48% | 1.307% | 1.977% |  | 12.5% | 12.5% | 11.88% |  | 173.2% | 50% | 42.15% | 173.2% | 60.27% | 17.32% | 4.13% | 0.1976% | 0.9564% | 0.5423% | 1.078% | 1.771% | 1.743% | 173.2% | 88.19% | 24.22% | 11.63%
3 | CAL BLANK |  | 6/8/2023 1:35:02 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CAL BLANK.358 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 173.2% |  | 1.855% | 173.2% | 4.223% | 5.59% | 40.52% | 173.2% | 173.2% | 45.83% | 51.49% | 2.03% | 22.91% | 3.049% | 91.65% | 60.27% | 16.67% | 28.09% | 9.032% | 10.31% | 173.2% | 27.7% | 27.7% | 6.931% | 24.74% | 86.6% | 114.6% | 24.73% |  | 33.33% | 36.74% | 0.7181% | 1.453% | 1.271% | 0.3834% | 1.229% | 0.3919% | 1.067% | 86.6% | 25% | 26.14% | 13.63%
4 | LOW 0.01 |  | 6/8/2023 1:40:50 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LOW 0.01.359 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 16.67% | 45.83% | 1.275% | 25% | 1.308% | 4.381% | 0% | 0% | 32.73% | 12.39% | 17.01% | 4.339% | 5.874% | 3.4% | 22.53% | 36.06% | 23.21% | 9.159% | 0.6181% | 5.55% | 100% | 17.03% | 17.03% | 3.889% | 114.6% | 32.73% | 42.14% | 54.64% | 69.28% | 4.847% | 3.186% | 1.278% | 1.13% | 1.195% | 1.267% | 0.8522% | 0.2996% | 1.297% | 12.33% | 5.155% | 5.458% | 4.154%
5 | LOW 0.05 |  | 6/8/2023 1:46:39 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LOW 0.05.360 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 11.76% | 14.19% | 0.817% | 29.67% | 1.048% | 1.414% | 14.61% | 14.78% | 15.03% | 10.99% | 14.61% | 1.455% | 4.452% | 6.59% | 6.523% | 31.33% | 4.086% | 4.595% | 4.124% | 4.161% | 35.25% | 3.912% | 3.912% | 5.354% | 24.12% | 20.83% | 30.75% | 22.95% | 10.83% | 1.173% | 5.743% | 0.9218% | 1.429% | 1.689% | 0.6189% | 0.4923% | 2.372% | 0.5563% | 9.876% | 4.066% | 2.019% | 1.803%
6 | LOW 0.1 |  | 6/8/2023 1:52:27 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LOW 0.1.361 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4.397% | 14% | 1.952% | 24.21% | 3.89% | 5.947% | 9.366% | 18.32% | 12.6% | 6.429% | 3.159% | 8.584% | 8.9% | 0.1031% | 10.87% | 1.961% | 7.607% | 7.864% | 2.613% | 5.21% | 24.98% | 6.001% | 6.001% | 3.643% | 61.39% | 35.25% | 20.2% | 39.57% | 8.333% | 3.42% | 5.259% | 0.8594% | 5.115% | 4.619% | 1% | 1.438% | 1.021% | 7.014% | 1.384% | 2.363% | 4.706% | 1.821%
7 | LOW 0.5 |  | 6/8/2023 1:58:15 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LOW 0.5.362 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4.326% | 4.931% | 0.7593% | 28.35% | 1.988% | 4.319% | 10.34% | 16.96% | 3.838% | 3.379% | 3.341% | 1.932% | 1.233% | 0.7453% | 3.493% | 1.838% | 3.89% | 1.952% | 1.055% | 96.27% | 6.863% | 1.834% | 1.834% | 3.427% | 13.91% | 9.629% | 4.752% | 7.358% | 7.391% | 1.428% | 1.927% | 0.2665% | 3.091% | 5.196% | 1.212% | 0.3302% | 0.4057% | 0.8037% | 1.34% | 1.019% | 0.506% | 0.5888%
8 | STD 1 |  | 6/8/2023 2:04:03 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 1.363 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 5.252% | 9.091% | 2.097% | 21.31% | 1.063% | 3.709% | 5.44% | 5.341% | 5.18% | 3.132% | 1.435% | 1.693% | 2.046% | 2.367% | 1.523% | 1.766% | 2.158% | 0.6509% | 0.4117% | 10.97% | 19.1% | 1.423% | 1.423% | 6.316% | 7.256% | 4.408% | 5.697% | 8.027% | 7.98% | 1.435% | 2.62% | 0.2087% | 2.187% | 0.8447% | 1.061% | 1.73% | 1.465% | 4.04% | 1.399% | 1.592% | 0.883% | 1.542%
9 | STD 5 |  | 6/8/2023 2:09:52 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 5.364 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.932% | 3.486% | 1.346% | 4.957% | 0.05396% | 4.599% | 3.95% | 3.235% | 3.536% | 2.358% | 0.4047% | 0.3914% | 1.112% | 2.377% | 2.333% | 2.533% | 2.282% | 0.511% | 0.3286% | 0.4229% | 4.734% | 0.358% | 0.358% | 2.475% | 2.203% | 0.9774% | 1.868% | 3.685% | 6.164% | 0.835% | 0.7926% | 4.676% | 0.3397% | 0.7155% | 0.5782% | 0.3337% | 2.109% | 0.3813% | 0.3482% | 0.9221% | 1.051% | 1.768%
10 | STD 10 |  | 6/8/2023 2:15:40 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 10.365 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.793% | 1.988% | 1.04% | 6.667% | 0.9964% | 1.313% | 0.2265% | 1.335% | 1.223% | 0.8185% | 1.122% | 2.796% | 1.4% | 0.9516% | 1.292% | 1.849% | 1.11% | 1.726% | 0.9491% | 0.8238% | 3.209% | 0.7406% | 0.7406% | 1.608% | 0.5591% | 3.829% | 2.184% | 3.517% | 1.882% | 0.697% | 0.9542% | 0.5486% | 2.039% | 0.9279% | 2.81% | 1.823% | 0.4268% | 1.998% | 1.249% | 1.434% | 1.003% | 2.463%
11 | STD 25 |  | 6/8/2023 2:21:28 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 25.366 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.908% | 0.6291% | 1.165% | 2.214% | 0.5021% | 2.811% | 0.515% | 1.883% | 2.615% | 1.021% | 1.32% | 0.3371% | 0.302% | 0.3169% | 2.992% | 3.04% | 1.722% | 1.017% | 0.5317% | 0.5726% | 1.882% | 0.6853% | 0.6853% | 0.7757% | 1.315% | 2.117% | 0.8545% | 1.315% | 1.318% | 1.43% | 0.4105% | 0.6048% | 1.045% | 0.5095% | 0.8364% | 4.03% | 1.321% | 4.63% | 0.4479% | 1.014% | 0.5469% | 2.385%
12 | STD 50 |  | 6/8/2023 2:27:16 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 50.367 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.798% | 3.044% | 0.4264% | 5.737% | 1.065% | 1.871% | 1.555% | 1.357% | 0.8734% | 1.533% | 0.8067% | 0.759% | 0.7798% | 2.101% | 1.644% | 2.205% | 1.137% | 0.7686% | 0.6988% | 0.8433% | 1.537% | 0.1838% | 0.1838% | 1.445% | 2.03% | 0.8822% | 0.8219% | 2.242% | 1.654% | 0.981% | 0.8348% | 1.27% | 3.464% | 2.381% | 1.638% | 1.147% | 0.6748% | 3.989% | 0.9674% | 0.3413% | 1.023% | 0.4339%
13 | STD 100 |  | 6/8/2023 2:33:03 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 100.368 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.7286% | 0.4948% | 0.3982% | 0.6236% | 1.754% | 1.368% | 1.236% | 1.305% | 1.047% | 1.741% | 0.2713% | 1.419% | 1.043% | 0.3924% | 0.8817% | 1.412% | 0.8791% | 0.9638% | 0.6484% | 0.5329% | 0.8737% | 0.4491% | 0.4491% | 1.146% | 1.233% | 0.4639% | 0.5485% | 1.753% | 0.9532% | 0.9191% | 1.15% | 0.6864% | 0.6608% | 0.6473% | 1.208% | 1.065% | 1.602% | 1.482% | 0.4244% | 1.493% | 1.295% | 0.5729%
15 | LBLANK |  | 6/8/2023 2:38:52 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.369 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 50% | 86.6% | 0.931% | 43.3% | 3.014% | 3.561% | 45.07% |  | 78.06% | 11.11% | 31.46% | 5.605% | 43.37% | 2.331% | 98.97% | 11.76% | 47.21% | 3.612% | 4.441% | 6.549% | 100% | 12.16% | 12.16% | 6.198% | 173.2% | 34.64% | 33.33% | 4.697% | 100% | 4.648% | 8.005% | 0.826% | 0.5315% | 2.954% | 0.7769% | 1.662% | 0.9978% | 3.796% | 39.36% | 7.512% | 9.474e-06% | 1.493%
16 | LBLANK |  | 6/8/2023 2:57:55 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.370 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 69.28% |  | 50.17% | 173.2% | 3.497% | 8.71% | 8.66% |  | 69.28% | 49.96% | 26.38% | 6.947% | 22.22% | 2.619% | 100% | 9.116% | 24.12% | 13.48% | 4.387% | 3.78% | 91.65% | 27.45% | 27.45% | 4.205% | 100% | 65.47% | 78.06% | 26.67% | 173.2% | 17.76% | 9.472% | 0.6473% | 0.2683% | 11.92% | 0.8077% | 1.582% | 0.5199% | 1.146% | 86.6% | 7.531% | 39.11% | 17.32%
17 | ICV PE |  | 6/8/2023 3:03:43 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\ICV PE.371 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.048% | 1.975% | 0.8687% | 6.276% | 2.352% | 3.254% | 2.552% | 1.816% | 0.3814% | 1.159% | 0.3911% | 1.345% | 0.9409% | 0.8106% | 1.312% | 2.151% | 1.982% | 0.157% | 0.3878% | 2.099% | 5.33% | 1.125% | 1.125% | 0.5406% | 2.044% | 2.656% | 0.7734% | 1.592% | 4.563% | 0.8487% | 1.164% | 1.695% | 0.7044% | 1.85% | 1.344% | 6.61% | 2.238% | 2.261% | 0.981% | 0.3684% | 0.2799% | 0.9021%
18 | ICV PE |  | 6/8/2023 3:09:59 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\ICV PE.372 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.9483% | 1.189% | 1.25% | 13.23% | 0.9539% | 1.636% | 1.644% | 3.068% | 3.792% | 2.026% | 0.5277% | 2.51% | 0.5906% | 2.28% | 0.5401% | 1.489% | 1.074% | 1.245% | 0.7248% | 1.994% | 6.797% | 0.2455% | 0.2455% | 1.944% | 4.221% | 0.7513% | 1.558% | 1.266% | 4.351% | 0.3418% | 0.9108% | 0.7618% | 0.9079% | 1.909% | 0.9095% | 1.115% | 2.288% | 2.925% | 1.451% | 1.079% | 0.7628% | 0.9735%
19 | CCV 1ppb TRACE |  | 6/8/2023 3:15:47 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.373 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.881% | 4.272% | 1.207% | 10.58% | 1.696% | 3.145% | 2.801% | 11.87% | 0.9036% | 4.821% | 0.9651% | 1.682% | 2.793% | 1.572% | 2.516% | 2.989% | 2.936% | 0.4957% | 1.789% | 4.708% | 12.44% | 1.366% | 1.366% | 2.581% | 12.47% | 7.174% | 3.117% | 7.009% | 7.692% | 2.441% | 1.086% | 1.514% | 2.122% | 0.9313% | 1.015% | 0.5324% | 1.54% | 1.954% | 2.683% | 3.085% | 0.8872% | 0.5771%
20 | CCV 1ppb TRACE |  | 6/8/2023 3:22:03 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.374 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.3845% | 2.518% | 0.8577% | 39.3% | 2.306% | 5.515% | 5.561% | 10.67% | 5.971% | 0.6422% | 3.827% | 2.578% | 2.675% | 3.699% | 2.004% | 6.855% | 3.169% | 0.9112% | 0.6261% | 5.718% | 15.84% | 0.7916% | 0.7916% | 3.743% | 8.347% | 7.101% | 3.638% | 1.719% | 13.73% | 2.119% | 2.01% | 0.9699% | 1.165% | 0.4048% | 2.209% | 0.5599% | 0.9466% | 3.065% | 2.06% | 2.222% | 1.704% | 1.641%
21 | LFB |  | 6/8/2023 3:27:51 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.375 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.437% | 3.913% | 1.376% | 3.518% | 0.5292% | 2.871% | 0.9415% | 3.356% | 1.734% | 1.721% | 0.6826% | 1.399% | 0.6836% | 1.677% | 1.689% | 3.768% | 0.5786% | 1.16% | 1.351% | 0.7885% | 3.962% | 1.693% | 1.693% | 2.209% | 0.9426% | 1.575% | 2.281% | 3.443% | 2.387% | 1.42% | 1.356% | 1.704% | 0.5031% | 0.6252% | 3.826% | 1.025% | 2.003% | 0.9338% | 1.715% | 0.4631% | 1.19% | 0.8962%
22 | LLOQ 0.01 |  | 6/8/2023 3:33:39 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LLOQ 0.01.376 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 44.3% | 32.73% | 1.377% | 13.32% | 0.9095% | 1.18% | 16.9% | 65.47% | 56.79% | 4.225% | 24.83% | 4.822% | 5.44% | 2.657% | 21.5% | 14.32% | 25.94% | 1.62% | 3.116% | 5.929% | 86.6% | 16.05% | 16.05% | 7.266% | 65.47% | 24.74% | 24.74% | 849.7% | 50% | 5.8% | 3.571% | 2.784% | 1.412% | 4.349% | 0.4254% | 1.262% | 0.3261% | 1.502% | 13.68% | 1.348% | 5.111% | 1.505%
23 | CCV 10 ppb MAJORS |  | 6/8/2023 3:39:28 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.377 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 100% | 173.2% | 2.899% | 173.2% | 9.967% | 13.85% | 34.64% | 173.2% | 173.2% | 28.87% | 22.53% | 3.931% | 20% | 1.479% | 92.49% | 34.64% | 39.74% | 12.03% | 5.078% | 5.158% |  | 37.65% | 37.65% | 13.7% | 86.6% | 100% | 43.3% | 72.33% | 86.6% | 16.37% | 13.32% | 2.953% | 3.68% | 2.815% | 2.368% | 3.804% | 4.976% | 6.926% | 0% | 43.3% | 10.19% | 45.83%
24 | CCV 10 ppb MAJORS |  | 6/8/2023 3:45:42 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.378 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.603% | 2.497% | 1.11% | 5.2% | 3.321% | 1.618% | 1.695% | 4.634% | 2.833% | 0.5455% | 0.8984% | 0.8868% | 0.7748% | 2.385% | 1.738% | 1.832% | 0.8955% | 0.9343% | 0.4926% | 1.209% | 6.752% | 0.3535% | 0.3535% | 1.585% | 2.081% | 3.219% | 1.698% | 1.687% | 6.77% | 0.9154% | 1.16% | 0.8399% | 0.918% | 2.645% | 1.254% | 0.8326% | 2.677% | 1.684% | 1.428% | 1.782% | 0.7024% | 1.533%
25 | LLOQ 0.1 |  | 6/8/2023 3:51:31 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LLOQ 0.1.379 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 8.104% | 13.22% | 1.844% | 18.55% | 0.962% | 6.84% | 21.29% | 2.665% | 6.928% | 2.054% | 4.021% | 3.709% | 2.563% | 6.028% | 6.815% | 16.88% | 9.323% | 1.022% | 0.2155% | 16.25% | 25% | 5.518% | 5.518% | 1.408% | 20% | 23.39% | 24.77% | 13.04% | 26.34% | 3.397% | 2.4% | 1.363% | 0.7574% | 1.414% | 0.5907% | 1.786% | 2.48% | 3.45% | 5.876% | 4.353% | 1.887% | 1.868%
26 | STD 0.5 |  | 6/8/2023 3:57:19 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 0.5.380 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.432% | 2.632% | 1.73% | 10.19% | 3.426% | 9.39% | 4.899% | 14.1% | 8.77% | 2.728% | 1.962% | 1.11% | 2.228% | 0.7453% | 4.782% | 8.988% | 6.84% | 3.617% | 1.41% | 20.19% | 13.35% | 3.919% | 3.919% | 4.803% | 8.173% | 8.824% | 6.845% | 19.44% | 12.48% | 1.096% | 3.533% | 0.5691% | 1.706% | 1.928% | 5.45% | 0.3688% | 1.667% | 2.241% | 1.768% | 2.54% | 0.8076% | 2.788%
27 | 8_2_21_UF_1 |  | 6/8/2023 4:03:07 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_UF_1.381 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.414% | 3.6% | 0.6617% | 8.171% | 2.783% | 2.843% | 0.6199% | 0.2873% | 9.109% | 10.87% | 10.7% | 2.229% | 1.179% | 6.682% | 26.46% | 10.05% | 3.976% | 1.544% | 1.899% | 1.641% | 3.318% | 1.932% | 1.932% | 8.921% | 70.5% | 20.4% | 3.846% | 10.73% | 45.83% | 1.266% | 0.6756% | 1.21% | 0.4188% | 1.712% | 1.335% | 0.7152% | 2.049% | 0.07906% | 26.52% | 4.534% | 3.017% | 4.415%
28 | 8_2_21_UF_6 |  | 6/8/2023 4:08:55 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_UF_6.382 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.078% | 1.51% | 3.828% | 14.7% | 1.306% | 5.343% | 8.02% | 2.865% | 3.655% | 16.42% | 5.863% | 4.378% | 1.661% | 3.146% | 3.646% | 0.6298% | 4.24% | 1.652% | 0.6546% | 1.215% | 11.05% | 2.113% | 2.113% | 3.488% | 6.928% | 11.6% | 21.55% | 11.37% | 31.11% | 1.261% | 0.9142% | 0.6371% | 0.02132% | 0.9061% | 1.816% | 1.39% | 0.8845% | 2.885% | 12.88% | 8.986% | 10.39% | 2.664%
29 | 8_2_21_UF_6_SP |  | 6/8/2023 4:14:43 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_UF_6_SP.383 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.041% | 3.446% | 0.9948% | 2.187% | 0.5297% | 2.044% | 2.925% | 2.157% | 1.457% | 2.167% | 1.13% | 1.634% | 0.7492% | 5.569% | 0.9176% | 1.08% | 1.445% | 0.5914% | 0.6317% | 1.092% | 3.252% | 1.046% | 1.046% | 3.042% | 4.727% | 2.748% | 3.07% | 5.515% | 3.974% | 1.374% | 1.9% | 0.5132% | 0.1651% | 0.8316% | 1.552% | 1.917% | 5.294% | 1.965% | 0.8519% | 0.823% | 0.983% | 1.376%
30 | STD 1 |  | 6/8/2023 4:20:31 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 1.384 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.8605% | 2.876% | 0.5294% | 9.625% | 2.653% | 5.137% | 4.64% | 7.328% | 4.88% | 6.337% | 0.7621% | 14.32% | 1.263% | 1.181% | 4.237% | 1.243% | 5.168% | 1.339% | 1.495% | 6.414% | 2.744% | 1.569% | 1.569% | 2.766% | 7.391% | 11.28% | 6.424% | 11.62% | 7.418% | 1.489% | 1.692% | 1.884% | 0.5413% | 2.148% | 1.09% | 1.041% | 1.707% | 5.861% | 2.752% | 1.148% | 1.804% | 1.121%
31 | 8_2_21_UF_6_SP_LD |  | 6/8/2023 4:26:19 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_UF_6_SP_LD.385 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.529% | 2.941% | 1.108% | 9.984% | 1.688% | 0.7126% | 3.72% | 2.036% | 1.843% | 0.6328% | 1.489% | 1.187% | 2.141% | 0.4877% | 2.529% | 1.791% | 1.835% | 1.904% | 0.6472% | 1.039% | 4.307% | 0.6621% | 0.6621% | 0.3616% | 7.149% | 3.094% | 0.9146% | 1.465% | 3.995% | 0.8796% | 1.401% | 0.5446% | 0.8162% | 1.402% | 1.325% | 1.871% | 2.392% | 2.015% | 0.266% | 0.7998% | 1.552% | 1.097%
32 | 8_2_21_UF_6_SP_LSP |  | 6/8/2023 4:32:06 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_UF_6_SP_LSP.386 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.2982% | 0.5618% | 1.054% | 9.317% | 1.506% | 1.072% | 0.3384% | 0.9991% | 0.7798% | 1.456% | 1.627% | 1.41% | 0.9141% | 1.597% | 0.9195% | 0.7426% | 0.6678% | 0.2953% | 0.6393% | 1.569% | 1.734% | 0.3588% | 0.3588% | 2.001% | 2.799% | 1.026% | 2.026% | 0.8969% | 3.879% | 1.483% | 0.2384% | 2.059% | 0.722% | 0.9421% | 0.9866% | 0.8148% | 0.9554% | 0.5294% | 0.8727% | 0.6725% | 0.8749% | 1.515%
33 | 8_2_21_UF_8 |  | 6/8/2023 4:37:54 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_UF_8.387 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.038% | 1.528% | 1.097% | 13.58% | 2.472% | 3.679% | 2.123% | 3.899% | 2.793% | 7.164% | 1.816% | 1.041% | 0.8596% | 1.442% | 12.9% | 2.41% | 2.072% | 0.9708% | 1.331% | 0.9427% | 9.575% | 0.8148% | 0.8148% | 2.8% | 9.116% | 13.36% | 13.22% | 4.052% | 12.37% | 1.734% | 1.255% | 0.7286% | 0.1808% | 1.148% | 1.363% | 1.293% | 2.475% | 3.389% | 7.943% | 7.123% | 1.171% | 0.8433%
34 | 8_2_21_UF_9 |  | 6/8/2023 4:43:42 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_UF_9.388 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4.025% | 1.62% | 1.303% | 33.25% | 0.951% | 1.062% | 2.565% | 2.534% | 6.611% | 10.94% | 6.497% | 0.7805% | 0.6461% | 4.9% | 25.81% | 9.333% | 0.4927% | 1.577% | 1.06% | 2.866% | 4.581% | 0.7976% | 0.7976% | 3.229% | 45.83% | 28.64% | 21.31% | 5.641% | 24.74% | 1.657% | 0.7159% | 0.7503% | 0.7306% | 2.566% | 0.6562% | 0.7723% | 1.037% | 1.547% | 6.864% | 7.132% | 2.612% | 4.005%
35 | LBLANK |  | 6/8/2023 4:49:31 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.389 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 66.67% |  | 3.996% | 100% | 9.403% | 4.928% | 70.89% | 173.2% | 0% | 74.18% | 23.81% | 6.25% | 19.25% | 0.7867% | 173.2% | 29.57% | 32.87% | 23.11% | 19.34% | 8.506% | 91.65% | 9.141% | 9.141% | 6.26% | 34.64% | 43.3% | 26.96% | 43.26% | 43.3% | 27.15% | 37.55% | 0.6767% | 1.232% | 7.964% | 1.232% | 0.2253% | 0.9105% | 0.5329% | 100% | 15.75% | 12.5% | 12.06%
36 | CCV 1ppb TRACE |  | 6/8/2023 4:55:19 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.390 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.251% | 5.161% | 2.004% | 11.46% | 3.057% | 6.206% | 6.063% | 15.68% | 5.005% | 3.619% | 3.422% | 0.4738% | 1.052% | 1.897% | 2.312% | 3.596% | 2.736% | 0.3748% | 1.5% | 5.888% | 12.48% | 1.396% | 1.396% | 2.632% | 19.97% | 2.244% | 2.567% | 10.7% | 18.38% | 1.24% | 0.3543% | 1.163% | 0.909% | 3.291% | 0.5874% | 3.356% | 1.97% | 2.573% | 2.586% | 2.409% | 2.776% | 4.087%
37 | CCV 1ppb TRACE |  | 6/8/2023 5:01:37 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.391 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.855% | 4.747% | 2.817% | 13.63% | 6.02% | 3.895% | 8.512% | 4.908% | 4.066% | 5.528% | 1.259% | 4.53% | 1.264% | 1.087% | 2.581% | 3.089% | 1.241% | 1.254% | 0.7158% | 3.242% | 9.109% | 1.97% | 1.97% | 0.5758% | 4.354% | 6.493% | 4.244% | 8.902% | 12.35% | 1.38% | 0.5469% | 1.534% | 0.7503% | 1.252% | 0.8184% | 0.7491% | 1.883% | 2.468% | 0.958% | 1.597% | 1.781% | 0.4497%
38 | CCV 10 ppb MAJORS |  | 6/8/2023 5:07:25 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.392 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.66% | 3.259% | 1.178% | 2.677% | 2.422% | 3.587% | 2.759% | 4.9% | 1.232% | 0.783% | 1.064% | 1.617% | 1.464% | 1.759% | 1.133% | 1.175% | 2.186% | 1.046% | 0.08356% | 0.995% | 7.253% | 0.645% | 0.645% | 2.958% | 5.757% | 0.8601% | 1.855% | 0.5496% | 1.789% | 1.771% | 0.4285% | 2.064% | 1.555% | 1.419% | 0.6118% | 1.09% | 1.704% | 1.342% | 1.126% | 0.8908% | 0.8016% | 1.617%
39 | CCV 10 ppb MAJORS |  | 6/8/2023 5:13:42 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.393 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.081% | 3.633% | 1.257% | 2.797% | 3.966% | 2.074% | 1.6% | 2.151% | 0.7975% | 1.942% | 1.96% | 2.758% | 0.7395% | 0.9136% | 0.9172% | 0.7266% | 0.6498% | 0.8355% | 0.435% | 1.239% | 3.982% | 1.17% | 1.17% | 1.487% | 4.633% | 1.79% | 3.38% | 1.944% | 0.296% | 0.2896% | 0.1344% | 1.514% | 0.9778% | 1.687% | 0.4846% | 0.4652% | 1.479% | 2.845% | 0.6598% | 0.4985% | 0.3291% | 3.3%
40 | 8_2_21_UF_10 |  | 6/8/2023 5:19:31 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_UF_10.394 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.5648% | 2.102% | 1.204% | 24.22% | 0.3332% | 1.595% | 0.6442% | 3.131% | 2.345% | 12.97% | 8.032% | 1.419% | 1.314% | 2.971% | 30.71% | 8.873% | 2.226% | 0.7347% | 1.301% | 2.037% | 2.263% | 0.6373% | 0.6373% | 3.263% | 9.116% | 36.92% | 19.16% | 20.67% | 11.17% | 1.796% | 1.295% | 1.627% | 0.9795% | 5.576% | 0.8848% | 0.3215% | 0.8424% | 1.95% | 5.589% | 13.32% | 4.397% | 2.129%
41 | 8_2_21_UF_11 |  | 6/8/2023 5:25:19 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_UF_11.395 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.168% | 4.39% | 1.377% | 9.75% | 0.371% | 1.862% | 4.848% | 5.038% | 4.13% | 2.789% | 1.865% | 1.852% | 0.7634% | 2.385% | 13.12% | 7.777% | 2.881% | 0.4953% | 1.578% | 0.8841% | 3.665% | 0.4166% | 0.4166% | 3.626% | 44.41% | 9.608% | 7.747% | 18.57% | 34.64% | 1.454% | 0.846% | 0.4756% | 1.187% | 17.91% | 0.5905% | 1.555% | 1.831% | 2.621% | 7.16% | 5.188% | 2.064% | 4.048%
42 | 8_2_21_UF_11_DUP |  | 6/8/2023 5:31:09 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_UF_11_DUP.396 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.77% | 1.87% | 1.212% | 44.61% | 1.767% | 3.125% | 6.099% | 3.121% | 3.169% | 3.446% | 9.141% | 4.778% | 2.042% | 0.8406% | 12.91% | 5.547% | 3.12% | 1.932% | 1.392% | 1.13% | 9.068% | 0.7259% | 0.7259% | 2.798% | 39.85% | 11.1% | 11.64% | 5.279% | 22.05% | 0.5387% | 1.322% | 0.4259% | 1.515% | 1.876% | 1.526% | 0.4075% | 1.775% | 2.95% | 14.24% | 12.74% | 2.035% | 0.9364%
43 | 8_2_21_F_1 |  | 6/8/2023 5:36:57 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_F_1.397 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.248% | 1.331% | 0.971% | 1.682% | 3.329% | 3.915% | 5.377% | 2.182% | 7.908% | 9.938% | 8.873% | 0.9704% | 0.8036% | 2.696% | 5.329% | 5.6% | 2.998% | 0.6847% | 0.1846% | 1.796% | 13.48% | 1.376% | 1.376% | 6.893% | 66.14% | 17.64% | 35.35% | 7.52% | 48.24% | 0.5095% | 1.783% | 1.102% | 0.7922% | 2.018% | 0.5353% | 0.2488% | 2.767% | 3.616% | 18.73% | 21.16% | 5.989% | 1.829%
44 | 8_2_21_F_6 |  | 6/8/2023 5:44:00 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_F_6.398 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.303% | 2.312% | 0.9327% | 26.23% | 1.368% | 1.468% | 2.324% | 2.466% | 1.511% | 4.552% | 0.9595% | 0.9236% | 3.484% | 1.833% | 19.49% | 0.885% | 1.049% | 0.2984% | 0.8737% | 0.4498% | 6.396% | 4.231% | 4.231% | 2.488% | 30% | 12.43% | 18.94% | 4.201% | 19.87% | 0.8062% | 2.985% | 2.107% | 16.96% | 1.116% | 0.5041% | 0.4939% | 0.7912% | 16.41% | 3.775% | 9.918% | 0.7512% | 0.5926%
45 | 8_2_21_F_8 |  | 6/8/2023 5:49:48 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_F_8.399 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.472% | 1.045% | 0.5882% | 15.47% | 2.287% | 3.388% | 4.274% | 0.9783% | 6.237% | 0% | 4.565% | 1.04% | 2.255% | 3.585% | 17.14% | 9.22% | 2.099% | 0.7598% | 0.8792% | 1.239% | 7.896% | 1.519% | 1.519% | 4.992% | 39.74% | 33.02% | 18.66% | 9.582% | 40.52% | 1.304% | 2.387% | 2.065% | 4.363% | 0.4728% | 0.7762% | 0.5153% | 1.324% | 2.711% | 13.92% | 11.07% | 4.483% | 2.849%
46 | 8_2_21_F_8_DUP |  | 6/8/2023 5:55:36 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_F_8_DUP.400 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.071% | 3.861% | 2.554% | 7.26% | 1.206% | 6.572% | 3.352% | 4.812% | 4.242% | 17.36% | 5.476% | 3.114% | 0.9213% | 3.809% | 12.03% | 8.776% | 1.979% | 0.8034% | 0.4836% | 1.65% | 9.114% | 1.147% | 1.147% | 5.587% | 33.99% | 19.09% | 3.482% | 9.791% | 33.33% | 2.117% | 2.25% | 1.061% | 1.065% | 16.67% | 0.6848% | 0.3242% | 1.959% | 1.914% | 6.216% | 9.841% | 2.457% | 0.1623%
47 | 8_2_21_F_8_DUP_LD |  | 6/8/2023 6:01:24 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_F_8_DUP_LD.401 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.691% | 1.617% | 0.7823% | 14.75% | 1.217% | 2.659% | 1.946% | 2.399% | 5.376% | 7.462% | 3.246% | 0.6385% | 0.8061% | 0.85% | 17.89% | 6.151% | 3.564% | 1.206% | 1.475% | 1.711% | 4.307% | 2.417% | 2.417% | 1.716% | 16.37% | 19.23% | 18.78% | 2.428% | 35.49% | 1.681% | 1.699% | 1.757% | 1.732% | 0.4066% | 7.31% | 1.44% | 1.493% | 2.337% | 17.44% | 13.98% | 1.81% | 1.899%
48 | 8_2_21_F_8_DUP_LD |  | 6/8/2023 6:07:40 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_F_8_DUP_LD.402 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.136% | 1.962% | 2.047% | 6.931% | 1.095% | 3.092% | 5.97% | 3.021% | 1.236% | 20.08% | 4.515% | 0.501% | 1.663% | 4.156% | 8.847% | 4.367% | 1.028% | 0.716% | 1.734% | 1.435% | 5.045% | 0.1765% | 0.1765% | 10.45% | 36.64% | 32.74% | 8.04% | 4.589% | 22.91% | 0.4906% | 1.49% | 0.4498% | 0.5478% | 1.953% | 2.404% | 0.8697% | 0.7976% | 0.9844% | 17.79% | 5.56% | 2.787% | 2.173%
49 | 8_2_21_F_8_DUP_LSP |  | 6/8/2023 6:13:28 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_F_8_DUP_LSP.403 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.125% | 0.8457% | 0.2561% | 6.396% | 1.318% | 1.724% | 0.4745% | 1.375% | 2.331% | 0.7047% | 1.259% | 0.4406% | 1.047% | 3.504% | 0.7851% | 1.559% | 1.841% | 0.3977% | 0.8581% | 0.4009% | 3.96% | 0.5302% | 0.5302% | 1.166% | 3.952% | 1.509% | 2.261% | 2.733% | 0.8427% | 0.5543% | 0.5571% | 1.162% | 0.03388% | 0.7448% | 0.4913% | 0.8345% | 1.718% | 2.177% | 2.158% | 1.21% | 1.942% | 1.801%
50 | LBLANK |  | 6/8/2023 6:19:16 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.404 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 36.74% | 43.3% | 0.5605% | 100% | 3.922% | 4.555% | 57.28% |  | 173.2% | 10.19% | 31.49% | 3.39% | 27.74% | 4.379% | 100% | 30.7% | 30.12% | 9.658% | 5.588% | 6.472% | 86.6% | 11.67% | 11.67% | 7.004% |  | 86.6% | 69.28% | 29.45% | 0% | 15.94% | 8.217% | 0.4256% | 0.915% | 0.2225% | 1.344% | 2.11% | 0.9712% | 2.467% | 100% | 39.16% | 7.531% | 8.579%
51 | CCV 1ppb TRACE |  | 6/8/2023 6:25:05 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.405 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.971% | 9.443% | 1.441% | 20.26% | 1.488% | 5.126% | 3.959% | 8.389% | 6.87% | 2.476% | 2.377% | 1.363% | 2.377% | 4.202% | 2.626% | 8.58% | 2.587% | 2.448% | 1.175% | 8.079% | 20.69% | 2.477% | 2.477% | 2.329% | 4.696% | 3.844% | 1.594% | 14.53% | 4.409% | 1.069% | 1.123% | 1.242% | 0.9058% | 3.012% | 2.135% | 1.906% | 1.345% | 4.724% | 1.839% | 1.178% | 0.8062% | 2.03%
52 | CCV 1ppb TRACE |  | 6/8/2023 6:31:23 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.406 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.823% | 4.146% | 3.943% | 13.39% | 3.272% | 7.117% | 6.493% | 14.92% | 3.57% | 3.364% | 3.616% | 0.9452% | 0.3% | 1.345% | 2.114% | 2.942% | 3.431% | 1.327% | 2.209% | 15.78% | 5.852% | 2.787% | 2.787% | 6.122% | 7.317% | 2.05% | 7.564% | 11.27% | 2.838% | 0.9743% | 1.717% | 6.156% | 1.851% | 1.638% | 0.8231% | 1.744% | 0.9822% | 0.5946% | 1.29% | 1.529% | 0.2706% | 0.9012%
53 | LFB |  | 6/8/2023 6:37:11 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.407 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.04% | 0.6165% | 2.292% | 5.484% | 1.008% | 2.74% | 2.885% | 3.14% | 1.9% | 2.9% | 1.643% | 0.9771% | 1.576% | 2.066% | 0.4818% | 0.5531% | 2.781% | 0.9086% | 0.9863% | 1.984% | 4.216% | 0.9994% | 0.9994% | 1.256% | 4.383% | 4.01% | 3.261% | 1.333% | 2.198% | 1.332% | 0.9408% | 0.9397% | 1.322% | 0.8294% | 0.7835% | 1.203% | 1.693% | 2.628% | 0.414% | 0.5435% | 1.422% | 1.542%
54 | CCV 10 ppb MAJORS |  | 6/8/2023 6:42:59 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.408 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.464% | 1.97% | 0.6447% | 11.42% | 1.645% | 5.118% | 1.647% | 1.927% | 1.5% | 0.7775% | 0.8713% | 1.18% | 0.9861% | 1.786% | 0.5727% | 2.964% | 0.9404% | 1.063% | 0.05468% | 0.3174% | 4.347% | 1.023% | 1.023% | 2.21% | 4.392% | 0.5383% | 1.524% | 4.228% | 1.418% | 1.099% | 1.749% | 1.526% | 0.3053% | 3.025% | 0.2478% | 0.4219% | 1.176% | 2.545% | 0.9846% | 0.4603% | 0.8232% | 1.305%
55 | CCV 10 ppb MAJORS |  | 6/8/2023 6:49:21 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.409 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.326% | 2.485% | 5.939% | 6.819% | 1.04% | 1.375% | 1.978% | 3.687% | 2.485% | 0.2842% | 2.976% | 0.8386% | 0.7901% | 0.56% | 0.7615% | 0.5188% | 0.2008% | 0.9491% | 0.8927% | 5.116% | 1.96% | 0.4644% | 0.4644% | 3.222% | 2.089% | 2.622% | 1.056% | 5.857% | 2.774% | 1.226% | 2.06% | 12.12% | 1.358% | 5.813% | 0.5542% | 0.6777% | 0.8079% | 1.634% | 1.051% | 0.1532% | 3.808% | 2.886%
56 | LBLANK |  | 6/8/2023 7:26:31 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.410 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 100% |  | 2.675% | 0% | 2.38% | 7.598% | 28.39% |  | 173.2% | 62.45% | 8.439% | 5.108% | 36.74% | 2.773% | 173.2% | 24.12% | 27.15% | 15.11% | 17.3% | 8.807% | 173.2% | 34.51% | 34.51% | 7.669% | 173.2% | 91.65% | 100% | 55.03% |  | 19.6% | 24.3% | 2.344% | 1.316% | 0.7243% | 1.651% | 0.5655% | 0.742% | 4.12% | 100% | 45.81% | 10% | 31.92%
57 | PE ICV |  | 6/8/2023 7:32:19 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\PE ICV.411 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.797% | 2.661% | 1.916% | 5.204% | 0.6929% | 1.211% | 0.5165% | 5.096% | 1.72% | 1.186% | 0.2746% | 3.12% | 1.115% | 1.345% | 0.3465% | 3.356% | 1.02% | 1.014% | 0.7386% | 2.281% | 2.713% | 1.39% | 1.39% | 2.896% | 4.226% | 0.9558% | 2.069% | 2.114% | 3.489% | 0.7157% | 1.687% | 0.6745% | 1.012% | 0.3535% | 0.5433% | 1.359% | 0.9893% | 5.248% | 1.949% | 0.6747% | 1.041% | 1.182%
58 | 8_2_21_F_9 |  | 6/8/2023 7:38:07 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_F_9.412 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.854% | 1.133% | 1.536% | 6.02% | 0.6225% | 1.427% | 2.077% | 1.213% | 5.162% | 7.859% | 1.823% | 2.696% | 1.166% | 1.614% | 13.32% | 1.205% | 1.877% | 2.039% | 0.7745% | 2.093% | 7.108% | 0.3524% | 0.3524% | 4.442% | 53.29% | 9.623% | 27.61% | 28.19% | 20.83% | 0.4683% | 1.191% | 0.8986% | 1.027% | 2.668% | 0.5173% | 1.548% | 0.4929% | 4.268% | 24.16% | 10.03% | 1.543% | 1.345%
59 | 8_2_21_F_10 |  | 6/8/2023 7:43:55 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_F_10.413 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.927% | 1.345% | 1.803% | 9.938% | 1.878% | 3.025% | 2.185% | 2.425% | 2.127% | 11.85% | 11.5% | 1.758% | 0.9291% | 3.145% | 6.522% | 12.26% | 0.8549% | 1.562% | 0.4438% | 0.89% | 5.352% | 0.4782% | 0.4782% | 2.191% | 60% | 21.05% | 14.5% | 14.09% | 24.35% | 0.3838% | 1.495% | 1.2% | 0.2992% | 8.127% | 0.7232% | 0.7729% | 1.136% | 1.7% | 28.91% | 11.03% | 2.487% | 2.527%
60 | 8_2_21_F_10_DUP |  | 6/8/2023 7:49:43 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_F_10_DUP.414 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.557% | 1.447% | 2.418% | 1.493% | 0.6804% | 2.47% | 2.917% | 2.253% | 3.662% | 5.103% | 1.553% | 2.186% | 1.865% | 1.231% | 22.78% | 3.115% | 1.015% | 1.311% | 1.059% | 1.134% | 4.444% | 0.1756% | 0.1756% | 2.329% | 22.91% | 24.66% | 11.76% | 23.97% | 43.53% | 1.487% | 0.7384% | 0.7766% | 7.909% | 1.151% | 0.4265% | 1.137% | 1.019% | 8.433% | 31.99% | 5.648% | 0.8319% | 1.569%
61 | 8_2_21_F_11 |  | 6/8/2023 7:55:32 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_F_11.415 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4.913% | 3.525% | 1.376% | 9.22% | 2.027% | 0.1259% | 8.97% | 1.158% | 4.733% | 5.12% | 7.281% | 1.965% | 0.3984% | 0.8309% | 21.54% | 11.22% | 1.328% | 0.8915% | 2.158% | 2.038% | 3.664% | 0.923% | 0.923% | 3.874% | 43.64% | 21.41% | 10.33% | 18.38% | 32.83% | 0.5595% | 0.9174% | 1.377% | 0.4379% | 0.7896% | 1.174% | 3.282% | 2.071% | 1.158% | 34.91% | 12.09% | 0.6711% | 2.311%
62 | 8_2_21_F_11_SP |  | 6/8/2023 8:01:20 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_F_11_SP.416 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.16% | 0.6466% | 0.54% | 3.312% | 1.584% | 0.8581% | 1.675% | 2.072% | 1.865% | 0.2915% | 1.051% | 1.936% | 1.165% | 1.926% | 0.6936% | 1.518% | 1.847% | 1.62% | 1.239% | 0.7753% | 1.847% | 0.4812% | 0.4812% | 0.6531% | 2.774% | 2.846% | 1.892% | 2.952% | 3.416% | 1.632% | 0.7652% | 2.49% | 1.337% | 2.096% | 0.2242% | 0.5709% | 1.169% | 1.045% | 2.652% | 1.168% | 0.2945% | 1.07%
63 | 8_2_21_W_1 |  | 6/8/2023 8:07:08 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_W_1.417 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.538% | 2.293% | 0.7824% | 6.217% | 2.107% | 4.227% | 6.184% | 4.6% | 6.013% | 6.667% | 4.802% | 1.207% | 0.6358% | 0.9591% | 27.71% | 20.5% | 4.195% | 0.6558% | 1.758% | 9.539% | 24.62% | 2.237% | 2.237% | 7.59% | 89.21% | 25% | 36.33% | 22.06% | 43.3% | 2.132% | 2.416% | 0.2326% | 1.946% | 0.5275% | 0.8372% | 0.6291% | 1.08% | 1.133% | 26.96% | 14.39% | 4.358% | 2.238%
64 | 8_2_21_W_6 |  | 6/8/2023 8:12:56 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_W_6.418 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 8.94% | 4.356% | 1.793% | 5.008% | 2.846% | 1.87% | 6.358% | 5.019% | 4.337% | 23.6% | 4.644% | 0.7754% | 0.4685% | 1.579% | 28.25% | 23.2% | 1.911% | 1.68% | 0.9179% | 5.189% | 14.18% | 3.141% | 3.141% | 2.422% | 57.28% | 44.1% | 11.18% | 56.15% | 45.83% | 1.862% | 2.294% | 0.8327% | 1.807% | 2.276% | 1.053% | 1.456% | 1.445% | 1.505% | 27.96% | 13.8% | 2.638% | 2.082%
65 | 8_2_21_W_6_DUP |  | 6/8/2023 8:18:44 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_W_6_DUP.419 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.7242% | 3.961% | 2.494% | 7.157% | 2.256% | 3.431% | 8.841% | 2.139% | 5.022% | 2.857% | 3.566% | 1.237% | 1.021% | 2.66% | 39.67% | 2.703% | 2.537% | 0.9068% | 3.182% | 2.692% | 23.17% | 0.8811% | 0.8811% | 13.64% | 173.2% | 17.63% | 27.73% | 40.35% | 86.6% | 2.796% | 2.074% | 2.209% | 0.3325% | 1.169% | 2.007% | 0.5503% | 0.5122% | 2.329% | 5.728% | 14.1% | 2.874% | 2.119%
66 | LBLANK |  | 6/8/2023 8:24:33 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.420 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 173.2% |  | 10.13% | 86.6% | 2.915% | 4.965% | 26.65% |  | 173.2% | 18.23% | 2.936% | 9.078% | 32.83% | 4.731% | 173.2% | 75.5% | 80% | 10.66% | 6.047% | 2.372% | 173.2% | 34.64% | 34.64% | 2.413% | 173.2% | 86.6% | 57.28% | 61.75% |  | 43.14% | 74.18% | 1.089% | 1.463% | 2.338% | 1.806% | 0.3488% | 1.454% | 1.954% | 100% | 34.32% | 43.3% | 5.871%
67 | CCV 1ppb TRACE |  | 6/8/2023 8:30:21 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.421 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4.75% | 1.712% | 1.368% | 27.97% | 1.476% | 5.155% | 4.896% | 9.79% | 4.929% | 0.991% | 0.4364% | 1.032% | 1.85% | 2.74% | 1.02% | 2.973% | 3.47% | 0.7428% | 2.689% | 6.279% | 1.786% | 0.4283% | 0.4283% | 3.397% | 12.72% | 7.844% | 8.25% | 4.032% | 2.817% | 1.768% | 0.7592% | 4.546% | 1.629% | 2.687% | 0.3267% | 0.4631% | 0.595% | 0.7419% | 4.023% | 0.6237% | 2.476% | 0.9011%
68 | CCV 1ppb TRACE |  | 6/8/2023 8:36:39 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.422 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.88% | 1.407% | 1.342% | 9.245% | 0.5277% | 8.435% | 4.167% | 6.158% | 1.484% | 4.4% | 4.024% | 1.586% | 0.6204% | 1.379% | 3.676% | 5.7% | 1.888% | 1.219% | 0.9618% | 8.007% | 11.74% | 1.554% | 1.554% | 5.355% | 8.331% | 4.311% | 2.029% | 6.247% | 2.569% | 1.78% | 0.8094% | 0.8958% | 1.013% | 1.382% | 0.6356% | 1.164% | 2.832% | 3.492% | 2.17% | 1.471% | 2.267% | 1.594%
69 | CCV 10 ppb MAJORS |  | 6/8/2023 8:42:27 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.423 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2% | 2.549% | 1.289% | 10.72% | 2.031% | 1.413% | 1.584% | 1.941% | 0.9055% | 2.153% | 1.135% | 3.044% | 0.9615% | 2.341% | 3.665% | 1.168% | 1.763% | 0.3947% | 0.8258% | 1.9% | 4.353% | 1.42% | 1.42% | 0.4231% | 3.394% | 2.579% | 1.519% | 1.948% | 3.365% | 1.46% | 1.148% | 1.809% | 0.6531% | 8.73% | 0.5137% | 1.389% | 1.196% | 1.013% | 1.813% | 0.768% | 0.3004% | 1.841%
70 | CCV 10 ppb MAJORS |  | 6/8/2023 8:48:44 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.424 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.8657% | 1.387% | 0.6518% | 3.138% | 1.413% | 1.261% | 0.9771% | 1.959% | 3.02% | 1.933% | 0.5989% | 2.483% | 1.011% | 2.993% | 1.233% | 2.546% | 1.209% | 0.6165% | 1.172% | 1.482% | 5.333% | 0.6211% | 0.6211% | 0.2169% | 4.36% | 0.8268% | 2.142% | 2.834% | 1.843% | 0.9022% | 1.201% | 2.857% | 0.5105% | 1.842% | 1.205% | 6.545% | 1.067% | 0.9986% | 0.2125% | 0.5842% | 1.235% | 4.49%
71 | 8_2_21_W_8 |  | 6/8/2023 8:54:33 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_W_8.425 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.476% | 4.659% | 2.378% | 3.393% | 0.9997% | 1.575% | 2.572% | 4.309% | 8.937% | 10.98% | 1.735% | 0.8646% | 1.332% | 1.523% | 3.535% | 21.03% | 0.505% | 1.52% | 1.006% | 2.588% | 9.939% | 0.4242% | 0.4242% | 4.11% | 34.64% | 38.49% | 24.98% | 34.67% | 21.65% | 0.3697% | 1.764% | 0.5854% | 0.6998% | 2.047% | 4.571% | 0.5382% | 1.08% | 3.225% | 38.57% | 19.1% | 3.253% | 2.21%
72 | 8_2_21_W_9 |  | 6/8/2023 9:00:21 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_W_9.426 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.189% | 3.931% | 1.921% | 9.24% | 2.435% | 5.396% | 10.53% | 2.485% | 2.97% | 11.5% | 1.219% | 1.592% | 1.835% | 1.011% | 19.87% | 4.33% | 4.355% | 0.9704% | 1.468% | 8.766% | 10.2% | 2.395% | 2.395% | 6.882% | 50% | 53.91% | 59.01% | 74.45% | 100% | 0.3872% | 1.181% | 0.9821% | 0.7714% | 2.336% | 2.308% | 2.456% | 1.057% | 2.274% | 20.83% | 19.44% | 2.395% | 0.7944%
73 | 8_2_21_W_9_DUP |  | 6/8/2023 9:06:09 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_W_9_DUP.427 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.636% | 2.939% | 0.8435% | 8.053% | 0.7079% | 8.936% | 3.268% | 4.369% | 15.85% | 21.05% | 22% | 0.1668% | 1.033% | 0.5874% | 12.37% | 19.44% | 2.13% | 1.689% | 1.337% | 7.309% | 17% | 2.757% | 2.757% | 3.389% | 86.6% | 9.116% | 33.99% | 57.33% | 100% | 0.6672% | 1.673% | 2.159% | 1.14% | 0.4551% | 0.1802% | 0.19% | 0.8813% | 5.538% | 27.55% | 22.68% | 8.133% | 2.707%
74 | 8_2_21_W_10 |  | 6/8/2023 9:11:57 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_W_10.428 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 5.869% | 2.89% | 0.5931% | 6.668% | 3.895% | 5.188% | 7.464% | 1.601% | 5.201% | 28.65% | 3.61% | 0.8684% | 1.146% | 2.894% | 4.33% | 27.37% | 6.118% | 1.578% | 1.305% | 2.704% | 13.73% | 2.732% | 2.732% | 3.745% | 65.47% | 53.29% | 13.29% | 36.15% | 43.3% | 0.5688% | 1.734% | 1.147% | 1.297% | 1.335% | 1.083% | 0.7649% | 2.394% | 2.133% | 16.61% | 7.865% | 4.301% | 3.644%
75 | 8_2_21_W_11 |  | 6/8/2023 9:17:45 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_W_11.429 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4.838% | 2.49% | 0.1383% | 4.233% | 3.748% | 9.084% | 8.512% | 3.357% | 3.441% | 8.794% | 10.89% | 1.927% | 1.351% | 0.9513% | 20.12% | 15.2% | 0.7342% | 1.942% | 0.9154% | 1.047% | 21.41% | 1.664% | 1.664% | 3.355% | 100% | 61.86% | 48.37% | 65.9% | 41.66% | 0.5402% | 2.8% | 0.6368% | 1.979% | 2.123% | 0.8546% | 2.034% | 1.398% | 2.777% | 15.02% | 24.42% | 1.041% | 1.984%
76 | 5_10_23_DB |  | 6/8/2023 9:23:33 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_10_23_DB.430 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 132.3% |  | 1.447% | 15.8% | 1.875% | 12.76% | 20.38% |  |  | 31.07% | 12.79% | 8.214% | 26.54% | 1.841% | 86.6% | 57.14% | 25.45% | 8.301% | 6.307% | 13.27% | 173.2% | 16.54% | 16.54% | 3.117% | 173.2% | 173.2% | 86.6% | 122.5% | 0% | 28.64% | 0% | 1.723% | 0.2643% | 1.344% | 1.887% | 0.5783% | 1.941% | 3.629% | 173.2% | 31.49% | 3.297% | 7.88%
77 | 5_10_23_MFB |  | 6/8/2023 9:29:21 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_10_23_MFB.431 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.8416% | 1.555% | 1.069% | 2.82% | 2.841% | 2.551% | 1.048% | 2.725% | 2.637% | 0.7845% | 1.19% | 2.275% | 0.2435% | 2.919% | 2.577% | 1.447% | 2.035% | 0.8562% | 0.7325% | 2.474% | 3.307% | 1.405% | 1.405% | 3.126% | 6.584% | 3.132% | 3.348% | 1.936% | 2.759% | 1.859% | 1.383% | 0.4281% | 0.5394% | 1.838% | 1.264% | 0.1646% | 1.932% | 1.84% | 0.2485% | 0.5478% | 0.9243% | 2.882%
78 | 5_10_23_T-239 |  | 6/8/2023 9:35:10 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_10_23_T-239.432 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.383% | 1.794% | 1.969% | 1.529% | 0.4292% | 6.668% | 1.605% | 1.06% | 3.206% | 3.244% | 1.985% | 0.7441% | 0.8569% | 2.54% | 2.85% | 0.9368% | 2.198% | 2.161% | 0.4287% | 1.67% | 16.08% | 1.558% | 1.558% | 3.015% | 5.422% | 5.583% | 1.905% | 5.657% | 6.328% | 0.3851% | 1.643% | 0.7626% | 1.058% | 3.463% | 0.8974% | 1.372% | 0.4926% | 2.339% | 0.5229% | 2.372% | 1.848% | 1.955%
79 | 5_10_23_DB |  | 6/8/2023 9:40:58 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_10_23_DB.433 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 66.67% |  | 0.9566% | 18.37% | 0.8956% | 8.167% | 24.12% | 0% | 173.2% | 22.91% | 33.72% | 6.686% | 5.413% | 2.067% | 86.6% | 8.66% | 9.84% | 9.609% | 6.233% | 9.86% | 86.6% | 18.9% | 18.9% | 1.719% |  | 173.2% | 75.5% | 46.98% | 173.2% | 4.225% | 68.88% | 3.917% | 2.06% | 3.543% | 0.8157% | 2.021% | 1.893% | 2.819% | 43.3% | 83.32% | 12.9% | 7.792%
80 | 8_16_21_UF_1 |  | 6/8/2023 9:46:46 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_UF_1.434 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4.516% | 2.82% | 2.39% | 9.391% | 1.245% | 3.564% | 3.762% | 2.178% | 3.873% | 15.4% | 10.34% | 2.447% | 0.7146% | 1.938% | 12.65% | 10.34% | 5.633% | 0.08954% | 2.62% | 1.461% | 8.11% | 0.926% | 0.926% | 1.968% | 62.98% | 33.8% | 23.67% | 2.492% | 20.38% | 1.272% | 1.183% | 2.166% | 0.2769% | 3.495% | 0.8223% | 1.747% | 0.8733% | 1.835% | 4.028% | 6.234% | 4.188% | 3.376%
81 | LBLANK |  | 6/8/2023 9:52:35 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.435 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 86.6% |  | 6.392% | 114.6% | 2.692% | 5.21% | 17.32% | 86.6% | 86.6% | 44.1% | 23.6% | 0.353% | 28.57% | 4% | 173.2% | 7.873% | 10.19% | 11.01% | 10.52% | 3.051% | 173.2% | 7.189% | 7.189% | 3.61% | 86.6% | 91.65% | 91.65% | 175.9% | 100% | 6.454% | 4.545% | 1.986% | 0.6458% | 1.782% | 1.358% | 0.7688% | 2.472% | 3.367% |  | 32.73% | 20.83% | 7.901%
82 | CCV 1ppb TRACE |  | 6/8/2023 9:58:24 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.436 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 5.67% | 6.475% | 1.433% | 12.38% | 5.823% | 3.799% | 6.71% | 6.041% | 6.116% | 4.312% | 0.9858% | 1.777% | 0.3935% | 1.656% | 2.484% | 3.01% | 5.679% | 1.886% | 1.994% | 9.216% | 12.14% | 3.305% | 3.305% | 1.544% | 12.59% | 7.26% | 4.331% | 15.75% | 4.648% | 0.5532% | 3.402% | 1.714% | 0.365% | 2.303% | 0.6137% | 0.3899% | 1.344% | 1.048% | 3.433% | 1.554% | 0.3022% | 0.5958%
83 | CCV 1ppb TRACE |  | 6/8/2023 10:04:41 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.437 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.719% | 3.207% | 2.16% | 12.06% | 2.185% | 7.78% | 14.93% | 6.872% | 2.124% | 5.926% | 0.4151% | 0.2956% | 2.364% | 2.747% | 2.048% | 6.985% | 2.713% | 2.654% | 2.003% | 2.259% | 13.73% | 1.356% | 1.356% | 4.575% | 3.711% | 16.24% | 4.567% | 7.056% | 21.76% | 0.868% | 0.9885% | 1.331% | 1.256% | 1.775% | 1.681% | 10.01% | 3.655% | 3.007% | 0.7667% | 0.9296% | 1.37% | 1.62%
84 | LFB |  | 6/8/2023 10:10:29 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.438 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.205% | 0.9538% | 0.371% | 3.597% | 2.219% | 2.91% | 3.443% | 2.013% | 0.5611% | 0.4448% | 0.6678% | 1.149% | 0.7938% | 1.385% | 1.275% | 1.364% | 0.153% | 0.3542% | 0.6018% | 1.086% | 4.385% | 0.9334% | 0.9334% | 2.394% | 1.851% | 2.212% | 0.8443% | 1.043% | 2.286% | 0.9776% | 0.9838% | 2.722% | 1.402% | 0.7825% | 0.78% | 0.2444% | 0.838% | 0.7367% | 0.5361% | 1.33% | 0.8472% | 1.076%
85 | CCV 10 ppb MAJORS |  | 6/8/2023 10:16:17 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.439 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.966% | 0.8131% | 2.31% | 3.732% | 1.534% | 3.68% | 1.99% | 4.389% | 1.919% | 0.6595% | 0.9888% | 1.259% | 2.397% | 0.8114% | 1.195% | 1.746% | 2.305% | 1.225% | 0.9308% | 2.001% | 7.376% | 1.664% | 1.664% | 2.226% | 1.64% | 0.9766% | 1.606% | 2.721% | 2.775% | 1.614% | 2.05% | 2.14% | 0.5004% | 2.31% | 0.6927% | 1.531% | 0.5834% | 2.649% | 0.1746% | 0.3386% | 1.457% | 1.971%
86 | CCV 10 ppb MAJORS |  | 6/8/2023 10:22:34 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.440 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.8141% | 1.652% | 1.79% | 6.991% | 2.501% | 1.889% | 1.852% | 2.423% | 1.086% | 1.303% | 1.339% | 1.369% | 0.7869% | 1.859% | 1.236% | 2.987% | 2.202% | 0.8391% | 0.4853% | 1.041% | 2.382% | 0.8588% | 0.8588% | 2.101% | 5.51% | 3.659% | 2.091% | 3.175% | 3.284% | 1.384% | 0.3816% | 0.2995% | 1.173% | 1.881% | 1.318% | 1.333% | 1.045% | 2.697% | 0.1571% | 0.4689% | 0.9472% | 2.609%
87 | 8_16_21_UF_6 |  | 6/8/2023 10:28:23 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_UF_6.441 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.104% | 1.675% | 1.439% | 12.43% | 0.5194% | 3.311% | 2.54% | 1.588% | 0.3467% | 17.32% | 8.11% | 1.16% | 0.7864% | 4.45% | 33.76% | 4.617% | 4.547% | 1.025% | 1.14% | 1.858% | 10.65% | 1.321% | 1.321% | 3.608% | 74.17% | 9.623% | 5% | 8.934% | 32.87% | 0.851% | 1.029% | 0.2729% | 1.009% | 0.316% | 1.029% | 0.999% | 1.135% | 3.033% | 29.57% | 4.98% | 2.803% | 1.549%
88 | 8_16_21_UF_8 |  | 6/8/2023 10:34:11 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_UF_8.442 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.883% | 2.84% | 0.7955% | 19.56% | 1.551% | 4.134% | 5.78% | 3.256% | 3.717% | 26.32% | 2.388% | 2.269% | 1.228% | 3.696% | 21.41% | 3.041% | 3.669% | 2.435% | 1.013% | 0.4624% | 3.767% | 0.9898% | 0.9898% | 4.779% | 20.38% | 18.43% | 33.33% | 10.11% | 29.04% | 1.104% | 1.331% | 0.3269% | 0.9418% | 1.159% | 0.3786% | 1.574% | 1.443% | 1.964% | 26.38% | 9.636% | 6.161% | 2.98%
89 | 8_16_21_UF_9 |  | 6/8/2023 10:39:59 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_UF_9.443 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.7067% | 3.463% | 0.5598% | 3.125% | 1.193% | 0.964% | 6.611% | 2.697% | 4.105% | 6.454% | 2.71% | 4.308% | 0.5587% | 2.939% | 29.88% | 0.7832% | 3.128% | 1.535% | 2.985% | 1.086% | 16.36% | 0.1232% | 0.1232% | 9.337% | 34.64% | 26.96% | 11.59% | 9.96% | 90.14% | 0.6984% | 0.8966% | 0.9286% | 1.333% | 6.494% | 2.079% | 1.027% | 0.6065% | 1.406% | 23.01% | 8.98% | 1.206% | 4.015%
90 | 8_16_21_UF_10 |  | 6/8/2023 10:45:47 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_UF_10.444 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4.377% | 1.435% | 2.786% | 14.24% | 0.7262% | 0.4408% | 3.013% | 3.401% | 7.44% | 15.91% | 2.291% | 2.664% | 1.106% | 2.101% | 13.04% | 3.226% | 2.58% | 2.085% | 3.315% | 0.9647% | 10.79% | 0.4084% | 0.4084% | 5.346% | 11.11% | 26.99% | 11.08% | 32.58% | 12.37% | 0.9727% | 0.8545% | 0.53% | 0.4721% | 4.474% | 0.6153% | 2.17% | 0.0812% | 2.17% | 10.56% | 8.816% | 4.192% | 2.179%
91 | 8_16_21_UF_11 |  | 6/8/2023 10:51:35 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_UF_11.445 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.163% | 2.926% | 2.765% | 17.83% | 2.154% | 5.166% | 4.442% | 5.278% | 3.604% | 11.78% | 3.718% | 3.779% | 1.328% | 2.87% | 6.928% | 2.839% | 1.541% | 1.466% | 1.228% | 1.165% | 8.912% | 0.2435% | 0.2435% | 0.9912% | 45.83% | 22.18% | 16.79% | 4.846% | 21.65% | 0.6825% | 0.6309% | 0.6087% | 1.42% | 8.435% | 0.9888% | 1.059% | 2.268% | 3.491% | 14.78% | 21.85% | 1.591% | 2.106%
92 | 8_16_21_F_1 |  | 6/8/2023 10:57:23 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_F_1.446 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.421% | 1.193% | 1.271% | 4.711% | 0.8915% | 1.473% | 3.669% | 2.661% | 9.954% | 9.742% | 4.906% | 1.578% | 1.191% | 2.51% | 14.47% | 4.091% | 3.542% | 0.4866% | 1.751% | 1.591% | 2.451% | 0.6469% | 0.6469% | 3.549% | 13.32% | 14.19% | 17.38% | 13.96% | 36.46% | 1.123% | 1.135% | 1.826% | 2.627% | 1.547% | 2.628% | 1.181% | 4.479% | 6.253% | 11.11% | 5.871% | 1.901% | 0.7519%
93 | 8_16_21_F_1_DUP |  | 6/8/2023 11:03:11 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_F_1_DUP.447 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.7633% | 1.588% | 1.313% | 8.163% | 0.6629% | 2.98% | 4.248% | 1.218% | 8.826% | 15.44% | 6.722% | 1.769% | 1.558% | 1.289% | 2.389% | 12.32% | 0.3189% | 1.098% | 1.266% | 0.5831% | 6.223% | 1.159% | 1.159% | 6.777% | 13.32% | 16.54% | 22.05% | 9.823% | 47.19% | 0.2946% | 0.4407% | 0.4696% | 0.4723% | 1.549% | 0.1372% | 0.9902% | 2.864% | 2.161% | 2.373% | 22.02% | 8.27% | 6.484%
94 | 8_16_21_F_1_DUP_LD |  | 6/8/2023 11:08:58 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_F_1_DUP_LD.448 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.825% | 3.305% | 1.582% | 13.29% | 1.484% | 0.9854% | 1.103% | 0.4309% | 6.168% | 12.47% | 5.474% | 3.708% | 1.323% | 1.185% | 25.56% | 9.684% | 3.284% | 1.478% | 2.158% | 1.472% | 10.74% | 2.012% | 2.012% | 4.841% | 15.75% | 17.84% | 9.123% | 17.76% | 34.64% | 0.5164% | 1.207% | 0.3947% | 0.9927% | 6.076% | 1.261% | 0.9871% | 0.8203% | 1.144% | 19.61% | 14.45% | 2.753% | 5.835%
95 | 8_16_21_F_1_DUP_LD |  | 6/8/2023 11:15:14 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_F_1_DUP_LD.449 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.934% | 2.299% | 1.346% | 5.268% | 0.2906% | 5.233% | 0.8349% | 0.9268% | 6.303% | 27% | 5.486% | 0.8446% | 1.527% | 2.257% | 28.82% | 3.881% | 5.143% | 1.373% | 3.217% | 0.4269% | 2.64% | 0.3187% | 0.3187% | 6.591% | 31.49% | 35.92% | 7.873% | 13% | 36.66% | 0.614% | 1.184% | 0.09173% | 1.145% | 3.099% | 2.351% | 0.7248% | 0.5667% | 0.773% | 11.71% | 9.384% | 2.529% | 3.55%
96 | 8_16_21_F_1_DUP_LSP |  | 6/8/2023 11:21:03 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_F_1_DUP_LSP.450 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.9437% | 2.032% | 0.9235% | 3.01% | 0.1781% | 1.797% | 0.8961% | 1.241% | 1.464% | 0.9353% | 0.7885% | 3.266% | 1.223% | 0.7356% | 0.7753% | 1.062% | 1.824% | 0.6584% | 1.525% | 1.172% | 1.36% | 0.6162% | 0.6162% | 2.292% | 4.965% | 2.016% | 1.016% | 2.39% | 7.16% | 0.6128% | 0.8731% | 0.7187% | 0.7526% | 1.516% | 1.123% | 1.947% | 1.145% | 0.8549% | 1.355% | 1.959% | 0.1747% | 1.887%
97 | LBLANK |  | 6/8/2023 11:26:52 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.451 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 66.67% | 173.2% | 3.196% | 173.2% | 8.284% | 6.86% | 12.5% | 173.2% | 173.2% | 48.04% | 23.68% | 1.786% | 29.86% | 2.609% | 43.3% | 26.96% | 44.44% | 4.493% | 5.374% | 5.366% | 173.2% | 12.05% | 12.05% | 9.099% | 86.6% | 43.3% | 25% | 43.37% | 100% | 14.21% | 8.66% | 9.862% | 1.758% | 1.371% | 0.7345% | 0.6461% | 0.6941% | 2.442% | 173.2% | 27.94% | 13.32% | 19.58%
98 | CCV 1ppb TRACE |  | 6/8/2023 11:32:40 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.452 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.669% | 4.383% | 1.63% | 28.58% | 2.628% | 3.142% | 3.732% | 10.99% | 0.5726% | 6.063% | 1.302% | 2.349% | 0.4283% | 7.777% | 4.269% | 9.129% | 1.627% | 1.739% | 1.849% | 5.517% | 12.18% | 1.22% | 1.22% | 2.08% | 23.4% | 7.25% | 6.738% | 9.029% | 7.33% | 0.5292% | 2.069% | 1.712% | 1.207% | 1.071% | 1.81% | 1.001% | 1.833% | 2.31% | 1.036% | 2.528% | 1.374% | 1.306%
99 | CCV 1ppb TRACE |  | 6/8/2023 11:38:58 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.453 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.185% | 3.289% | 2.783% | 19.1% | 3.644% | 5.358% | 0.1407% | 9.512% | 2.263% | 2.676% | 3.534% | 0.7063% | 2.239% | 3.053% | 1.155% | 4.333% | 2.137% | 0.9058% | 2.088% | 8.088% | 21.17% | 1.695% | 1.695% | 5.251% | 9.542% | 4.241% | 7.277% | 2.837% | 2.054% | 0.7333% | 0.2512% | 1.523% | 0.9683% | 1.687% | 1.931% | 0.5873% | 1.156% | 1.603% | 1.472% | 1.106% | 1.937% | 1.416%
100 | CCV 10 ppb MAJORS |  | 6/8/2023 11:44:46 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.454 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.956% | 3.42% | 1.404% | 1.323% | 1.75% | 2.485% | 1.425% | 2.597% | 0.8568% | 1.306% | 1.319% | 1.38% | 1.522% | 1.072% | 0.8823% | 0.7742% | 0.1509% | 0.9572% | 0.628% | 0.8742% | 6.193% | 0.8808% | 0.8808% | 2.496% | 4.273% | 1.045% | 4.024% | 1.909% | 1.303% | 0.445% | 1.301% | 1.337% | 0.7502% | 1.837% | 1.87% | 0.8617% | 1.188% | 3.266% | 1.052% | 0.688% | 0.7626% | 0.7889%
101 | CCV 10 ppb MAJORS |  | 6/8/2023 11:51:04 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.455 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.899% | 1.018% | 0.9716% | 4.569% | 0.7716% | 2.71% | 2.498% | 2.879% | 2.19% | 1.587% | 0.8965% | 1.365% | 0.5713% | 1.014% | 1.2% | 2.583% | 1.438% | 0.5202% | 0.7631% | 0.473% | 3.309% | 1.409% | 1.409% | 1.583% | 4.808% | 3.739% | 0.9232% | 1.772% | 3.842% | 0.9116% | 1.306% | 1.83% | 0.8537% | 4.35% | 1.445% | 1.142% | 0.9676% | 0.6513% | 0.9203% | 1.51% | 0.6759% | 1.331%
102 | 8_16_21_F_6 |  | 6/8/2023 11:56:52 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_F_6.456 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.715% | 1.426% | 0.3871% | 7.482% | 0.2551% | 1.857% | 6.281% | 2.24% | 4.631% | 7.37% | 7.433% | 1.287% | 0.6596% | 3.211% | 16.41% | 3.952% | 1.357% | 2.304% | 1.786% | 0.2017% | 4.887% | 0.6166% | 0.6166% | 3.787% | 24.74% | 44.1% | 6.522% | 9.354% | 16.67% | 1.138% | 0.8501% | 4.12% | 1.066% | 0.5566% | 1.839% | 0.4059% | 0.9384% | 3.185% | 3.808% | 5.003% | 1.368% | 1.199%
103 | 8_16_21_F_6_SP |  | 6/9/2023 12:02:41 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_F_6_SP.457 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.132% | 1.58% | 2.295% | 13.49% | 1.213% | 1.942% | 0.8133% | 1.155% | 2.697% | 0.5644% | 0.6853% | 0.3926% | 0.7426% | 2.066% | 0.334% | 2.965% | 3.17% | 0.5545% | 1.576% | 0.8274% | 4.076% | 0.464% | 0.464% | 1.469% | 7.755% | 4.641% | 1.446% | 2.732% | 6.676% | 1.033% | 0.6432% | 0.9558% | 1.787% | 3.645% | 1.124% | 0.8169% | 1.191% | 1.719% | 1.485% | 1.486% | 0.8023% | 0.1661%
104 | 8_16_21_F_8 |  | 6/9/2023 12:08:29 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_F_8.458 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.9097% | 2.726% | 2.526% | 4.97% | 2.914% | 1.321% | 5.05% | 1.049% | 2.9% | 17.71% | 4.4% | 0.6309% | 2.394% | 1.602% | 22.79% | 11.37% | 1.891% | 0.8822% | 0.372% | 1.359% | 5.309% | 0.6859% | 0.6859% | 6.514% | 100% | 5.728% | 9.999% | 7.533% | 18.23% | 0.5396% | 0.5861% | 0.7318% | 1.297% | 4.947% | 1.349% | 1.596% | 0.583% | 1.253% | 12.88% | 6.524% | 2.21% | 1.775%
105 | 8_16_21_F_8_DUP |  | 6/9/2023 12:14:18 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_F_8_DUP.459 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.6378% | 3.959% | 1.863% | 3.719% | 0.9483% | 1.386% | 5.816% | 1.472% | 3.319% | 10.32% | 3.936% | 0.775% | 0.2701% | 1.622% | 3.696% | 8.907% | 1.44% | 0.3247% | 0.7505% | 0.9579% | 7.015% | 1.547% | 1.547% | 3.143% | 32.83% | 18.65% | 8.873% | 10.98% | 39.16% | 0.4007% | 1.212% | 0.7788% | 0.937% | 1.43% | 1.107% | 2.573% | 1.037% | 2.029% | 4.461% | 4.625% | 2.179% | 3.458%
106 | 8_16_21_F_9 |  | 6/9/2023 12:20:06 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_F_9.460 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.3475% | 0.8477% | 0.8819% | 2.894% | 1.1% | 3.563% | 1.886% | 3.413% | 1.843% | 8.139% | 1.479% | 1.496% | 0.2837% | 1.637% | 27.36% | 1.961% | 1.654% | 0.4919% | 2.059% | 1.402% | 9.717% | 1.694% | 1.694% | 1.639% | 6.928% | 10.53% | 15.38% | 7.427% | 45.43% | 1.407% | 1.103% | 1.521% | 0.5233% | 0.8233% | 0.204% | 0.2324% | 1.431% | 3.191% | 19.44% | 17.63% | 3.755% | 1.838%
107 | 8_16_21_F_10 |  | 6/9/2023 12:25:54 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_F_10.461 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.718% | 1.714% | 1.172% | 5.746% | 1.844% | 2.049% | 3.985% | 1.759% | 3.475% | 15.29% | 6.72% | 0.9575% | 1.717% | 2.405% | 23.39% | 7.825% | 1.142% | 1.374% | 0.2777% | 0.9455% | 8.555% | 0.279% | 0.279% | 7.223% | 32.73% | 11.95% | 33.38% | 15.05% | 30.12% | 2.001% | 0.7587% | 0.358% | 2.012% | 1.344% | 1.583% | 1.291% | 0.6617% | 3.59% | 10.19% | 15.48% | 0.9212% | 1.839%
108 | 8_16_21_F_10_DUP |  | 6/9/2023 12:31:42 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_F_10_DUP.462 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.146% | 1.762% | 1.151% | 19.07% | 0.2909% | 2.103% | 4.157% | 3.066% | 6.926% | 22.97% | 2.134% | 1.454% | 0.7784% | 2.246% | 4.225% | 9.609% | 0.6599% | 0.4687% | 1.396% | 1.518% | 7.382% | 1.597% | 1.597% | 2.872% | 53.29% | 11.54% | 3.851% | 14.43% | 22.91% | 2.073% | 1.757% | 0.8373% | 4.53% | 1.341% | 3.78% | 1.365% | 2.745% | 4.607% | 3.464% | 8.379% | 1.816% | 2.426%
109 | 8_16_21_F_11 |  | 6/9/2023 12:37:30 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_F_11.463 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.061% | 2.787% | 2.143% | 5.015% | 2.263% | 3.079% | 1.809% | 1.888% | 4.474% | 9.245% | 9.196% | 1.173% | 1.447% | 1.291% | 8.837% | 10.35% | 1.844% | 1.322% | 0.5501% | 2.079% | 7.006% | 0.8314% | 0.8314% | 3.799% | 53.29% | 37.5% | 6.585% | 4.132% | 24.74% | 0.8398% | 1.478% | 1.65% | 1.315% | 4.767% | 1.686% | 0.7693% | 1.186% | 0.6053% | 9.962% | 23.09% | 0.876% | 0.4115%
110 | 8_16_21_W_1 |  | 6/9/2023 12:43:18 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_W_1.464 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4.661% | 4.268% | 0.4051% | 6.638% | 0.2116% | 1.751% | 4.797% | 2.216% | 19.42% | 29.06% | 7.616% | 1.723% | 0.5402% | 4.123% | 42.48% | 4.055% | 3.709% | 0.9353% | 5.238% | 4.758% | 14.01% | 1.832% | 1.832% | 4.749% | 173.2% | 28.64% | 26.96% | 30.36% | 86.6% | 1.601% | 0.5263% | 1.295% | 0.2771% | 5.899% | 10.25% | 0.9086% | 2.217% | 1.242% | 26.46% | 18.65% | 7.525% | 7.066%
111 | 8_16_21_W_6 |  | 6/9/2023 12:49:05 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_W_6.465 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.866% | 2.042% | 0.7856% | 1.343% | 0.263% | 1.511% | 7.626% | 2.829% | 7.499% | 7.217% | 13.18% | 2.382% | 0.3264% | 2.567% | 16.67% | 5.036% | 3.496% | 0.8103% | 0.8099% | 2.934% | 15.17% | 1.183% | 1.183% | 4.325% | 43.3% | 19.92% | 5.413% | 34.91% | 24.74% | 2.338% | 1.426% | 2.28% | 0.9384% | 16.91% | 1.167% | 0.8165% | 1.573% | 1.469% | 28.49% | 19.42% | 5.345% | 2.324%
112 | LBLANK |  | 6/9/2023 12:54:53 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.466 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 173.2% |  | 1.685% | 43.3% | 0.8053% | 8.419% | 50.94% |  |  | 25% | 10.53% | 8.529% | 18.33% | 1.502% | 173.2% | 48.04% | 45.58% | 4.558% | 13.21% | 7.192% |  | 37.5% | 37.5% | 8.788% |  | 86.6% | 124.9% | 205.8% | 100% | 15.61% | 26.65% | 1.964% | 1.029% | 1.548% | 1.078% | 1.524% | 2.507% | 2.19% | 173.2% | 32.87% | 30.05% | 14.32%
113 | CCV 1ppb TRACE |  | 6/9/2023 1:00:42 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.467 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.437% | 11.03% | 0.521% | 21.17% | 2.254% | 9.288% | 1.642% | 7.966% | 1.667% | 2.087% | 1.419% | 0.9827% | 1.581% | 2.097% | 1.162% | 4.454% | 0.2763% | 1.797% | 1.467% | 6.165% | 7.674% | 0.8398% | 0.8398% | 4.959% | 6.375% | 6.08% | 5.603% | 13.74% | 1.337% | 1.8% | 0.4676% | 1.548% | 0.9687% | 3.05% | 1.976% | 0.9733% | 0.9663% | 3.391% | 2.541% | 2.172% | 0.7105% | 0.04877%
114 | CCV 1ppb TRACE |  | 6/9/2023 1:06:59 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.468 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 5.52% | 7.63% | 1.267% | 13.88% | 1.223% | 5.579% | 9.945% | 10.24% | 3.619% | 3.657% | 2.735% | 4.007% | 0.803% | 2.306% | 3.905% | 1.295% | 5.327% | 1.378% | 1.53% | 7.308% | 7.723% | 2.663% | 2.663% | 2.38% | 7.616% | 2.301% | 3.03% | 7.473% | 4.949% | 0.4418% | 1.615% | 1.173% | 1.962% | 2.863% | 3.339% | 1.21% | 1.275% | 3.203% | 1.598% | 1.085% | 0.4865% | 1.216%
115 | LFB |  | 6/9/2023 1:12:47 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.469 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.278% | 1.628% | 1.665% | 12.57% | 1.935% | 3.685% | 0.7638% | 3.446% | 1.817% | 2.793% | 0.8454% | 1.093% | 1.288% | 3.021% | 2.102% | 1.713% | 1.838% | 0.4698% | 0.928% | 2.155% | 4.729% | 1.375% | 1.375% | 1.521% | 1.957% | 3.393% | 0.5904% | 1.144% | 4.422% | 1.19% | 0.6947% | 2.58% | 1.871% | 1.793% | 0.279% | 1.137% | 2.477% | 2.268% | 0.429% | 0.7035% | 2.106% | 1.401%
116 | CCV 10 ppb MAJORS |  | 6/9/2023 1:18:35 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.470 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.9914% | 3.544% | 1.62% | 14.09% | 3.026% | 2.445% | 2.601% | 2.435% | 1.61% | 0.1364% | 1.425% | 1.367% | 1.188% | 0.7575% | 0.5046% | 1.51% | 0.9319% | 0.3861% | 0.4919% | 3.276% | 4.191% | 1.157% | 1.157% | 3.141% | 2.109% | 1.44% | 2.099% | 3.931% | 4.655% | 0.1163% | 0.1772% | 1.314% | 0.7835% | 1.261% | 1.627% | 0.504% | 0.3669% | 2.207% | 0.8535% | 1.769% | 1.989% | 0.9551%
117 | CCV 10 ppb MAJORS |  | 6/9/2023 1:24:53 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.471 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.241% | 1.866% | 1.353% | 2.065% | 1.669% | 4.418% | 2.577% | 0.4601% | 1.495% | 1.492% | 0.1619% | 0.6461% | 0.9131% | 1.831% | 1.628% | 1.11% | 1.833% | 0.8018% | 0.8843% | 2.401% | 1.481% | 0.3886% | 0.3886% | 1.632% | 2.242% | 1.165% | 2.922% | 2.222% | 3.299% | 0.6405% | 0.6923% | 0.4798% | 0.7623% | 1.984% | 0.6755% | 1.288% | 1.023% | 0.9139% | 1.747% | 1.035% | 0.3949% | 3.018%
118 | 8_16_21_W_6_DUP |  | 6/9/2023 1:30:41 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_W_6_DUP.472 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.921% | 5.473% | 1.694% | 1.582% | 2.236% | 0.4153% | 10.06% | 3.642% | 11.03% | 6.439% | 3.136% | 2.179% | 0.9821% | 0.4944% | 8.646% | 31.58% | 3.919% | 1.597% | 2.684% | 3.382% | 17.68% | 2.59% | 2.59% | 5.415% | 173.2% | 65.74% | 19.87% | 18.46% | 50% | 0.7188% | 0.6546% | 0.8148% | 1.833% | 2.474% | 1.861% | 2.343% | 5.142% | 3.532% | 31.13% | 20.65% | 4.96% | 3.183%
119 | 8_16_21_W_8 |  | 6/9/2023 1:36:29 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_W_8.473 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.235% | 6.721% | 2.166% | 0.5698% | 1.41% | 2.607% | 4.459% | 5.313% | 3.218% | 20.03% | 3.09% | 0.5783% | 1.075% | 2.151% | 14.24% | 16.38% | 3.866% | 0.8614% | 1.236% | 3.648% | 14.53% | 0.7688% | 0.7688% | 1.107% | 45.83% | 62.5% | 22.79% | 18.13% | 91.65% | 1.901% | 0.1789% | 0.8363% | 0.7829% | 4.818% | 1.071% | 1.594% | 1.305% | 3.35% | 14.7% | 9.476% | 0.9823% | 1.715%
120 | 8_16_21_W_9 |  | 6/9/2023 1:42:18 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_W_9.474 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4.025% | 2.545% | 1.473% | 1.002% | 0.5912% | 2.085% | 5.59% | 1.552% | 4.919% | 10.48% | 6.222% | 0.8653% | 1.132% | 1.546% | 16.9% | 39.54% | 6.368% | 2.808% | 1.457% | 8.586% | 11.22% | 3.154% | 3.154% | 5.558% | 66.67% | 12.37% | 13.09% | 25.44% | 34.64% | 0.8518% | 2.705% | 0.6867% | 1.875% | 1.902% | 8.411% | 0.327% | 0.898% | 3.381% | 19.29% | 26.57% | 1.594% | 1.048%
121 | 8_16_21_W_9_DUP |  | 6/9/2023 1:48:06 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_W_9_DUP.475 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.044% | 2.662% | 1.345% | 5.284% | 1.799% | 7.501% | 6.952% | 1.348% | 1.553% | 17.54% | 3.571% | 1.599% | 0.6743% | 1.073% | 13.02% | 29.04% | 2.761% | 0.5384% | 3.126% | 0.3307% | 17.78% | 2.438% | 2.438% | 3.131% |  | 47.89% | 14.32% | 162.6% | 49.49% | 1.659% | 1.002% | 2.945% | 0.568% | 2.032% | 1.667% | 1.656% | 0.7095% | 3.362% | 18.37% | 11.66% | 0.9729% | 1.129%
122 | 8_16_21_W_10 |  | 6/9/2023 1:53:55 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_W_10.476 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.325% | 6.089% | 2.216% | 7.425% | 2.256% | 4.859% | 6.534% | 4.741% | 10.15% | 8.278% | 4.264% | 1.769% | 0.3925% | 1.119% | 43.64% | 13.78% | 2.542% | 1.798% | 2.895% | 2.7% | 2.794% | 1.151% | 1.151% | 7.491% | 57.28% | 66.67% | 19.29% | 38.48% | 100% | 2.147% | 0.4792% | 0.4395% | 1.074% | 0.4996% | 1.338% | 1.091% | 1.271% | 2.239% | 8.449% | 13.85% | 3.622% | 1.746%
123 | 8_16_21_W_11 |  | 6/9/2023 1:59:43 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_W_11.477 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.311% | 1.937% | 1.75% | 8.119% | 2.987% | 8.577% | 12.12% | 3.564% | 5.501% | 25.38% | 8.832% | 2.489% | 1.471% | 1.006% | 57.74% | 32.29% | 5.071% | 1.314% | 3.663% | 4.94% | 18.87% | 2.348% | 2.348% | 2.599% | 34.64% | 15.75% | 19.68% | 33.2% | 24.74% | 0.831% | 3.632% | 1.254% | 1.455% | 0.7518% | 0.8231% | 13.11% | 1.793% | 3.945% | 14.29% | 25.08% | 5.44% | 2.171%
124 | 5_10_23_DB |  | 6/9/2023 2:05:32 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_10_23_DB.478 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth |  |  | 2.889% | 38.57% | 2.003% | 10.55% | 40% |  |  | 19.92% | 25.28% | 4.023% | 24.12% | 4.03% | 86.6% | 45.83% | 17.83% | 3.819% | 13.98% | 4.061% | 173.2% | 43.3% | 43.3% | 2.652% | 173.2% | 173.2% | 43.3% | 56.66% | 86.6% | 50% | 25% | 0.6514% | 0.3468% | 4.246% | 1.196% | 0.8099% | 1.351% | 2.522% |  | 62.45% | 38.01% | 18.3%
125 | 5_10_23_MFB |  | 6/9/2023 2:11:19 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_10_23_MFB.479 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.083% | 2.742% | 1.456% | 4.668% | 0.4482% | 0.3038% | 3.751% | 6.053% | 1.36% | 2.198% | 0.6201% | 1.385% | 1.087% | 3.466% | 3.362% | 3.274% | 2.778% | 1.13% | 0.8614% | 1.025% | 4.661% | 0.5554% | 0.5554% | 2.326% | 5.728% | 2.003% | 1.011% | 2.106% | 3.101% | 0.5028% | 1.19% | 0.9238% | 1.177% | 0.4866% | 2.302% | 2.018% | 1.174% | 3.21% | 1.355% | 0.4066% | 0.8778% | 0.7907%
126 | 5_10_23_T-239 |  | 6/9/2023 2:17:07 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_10_23_T-239.480 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.297% | 2.88% | 0.7111% | 1.661% | 1.163% | 3.126% | 1.025% | 0.9701% | 1.42% | 2.776% | 0.7535% | 0.2601% | 1.026% | 0.812% | 2.332% | 5.959% | 1.873% | 0.2385% | 0.2207% | 3.345% | 14.53% | 1.393% | 1.393% | 0.3112% | 5.092% | 6.301% | 3.293% | 5.082% | 5.578% | 3.268% | 1.222% | 3.786% | 0.4834% | 1.372% | 0.2909% | 1.177% | 0.1404% | 1.484% | 0.5384% | 0.9014% | 0.4735% | 1.392%
127 | 5_10_23_DB |  | 6/9/2023 2:22:55 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_10_23_DB.481 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 65.47% | 173.2% | 2.112% | 8.449% | 0.7964% | 2.972% | 45.07% | 86.6% | 43.3% | 28.28% | 21.31% | 4.684% | 0% | 0.9128% | 173.2% | 33.33% | 5.036% | 6.322% | 8.737% | 6.907% |  | 53.91% | 53.91% | 0.6197% | 173.2% |  | 43.3% | 382.9% | 86.6% | 24.05% | 32.73% | 5.084% | 1.435% | 3.572% | 0.4082% | 0.9682% | 1.095% | 3.749% | 173.2% | 17.32% | 6.69% | 6.157%
128 | LBLANK |  | 6/9/2023 2:28:44 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.482 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 86.6% |  | 1.974% | 173.2% | 1.357% | 8.783% | 39.74% | 86.6% | 173.2% | 31.23% | 25.52% | 5.094% | 56.79% | 2.548% |  | 25% | 65.85% | 13.88% | 19.04% | 4.541% |  | 31.23% | 31.23% | 6.941% | 173.2% | 173.2% | 33.33% | 174.5% |  | 17.64% | 6.662% | 1.952% | 0.4323% | 1.415% | 1.112% | 1.548% | 0.4756% | 4.293% |  | 66.14% | 22.21% | 15.46%
129 | CCV 1ppb TRACE |  | 6/9/2023 2:34:32 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.483 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4.989% | 3.829% | 1.459% | 4.028% | 1.441% | 2.498% | 3.544% | 8.209% | 5.691% | 1.279% | 1.412% | 4.781% | 0.8348% | 2.729% | 1.344% | 1.493% | 0.9143% | 1.499% | 1.628% | 9.596% | 8.861% | 2.342% | 2.342% | 2.234% | 15.51% | 2.497% | 3.815% | 14.68% | 8.526% | 0.5625% | 1.022% | 8.779% | 1.578% | 2.068% | 1.398% | 0.705% | 0.3807% | 2.378% | 0.7154% | 1.989% | 2.341% | 0.821%
130 | CCV 1ppb TRACE |  | 6/9/2023 2:40:49 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.484 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.248% | 3.218% | 1.614% | 6.681% | 1.765% | 6.957% | 4.048% | 12.35% | 7.755% | 2.605% | 3.223% | 1.609% | 1.705% | 1.517% | 1.529% | 8.064% | 3.014% | 1.209% | 0.4979% | 1.744% | 7.531% | 0.938% | 0.938% | 1.343% | 17.59% | 10.43% | 7.151% | 7.395% | 10.52% | 2.67% | 0.4298% | 0.9302% | 2.163% | 1.547% | 1.608% | 0.2209% | 0.6941% | 5.292% | 0.4129% | 1.55% | 1.261% | 0.9917%
131 | CCV 10 ppb MAJORS |  | 6/9/2023 2:46:38 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.485 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.824% | 1.025% | 1.373% | 6.922% | 1.349% | 3.931% | 0.848% | 5.324% | 1.39% | 1.482% | 1.772% | 1.396% | 1.429% | 2.536% | 0.7421% | 2.028% | 1.393% | 1.013% | 1.065% | 0.6691% | 5.396% | 0.1708% | 0.1708% | 1.611% | 1.785% | 1.778% | 0.5509% | 1.902% | 1.832% | 1.94% | 0.45% | 1.519% | 0.8867% | 3.707% | 0.735% | 0.6177% | 1.684% | 3.675% | 0.5487% | 1.646% | 0.7981% | 1.054%
132 | CCV 10 ppb MAJORS |  | 6/9/2023 2:52:55 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.486 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.329% | 3.185% | 0.3045% | 3.179% | 0.9891% | 4.147% | 1.447% | 3.385% | 1.957% | 1.773% | 1.655% | 0.7325% | 1.026% | 0.6687% | 1.761% | 0.8154% | 2.241% | 0.6811% | 0.8469% | 1.492% | 3.016% | 1.138% | 1.138% | 1.045% | 3.346% | 3.469% | 0.6297% | 1.673% | 5.199% | 0.9133% | 0.4129% | 7.42% | 0.5892% | 0.8202% | 0.7148% | 0.689% | 0.8504% | 2.359% | 0.9854% | 0.0905% | 1.212% | 1.507%
133 | MT 34 D |  | 6/9/2023 2:58:43 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\MT 34 D.487 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 98.97% |  | 1.42% | 21.17% | 1.008% | 7.467% | 44.1% | 89.21% | 91.65% | 15.38% | 24.62% | 2.174% | 19.89% | 6.929% | 43.3% | 39.74% | 15.68% | 11.15% | 6.98% | 2.193% |  | 20.35% | 20.35% | 6.982% |  | 86.6% | 100% | 9.546% |  | 5.191% | 8.231% | 0.8334% | 1.127% | 2.42% | 1.282% | 1.715% | 2.209% | 1.242% | 173.2% | 27.73% | 5.866% | 5.719%
134 | MT 35 D |  | 6/9/2023 3:04:31 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\MT 35 D.488 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 86.6% |  | 1.429% | 6.928% | 1.148% | 4.459% | 32.73% | 173.2% | 173.2% | 23.89% | 2.083% | 5.037% | 28.97% | 3.479% | 173.2% | 26.96% | 19.92% | 6.321% | 9.588% | 6.294% | 173.2% | 52.07% | 52.07% | 5.963% | 173.2% |  | 43.3% | 489.8% | 114.6% | 20.74% | 64.58% | 1.242% | 0.4161% | 1.153% | 1.601% | 0.6902% | 1.27% | 3.417% | 173.2% | 32.78% | 5.61% | 3.554%
135 | MT 36 D |  | 6/9/2023 3:10:18 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\MT 36 D.489 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 69.28% |  | 0.9601% | 23.09% | 1.182% | 7.186% | 14.29% | 86.6% |  | 8.646% | 1.642% | 2.471% | 4.167% | 4.664% |  | 18.23% | 8.439% | 7.915% | 2.606% | 8.745% |  | 32.78% | 32.78% | 2.817% | 173.2% | 173.2% | 50% | 34.73% | 173.2% | 20.84% | 14.84% | 0.8744% | 0.1385% | 1.874% | 1.032% | 1.176% | 1.504% | 3.88% | 43.3% | 25% | 4.761% | 4.274%
136 | MT 37 D |  | 6/9/2023 3:16:06 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\MT 37 D.490 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 31.49% | 173.2% | 1.95% | 19.52% | 1.846% | 5.751% | 28.64% | 114.6% |  | 14.78% | 11.32% | 5.228% | 34.64% | 2.143% | 86.6% | 31.23% | 42.72% | 4.939% | 2.249% | 2.813% |  | 17.32% | 17.32% | 3.653% |  | 86.6% | 78.06% | 265.8% |  | 23.83% | 9.352% | 0.5003% | 1.251% | 2.686% | 2.011% | 1.692% | 2.203% | 1.899% | 86.6% | 27.15% | 1.763% | 7.603%
137 | MT 38 D |  | 6/9/2023 3:21:55 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\MT 38 D.491 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 32.87% | 173.2% | 0.2445% | 5.384% | 1.908% | 3.899% | 16.37% | 50% | 173.2% | 22.05% | 8.789% | 2.677% | 4.705% | 1.378% | 49.49% | 33.76% | 5.961% | 2.259% | 1.282% | 9.357% | 0% | 39.16% | 39.16% | 7.946% |  | 132.3% | 132.3% | 65.89% |  | 29.4% | 15.68% | 0.7602% | 0.7168% | 1.825% | 0.895% | 1.248% | 1.445% | 0.5458% | 114.6% | 6.928% | 9.632% | 2.257%
138 | MT 39 D |  | 6/9/2023 3:27:43 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\MT 39 D.492 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 34.64% |  | 1.857% | 11.62% | 0.5102% | 5.965% | 39.74% | 0% | 43.3% | 21.92% | 14.22% | 11.93% | 17.67% | 6.237% | 173.2% | 56.73% | 22.18% | 8.426% | 3.374% | 10.53% |  | 24.12% | 24.12% | 4.129% |  | 173.2% | 173.2% | 142.2% |  | 15.05% | 19.66% | 0.8961% | 1.76% | 1.431% | 6.407% | 0.4438% | 0.8516% | 0.5133% | 114.6% | 32.77% | 1.441% | 1.886%
139 | MT 40 D |  | 6/9/2023 3:33:31 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\MT 40 D.493 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 100% |  | 0.4405% | 13.58% | 0.8963% | 9.046% | 28.64% | 100% |  | 23.89% | 10.57% | 2.233% | 6.928% | 2.508% | 88.19% | 33.07% | 28.57% | 16.47% | 2.163% | 3.44% |  | 21.65% | 21.65% | 3.1% |  | 173.2% | 34.64% | 6.976% |  | 29.17% | 11.62% | 0.1994% | 1.359% | 2.365% | 0.7701% | 4.341% | 1.556% | 2.984% | 50% | 21.65% | 5.282% | 0.7895%
140 | MT 32 D |  | 6/9/2023 3:39:20 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\MT 32 D.494 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 69.28% |  | 1.058% | 22.58% | 1.44% | 7.573% | 47.19% | 34.64% | 173.2% | 51.7% | 12.77% | 1.284% | 16.37% | 9.249% |  | 60.09% | 1.979% | 2.017% | 6.649% | 11.86% | 173.2% | 12.39% | 12.39% | 5.679% | 86.6% | 86.6% | 124.9% | 125% |  | 19.79% | 4.762% | 1.434% | 1.689% | 5.538% | 0.9985% | 0.9056% | 2.31% | 2.124% | 86.6% | 16.37% | 10.01% | 1.925%
141 | MT 32 D LD |  | 6/9/2023 3:45:07 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\MT 32 D LD.495 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 34.64% |  | 2.8% | 12.6% | 1.823% | 4.042% | 41.66% | 173.2% |  | 49.81% | 13.8% | 2.979% | 15.06% | 5.359% | 173.2% | 63.81% | 16.33% | 5.408% | 4.367% | 6.674% |  | 33.31% | 33.31% | 1.344% | 173.2% | 173.2% | 114.6% | 70.16% |  | 12.39% | 2.547% | 1.454% | 1.796% | 0.07034% | 0.9252% | 11.27% | 1.47% | 1.56% |  | 26.65% | 5.46% | 3.565%
142 | MT 32 D LD |  | 6/9/2023 3:51:24 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\MT 32 D LD.496 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 114.6% | 173.2% | 1.387% | 24.3% | 1.348% | 1.437% | 20.83% | 86.6% | 173.2% | 11.46% | 7.92% | 2.97% | 14.78% | 2.193% |  | 43.3% | 23.2% | 4.84% | 2.941% | 1.748% |  | 12.37% | 12.37% | 2.433% | 173.2% | 173.2% | 65.47% | 80.33% | 114.6% | 15.75% | 25.42% | 0.9039% | 4.251% | 0.6272% | 1.013% | 0.4387% | 2.442% | 7.927% | 86.6% | 44.61% | 2.92% | 1.618%
143 | LBLANK |  | 6/9/2023 3:57:12 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.497 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 100% |  | 2.791% | 173.2% | 1.983% | 5.442% | 21.65% | 173.2% | 173.2% | 12.37% | 17.32% | 4.675% | 22.91% | 7.95% |  | 54.13% | 26.19% | 2.325% | 5.474% | 8.709% |  | 16.67% | 16.67% | 5.974% | 86.6% | 173.2% | 86.6% | 114.2% |  | 26.96% | 15.75% | 0.8985% | 12.04% | 1.277% | 5.761% | 1.625% | 2.086% | 11% | 173.2% | 56.77% | 18.18% | 24.03%
144 | CCV 1ppb TRACE |  | 6/9/2023 4:03:01 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.498 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.8918% | 3.107% | 0.4124% | 14.32% | 2.633% | 3.156% | 3.188% | 7.115% | 5.34% | 3.512% | 1.026% | 3.286% | 1.801% | 4.804% | 2.401% | 0.4316% | 3.737% | 0.4973% | 1.441% | 8.161% | 12.41% | 2.852% | 2.852% | 2.701% | 5.582% | 2.903% | 2.172% | 20.08% | 16.71% | 1.218% | 0.8638% | 2.215% | 0.591% | 1.411% | 0.5086% | 0.8483% | 1.025% | 1.69% | 0.6198% | 0.8992% | 1.595% | 0.3966%
145 | CCV 1ppb TRACE |  | 6/9/2023 4:09:18 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.499 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.804% | 4.969% | 1.752% | 19.9% | 2.251% | 7.138% | 2.806% | 11.5% | 4.149% | 2.087% | 3.522% | 2.542% | 1.125% | 2.687% | 1.82% | 1.042% | 2.266% | 2.563% | 3.003% | 8.532% | 11.93% | 0.7972% | 0.7972% | 1.702% | 12.73% | 6.316% | 10.52% | 23.99% | 1.975% | 1.513% | 1.817% | 2.728% | 0.5777% | 2.163% | 0.6755% | 0.9018% | 1.448% | 3.282% | 1.183% | 1.037% | 0.6416% | 0.1855%
146 | LFB |  | 6/9/2023 4:15:06 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.500 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.3891% | 1.302% | 3.282% | 3.632% | 0.3417% | 2.3% | 1.537% | 3.005% | 1.456% | 2.727% | 0.7926% | 0.6301% | 1.504% | 0.7657% | 1.255% | 1.663% | 0.8718% | 0.6907% | 0.4434% | 1.998% | 5.581% | 1.965% | 1.965% | 2.435% | 6.267% | 3.214% | 0.8317% | 3.368% | 7.166% | 1.142% | 0.4762% | 1.636% | 2.248% | 1.032% | 0.7409% | 1.497% | 1.141% | 1.784% | 1.687% | 0.3392% | 0.1801% | 2.183%
147 | LFB |  | 6/9/2023 4:21:22 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.501 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.2485% | 0.9917% | 1.141% | 11.44% | 1.024% | 3.518% | 0.4828% | 3.608% | 1.372% | 1.108% | 1.196% | 1.812% | 0.1008% | 2.623% | 1.842% | 2.047% | 1.127% | 0.9929% | 0.5389% | 0.449% | 1.528% | 0.2431% | 0.2431% | 2.072% | 4.81% | 2.173% | 3.196% | 2.322% | 0.9175% | 1.495% | 0.5759% | 2.188% | 1.864% | 1.661% | 0.4768% | 0.788% | 0.5475% | 3.237% | 1.098% | 1.645% | 1.097% | 1.442%
148 | CCV 10 ppb MAJORS |  | 6/9/2023 4:27:10 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.502 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.593% | 3.251% | 1.217% | 7.895% | 0.8363% | 2.868% | 1.245% | 1.521% | 1.623% | 2.843% | 1.149% | 2.009% | 1.119% | 3.693% | 0.545% | 0.8083% | 0.8824% | 1.776% | 1.058% | 1.024% | 1.414% | 1.242% | 1.242% | 1.609% | 1.463% | 0.1632% | 1.289% | 3.988% | 3.106% | 1.334% | 0.6989% | 1.966% | 1.544% | 1.217% | 2.329% | 1.18% | 0.3998% | 3.078% | 1.002% | 1.21% | 0.3177% | 1.475%
149 | CCV 10 ppb MAJORS |  | 6/9/2023 4:33:27 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.503 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.6756% | 0.8752% | 1.284% | 11.24% | 2.015% | 5.157% | 1.705% | 1.74% | 0.5877% | 0.6383% | 3.589% | 2.66% | 0.4765% | 0.6384% | 1.101% | 2.136% | 1.275% | 3.566% | 3.522% | 1.779% | 11.5% | 0.5479% | 0.5479% | 0.999% | 2.459% | 0.33% | 1.285% | 2.127% | 2.882% | 2.281% | 1.977% | 1.114% | 0.5826% | 2.441% | 7.878% | 0.8698% | 0.4604% | 2.582% | 3.316% | 3.475% | 0.07378% | 1.806%
150 | MT 33 D |  | 6/9/2023 4:39:15 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\MT 33 D.504 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 45.83% |  | 3.284% | 48.16% | 0.8067% | 1.541% | 19.92% | 173.2% | 173.2% | 24.05% | 17% | 7.504% | 32.01% | 3.368% | 114.6% | 55.45% | 26.22% | 3.431% | 7.93% | 3.807% |  | 7.37% | 7.37% | 1.767% |  | 173.2% | 114.6% | 44.53% | 86.6% | 2.965% | 7.293% | 1.579% | 0.7729% | 2.11% | 0.5099% | 0.5528% | 1.372% | 2.005% | 33.33% | 29.57% | 4.514% | 5.896%
151 | MT 33 D LSP |  | 6/9/2023 4:45:03 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\MT 33 D LSP.505 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.355% | 1.617% | 1.141% | 4.903% | 1.472% | 1.355% | 1.832% | 1.174% | 0.8321% | 0.2926% | 0.2186% | 1.921% | 0.438% | 0.654% | 1.679% | 1.894% | 0.8744% | 1.136% | 1.066% | 1.338% | 3.777% | 0.5169% | 0.5169% | 1.852% | 2.779% | 2.074% | 0.3471% | 1.026% | 1.435% | 1.445% | 0.7924% | 3.993% | 0.7362% | 2.683% | 1.168% | 0.9815% | 1.674% | 3.671% | 0.9864% | 0.8669% | 0.1743% | 0.7346%
152 | cal blank |  | 6/9/2023 4:50:51 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\cal blank.506 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 91.65% |  | 2.946% |  | 5.252% | 11.04% | 87.18% |  | 173.2% | 94.37% | 30.2% | 8.474% | 34.32% | 0.8673% | 86.6% | 37.75% | 65.47% | 8.497% | 6.765% | 1.101% | 173.2% | 2.794% | 2.794% | 5.516% | 173.2% | 65.47% | 86.6% | 40.55% | 86.6% | 7.569% | 6.667% | 2.059% | 0.871% | 1.863% | 1.028% | 1.469% | 0.6951% | 2.1% | 69.28% | 52.92% | 36.3% | 19.87%
153 | STD 0.01 |  | 6/9/2023 4:56:40 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 0.01.507 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 15.8% | 53.91% | 2.477% | 26.65% | 1.708% | 3.489% | 13.86% | 34.64% | 74.18% | 21.07% | 7.407% | 3.367% | 6.773% | 4.44% | 17.47% | 6.928% | 24.97% | 3.363% | 1.235% | 3.424% |  | 6.926% | 6.926% | 6.698% | 86.6% | 33.33% | 52.92% | 99.87% | 100% | 3.837% | 4.672% | 1.81% | 1.032% | 0.8952% | 1.351% | 1.093% | 1.372% | 3.022% | 8.297% | 13.52% | 8.973% | 1.582%
154 | std 0.05 |  | 6/9/2023 5:02:29 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\std 0.05.508 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4.124% | 13.53% | 3.35% | 28.28% | 2.145% | 8.727% | 26.87% | 22.35% | 32.89% | 12.49% | 1.653% | 1.747% | 1.336% | 6.048% | 11.77% | 15.08% | 17.03% | 4.967% | 0.638% | 3.757% | 41.66% | 5.036% | 5.036% | 9.018% | 41.66% | 14.52% | 7.783% | 58.93% | 53.29% | 4.452% | 6.328% | 1.582% | 2.929% | 1.151% | 1.193% | 15.76% | 1.92% | 4.475% | 4.842% | 3.417% | 1.263% | 1.436%
155 | std 0.1 |  | 6/9/2023 5:08:17 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\std 0.1.509 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 6.96% | 25.8% | 1.345% | 52.92% | 2.061% | 7.346% | 6.784% | 7.07% | 10.07% | 9.799% | 3.964% | 4.185% | 1.925% | 2.174% | 2.604% | 32.9% | 3.182% | 2.336% | 5.193% | 19.64% | 18.23% | 5.514% | 5.514% | 3.177% | 29.4% | 6.84% | 12.83% | 73.95% | 22.22% | 3.895% | 2.273% | 0.3687% | 0.8705% | 1.407% | 0.9109% | 2.12% | 1.988% | 2.297% | 4.486% | 1.444% | 3.1% | 3.942%
156 | std 0.5 |  | 6/9/2023 5:14:05 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\std 0.5.510 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4.096% | 10.71% | 1.959% | 20.25% | 2.099% | 2.332% | 5.13% | 13.64% | 2.363% | 3.605% | 2.993% | 0.9498% | 0.764% | 1.962% | 2.25% | 3.654% | 2.156% | 0.819% | 2.784% | 23.25% | 17.32% | 1.621% | 1.621% | 5.571% | 9.387% | 11.63% | 11.15% | 8.541% | 7.922% | 0.9086% | 1.24% | 0.203% | 1.675% | 3.227% | 1.293% | 0.5741% | 2.32% | 1.435% | 1.755% | 1.083% | 2.52% | 0.2415%
157 | std 1 |  | 6/9/2023 5:19:54 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\std 1.511 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.5337% | 7.608% | 0.9484% | 4.092% | 3.667% | 5.521% | 7.471% | 3.616% | 2.857% | 2.431% | 1.826% | 0.5863% | 1.299% | 1.423% | 2.326% | 2.288% | 1.755% | 1.597% | 1.121% | 2.374% | 21.05% | 3.513% | 3.513% | 1.369% | 17.73% | 6.911% | 7.924% | 4.954% | 9.434% | 0.677% | 3.378% | 0.515% | 1.524% | 2.639% | 2.548% | 1.507% | 1.548% | 1.744% | 0.9674% | 1.021% | 1.323% | 1.029%
158 | std 5 |  | 6/9/2023 5:25:42 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\std 5.512 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.277% | 4.372% | 2.369% | 12.71% | 0.9776% | 2.288% | 1.051% | 5.144% | 0.9801% | 0.9458% | 1.216% | 2.185% | 1.423% | 0.7104% | 2.496% | 4.591% | 1.881% | 1.707% | 1.927% | 2.894% | 7.362% | 0.9776% | 0.9776% | 3.101% | 6.753% | 0.576% | 2.928% | 3.24% | 5.005% | 1.85% | 0.4519% | 1.261% | 0.3108% | 1.144% | 5.056% | 0.4924% | 2.468% | 1.677% | 1.858% | 1.716% | 0.5391% | 0.7637%
159 | std 10 |  | 6/9/2023 5:31:30 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\std 10.513 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.92% | 3.096% | 1.332% | 13.75% | 1.085% | 4.256% | 1.829% | 3.555% | 1.847% | 2.502% | 1.077% | 1.325% | 1.485% | 1.794% | 0.6996% | 1.37% | 1.535% | 0.8583% | 0.9367% | 2.295% | 3.413% | 0.6339% | 0.6339% | 3.075% | 0.8081% | 1.741% | 0.5078% | 2.078% | 3.407% | 0.8711% | 0.7829% | 1.543% | 0.7209% | 1.525% | 1.776% | 1.106% | 0.3193% | 2.394% | 0.8146% | 0.6726% | 0.2197% | 1.969%
160 | LBLANK |  | 6/9/2023 5:37:19 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.514 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 86.6% | 173.2% | 1.555% | 0% | 4.509% | 3.133% | 75.25% |  | 173.2% | 28.64% | 26.02% | 2.827% | 21.65% | 0.4405% | 91.65% | 39.03% | 33.33% | 20.52% | 6.233% | 10.14% | 173.2% | 32.78% | 32.78% | 6.272% |  | 173.2% | 100% | 55.45% |  | 5.268% | 20.76% | 1.101% | 1.003% | 1.789% | 1.206% | 4.265% | 1.353% | 1.771% | 173.2% | 33.33% | 40.81% | 19.25%
161 | CCV 1ppb TRACE |  | 6/9/2023 5:43:07 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.515 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 5.524% | 2.747% | 0.8973% | 32.81% | 5.526% | 7.146% | 4.546% | 14.47% | 1.765% | 1.576% | 1.864% | 0.6353% | 0.5123% | 2.872% | 1.984% | 0.9993% | 6.09% | 2.671% | 0.1613% | 3.028% | 8.98% | 1.766% | 1.766% | 4.138% | 17.38% | 2.736% | 5.053% | 11.1% | 8.875% | 1.703% | 3.283% | 1.367% | 0.4955% | 6.739% | 0.1628% | 1.393% | 1.109% | 0.5009% | 2.568% | 0.9815% | 1.801% | 1.292%
162 | CCV 1ppb TRACE |  | 6/9/2023 5:49:25 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.516 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.858% | 1.315% | 0.9767% | 5.161% | 3.894% | 2.294% | 4.083% | 5.707% | 8.774% | 3.926% | 1.532% | 2.55% | 1.33% | 1.398% | 5.492% | 2.891% | 2.799% | 0.4369% | 2.048% | 5.398% | 12.14% | 0.8456% | 0.8456% | 5.566% | 19.47% | 4.846% | 3.367% | 8.254% | 12.86% | 0.9096% | 0.4958% | 1.563% | 5.064% | 0.8824% | 1.512% | 1.7% | 1.828% | 3.902% | 0.9057% | 1.79% | 0.9313% | 2.096%
163 | CCV 10 ppb MAJORS |  | 6/9/2023 5:55:13 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.517 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.3178% | 2.257% | 2.04% | 13.5% | 1.145% | 1.863% | 1.3% | 1.37% | 1.572% | 0.5472% | 0.4889% | 1.121% | 1.045% | 1.565% | 0.7328% | 1.54% | 0.2956% | 0.7801% | 0.9463% | 1.99% | 2.533% | 0.4644% | 0.4644% | 2.328% | 3.34% | 1.356% | 0.5746% | 2.267% | 6.645% | 1.088% | 0.5627% | 0.7089% | 0.2961% | 1.216% | 1.158% | 23.71% | 1.911% | 4.489% | 0.5374% | 1.138% | 1.942% | 1.155%
164 | CCV 10 ppb MAJORS |  | 6/9/2023 6:01:30 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.518 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.127% | 1.424% | 1.661% | 9.115% | 2.554% | 3.557% | 1.436% | 2.809% | 1.736% | 1.377% | 0.4748% | 0.1414% | 1.065% | 0.7126% | 1.827% | 2.1% | 1.485% | 1.786% | 0.4402% | 0.2274% | 3.765% | 0.7112% | 0.7112% | 2.491% | 3.296% | 2.412% | 1.363% | 1.401% | 1.585% | 0.5074% | 1.736% | 1.532% | 0.8206% | 4.924% | 0.5745% | 0.7944% | 1.476% | 1.656% | 0.3266% | 0.4769% | 0.5185% | 2.201%
165 | std 25 |  | 6/9/2023 6:07:19 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\std 25.519 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.2015% | 1.614% | 0.7492% | 4.072% | 0.6029% | 5.25% | 1.086% | 4.019% | 1.033% | 0.359% | 0.05832% | 0.631% | 0.6506% | 1.95% | 0.4151% | 1.479% | 1.892% | 1.655% | 0.7727% | 0.908% | 3.692% | 0.5747% | 0.5747% | 1.464% | 1.317% | 1.076% | 0.9339% | 1.091% | 2.152% | 0.2259% | 1.097% | 1.375% | 1.127% | 3.284% | 0.9746% | 0.9442% | 0.6637% | 0.3031% | 0.7917% | 0.9857% | 1.017% | 1.754%
166 | std 50 |  | 6/9/2023 6:13:06 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\std 50.520 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.563% | 1.68% | 2.178% | 2.602% | 2.025% | 0.6115% | 0.8899% | 1.078% | 0.5236% | 0.7398% | 0.7454% | 2.502% | 0.885% | 1.143% | 0.647% | 1.091% | 1.11% | 1.524% | 1.114% | 1.306% | 2.708% | 0.7741% | 0.7741% | 0.5575% | 1.849% | 1.031% | 1.061% | 1.325% | 1.557% | 1.406% | 0.08422% | 0.5775% | 2.68% | 1.784% | 1.458% | 0.6424% | 1.353% | 2.914% | 1.479% | 0.7621% | 1.965% | 0.995%
167 | std 50 |  | 6/9/2023 6:19:23 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\std 50.521 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.322% | 0.8697% | 2.075% | 2.409% | 1.35% | 3.208% | 1.803% | 1.677% | 1.563% | 1.989% | 1.331% | 0.7998% | 0.7598% | 1.719% | 0.9587% | 0.677% | 0.5935% | 0.4423% | 1.202% | 1.413% | 1.624% | 0.8888% | 0.8888% | 0.3424% | 1.199% | 2.404% | 0.4146% | 2.049% | 1.873% | 0.6356% | 0.2121% | 1.117% | 1.113% | 3.678% | 7.884% | 1.085% | 1.03% | 2.887% | 0.7418% | 0.8757% | 0.3789% | 1.658%
168 | std 100 |  | 6/9/2023 6:25:11 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\std 100.522 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.652% | 0.6616% | 1.852% | 0.2062% | 1.618% | 1.119% | 1.14% | 2.519% | 1.304% | 1.021% | 0.9771% | 0.5669% | 0.6127% | 0.9544% | 1.702% | 1.174% | 1.044% | 0.06528% | 0.8399% | 0.9987% | 1.892% | 0.1195% | 0.1195% | 0.8049% | 2.014% | 1.322% | 0.9001% | 0.8483% | 1.557% | 1.502% | 0.5718% | 1.628% | 0.7071% | 2.297% | 1.782% | 0.6804% | 2.241% | 1.334% | 0.5724% | 1.23% | 1.14% | 0.4018%
169 | std 100 |  | 6/9/2023 6:31:28 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\std 100.523 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.126% | 2.017% | 0.8288% | 1.003% | 0.6814% | 1.233% | 0.726% | 1.162% | 1.607% | 1.42% | 0.9255% | 0.3968% | 1.004% | 1.068% | 0.9067% | 0.7574% | 0.5841% | 0.3233% | 0.1656% | 0.7848% | 1.325% | 1.202% | 1.202% | 1.82% | 1.902% | 0.7429% | 0.3064% | 1.185% | 1.044% | 1.266% | 1.306% | 1.481% | 1.029% | 1.933% | 0.4558% | 0.8705% | 1.048% | 0.8905% | 0.9042% | 0.6599% | 0.5046% | 10.35%
170 | LBLANK |  | 6/9/2023 6:37:16 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.524 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 66.14% | 86.6% | 0.984% | 34.64% | 3.745% | 3.974% | 76.03% |  | 50% | 16.67% | 22.59% | 3.39% | 17.64% | 2.715% | 43.83% | 20% | 43.07% | 13.01% | 5.886% | 7.532% | 57.74% | 10.57% | 10.57% | 3.153% | 100% | 43.3% | 26.96% | 83.89% |  | 2.118% | 3.17% | 2.775% | 1.653% | 5.343% | 0.9078% | 1.72% | 0.4196% | 2.227% | 78.06% | 18.9% | 19.87% | 25.78%
171 | PE ICV |  | 6/9/2023 6:43:04 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\PE ICV.525 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.045% | 2.151% | 1.059% | 0.8111% | 1.102% | 3.967% | 2.499% | 2.47% | 2.908% | 1.15% | 1.119% | 0.9269% | 1.519% | 1.468% | 0.4249% | 1.193% | 0.9551% | 0.3444% | 1.01% | 0.9902% | 0.9346% | 1.182% | 1.182% | 2.223% | 0.9368% | 2.366% | 2.129% | 2.455% | 5.532% | 0.2422% | 1.417% | 1.004% | 1.806% | 5.588% | 1.091% | 0.3502% | 1.686% | 1.103% | 0.5061% | 0.4868% | 1.338% | 2.751%
172 | 5ppb CCV |  | 6/9/2023 6:48:52 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5ppb CCV.526 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.5674% | 2.65% | 1.404% | 8.761% | 2.371% | 2.489% | 1.142% | 6.393% | 1.495% | 0.7893% | 0.3131% | 4.538% | 0.9602% | 0.2281% | 1.325% | 0.3265% | 1.682% | 0.9424% | 1.353% | 3.127% | 5.428% | 1.617% | 1.617% | 1.549% | 4.12% | 1.701% | 0.4588% | 6.564% | 1.584% | 0.1375% | 1.943% | 4.711% | 2.128% | 0.7945% | 0.3058% | 1.024% | 1.975% | 3.108% | 0.8455% | 1.25% | 1.067% | 2.203%
173 | 25ppb CCV |  | 6/9/2023 6:54:40 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\25ppb CCV.527 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.031% | 2.554% | 1.229% | 6.924% | 1.223% | 2.851% | 1.004% | 4.158% | 0.854% | 1.337% | 0.6294% | 1.302% | 1.547% | 3.143% | 1.222% | 1.35% | 0.628% | 1.473% | 1.628% | 0.9845% | 4.874% | 0.2925% | 0.2925% | 0.3271% | 2.607% | 2.284% | 0.8494% | 1.337% | 0.8597% | 1.813% | 0.6495% | 1.474% | 1.228% | 1.598% | 1.323% | 0.5134% | 1.473% | 0.6235% | 0.9997% | 0.8408% | 0.3417% | 0.533%
174 | 8_30_21_W_1 |  | 6/9/2023 7:00:27 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_W_1.528 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 7.59% | 2.402% | 2.514% | 7.148% | 16.2% | 2.021% | 12.85% | 3.627% | 18.49% | 3.571% | 13.22% | 0.2303% | 0.9202% | 2.201% | 37.12% | 8.108% | 1.89% | 0.4305% | 0.9123% | 6.629% | 14.99% | 2.128% | 2.128% | 3.439% | 100% | 69.28% | 47.19% | 118.8% | 114.6% | 1.389% | 0.01613% | 1.916% | 1.18% | 2.608% | 2.37% | 1.791% | 1.311% | 1.401% | 19.92% | 12.6% | 3.186% | 2.628%
175 | LBLANK |  | 6/9/2023 7:06:16 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.529 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 173.2% | 173.2% | 5.103% | 43.3% | 3.982% | 6.261% | 63.55% |  | 86.6% | 81.51% | 19.92% | 5.366% | 35.25% | 1.073% | 173.2% | 33.07% | 36.74% | 10.66% | 7.845% | 12.07% | 173.2% | 20.28% | 20.28% | 5.62% | 173.2% | 173.2% | 33.33% | 113.7% | 173.2% | 6.701% | 33.54% | 4.98% | 1.15% | 2.6% | 0.7163% | 1.393% | 1.442% | 2.139% | 86.6% | 10.83% | 38.42% | 5.413%
176 | CCV 1ppb TRACE |  | 6/9/2023 7:12:04 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.530 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.031% | 8.46% | 1.335% | 22.19% | 2.23% | 4.384% | 9.322% | 3.977% | 5.269% | 5.161% | 2.314% | 4.69% | 3.434% | 5.364% | 2.216% | 1.03% | 2.135% | 2.71% | 0.3507% | 12.06% | 18.04% | 1.083% | 1.083% | 3.494% | 8.61% | 6.757% | 5.736% | 4.393% | 5.951% | 2.163% | 1.874% | 1.685% | 8.847% | 0.06391% | 0.4677% | 1.674% | 3.025% | 5.982% | 2.116% | 1.825% | 1.978% | 1.444%
177 | CCV 1ppb TRACE |  | 6/9/2023 7:18:22 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.531 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 43.3% |  | 7.192% |  | 11.88% | 21.11% | 120.2% |  |  | 78.73% | 18.41% | 10.01% | 45.83% | 8.434% | 88.19% | 57.28% | 41.66% | 5.413% | 2.207% | 10.42% | 173.2% | 43.3% | 43.3% | 9.534% | 173.2% | 86.6% | 173.2% | 566.8% |  | 3.953% | 91.65% | 2.553% | 1.899% | 5.267% | 2.289% | 4.948% | 7.384% | 4.465% | 43.3% | 65.47% | 63.81% | 5.973%
178 | LFB |  | 6/9/2023 7:24:08 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.532 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.6308% | 3.262% | 1.641% | 1.613% | 0.2127% | 4.46% | 1.99% | 3.834% | 1.796% | 0.4975% | 1.123% | 0.3874% | 0.1663% | 2.137% | 0.5966% | 0.8823% | 2.215% | 0.8708% | 0.5523% | 0.5959% | 2.954% | 1.263% | 1.263% | 3.03% | 1.993% | 0.7484% | 1.51% | 1.378% | 1.52% | 2.278% | 0.6952% | 3.142% | 1.134% | 2.348% | 1.089% | 0.4019% | 1.092% | 4.363% | 0.6803% | 1.418% | 0.2993% | 1.001%
179 | LFB |  | 6/9/2023 7:30:24 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.533 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.273% | 0.5435% | 0.7844% | 4.189% | 0.4152% | 5.019% | 2.118% | 0.6764% | 2.39% | 3.341% | 1.257% | 0.6775% | 0.7599% | 2.311% | 0.4615% | 1.53% | 1.519% | 1.294% | 0.8729% | 1.294% | 3.397% | 0.9944% | 0.9944% | 3.354% | 4.591% | 0.4282% | 1.701% | 1.847% | 1.547% | 0.6382% | 1.797% | 0.9668% | 1.803% | 2.092% | 0.8352% | 0.1633% | 0.9283% | 0.8912% | 1.569% | 0.5947% | 0.878% | 0.6679%
180 | CCV 10 ppb MAJORS |  | 6/9/2023 7:36:12 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.534 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.397% | 0.242% | 2.059% | 1.691% | 0.3166% | 3.336% | 1.294% | 3.047% | 2.241% | 0.8746% | 0.3311% | 0.8657% | 0.9418% | 1.338% | 0.8717% | 1.837% | 1.481% | 0.1778% | 1.176% | 0.9921% | 4.201% | 0.7788% | 0.7788% | 1.553% | 5.039% | 0.4281% | 1.246% | 2.659% | 5.957% | 0.2792% | 0.6889% | 1.766% | 0.9397% | 3.918% | 1.11% | 1.656% | 0.3% | 2.833% | 0.4533% | 0.1679% | 0.5307% | 1.274%
181 | CCV 10 ppb MAJORS |  | 6/9/2023 7:42:30 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.535 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.257% | 0.897% | 0.8936% | 1.239% | 1.109% | 2.126% | 2.357% | 1.011% | 1.84% | 0.8572% | 1.664% | 2.158% | 1.646% | 2.68% | 4.537% | 0.717% | 2.754% | 1.62% | 0.8516% | 1.229% | 7.096% | 0.8015% | 0.8015% | 2.226% | 4.414% | 1.176% | 0.1032% | 2.989% | 3.565% | 1.537% | 0.6915% | 0.6312% | 1.746% | 4.298% | 1.738% | 4.305% | 1.668% | 2.034% | 0.4949% | 0.7409% | 0.2661% | 1.666%
182 | 8_30_21_W_6 |  | 6/9/2023 7:48:18 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_W_6.536 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 7.343% | 10.23% | 1.278% | 5.207% | 1.524% | 3.766% | 10.38% | 3.823% | 14.6% | 17.8% | 3.811% | 1.445% | 1.683% | 3.014% | 9.897% | 4.879% | 3.74% | 1.835% | 3.353% | 14.57% | 18.56% | 1.212% | 1.212% | 2.959% | 33.33% | 9.116% | 29.4% | 123.9% | 13.32% | 1.467% | 0.8006% | 1.034% | 0.6132% | 9.924% | 3.099% | 2.839% | 1.362% | 2.763% | 58.47% | 11.95% | 3.093% | 1.762%
183 | 8_30_21_W_6_LD |  | 6/9/2023 7:54:05 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_W_6_LD.537 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.369% | 4.56% | 2.272% | 8.226% | 0.1878% | 5.865% | 16.4% | 3.398% | 15.93% | 10.77% | 4.406% | 1.148% | 1.121% | 0.9101% | 10.83% | 5.393% | 3.298% | 1.352% | 2.698% | 9.81% | 16.82% | 1.138% | 1.138% | 5.88% | 86.6% | 24.12% | 14.78% | 31.39% | 50% | 1.185% | 1.95% | 1.231% | 5.461% | 1.367% | 2.412% | 2.559% | 1.457% | 7.367% | 24.44% | 9.622% | 4.349% | 1.937%
184 | 8_30_21_W_6_LD |  | 6/9/2023 8:00:20 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_W_6_LD.538 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.858% | 6.129% | 1.529% | 3.147% | 0.9829% | 6.868% | 13.14% | 3.588% | 14.83% | 34.64% | 12.14% | 0.6496% | 0.9522% | 1.577% | 15.75% | 10.57% | 3.261% | 0.6367% | 0.7968% | 1.835% | 5% | 2.448% | 2.448% | 6.646% | 50% | 43.3% | 32.78% | 31.53% | 66.67% | 2.213% | 1.262% | 1.037% | 0.2317% | 5.866% | 0.7221% | 11.92% | 1.205% | 0.8942% | 14.7% | 17.76% | 4.554% | 3.449%
185 | 8_30_21_W_6_DUP |  | 6/9/2023 8:06:08 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_W_6_DUP.539 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.229% | 3.942% | 1.685% | 5.914% | 2.221% | 3.391% | 3.199% | 6.153% | 19.95% | 31.02% | 7.531% | 2.424% | 0.617% | 3.013% | 24.05% | 9.938% | 0.9626% | 2.668% | 1.08% | 2.593% | 9.324% | 0.8746% | 0.8746% | 3.641% | 173.2% | 49.49% | 36.74% | 44.7% | 114.6% | 2.373% | 1.348% | 3.656% | 1.327% | 0.5923% | 1.53% | 1.609% | 1.009% | 4.07% | 21.31% | 8.707% | 2.821% | 4.589%
186 | 8_30_21_W_6 |  | 6/9/2023 8:11:56 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_W_6.540 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4.203% | 3.254% | 3.333% | 3.555% | 3.404% | 4.636% | 19.35% | 2.71% | 7.855% | 27.27% | 10.43% | 1.083% | 0.8844% | 2.367% | 16.37% | 40.43% | 4.16% | 1.564% | 0.7982% | 3.725% | 6.744% | 1.623% | 1.623% | 1.865% | 173.2% | 86.6% | 45.07% | 59.63% | 86.6% | 0.297% | 0.9976% | 1.139% | 1.143% | 2.507% | 0.671% | 29.4% | 8.17% | 3.673% | 13.32% | 24.12% | 4.339% | 4.77%
187 | 8_30_21_W_6_DUP_LSP |  | 6/9/2023 8:17:43 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_W_6_DUP_LSP.541 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.086% | 1.529% | 2.942% | 6.051% | 2.286% | 2.481% | 2.52% | 2.181% | 1.365% | 2.995% | 0.6558% | 0.2082% | 0.7051% | 0.7606% | 1.881% | 1.709% | 1.265% | 1.021% | 0.2169% | 3.341% | 4.607% | 0.1516% | 0.1516% | 5.634% | 1.692% | 1.414% | 2.076% | 1.35% | 2.331% | 0.5223% | 1.275% | 23.01% | 0.3699% | 3.282% | 0.9642% | 1.085% | 2.727% | 0.9985% | 1.296% | 0.5843% | 2.593% | 0.7684%
188 | 8_30_21_W_6_DUP_LSP |  | 6/9/2023 8:24:01 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_W_6_DUP_LSP.542 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.8693% | 0.6847% | 2.931% | 5.698% | 1.166% | 3.082% | 1.685% | 2.339% | 0.5294% | 0.389% | 0.841% | 1.679% | 1.068% | 0.6996% | 1.73% | 1.206% | 1.412% | 0.8739% | 0.3245% | 1.048% | 3.131% | 0.8292% | 0.8292% | 0.7478% | 0.4369% | 1.77% | 1.092% | 1.751% | 2.601% | 1.297% | 0.4507% | 0.3017% | 0.6686% | 3.76% | 0.3664% | 0.1975% | 1.051% | 1.714% | 1.668% | 0.9308% | 1.138% | 1.423%
189 | 8_30_21_W_8 |  | 6/9/2023 8:29:49 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_W_8.543 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4.599% | 4.229% | 0.7403% | 2.179% | 2.743% | 0.1527% | 3.287% | 3.606% | 7.001% | 16.87% | 10.84% | 4.043% | 1.174% | 0.7274% | 23.08% | 9.799% | 5.264% | 0.9361% | 0.6051% | 7.129% | 18.09% | 0.6893% | 0.6893% | 1.891% | 173.2% | 20% | 22.21% | 16.62% | 50% | 0.7066% | 1.2% | 0.3451% | 0.3964% | 8.808% | 3.149% | 0.07555% | 5.605% | 3.734% | 38.42% | 18.2% | 3.279% | 3.235%
190 | 8_30_21_W_8 |  | 6/9/2023 8:36:07 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_W_8.544 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.785% | 9.639% | 3.055% | 11.7% | 1.845% | 7.465% | 14.31% | 2.2% | 2.687% | 11.95% | 17.38% | 1.04% | 0.8709% | 1.699% | 40.52% | 20.76% | 5.953% | 2.043% | 0.4246% | 1.21% | 19.81% | 2.042% | 2.042% | 4.624% | 107.9% | 12.37% | 27.71% | 30.83% | 86.6% | 1.308% | 0.9543% | 0.8612% | 0.6009% | 1.655% | 1.862% | 1.852% | 1.763% | 2.611% | 46.1% | 11.18% | 3.234% | 5.019%
191 | 8_30_21_W_8_SP |  | 6/9/2023 8:41:55 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_W_8_SP.545 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.387% | 5.748% | 0.6506% | 7.188% | 2.095% | 4.769% | 4.551% | 6.073% | 11.48% | 5.041% | 1.821% | 1.382% | 0.354% | 2.562% | 1.247% | 1.342% | 6.063% | 1.737% | 0.3342% | 1.271% | 12.74% | 1.3% | 1.3% | 3.698% | 14.88% | 7.151% | 5.727% | 13.81% | 11.13% | 0.7585% | 0.538% | 0.8077% | 0.8928% | 5.821% | 4.121% | 1.318% | 1.526% | 1.585% | 2.078% | 1.375% | 1.15% | 2.437%
192 | LBLANK |  | 6/9/2023 8:47:44 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.546 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 173.2% |  | 1.736% | 86.6% | 6.46% | 5.531% | 12.37% |  |  | 24.74% | 20.55% | 6.916% | 44.44% | 3.877% | 173.2% | 18.37% | 66.14% | 12.38% | 6.349% | 8.006% | 173.2% | 76.36% | 76.36% | 1.822% | 173.2% | 173.2% | 49.49% | 120.7% |  | 14.47% | 8.449% | 0.3913% | 7.999% | 4.817% | 2.381% | 0.7951% | 2.896% | 7.852% | 100% | 75% | 39.03% | 10.83%
193 | CCV 1ppb TRACE |  | 6/9/2023 8:53:33 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.547 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.3259% | 11.21% | 0.8925% | 8.332% | 6.223% | 3.562% | 1.906% | 14.33% | 4.405% | 2.656% | 2.488% | 1.579% | 2.298% | 2.158% | 2.84% | 2.657% | 3.116% | 1.867% | 1.408% | 7.862% | 7.217% | 1.888% | 1.888% | 2.427% | 10.89% | 11.48% | 1.73% | 10.21% | 3.326% | 1.14% | 0.6915% | 4.213% | 0.6761% | 6.604% | 0.6657% | 2.555% | 2.122% | 1.8% | 2% | 2.442% | 0.9426% | 1.078%
194 | CCV 1ppb TRACE |  | 6/9/2023 8:59:50 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.548 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4.201% | 5.983% | 2.306% | 12.5% | 2.794% | 7.524% | 1.38% | 2.206% | 1.893% | 0.6607% | 2.532% | 1.923% | 0.6729% | 1.952% | 1.214% | 6.95% | 0.8653% | 1.282% | 1.968% | 8.462% | 20.34% | 2.872% | 2.872% | 8.174% | 12.76% | 4.766% | 3.228% | 9.267% | 11.21% | 1.724% | 2.588% | 0.51% | 0.9168% | 4.647% | 1.172% | 0.8966% | 0.5803% | 3.044% | 0.8494% | 2.098% | 0.1957% | 0.3113%
195 | CCV 10 ppb MAJORS |  | 6/9/2023 9:05:39 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.549 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.739% | 2.576% | 2.306% | 2.705% | 0.2395% | 4.172% | 1.381% | 1.37% | 0.5338% | 3.227% | 1.002% | 2.836% | 0.6541% | 0.8215% | 2.604% | 3.869% | 2.089% | 0.5963% | 0.3665% | 0.8603% | 4.37% | 0.3828% | 0.3828% | 2.484% | 3.667% | 2.27% | 1.701% | 1.107% | 4.222% | 1.395% | 1.631% | 2.923% | 2.596% | 5.936% | 1.276% | 0.3865% | 0.8899% | 2.217% | 0.5618% | 0.8855% | 0.145% | 1.924%
196 | CCV 10 ppb MAJORS |  | 6/9/2023 9:11:56 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.550 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.325% | 4.218% | 2.3% | 6.986% | 1.093% | 2.751% | 2.398% | 2.179% | 0.8675% | 0.7552% | 0.7192% | 1.182% | 1.686% | 1.381% | 1.104% | 2.183% | 1.06% | 0.5608% | 1.168% | 0.3495% | 7.036% | 0.9149% | 0.9149% | 1.379% | 5.138% | 4.369% | 1.624% | 3.667% | 8% | 2.339% | 1.515% | 0.8004% | 1.302% | 0.9883% | 1.996% | 1.524% | 1.028% | 4.687% | 1.28% | 0.8352% | 0.844% | 1.673%
197 | 8_30_21_W_9 |  | 6/9/2023 9:17:44 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_W_9.551 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.855% | 5.563% | 2.078% | 10.06% | 0.866% | 6.426% | 10.47% | 6.125% | 10.95% | 22.15% | 15.83% | 1.089% | 0.8402% | 0.4894% | 31.17% | 31.23% | 5.247% | 1.112% | 2.56% | 9.223% | 14.31% | 3.148% | 3.148% | 8.045% | 86.6% | 96.44% | 46.81% | 63.66% | 91.65% | 0.3951% | 1.445% | 4.426% | 1.337% | 1.676% | 0.335% | 1.3% | 0.8568% | 1.434% | 27.06% | 25.32% | 3.299% | 3.393%
198 | 8_30_21_W_10 |  | 6/9/2023 9:23:32 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_W_10.552 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 5.183% | 1.426% | 1.253% | 6.862% | 1.176% | 6.315% | 10.49% | 1.683% | 15.06% | 14.93% | 9.947% | 1.987% | 2.241% | 1.572% | 21.31% | 22.93% | 6.523% | 2.355% | 3.859% | 10.43% | 22.91% | 2.223% | 2.223% | 0.9771% | 50% | 16.67% | 28.64% | 11.97% | 69.28% | 0.3592% | 2.733% | 3.153% | 1.325% | 3.178% | 2.383% | 0.4817% | 3.266% | 1.412% | 50% | 4.558% | 2.801% | 1.015%
199 | 8_30_21_W_11 |  | 6/9/2023 9:29:20 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_W_11.553 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.687% | 0.2212% | 2.2% | 3.999% | 0.6313% | 4.546% | 3.632% | 4.311% | 6.703% | 39.05% | 6.614% | 3.189% | 2.134% | 2.537% | 21.57% | 30.05% | 0.5695% | 1.372% | 4.154% | 11.51% | 16.4% | 1.374% | 1.374% | 2.658% | 173.2% | 31.49% | 27.27% | 39.3% | 34.64% | 2.737% | 1.948% | 1.405% | 0.9277% | 0.7839% | 6.11% | 3.35% | 1.79% | 0.9791% | 49.81% | 10.39% | 1.077% | 0.3388%
200 | 8_30_21_F_1 |  | 6/9/2023 9:35:08 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_F_1.554 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.605% | 1.773% | 1.838% | 8.526% | 1.576% | 0.2814% | 3.682% | 1.797% | 9.06% | 12.78% | 3.475% | 1.508% | 1.23% | 5.414% | 17.5% | 7.426% | 3.095% | 1.128% | 0.2734% | 1.507% | 5.036% | 0.06486% | 0.06486% | 3.457% | 56.77% | 24.98% | 29.64% | 12.28% | 50% | 1.661% | 1.564% | 1.713% | 0.5226% | 2.909% | 0.521% | 0.8469% | 0.8928% | 3.183% | 16.95% | 8.486% | 4.671% | 2.214%
201 | 8_30_21_F_1 |  | 6/9/2023 9:41:25 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_F_1.555 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.311% | 3.234% | 1.956% | 14.19% | 0.3596% | 2.901% | 4.646% | 3.875% | 8.356% | 10.39% | 5.482% | 2.559% | 0.7252% | 0.9932% | 4.917% | 11.21% | 1.819% | 0.8518% | 0.6688% | 0.6354% | 10.01% | 1.189% | 1.189% | 4.415% | 13.32% | 28.87% | 31.49% | 16.48% | 22.91% | 0.9262% | 0.8023% | 0.9547% | 0.4684% | 7.26% | 0.4881% | 0.6859% | 2.913% | 1.754% | 3.464% | 4.562% | 2.399% | 1.816%
202 | 8_30_21_F_1_DUP |  | 6/9/2023 9:47:13 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_F_1_DUP.556 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.241% | 2.951% | 2.2% | 9.872% | 0.5452% | 0.8866% | 3.825% | 0.4882% | 6.113% | 4.58% | 1.776% | 1.025% | 0.3769% | 3.642% | 2.839% | 6.99% | 1.165% | 1.391% | 0.2717% | 1.733% | 6.311% | 1.097% | 1.097% | 7.25% | 45.83% | 37.89% | 8.646% | 24.34% | 68.63% | 0.9717% | 0.8407% | 13.09% | 0.6088% | 4.189% | 0.942% | 1.752% | 1.458% | 1.68% | 25.31% | 18.78% | 4.693% | 0.4582%
203 | 8_30_21_F_1_DUP |  | 6/9/2023 9:53:30 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_F_1_DUP.557 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.983% | 2.231% | 2.425% | 4.322% | 1.544% | 2.875% | 4.473% | 3.265% | 6.744% | 15.38% | 9.21% | 0.8198% | 1.376% | 1.425% | 32.01% | 4.814% | 0.5372% | 0.8959% | 0.3488% | 2.492% | 5.041% | 0.4907% | 0.4907% | 2.381% | 24.74% | 12.5% | 48.06% | 5.822% | 32.73% | 1.46% | 0.8546% | 4.542% | 1.189% | 2.739% | 0.4843% | 0.9174% | 1.666% | 1.62% | 28.54% | 9.556% | 3.62% | 4.326%
204 | 8_30_21_F_6 |  | 6/9/2023 9:59:18 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_F_6.558 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.9726% | 4.386% | 0.91% | 4.364% | 1.128% | 0.4357% | 3.36% | 1.364% | 3.33% | 10.5% | 5.555% | 0.3408% | 0.6931% | 2.199% | 17.84% | 11.45% | 1.674% | 0.3757% | 0.3612% | 0.929% | 6.248% | 0.6637% | 0.6637% | 2.039% | 58.08% | 21.65% | 18.92% | 6.799% | 22.22% | 0.5618% | 0.6289% | 0.9022% | 1.272% | 3.06% | 1.354% | 5.655% | 1.483% | 5.335% | 12.07% | 10.47% | 1.997% | 1.164%
205 | 8_30_21_F_6 |  | 6/9/2023 10:05:38 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_F_6.559 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.399% | 1.802% | 0.8121% | 22.81% | 1.239% | 1.203% | 2.043% | 1.07% | 3.581% | 8.794% | 3.252% | 0.3729% | 1.106% | 2.794% | 16.03% | 5.149% | 1.741% | 0.9234% | 0.7919% | 1.993% | 8.813% | 0.5007% | 0.5007% | 3.046% | 73.47% | 19.25% | 8.408% | 16.59% | 0% | 0.8658% | 0.03303% | 0.2916% | 0.8758% | 12.23% | 0.794% | 1.8% | 0.2391% | 5.492% | 11.35% | 11.76% | 3.152% | 3.061%
206 | 8_30_21_F_8 |  | 6/9/2023 10:11:26 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_F_8.560 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.349% | 0.2547% | 1.293% | 4.651% | 1.287% | 3.814% | 9.199% | 1.854% | 7.513% | 23.09% | 3.783% | 1.922% | 1.717% | 2.046% | 21.72% | 17.25% | 0.7229% | 1.249% | 3.065% | 0.5469% | 7.155% | 2.26% | 2.26% | 8.544% | 45.83% | 28.35% | 14.47% | 5.192% | 53.91% | 0.5679% | 2.42% | 1.055% | 6.394% | 3.857% | 0.8661% | 1.578% | 0.2328% | 5.7% | 13.41% | 2.525% | 4.019% | 0.2449%
207 | 8_30_21_F_8 |  | 6/9/2023 10:17:43 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_F_8.561 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.863% | 4.269% | 1.651% | 16.54% | 1.534% | 2.218% | 2.262% | 1.126% | 3.316% | 5.898% | 1.202% | 0.6831% | 0.5521% | 2.562% | 6.183% | 12.78% | 1.037% | 1.905% | 1.746% | 3.452% | 5.889% | 0.6046% | 0.6046% | 2.338% | 9.116% | 19.25% | 28.6% | 11.9% | 0% | 1.044% | 0.9489% | 0.171% | 4.617% | 0.9205% | 1.178% | 1.598% | 1.177% | 3.342% | 22.73% | 8.217% | 1.917% | 2.297%
208 | LBLANK |  | 6/9/2023 10:23:31 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.562 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 86.6% |  | 6.01% | 34.64% | 5.211% | 9.079% | 0% |  | 173.2% | 17.32% | 29.05% | 5.346% | 6.662% | 6.357% |  | 8.66% | 69.28% | 4.651% | 1.386% | 8.958% |  | 9.75% | 9.75% | 0.8735% |  | 0% | 33.33% | 80.5% | 86.6% | 10.83% | 40.03% | 2.063% | 1.995% | 0.6249% | 12.27% | 2.261% | 1.466% | 1.347% | 173.2% | 26.65% | 39.74% | 3.149%
209 | CCV 1ppb TRACE |  | 6/9/2023 10:29:20 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.563 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 6.411% | 9.167% | 1.468% | 22.72% | 2.473% | 3.464% | 6.402% | 5.491% | 1.008% | 3.326% | 2.052% | 0.9794% | 0.9572% | 2.635% | 3.941% | 1.569% | 2.544% | 0.8088% | 2.083% | 13.97% | 9.75% | 1.487% | 1.487% | 6.431% | 13.44% | 2.332% | 1.03% | 10.35% | 14.59% | 0.5918% | 0.82% | 0.9334% | 0.8896% | 1.613% | 1.208% | 0.4991% | 1.182% | 0.9958% | 2.232% | 1.127% | 1.431% | 0.1964%
210 | CCV 1ppb TRACE |  | 6/9/2023 10:35:37 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.564 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.26% | 1.472% | 2.343% | 15.16% | 3.449% | 6.661% | 2.265% | 9.467% | 4.406% | 2.471% | 1.1% | 3.095% | 1.438% | 2.943% | 2.856% | 1.378% | 0.9653% | 0.3229% | 1.572% | 8.093% | 12.92% | 1.082% | 1.082% | 1.051% | 10.97% | 6.075% | 4.406% | 5.044% | 4.167% | 0.6524% | 1.536% | 0.7658% | 1.496% | 5.835% | 0.5145% | 0.7479% | 1.299% | 2.912% | 0.817% | 1.881% | 1.722% | 0.4354%
211 | LFB |  | 6/9/2023 10:41:25 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.565 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.4268% | 0.9701% | 1.018% | 2.026% | 0.8966% | 2.252% | 2.489% | 2.901% | 1.948% | 1.054% | 1.128% | 2.565% | 0.6167% | 0.1528% | 0.7468% | 3.262% | 1.795% | 1.313% | 1.487% | 1.337% | 0.5012% | 0.7653% | 0.7653% | 1.404% | 0.8747% | 1.753% | 2.911% | 2.788% | 3.49% | 1.068% | 1.097% | 0.6444% | 0.7507% | 0.3039% | 0.3871% | 0.5654% | 1.014% | 1.665% | 1.166% | 0.7668% | 0.4033% | 1.415%
212 | LFB |  | 6/9/2023 10:47:41 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.566 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.5359% | 1.137% | 1.027% | 9.963% | 1.597% | 2.365% | 2.52% | 2.587% | 2.116% | 2.137% | 0.8405% | 1.981% | 1.705% | 2.16% | 1.103% | 1.706% | 0.7416% | 1.651% | 1.008% | 2.037% | 5.313% | 0.9186% | 0.9186% | 0.2088% | 1.935% | 1.806% | 1.571% | 2.007% | 4.778% | 1.112% | 0.4937% | 1.263% | 1.28% | 1.553% | 1.24% | 0.9009% | 1.28% | 3.83% | 1.186% | 0.654% | 0.2206% | 2.289%
213 | CCV 10 ppb MAJORS |  | 6/9/2023 10:53:35 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.567 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.278% | 1.124% | 1.425% | 9.936% | 2.377% | 1.258% | 1.274% | 0.9908% | 1.781% | 0.813% | 0.7083% | 0.1338% | 1.701% | 2.086% | 0.4509% | 4.274% | 1.729% | 0.6387% | 0.5895% | 2.056% | 3.73% | 0.6261% | 0.6261% | 2.795% | 2.32% | 2.329% | 1.157% | 6.754% | 4.243% | 0.9673% | 0.2929% | 0.9541% | 1.625% | 2.554% | 1.471% | 1.049% | 0.7704% | 1.192% | 0.931% | 0.4358% | 0.7185% | 0.7968%
214 | LBLANK |  | 6/9/2023 11:28:06 AM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.568 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 173.2% |  | 6.416% | 24.74% | 2.033% | 7.489% | 68.63% |  | 173.2% | 24.74% | 34.64% | 1.438% | 44.41% | 3.79% | 173.2% | 26.96% | 56.73% | 22.15% | 12.55% | 10.53% |  | 12.37% | 12.37% | 1.812% |  | 100% |  | 71.75% |  | 40% | 45.07% | 1.137% | 0.473% | 1.171% | 0.9224% | 0.9045% | 1.856% | 3.571% | 173.2% | 65.61% | 38.49% | 24.03%
215 | Cal blank |  | 6/9/2023 11:33:55 AM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\Cal blank.569 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth |  |  | 8.224% | 45.83% | 4.347% | 5.57% | 70.5% |  | 173.2% | 0% | 9.091% | 1.152% | 10.83% | 5.953% | 34.64% | 26.96% | 36.74% | 5.036% | 8.533% | 3.658% | 173.2% | 26.65% | 26.65% | 7.989% |  | 173.2% | 0% | 77.02% | 86.6% | 92.82% | 17.32% | 1.732% | 1.833% | 2.313% | 0.9524% | 0.5893% | 1.579% | 3.016% | 43.3% | 56.77% | 42.43% | 16.76%
216 | CAL BLANK |  | 6/9/2023 11:39:43 AM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CAL BLANK.570 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 173.2% | 173.2% | 8.658% | 26.65% | 4.36% | 11.54% | 40% |  |  | 15.75% | 21.65% | 1.481% | 7.531% | 1.618% | 100% | 36.74% | 17.32% | 18.37% | 34.42% | 4.498% |  | 34.64% | 34.64% | 6.695% |  | 86.6% | 43.3% | 40.36% |  | 26.65% | 86.6% | 0.6623% | 0.5934% | 0.8174% | 1.564% | 0.7852% | 1.523% | 0.8993% |  | 56.77% | 53.3% | 25.04%
217 | LOW 0.01 |  | 6/9/2023 11:45:32 AM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LOW 0.01.571 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 42.86% | 66.14% | 2.305% | 19.68% | 3.136% | 8.952% | 33.33% | 0% | 86.6% | 4.028% | 20.94% | 2.928% | 5.404% | 8.491% | 20.52% | 43.3% | 22.49% | 2.407% | 4.7% | 11.53% |  | 28.93% | 28.93% | 1.454% | 69.28% | 26.96% | 16.67% | 83.27% | 100% | 7.563% | 6.879% | 0.3673% | 1.84% | 0.9474% | 0.7389% | 1.658% | 0.8092% | 1.431% | 3.646% | 7.864% | 8.541% | 3.913%
218 | LOW 0.05 |  | 6/9/2023 11:51:20 AM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LOW 0.05.572 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 11.46% | 21.31% | 4.319% | 19.92% | 1.696% | 5.078% | 16.21% | 32.73% | 8.779% | 15.55% | 8.251% | 3.659% | 7.2% | 1.691% | 4.277% | 33.08% | 3.482% | 3.077% | 9.805% | 8.108% | 50% | 15.34% | 15.34% | 2.413% | 78.06% | 36.74% | 26.59% | 17.52% | 45.07% | 10.04% | 1.29% | 1.424% | 1.024% | 0.635% | 0.6804% | 0.3827% | 1.29% | 1.942% | 4.8% | 5.838% | 1.705% | 1.917%
219 | LOW 0.1 |  | 6/9/2023 11:57:09 AM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LOW 0.1.573 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 8.332% | 9.799% | 1.461% | 60.27% | 0.7321% | 5.008% | 6.522% | 16.54% | 14.9% | 4.713% | 3.638% | 6.016% | 2.161% | 4.257% | 6.027% | 3.273% | 4.803% | 4.013% | 2.413% | 7.889% | 33.33% | 5.03% | 5.03% | 3.359% | 20.15% | 39.31% | 6.013% | 20.39% | 12.37% | 4.909% | 0.6231% | 1.825% | 1.283% | 1.076% | 0.2006% | 1.706% | 1.319% | 1.875% | 5.18% | 0.636% | 3.197% | 0.2099%
220 | LOW 0.5 |  | 6/9/2023 12:02:57 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LOW 0.5.574 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.739% | 9.581% | 3.657% | 23.4% | 1.452% | 3.929% | 6.345% | 5.973% | 10.55% | 6.283% | 3.912% | 2.835% | 1.931% | 2.014% | 7.537% | 4.347% | 0.7635% | 3.363% | 1.3% | 155.4% | 9.934% | 3.731% | 3.731% | 2.779% | 5.714% | 1.546% | 4.464% | 25.25% | 21.27% | 4.142% | 3.434% | 2.946% | 0.292% | 0.5819% | 7.365% | 14.01% | 0.7157% | 2.644% | 3.879% | 1.292% | 2.48% | 2.817%
221 | STD 1 |  | 6/9/2023 12:08:45 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 1.575 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.582% | 2.208% | 4.624% | 36.84% | 3.021% | 2.678% | 7.119% | 11.84% | 3.134% | 3.137% | 2.411% | 1.283% | 0.5544% | 3.931% | 0.7491% | 5.446% | 2.129% | 1.692% | 1.812% | 4.889% | 10.62% | 1.443% | 1.443% | 2.234% | 20.26% | 3.352% | 7.31% | 6.105% | 8.979% | 1.729% | 0.6311% | 0.6116% | 0.7099% | 2.769% | 0.7788% | 1.58% | 2.355% | 2.789% | 2.987% | 0.09693% | 0.4495% | 1.027%
222 | STD 5 |  | 6/9/2023 12:14:33 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 5.576 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.386% | 1.539% | 1.922% | 6.508% | 2.371% | 4.717% | 2.495% | 3.598% | 0.3686% | 0.5161% | 0.1629% | 2.875% | 1.523% | 2.774% | 0.4769% | 0.3646% | 0.3266% | 0.8585% | 0.9612% | 1.771% | 12.41% | 0.6807% | 0.6807% | 2.319% | 4.246% | 5.31% | 5.005% | 4.448% | 4.497% | 0.7185% | 0.5339% | 0.6203% | 2.11% | 1.185% | 0.5942% | 1.504% | 1.625% | 2.701% | 1.646% | 1.165% | 0.8495% | 0.8291%
223 | STD 10 |  | 6/9/2023 12:20:21 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 10.577 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.8466% | 1.312% | 1.302% | 6.448% | 1.984% | 5.381% | 1.716% | 0.7226% | 0.4541% | 0.9115% | 1.285% | 2.175% | 1.016% | 1.185% | 1.187% | 1.529% | 0.5695% | 1.292% | 1.283% | 1.203% | 2.669% | 1.257% | 1.257% | 3.123% | 1.6% | 0.9435% | 1.854% | 1.818% | 4.954% | 0.9171% | 0.6624% | 0.9618% | 0.9593% | 3.411% | 0.9061% | 0.4659% | 0.1279% | 1.906% | 0.6047% | 0.3581% | 0.3966% | 0.7369%
224 | STD 25 |  | 6/9/2023 12:26:09 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 25.578 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.523% | 1.354% | 3.848% | 6.549% | 1.892% | 1.178% | 2.085% | 2.344% | 1.996% | 1.305% | 1.109% | 1.001% | 1.01% | 0.5627% | 1.039% | 1.145% | 1.924% | 1.316% | 0.2188% | 2.901% | 2.63% | 0.4879% | 0.4879% | 3.994% | 5.027% | 1.29% | 0.1598% | 2.853% | 1.093% | 1.122% | 0.3567% | 4.113% | 1.222% | 3.954% | 0.7926% | 1.09% | 1.664% | 2.126% | 0.5282% | 0.8927% | 1.3% | 1.613%
225 | STD 50 |  | 6/9/2023 12:31:57 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 50.579 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.8209% | 1.502% | 1.342% | 3.227% | 0.9976% | 1.311% | 2.426% | 0.827% | 0.4821% | 0.2543% | 1.635% | 1.984% | 0.5617% | 1.852% | 0.3026% | 1.075% | 0.4177% | 1.464% | 1.524% | 0.7162% | 1.648% | 1.094% | 1.094% | 2.972% | 1.175% | 0.8982% | 1.622% | 1.835% | 1.276% | 1.237% | 0.3869% | 0.4796% | 1.006% | 1.65% | 0.6768% | 6.87% | 1.442% | 2.755% | 1.532% | 1.659% | 0.9381% | 2.484%
226 | STD 100 |  | 6/9/2023 12:37:44 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 100.580 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.738% | 2.039% | 2.591% | 2.332% | 0.1248% | 1.31% | 2.871% | 1.966% | 2.462% | 2.448% | 1.448% | 0.766% | 1.391% | 2.513% | 2.117% | 2.395% | 2.045% | 0.7567% | 0.8525% | 2.719% | 1.234% | 0.2432% | 0.2432% | 3.488% | 2.126% | 0.344% | 1.44% | 1.8% | 0.8696% | 1.711% | 0.4826% | 13.82% | 0.7724% | 7.11% | 1.121% | 0.3927% | 0.5599% | 1.78% | 0.589% | 0.787% | 2.553% | 0.624%
228 | LBLANK |  | 6/9/2023 12:43:32 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.581 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 24.74% | 173.2% | 9.022% | 49.49% | 1.918% | 6.952% | 50.76% |  | 108.3% | 52.68% | 20% | 7.97% | 14.29% | 6.625% | 10.19% | 35.25% | 14.78% | 7.663% | 8.318% | 2.591% | 86.6% | 34.34% | 34.34% | 9.005% |  | 50% | 33.33% | 49.19% | 100% | 7.505% | 4.762% | 15.71% | 0.7355% | 0.8566% | 0.6528% | 0.7962% | 0.8173% | 2.215% | 137.8% | 30.05% | 33.63% | 12.43%
229 | LBLANK |  | 6/9/2023 1:13:17 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.583 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 44.1% |  | 41.84% | 12.37% | 4.072% | 28.6% | 8.66% |  | 114.6% | 50.94% | 16.65% | 4.848% | 54.53% | 11.62% | 57.28% | 22.91% | 45.07% | 6.977% | 43.48% | 3.038% | 173.2% | 31.23% | 31.23% | 4.378% | 173.2% | 50% | 173.2% | 31.88% |  | 1.011% | 22.91% | 15.24% | 0.4477% | 3.995% | 0.9617% | 0.8764% | 0.6516% | 3.854% | 173.2% | 44.41% | 48.71% | 14.84%
230 | LBLANK |  | 6/9/2023 1:19:06 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.584 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 34.64% |  | 2.639% | 43.3% | 4.995% | 5.121% | 22.91% |  | 173.2% | 0% | 33.82% | 3.375% | 34.32% | 5.564% | 50% | 45.83% | 20.38% | 14.72% | 15.05% | 2.946% |  | 52.92% | 52.92% | 6.308% |  | 50% | 34.64% | 101.5% |  | 9.369% | 14.78% | 1.166% | 0.3643% | 2.559% | 1.811% | 0.9652% | 2.079% | 0.8685% | 173.2% | 34.64% | 28.36% | 21.43%
231 | ICV PE |  | 6/9/2023 1:24:53 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\ICV PE.585 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.265% | 3.629% | 1.365% | 10.43% | 1.177% | 2.34% | 2.803% | 1.532% | 1.336% | 0.2524% | 1.758% | 1.873% | 0.9233% | 2.583% | 3.04% | 1.299% | 0.806% | 0.3543% | 1.287% | 1.013% | 7.186% | 0.4243% | 0.4243% | 1.098% | 1.871% | 0.7698% | 1.581% | 1.138% | 3.545% | 0.7911% | 1.109% | 2.723% | 1.423% | 4.253% | 0.5821% | 0.7338% | 0.491% | 2.63% | 0.2308% | 0.43% | 0.5957% | 1.315%
232 | ICV PE |  | 6/9/2023 1:31:10 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\ICV PE.586 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.7826% | 1.269% | 2.681% | 4.501% | 0.2219% | 0.7935% | 1.255% | 2.007% | 1.463% | 0.4055% | 3.436% | 4.058% | 0.3809% | 0.7396% | 1.828% | 1.551% | 0.1224% | 2.353% | 1.138% | 1.235% | 5.672% | 1.124% | 1.124% | 0.1948% | 2.922% | 1.112% | 1.655% | 1.395% | 2.174% | 2.897% | 0.547% | 1.911% | 1.198% | 2.819% | 14.57% | 0.7403% | 0.8267% | 1.843% | 3.609% | 1.259% | 0.6054% | 0.8791%
233 | CCV 1ppb TRACE |  | 6/9/2023 1:36:58 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.587 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.397% | 10.56% | 1.646% | 1.431% | 3.194% | 9.465% | 4.189% | 1.8% | 2.685% | 1.676% | 0.8708% | 0.9592% | 1.659% | 5.173% | 2.411% | 4.828% | 5.773% | 1.92% | 2.901% | 10.21% | 9.805% | 0.7081% | 0.7081% | 6.548% | 10.73% | 3.409% | 1.846% | 4.448% | 13.35% | 1.59% | 0.9142% | 0.7307% | 0.5671% | 9.904% | 1.524% | 0.368% | 0.9798% | 1.813% | 1.185% | 1.471% | 1.224% | 2.108%
234 | CCV 1ppb TRACE |  | 6/9/2023 1:43:15 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.588 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.539% | 7.525% | 0.8529% | 13.92% | 2.65% | 3.966% | 3.37% | 10.23% | 1.574% | 4.301% | 1.495% | 0.5585% | 1.859% | 2.979% | 2.024% | 7.369% | 4.045% | 1.526% | 2.698% | 4.238% | 4.809% | 1.756% | 1.756% | 7.149% | 5.679% | 4.854% | 5.096% | 12.4% | 5.315% | 1.193% | 0.5492% | 5.78% | 0.6314% | 2.097% | 1.011% | 0.4695% | 1.109% | 1.02% | 1.84% | 0.8897% | 2.724% | 0.7695%
235 | LFB |  | 6/9/2023 1:49:03 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.589 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.355% | 1.442% | 3.139% | 2.99% | 1.224% | 2.751% | 1.214% | 4.913% | 2.836% | 2.132% | 0.7214% | 1.516% | 1.529% | 1.389% | 0.8132% | 1.498% | 1.664% | 0.3031% | 1.941% | 2.071% | 5.31% | 0.4289% | 0.4289% | 2.171% | 4.062% | 1.515% | 0.8078% | 2.46% | 5.842% | 0.4368% | 1.085% | 1.275% | 1.445% | 1.39% | 0.4243% | 1.086% | 1.012% | 2.508% | 0.7153% | 0.492% | 0.3728% | 1.773%
236 | LLOQ 0.01 |  | 6/9/2023 1:54:51 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LLOQ 0.01.590 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0% | 31.49% | 1.464% | 40.88% | 3.952% | 3.939% | 24.77% | 69.28% | 18.18% | 35.69% | 11.68% | 2.99% | 6.604% | 2.491% | 11.55% | 72.63% | 14.84% | 6.033% | 10.75% | 3.922% | 173.2% | 15.02% | 15.02% | 7.419% | 100% | 96.44% | 75.5% | 104.3% | 0% | 5.695% | 7.911% | 1.702% | 1.889% | 1.14% | 1.595% | 1.349% | 0.7058% | 4.215% | 6.963% | 12.02% | 1.654% | 2.477%
237 | CCV 10 ppb MAJORS |  | 6/9/2023 2:00:40 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.591 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.495% | 4.085% | 0.3285% | 9.321% | 1.217% | 1.862% | 3.306% | 3.013% | 1.744% | 2.129% | 1.255% | 2.307% | 1.269% | 1.012% | 1.036% | 2.215% | 0.6349% | 0.9508% | 0.7969% | 1.296% | 3.448% | 0.9911% | 0.9911% | 2.409% | 7.784% | 3.039% | 1.036% | 2.197% | 5.899% | 1.14% | 0.5821% | 1.26% | 0.4545% | 3.227% | 3.636% | 1.393% | 1.728% | 2.451% | 1.84% | 0.8656% | 1.126% | 1.999%
238 | CCV 10 ppb MAJORS |  | 6/9/2023 2:06:57 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.592 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.96% | 4.243% | 0.2451% | 9.558% | 2.206% | 2.825% | 3.139% | 1.295% | 1.333% | 1.044% | 0.2491% | 0.8804% | 0.6063% | 1.71% | 1.151% | 1.409% | 1.385% | 0.539% | 1.459% | 0.6396% | 2.979% | 0.4272% | 0.4272% | 1.247% | 6.456% | 0.5625% | 2.314% | 0.992% | 1.854% | 0.3972% | 0.3415% | 0.2296% | 0.8037% | 1.232% | 0.336% | 0.8345% | 0.7409% | 1.433% | 0.8466% | 0.6393% | 1.108% | 1.412%
239 | LLOQ 0.1 |  | 6/9/2023 2:12:45 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LLOQ 0.1.593 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 13.31% | 1.665% | 1.5% | 6.186% | 5.468% | 2.972% | 12.08% | 41.67% | 6.415% | 6.662% | 2.536% | 0.7044% | 3.172% | 6.878% | 6.667% | 20.09% | 3.544% | 4.094% | 31.24% | 14.03% | 62.45% | 2.098% | 2.098% | 1.457% | 47.78% | 10.58% | 22.59% | 84.67% | 29.4% | 1.12% | 4.028% | 2.084% | 0.8149% | 1.134% | 0.3456% | 1.352% | 0.6818% | 1.147% | 5.348% | 1.675% | 3.491% | 1.095%
240 | 8_30_21_F_9 |  | 6/9/2023 2:19:27 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_F_9.594 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.952% | 2.906% | 3.101% | 7.362% | 0.8301% | 4.456% | 5.492% | 3.438% | 3.797% | 17.56% | 4.825% | 2.081% | 1.278% | 1.278% | 31.39% | 4.412% | 1.487% | 1.176% | 2.252% | 3.959% | 4.094% | 1.209% | 1.209% | 6.374% | 33.33% | 22.43% | 16.52% | 6.853% | 37.8% | 1.524% | 0.7876% | 1.214% | 0.5334% | 0.7421% | 0.7146% | 0.5728% | 1.536% | 2.363% | 9.644% | 3.614% | 4.324% | 3.634%
241 | 8_30_21_F_10 |  | 6/9/2023 2:25:14 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_F_10.595 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.53% | 1.547% | 1.683% | 9.252% | 0.8479% | 1.457% | 5.151% | 2.24% | 4.294% | 22.37% | 7.637% | 2.052% | 0.379% | 3.8% | 24.37% | 15.98% | 2.039% | 2.518% | 3.968% | 2.145% | 1.478% | 0.7195% | 0.7195% | 5.472% | 28.87% | 33.05% | 20.28% | 15.07% | 28.64% | 2.233% | 0.5733% | 0.6525% | 0.3767% | 3.578% | 4.496% | 1.659% | 3.656% | 1.553% | 29.06% | 12.89% | 3.684% | 4.021%
242 | 8_30_21_F_11 |  | 6/9/2023 2:31:03 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_F_11.596 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.989% | 1.216% | 1.46% | 10.78% | 1.871% | 1.389% | 5.286% | 1.96% | 3.407% | 8.698% | 0.8792% | 2.246% | 0.9516% | 0.6018% | 13.86% | 8.983% | 1.537% | 2.119% | 0.1882% | 1.134% | 6.039% | 1.265% | 1.265% | 2.895% | 7.531% | 17.32% | 12.92% | 24.14% | 33.07% | 1.526% | 0.8907% | 0.5886% | 0.6527% | 1.832% | 1.149% | 9.273% | 1.648% | 2.532% | 8% | 9.811% | 1.386% | 2.93%
243 | 8_30_21_UF_1 |  | 6/9/2023 2:36:51 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_UF_1.597 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.245% | 3.802% | 1.148% | 15.11% | 0.7251% | 3.137% | 6.822% | 2.167% | 3.382% | 16.57% | 4.052% | 0.4174% | 1.393% | 4.218% | 12.76% | 2.102% | 0.773% | 1.186% | 0.775% | 0.1389% | 1.928% | 1.067% | 1.067% | 0.3356% | 24.74% | 17.32% | 9.962% | 19.29% | 78.06% | 0.4268% | 0.3701% | 1.437% | 0.5175% | 5.774% | 1.353% | 1.215% | 1.494% | 4.192% | 14.84% | 6.07% | 3.704% | 3.149%
244 | 8_30_21_UF_6 |  | 6/9/2023 2:42:39 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_UF_6.598 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.631% | 1.253% | 3.554% | 62.48% | 1.397% | 0.7403% | 2.076% | 1.476% | 2.45% | 17.16% | 8.126% | 4.901% | 1.612% | 5.794% | 16.87% | 11.89% | 1.001% | 2.353% | 3.511% | 1.145% | 10.42% | 1.601% | 1.601% | 2.989% | 57.28% | 28.36% | 9.7% | 12.03% | 24.74% | 3.159% | 1.205% | 0.5587% | 10.53% | 0.8394% | 11.8% | 0.9709% | 1.227% | 13.27% | 8.296% | 13.51% | 2.853% | 3.075%
245 | 8_30_21_UF_8 |  | 6/9/2023 2:48:27 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_UF_8.599 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.104% | 1.525% | 1.356% | 19.93% | 1.513% | 4.8% | 5.246% | 3.023% | 5.137% | 9.799% | 8.759% | 5.349% | 1.095% | 2.259% | 36.25% | 6.344% | 0.9478% | 0.8859% | 2.338% | 1.686% | 1.555% | 0.07468% | 0.07468% | 3.181% | 25% | 20.9% | 16.06% | 33.97% | 39.03% | 0.6578% | 0.08196% | 0.9796% | 0.3013% | 1.754% | 0.4663% | 0.7182% | 1.557% | 1.112% | 27.04% | 14.87% | 6.154% | 2.569%
246 | 8_30_21_UF_9 |  | 6/9/2023 2:54:15 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_UF_9.600 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.76% | 3.158% | 2.868% | 12.6% | 0.9831% | 1.969% | 5.128% | 2.826% | 0.9295% | 17.63% | 2% | 1.34% | 0.3817% | 4.472% | 7.905% | 4.703% | 4.266% | 0.7263% | 17.37% | 2.103% | 14.94% | 1.503% | 1.503% | 4.021% | 25% | 16.9% | 8.056% | 10.51% | 13.32% | 0.5286% | 0.9492% | 3.237% | 2.818% | 5.445% | 0.6707% | 0.5695% | 1.097% | 2.003% | 4.028% | 13.8% | 1.367% | 1.079%
247 | 8_30_21_UF_9_LD |  | 6/9/2023 3:00:03 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_UF_9_LD.601 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.5444% | 0.9975% | 0.8695% | 9.488% | 1.289% | 3.249% | 2.419% | 3.47% | 8.47% | 8.534% | 12.48% | 4.952% | 1.121% | 0.5824% | 25.82% | 9.216% | 1.689% | 1.547% | 16.79% | 1.863% | 4.807% | 0.5867% | 0.5867% | 4.463% | 89.21% | 4.33% | 38.87% | 10.58% | 35.25% | 1.137% | 0.5806% | 2.52% | 0.3991% | 2.459% | 6.431% | 0.4269% | 1.181% | 1.131% | 31.6% | 16.94% | 3.863% | 0.9636%
248 | 8_30_21_UF_9_LD |  | 6/9/2023 3:06:19 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_UF_9_LD.602 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.7374% | 2.908% | 2.6% | 14.12% | 1.839% | 3.267% | 5.205% | 0.8985% | 4.632% | 7.862% | 4.263% | 2.643% | 1.142% | 2.294% | 27.24% | 11.29% | 4.221% | 0.4443% | 3.707% | 1.257% | 10.47% | 1.248% | 1.248% | 7.72% | 61.97% | 27.15% | 13.18% | 5.809% | 18.33% | 0.4552% | 0.5729% | 1.372% | 0.204% | 3.743% | 1.448% | 0.7068% | 0.9383% | 0.7132% | 4.499% | 11.3% | 3.548% | 3.459%
249 | 8_30_21_UF_10 |  | 6/9/2023 3:12:07 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_UF_10.603 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.3272% | 0.9996% | 0.7398% | 17.54% | 0.6286% | 1.371% | 4.842% | 1.013% | 3.856% | 10.87% | 8.438% | 4.15% | 1.275% | 1.513% | 15.8% | 2.706% | 3.458% | 0.5801% | 2.458% | 0.6859% | 9.512% | 1.416% | 1.416% | 3.35% | 65.47% | 22.79% | 21.12% | 8.431% | 32.83% | 1.146% | 0.1605% | 0.2173% | 1.153% | 1.461% | 1.049% | 1.291% | 0.1052% | 2.285% | 15.28% | 1.69% | 2.926% | 2.682%
250 | LBLANK |  | 6/9/2023 3:17:56 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.604 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 173.2% | 173.2% | 4.007% | 43.3% | 1.552% | 3.096% | 25% | 86.6% |  | 66.67% | 19.68% | 1.229% | 20.83% | 3.956% | 86.6% | 39.74% | 35.25% | 15.94% | 6.479% | 17.2% | 43.3% | 6.387% | 6.387% | 3.538% | 173.2% | 173.2% | 124.9% | 86.22% | 173.2% | 18.33% | 28.87% | 8.876% | 0.5208% | 4.792% | 0.2839% | 0.8184% | 1.055% | 2.787% | 173.2% | 33.33% | 47.19% | 19.81%
251 | CCV 1ppb TRACE |  | 6/9/2023 3:23:44 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.605 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4.388% | 4.118% | 2.446% | 14.04% | 2.16% | 10.2% | 2.134% | 15.81% | 6.627% | 7.238% | 1.049% | 3.45% | 2.063% | 3.419% | 3.825% | 1.533% | 3.11% | 1.393% | 1.752% | 4.404% | 13.56% | 1.756% | 1.756% | 0.9862% | 9.375% | 1.861% | 3.846% | 9.782% | 5.735% | 0.5764% | 0.4589% | 1.358% | 0.5748% | 3.713% | 1.627% | 5.649% | 0.6799% | 1.078% | 3.661% | 1.476% | 1.677% | 1.57%
252 | CCV 1ppb TRACE |  | 6/9/2023 3:30:02 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.606 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 5.626% | 7.001% | 1.022% | 14.1% | 1.987% | 7.619% | 4.176% | 7.719% | 6.277% | 4.67% | 2.712% | 1.053% | 1.86% | 1.991% | 1.467% | 5.12% | 4.931% | 1.292% | 2.394% | 6.14% | 4.398% | 1.475% | 1.475% | 4.87% | 0% | 8.109% | 3.126% | 1.621% | 5.634% | 1.714% | 1.057% | 1.574% | 0.211% | 0.6129% | 1.745% | 0.8026% | 1.759% | 2.836% | 0.1542% | 2.228% | 1.114% | 1.089%
253 | CCV 10 ppb MAJORS |  | 6/9/2023 3:35:51 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.607 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.255% | 1.173% | 1.798% | 8.852% | 1.796% | 2.511% | 2.39% | 3.183% | 1.606% | 2.204% | 0.4784% | 3.18% | 2.013% | 1.433% | 0.6198% | 0.7729% | 2.515% | 1.06% | 0.5482% | 1.525% | 4.295% | 1.32% | 1.32% | 4.052% | 3.021% | 2.417% | 0.3354% | 1.354% | 3.905% | 0.9849% | 1.25% | 0.8051% | 9.267% | 1.287% | 1.042% | 0.7017% | 2.35% | 10.17% | 0.2317% | 0.8021% | 0.09631% | 0.3579%
254 | CCV 10 ppb MAJORS |  | 6/9/2023 3:42:08 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.608 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.42% | 1.639% | 2.689% | 7.021% | 4.31% | 2.797% | 1.684% | 2.782% | 1.433% | 0.8159% | 0.6667% | 0.3551% | 1.213% | 2.035% | 0.716% | 1.84% | 0.9378% | 0.7138% | 0.1262% | 0.563% | 2.977% | 0.5493% | 0.5493% | 3.466% | 2.613% | 2.925% | 0.6658% | 1.506% | 5.472% | 0.6427% | 1.216% | 11.06% | 0.7684% | 1.187% | 0.3132% | 1.553% | 18.64% | 1.259% | 1.557% | 0.3786% | 1.948% | 1.185%
255 | 8_30_21_UF_10_LSP |  | 6/9/2023 3:47:56 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_UF_10_LSP.609 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.349% | 0.1411% | 0.7103% | 5.67% | 1.304% | 1.263% | 0.7026% | 0.738% | 1.961% | 0.7498% | 1.686% | 0.5431% | 1.11% | 0.4006% | 0.2955% | 1.401% | 0.7608% | 0.8238% | 1.537% | 1.262% | 3.753% | 1.131% | 1.131% | 2.174% | 3.055% | 1.497% | 1.803% | 1.691% | 4.512% | 1.109% | 0.9776% | 6.686% | 2.493% | 1.972% | 0.7269% | 4.064% | 0.2291% | 3.865% | 1.522% | 1.06% | 2.286% | 0.8319%
256 | 8_30_21_UF_10_DUP_SP |  | 6/9/2023 3:53:45 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_UF_10_DUP_SP.610 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.8869% | 1.649% | 1.186% | 3.527% | 1.847% | 3.179% | 1.67% | 0.5111% | 1.821% | 2.559% | 0.6054% | 1.599% | 0.5155% | 0.7491% | 0.7411% | 2.819% | 0.1602% | 0.7966% | 1.832% | 1.273% | 4.902% | 0.6178% | 0.6178% | 0.7817% | 2.822% | 2.517% | 0.4921% | 5.819% | 5.601% | 2.217% | 0.6467% | 0.5516% | 0.4731% | 1.335% | 2.533% | 1.251% | 0.691% | 1.585% | 1.649% | 1.092% | 0.9389% | 0.4365%
257 | 8_30_21_UF_11 |  | 6/9/2023 3:59:33 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_UF_11.611 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.4569% | 3.189% | 3.296% | 10.18% | 1.381% | 1.229% | 2.777% | 0.4226% | 10.5% | 7.211% | 9.901% | 0.8433% | 0.6681% | 0.6114% | 40.69% | 5.951% | 2.486% | 2.72% | 2.053% | 2.991% | 4.705% | 1.554% | 1.554% | 1.042% | 10.19% | 25% | 3.851% | 40.3% | 28.57% | 0.3891% | 1.368% | 2.354% | 0.2224% | 0.6777% | 0.3621% | 0.8093% | 0.6421% | 2.235% | 10.18% | 5.879% | 2.063% | 0.8739%
258 | 8_30_21_BB |  | 6/9/2023 4:05:21 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_BB.612 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 24.74% | 86.6% | 1.447% | 2.193% | 1.906% | 4.075% | 20.15% | 100% | 50% | 27.88% | 2.941% | 2.476% | 12.59% | 0.8053% | 114.6% | 43.83% | 4.683% | 3.41% | 4.089% | 8.664% | 0% | 28.89% | 28.89% | 2.526% |  | 45.83% | 124.9% | 77.07% |  | 15.23% | 31.17% | 0.814% | 0.7483% | 0.7162% | 9.986% | 0.3897% | 0.8353% | 0.4636% | 86.6% | 22.91% | 1.882% | 0.5809%
259 | 8_30_21_FB |  | 6/9/2023 4:11:09 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_FB.613 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 28.35% | 69.28% | 1.665% | 1.14% | 5.681% | 2.905% | 41.66% | 39.03% | 114.6% | 9% | 7.956% | 1.41% | 3.024% | 3.254% | 43.59% | 0.6242% | 2.654% | 0.7428% | 1.067% | 6.646% |  | 12.02% | 12.02% | 4.457% | 86.6% | 91.65% | 43.3% | 38.38% | 173.2% | 12.88% | 8.183% | 1.721% | 0.9475% | 0.1432% | 1.18% | 0.5146% | 3.437% | 2.342% | 13.32% | 4.949% | 6.02% | 3.028%
260 | 8_30_21_LAB_A |  | 6/9/2023 4:16:56 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_LAB_A.614 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 100% |  | 3.083% | 30.77% | 5.022% | 3.344% | 44.41% |  | 34.64% | 21.44% | 4.646% | 3.792% | 14% | 2.556% | 100% | 15.3% | 7.924% | 1.319% | 4.293% | 3.624% |  | 17.22% | 17.22% | 0.9652% |  | 173.2% | 173.2% | 70.37% |  | 5.447% | 4.681% | 1.746% | 1.115% | 1.406% | 0.9107% | 2.348% | 2.376% | 1.211% | 86.6% | 50% | 0.4738% | 4.788%
261 | 8_30_21_LAB_B |  | 6/9/2023 4:22:44 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_LAB_B.615 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 91.65% |  | 2.742% | 26.96% | 1.818% | 8.349% | 24.74% | 100% |  | 28.97% | 10.79% | 4.416% | 17.58% | 3.874% | 86.6% | 10.28% | 7.977% | 4.188% | 5.327% | 11.67% | 173.2% | 26.41% | 26.41% | 2.529% |  | 100% | 86.6% | 106.8% | 173.2% | 40.06% | 50.76% | 1.641% | 1.699% | 1.699% | 3.759% | 0.2761% | 0.878% | 3.94% |  | 33.33% | 2.603% | 0.7241%
262 | 9_13_21_W_1 |  | 6/9/2023 4:28:32 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_W_1.616 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.197% | 4.921% | 1.01% | 8.971% | 1.419% | 2.308% | 6.142% | 1.278% | 15.57% | 34.04% | 8.45% | 0.5561% | 0.6472% | 1.968% | 35.44% | 15.74% | 3.876% | 2.93% | 3.91% | 8.893% | 12.5% | 0.3159% | 0.3159% | 4.26% | 34.64% | 49.49% | 65.47% | 60.8% | 173.2% | 1.241% | 0.5189% | 1.711% | 0.1605% | 4.198% | 1.216% | 0.3187% | 0.7964% | 3.785% | 49.49% | 25.19% | 3.925% | 1.674%
263 | 9_13_21_W_6 |  | 6/9/2023 4:34:20 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_W_6.617 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.664% | 6.315% | 1.999% | 4.519% | 1.93% | 2.555% | 19.02% | 3.321% | 8.483% | 12.88% | 7.756% | 1.418% | 0.6458% | 0.1973% | 18.33% | 3.704% | 4.286% | 2.474% | 1.675% | 11.33% | 24.46% | 0.6023% | 0.6023% | 2.87% | 86.6% | 0% | 9.116% | 20.55% | 49.49% | 0.4736% | 1.465% | 0.7042% | 1.049% | 0.39% | 0.4836% | 0.6744% | 2.204% | 2.677% | 40% | 12.5% | 1.483% | 1.331%
264 | 9_13_21_W_8 |  | 6/9/2023 4:40:07 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_W_8.618 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.662% | 3.985% | 1.825% | 9.522% | 1.148% | 8.588% | 6.06% | 3.106% | 20.37% | 16.26% | 5.936% | 4.611% | 0.6094% | 0.7349% | 10% | 18.36% | 4.217% | 6.097% | 2.755% | 6.808% | 19.68% | 3.461% | 3.461% | 1.891% | 43.3% | 20% | 33.33% | 36.73% | 33.33% | 5.641% | 0.2826% | 2.482% | 2.347% | 1.263% | 23.81% | 0.8854% | 1.545% | 3.874% | 24.05% | 13.35% | 3.701% | 2.295%
265 | LBLANK |  | 6/9/2023 4:45:55 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.619 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 114.6% |  | 4.326% | 86.6% | 2.083% | 9.927% | 21.65% |  | 173.2% | 66.14% | 11.55% | 10.4% | 10.83% | 4.199% | 173.2% | 50% | 25% | 6.061% | 3.675% | 6.297% |  | 50% | 50% | 3.917% |  | 173.2% | 100% | 454.9% |  | 25.47% | 33.07% | 0.8423% | 0.8979% | 2.629% | 1.447% | 0.8619% | 0.213% | 2.021% | 100% | 25% | 28.87% | 25.47%
266 | CCV 1ppb TRACE |  | 6/9/2023 4:51:44 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.620 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 5.551% | 8.382% | 0.9446% | 2.341% | 3.13% | 3.413% | 4.186% | 10.53% | 5.049% | 4.691% | 0.8214% | 2.049% | 2.329% | 5.219% | 1.962% | 8.394% | 1.477% | 0.8802% | 2.408% | 22.55% | 18.61% | 1.15% | 1.15% | 3.737% | 9.913% | 8.497% | 8.942% | 23.94% | 8.411% | 0.9236% | 1.311% | 1.05% | 0.6168% | 2.365% | 0.7681% | 1.512% | 0.7936% | 1.078% | 2.359% | 0.4576% | 0.3197% | 1.31%
267 | CCV 1ppb TRACE |  | 6/9/2023 4:58:02 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.621 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.345% | 5.523% | 0.9033% | 13.35% | 1.342% | 4.592% | 2.006% | 6.913% | 2.092% | 1.461% | 2.524% | 3.9% | 1.508% | 2.173% | 0.8537% | 4.628% | 1.048% | 1.524% | 0.6535% | 3.23% | 15.15% | 1.401% | 1.401% | 7.878% | 4.467% | 3.812% | 3.868% | 11.92% | 15.83% | 1.491% | 0.6619% | 1.26% | 0.9046% | 1.009% | 0.4907% | 0.8966% | 1.375% | 1.213% | 0.5327% | 0.5985% | 0.6942% | 1.307%
268 | LFB |  | 6/9/2023 5:03:50 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.622 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.149% | 1.103% | 2.837% | 2.632% | 2.074% | 0.4168% | 4.833% | 4.967% | 1.446% | 0.9393% | 1.107% | 0.4579% | 0.2746% | 1.789% | 2.171% | 0.7783% | 0.5115% | 0.2112% | 0.2863% | 0.4274% | 6.888% | 0.3746% | 0.3746% | 1.369% | 3.891% | 0.7262% | 1.422% | 1.077% | 3.084% | 1.443% | 0.5929% | 2.218% | 0.7422% | 1.246% | 1.692% | 1.31% | 0.7621% | 1.379% | 1.224% | 0.9486% | 0.9932% | 0.5835%
269 | CCV 10 ppb MAJORS |  | 6/9/2023 5:09:38 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.623 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.102% | 3.318% | 2.01% | 9.185% | 1.007% | 1.179% | 1.722% | 1.373% | 1.045% | 2.531% | 2.085% | 1.823% | 2.11% | 0.9347% | 2.905% | 2.336% | 1.651% | 1.44% | 1.329% | 0.2386% | 5.947% | 1.602% | 1.602% | 3.232% | 7.815% | 3.458% | 1.738% | 0.4075% | 7.934% | 1.803% | 1.582% | 3.166% | 4.009% | 2.617% | 8.037% | 3.562% | 0.3541% | 2.135% | 0.4203% | 1.261% | 1.217% | 1.503%
270 | CCV 10 ppb MAJORS |  | 6/9/2023 5:15:55 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.624 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.027% | 4.489% | 0.9322% | 4% | 1.431% | 3.229% | 2.78% | 2.511% | 2.397% | 1.429% | 0.3032% | 0.6339% | 0.3249% | 1.345% | 1.09% | 1.237% | 1.665% | 0.7746% | 0.2564% | 2.271% | 1.342% | 0.673% | 0.673% | 1.201% | 3.145% | 1.34% | 2.056% | 3.393% | 0.9315% | 0.8967% | 0.8627% | 1.002% | 0.8197% | 0.786% | 0.3963% | 0.5005% | 1.964% | 1.09% | 1.128% | 0.6594% | 0.393% | 0.953%
271 | 9_13_21_W_9 |  | 6/9/2023 5:21:43 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_W_9.625 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.988% | 1.989% | 1.11% | 16.04% | 1.017% | 3.825% | 1.474% | 5.907% | 20.29% | 18.23% | 6.256% | 0.5063% | 1.198% | 1.587% | 45.07% | 36.47% | 1.33% | 2.034% | 1.24% | 11.16% | 19.79% | 2.793% | 2.793% | 4.779% | 86.6% | 45.83% | 17.32% | 88.81% | 24.74% | 0.9685% | 2.29% | 13.73% | 0.5977% | 1.193% | 0.7479% | 0.8537% | 0.6691% | 2.345% | 8.66% | 27.09% | 4.029% | 1.297%
272 | 9_13_21_W_10 |  | 6/9/2023 5:27:31 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_W_10.626 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.171% | 3.125% | 1.254% | 9.052% | 1.559% | 2.007% | 8.672% | 3.2% | 7.138% | 25.81% | 15.5% | 2.103% | 0.7209% | 1.442% | 15.75% | 11.62% | 1.635% | 1.947% | 0.411% | 14.55% | 8.268% | 1.508% | 1.508% | 6.211% | 173.2% | 10.19% | 22.3% | 38.91% | 173.2% | 0.8207% | 2.215% | 0.587% | 0.4249% | 0.1409% | 0.7192% | 1.239% | 0.4237% | 2.127% | 40.52% | 9.557% | 1.478% | 2.573%
273 | 9_13_21_W_10_DUP |  | 6/9/2023 5:33:20 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_W_10_DUP.627 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 5.11% | 0.9124% | 2.599% | 12.37% | 2.547% | 3.666% | 2.521% | 5.45% | 6.681% | 41.1% | 14.24% | 1.755% | 0.4748% | 2.254% | 43.99% | 11.11% | 2.766% | 0.8665% | 4.138% | 2.301% | 27.24% | 1.493% | 1.493% | 4.735% | 91.65% | 13.32% | 54.11% | 19.87% | 91.65% | 1.755% | 1.443% | 0.5151% | 1.36% | 1.166% | 0.4041% | 2.168% | 1.155% | 1.937% | 15.8% | 22.63% | 1.985% | 3.23%
274 | 9_13_21_W_11 |  | 6/9/2023 5:39:08 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_W_11.628 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.695% | 9.178% | 1.217% | 1.004% | 0.7591% | 5.657% | 4.728% | 2.338% | 11.99% | 17.56% | 5.661% | 0.2278% | 1.08% | 0.08571% | 12.06% | 20.38% | 2.471% | 1.761% | 0.227% | 6.864% | 13.95% | 3.328% | 3.328% | 1.301% | 114.6% | 12.37% | 44.44% | 50.6% | 43.3% | 1.385% | 2.606% | 8.16% | 1.464% | 0.6951% | 0.5997% | 1.677% | 0.1948% | 1.709% | 47.04% | 9.363% | 2.712% | 0.4963%
275 | 9_13_21_F_1 |  | 6/9/2023 5:44:56 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_F_1.629 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.7306% | 4.646% | 1.918% | 12.37% | 0.1967% | 1.682% | 3.109% | 2.006% | 11.5% | 13.68% | 2.643% | 0.2226% | 0.4739% | 4.084% | 15.68% | 4.377% | 3.694% | 0.7204% | 2.175% | 1.42% | 7.375% | 1.052% | 1.052% | 1.858% | 78.06% | 26.46% | 23.32% | 23.81% | 25% | 1.539% | 0.6074% | 4.279% | 1.298% | 0.9488% | 0.6295% | 0.8144% | 1.132% | 2.898% | 32.87% | 10.98% | 4.06% | 1.641%
276 | 9_13_21_F_6 |  | 6/9/2023 5:50:44 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_F_6.630 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.022% | 0.9012% | 0.9876% | 3.808% | 0.7188% | 1.91% | 1.775% | 2.83% | 1.735% | 8.69% | 4.669% | 1.066% | 1.721% | 1.21% | 8.108% | 6.681% | 3.906% | 2.191% | 1.314% | 0.2428% | 8.5% | 0.4808% | 0.4808% | 3.785% | 72.11% | 17.64% | 14.75% | 8.772% | 60.09% | 0.5972% | 1.415% | 0.2528% | 1.13% | 6.022% | 2.067% | 1.781% | 1.131% | 3.901% | 16.57% | 5.822% | 5.143% | 3.201%
277 | 9_13_21_F_8 |  | 6/9/2023 5:56:32 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_F_8.631 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.478% | 4.807% | 1.162% | 11.85% | 0.9714% | 2.93% | 5.239% | 2.083% | 3.195% | 4.696% | 7.094% | 0.9009% | 0.9603% | 1.333% | 26.65% | 11.01% | 1.269% | 0.4038% | 0.6958% | 0.5303% | 8.952% | 0.7888% | 0.7888% | 3.769% | 31.49% | 25.7% | 6.788% | 12.38% | 52.92% | 0.3371% | 0.9908% | 1.081% | 0.1801% | 5.054% | 0.8072% | 0.9484% | 1.494% | 1.856% | 4.949% | 4.663% | 1.074% | 2.757%
278 | 9_13_21_F_9 |  | 6/9/2023 6:02:20 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_F_9.632 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.52% | 0.4797% | 1.016% | 10.64% | 0.7027% | 6.119% | 1.012% | 2.167% | 1.745% | 4.879% | 1.779% | 1.514% | 0.5094% | 1.984% | 30.05% | 14.65% | 5.81% | 1.555% | 1.611% | 0.3641% | 5.587% | 1.638% | 1.638% | 2.006% | 65.47% | 22.35% | 24.87% | 24.9% | 39.03% | 1.913% | 1.836% | 0.3624% | 3.666% | 5.978% | 1.724% | 0.8754% | 1.262% | 6.462% | 24.05% | 3.035% | 2.014% | 1.577%
279 | 9_13_21_F_10 |  | 6/9/2023 6:08:08 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_F_10.633 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.3453% | 0.5972% | 0.7557% | 6.998% | 2.323% | 1.8% | 4.916% | 1.506% | 3.462% | 8.351% | 10.58% | 1.392% | 1.246% | 3.88% | 43.3% | 5.892% | 3.507% | 2.253% | 2.349% | 2.061% | 3.354% | 0.6148% | 0.6148% | 3.388% | 13.32% | 28.35% | 40.51% | 15.67% | 47.19% | 2.51% | 1.247% | 0.383% | 0.8397% | 3.106% | 7.523% | 0.6527% | 0.4728% | 0.09927% | 14.84% | 26.02% | 4.981% | 2.236%
280 | 9_13_21_F_11 |  | 6/9/2023 6:13:56 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_F_11.634 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.427% | 0.879% | 1.057% | 8.614% | 1.086% | 2.431% | 5.251% | 0.6368% | 4.24% | 17.83% | 8.38% | 1.355% | 0.8832% | 0.3927% | 25.08% | 19.96% | 3.812% | 1.368% | 3.05% | 1.087% | 11.67% | 0.7196% | 0.7196% | 4.733% | 50% | 6.667% | 24.47% | 9.142% | 48.22% | 1.392% | 0.5607% | 1.915% | 0.2078% | 3.012% | 3.366% | 2.789% | 1.788% | 3.673% | 18.87% | 5.328% | 3.421% | 2.23%
281 | 9_13_21_F_11_LD |  | 6/9/2023 6:19:44 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_F_11_LD.635 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.081% | 1.445% | 2.342% | 15.36% | 0.8802% | 3.631% | 1.438% | 1.9% | 2.967% | 11.63% | 2.857% | 2.344% | 0.7315% | 2.461% | 12.56% | 12.55% | 1.204% | 1.503% | 1.587% | 1.529% | 5.677% | 2.2% | 2.2% | 3.093% | 86.6% | 19.5% | 24.74% | 7.231% | 32.73% | 2.256% | 1.175% | 2.377% | 2.108% | 5.02% | 0.06788% | 1.014% | 1.819% | 0.9668% | 4% | 5.914% | 3.165% | 4.08%
282 | 9_13_21_F_11_LD |  | 6/9/2023 6:26:00 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_F_11_LD.636 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.5812% | 0.9836% | 2.489% | 2.402% | 4.406% | 2.777% | 3.663% | 1.298% | 5.128% | 20.3% | 6.771% | 1.802% | 1.428% | 1.129% | 24.37% | 9.561% | 0.5308% | 2.224% | 1.982% | 1.423% | 4.84% | 0.8469% | 0.8469% | 1.562% | 26.96% | 21.5% | 20.99% | 26.61% | 20% | 0.7708% | 1.537% | 1.663% | 0.1337% | 0.9135% | 1.239% | 1.175% | 1.141% | 4.819% | 23.89% | 1.056% | 3.355% | 1.056%
283 | LBLANK |  | 6/9/2023 6:31:48 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.637 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 173.2% |  | 5.384% | 124.9% | 9.569% | 6.828% | 15.06% |  | 173.2% | 86.6% | 43.95% | 6.576% | 20% | 1.389% | 173.2% | 18.23% | 35.25% | 18.98% | 6.356% | 6.413% |  | 45.83% | 45.83% | 6.121% | 173.2% | 100% | 86.6% | 97.26% | 173.2% | 30.2% | 12.5% | 5.394% | 1.186% | 1.318% | 0.04205% | 1.383% | 1.301% | 3.179% | 173.2% | 56.77% | 21.65% | 26.96%
284 | CCV 1ppb TRACE |  | 6/9/2023 6:37:37 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.638 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.8956% | 11.68% | 1.108% | 28% | 0.7188% | 11.66% | 3.402% | 3.2% | 1.792% | 2.703% | 1.543% | 1.955% | 1.631% | 0.641% | 2.819% | 4.166% | 4.589% | 1.816% | 2.635% | 5.543% | 12.42% | 0.76% | 0.76% | 3.21% | 5.12% | 1.226% | 6.609% | 12.05% | 12.02% | 0.8049% | 2.045% | 0.604% | 0.5305% | 1.872% | 0.5734% | 3.518% | 0.521% | 3.35% | 1.289% | 1.307% | 0.8605% | 0.9733%
285 | CCV 1ppb TRACE |  | 6/9/2023 6:43:54 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.639 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 6.287% | 11.15% | 2.461% | 37.6% | 5.472% | 6.225% | 9.205% | 8.526% | 3.831% | 1.639% | 2.067% | 1.887% | 0.4454% | 3.866% | 0% | 7.287% | 0.5234% | 1.7% | 2.588% | 15.8% | 18.46% | 2.027% | 2.027% | 0.9473% | 4.642% | 7.887% | 4.981% | 10.78% | 4.744% | 1.461% | 0.7845% | 1.363% | 1.239% | 3.178% | 0.7482% | 14.15% | 0.2544% | 2.792% | 1.193% | 2.691% | 2.028% | 4.119%
286 | CCV 10 ppb MAJORS |  | 6/9/2023 6:49:43 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.640 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.38% | 4.689% | 3.28% | 4.618% | 1.182% | 0.9557% | 1.231% | 1.616% | 0.8032% | 1.009% | 1.364% | 0.6435% | 0.6342% | 1.561% | 1.942% | 1.673% | 1.468% | 0.8515% | 1.213% | 2.206% | 5.084% | 0.6233% | 0.6233% | 1.136% | 3.915% | 2.181% | 2.256% | 2.122% | 2.518% | 0.7988% | 0.1385% | 9.981% | 6.897% | 0.4346% | 0.357% | 0.6787% | 0.253% | 8.823% | 0.7828% | 0.4983% | 1.93% | 0.9864%
287 | CCV 10 ppb MAJORS |  | 6/9/2023 6:56:01 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.641 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.808% | 2.902% | 1.468% | 12.52% | 0.9275% | 0.4348% | 2.49% | 2.482% | 1.256% | 0.875% | 2.991% | 1.528% | 1.295% | 0.5024% | 1.852% | 1.788% | 0.4828% | 1.68% | 1.847% | 0.2166% | 2.594% | 0.2162% | 0.2162% | 1.137% | 4.51% | 1.253% | 3.577% | 3.66% | 2.443% | 1.64% | 0.7756% | 1.495% | 0.8839% | 1.35% | 1.917% | 10.99% | 0.7604% | 0.3437% | 2.149% | 1.715% | 1.127% | 1.161%
288 | 9_13_21_F_11_DUP |  | 6/9/2023 7:01:49 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_F_11_DUP.642 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.2125% | 3.57% | 3.54% | 2.865% | 1.735% | 2.158% | 6.528% | 2.715% | 6.463% | 14.69% | 3.723% | 0.9882% | 1.018% | 0.2936% | 12.06% | 4.009% | 4.297% | 0.2023% | 1.051% | 3.391% | 16.67% | 0.9199% | 0.9199% | 4.594% | 41.66% | 23.52% | 21.85% | 1.749% | 19.92% | 2.219% | 1.957% | 13.54% | 0.3303% | 0.9009% | 0.3685% | 0.5912% | 2.615% | 2.573% | 3.571% | 11.49% | 6.406% | 0.8488%
289 | 9_13_21_F_11_DUP_LSP |  | 6/9/2023 7:07:37 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_F_11_DUP_LSP.643 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.4243% | 0.586% | 2.185% | 8.708% | 0.6757% | 2.96% | 2.087% | 1.687% | 1.466% | 1.3% | 1.482% | 0.722% | 0.7694% | 0.4771% | 0.3995% | 2.264% | 0.6327% | 0.5021% | 0.781% | 0.7156% | 0.9164% | 0.8987% | 0.8987% | 0.1344% | 5.006% | 1.717% | 2.536% | 0.7272% | 0.621% | 0.6508% | 0.1403% | 1.282% | 6.093% | 1.678% | 1.42% | 2.223% | 0.2122% | 2.894% | 1.723% | 0.8831% | 0.3014% | 1.391%
290 | 9_13_21_UF_1 |  | 6/9/2023 7:13:24 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_UF_1.644 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.047% | 2.817% | 2.935% | 25.6% | 1.144% | 3.477% | 4.217% | 0.9733% | 5.505% | 27.85% | 6.863% | 2.942% | 0.2905% | 0.8181% | 6.614% | 2.6% | 8.944% | 0.4596% | 2.989% | 2.677% | 8.124% | 0.8257% | 0.8257% | 3.916% | 24.74% | 15.75% | 20.38% | 23.85% | 88.19% | 0.9714% | 1.083% | 9.958% | 1.597% | 2.334% | 0.3272% | 2.583% | 0.5097% | 2.582% | 11.5% | 20.97% | 7.22% | 3.456%
291 | 9_13_21_UF_6 |  | 6/9/2023 7:19:12 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_UF_6.645 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.923% | 4.075% | 2.093% | 19.87% | 0.6681% | 2.252% | 1.743% | 1.745% | 6.381% | 9.527% | 6.143% | 3.841% | 0.9474% | 2.209% | 14.42% | 8.296% | 0.3507% | 1.145% | 0.2347% | 1.836% | 7.566% | 1.917% | 1.917% | 6.039% | 12.37% | 13.29% | 24.07% | 7.888% | 72.11% | 0.9144% | 1.006% | 0.8797% | 2.995% | 0.8532% | 2.878% | 1.589% | 1.367% | 4.814% | 6.153% | 3.846% | 8.248% | 2.05%
292 | 9_13_21_UF_6_DUP |  | 6/9/2023 7:25:00 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_UF_6_DUP.646 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.6153% | 1.252% | 2.848% | 20.35% | 0.3814% | 2.811% | 3.083% | 0.9943% | 10.37% | 6.662% | 4.312% | 1.143% | 1.01% | 7.236% | 8.882% | 9.391% | 2.809% | 1.889% | 2.13% | 0.1307% | 9.316% | 1.9% | 1.9% | 8.259% | 75.78% | 26.96% | 14.32% | 15.51% | 49.49% | 1.854% | 1.38% | 0.7562% | 7.21% | 1.161% | 0.6748% | 1.753% | 0.9393% | 3.764% | 5.808% | 15.15% | 2.287% | 2.859%
293 | 9_13_21_UF_8 |  | 6/9/2023 7:30:49 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_UF_8.647 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.5306% | 0.823% | 2.835% | 21.79% | 1.374% | 1.79% | 2.896% | 2.772% | 7.391% | 20.81% | 5.045% | 2.989% | 0.5803% | 2.354% | 12.21% | 9.791% | 0.6929% | 1.101% | 2.581% | 2.384% | 4.335% | 1.355% | 1.355% | 7.256% | 32.73% | 26.65% | 10.93% | 4.867% | 105.4% | 1.518% | 1.119% | 3.854% | 1.74% | 1.457% | 0.741% | 0.6348% | 2.173% | 1.886% | 24.5% | 10.8% | 1.233% | 2.162%

## Net Intensities
| Sample Id | R | Acquisition Time | QC Status | Dataset File | Method File | Na KED 23/23 Helium KED_Low (cps) | Mg KED 24/24 Helium KED_Low (cps) | Al 27/27 (cps) | Al KED 27/27 Helium KED_Low (cps) | P Oshift 31/47 Oxygen DRC (cps) | S (34) -1 34/34 Helium KED_Low (cps) | K KED 39/39 Helium KED_Low (cps) | Ca KED 43/43 Helium KED_Low (cps) | V KED 51/51 Helium KED_Low (cps) | Cr (52) 52/52 Helium KED_Low (cps) | Cr (53) NH3 53/53 Ammonia DRC (cps) | Fe (54) NH3 54/54 Ammonia DRC (cps) | Mn Oshift 55/71 Oxygen DRC (cps) | Fe (56) KED 56/56 Helium KED_Low (cps) | Co KED 59/59 Helium KED_Low (cps) | Ni KED 60/60 Helium KED_Low (cps) | Cu (63) KED 63/63 Helium KED_Low (cps) | Cu (65) NH3 65/65 Ammonia DRC (cps) | Zn NH3 66/66 Ammonia DRC (cps) | As 75/75 (cps) | As KED 75/75 Helium KED_Low (cps) | As RH Oshift 75/91 Oxygen DRC (cps) | As Oshift 75/91 Oxygen DRC (cps) | Se (77) 77/77 (cps) | Se (77) Oshift 77/93 Oxygen DRC (cps) | Se (78) Oshift 78/94 Oxygen DRC (cps) | Se (80) Oshift 80/96 Oxygen DRC (cps) | Se (82) 82/82 (cps) | Se (82) Oshift 82/98 Oxygen DRC (cps) | Mo (95) NH3 95/95 Ammonia DRC (cps) | Mo (95) OShift 95/127 Oxygen DRC (cps) | Rh (IS1) 103/103 (IS) (cps) | Rh (IS2) 103/103 (IS) Oxygen DRC (cps) | Rh (IS3) 103/103 (IS) Helium KED_Low (cps) | Rh (IS4) 103/103 (IS) Ammonia DRC (cps) | Rh (IS5) 103/103 (IS) Ammonia DRC (cps) | Rh (IS6) 103/103 (IS) Oxygen DRC (cps) | Rh OShift (IS1) 103/119 (IS) Oxygen DRC (cps) | Cd (111) NH3 111/111 Ammonia DRC (cps) | Cd (114) NH3 114/114 Ammonia DRC (cps) | Pb 208/208 (cps) | Pb NH3 208/208 Ammonia DRC (cps)
1 | Sac Blank |  | 06/08/2023 13:23:24 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\Sac Blank.356 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0 | 0 | 0.028209 | 0.0001822 | 0.01101 | 0.045017 | 0.0010783 | 0 | 9.1609e-05 | 0.0010236 | 0.00014692 | 0.0016664 | 6.2867e-05 | 0.098303 | 0.00013856 | 0.00041707 | 0.0007405 | 0.00039525 | 0.0014077 | -0.016638 | 0 | 0.0071544 | 0.00013838 | 0.0060186 | 5.0373e-06 | 1.509e-05 | 2.5179e-05 | 0.00029192 | 7.545e-06 | 5.6456e-05 | 3.781e-05 | 182769.2655 | 265085.5927 | 14337.1948 | 272750.8925 | 224928.416 | 238625.9044 | 5129.5876 | 7.4074e-06 | 4.1489e-05 | 0.00010212 | 0.00022251
2 | sac blank cal |  | 06/08/2023 13:29:13 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\sac blank cal.357 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9.4097e-05 | 0 | 0.028615 | 0.0041548 | 0.0090662 | 0.045542 | 0.00098978 | 0 | 4.6866e-05 | 0.00066101 | 7.8428e-05 | 0.0018178 | 3.614e-05 | 0.10119 | 4.6866e-05 | 0.00047229 | 0.00028336 | 0.00028793 | 0.0010918 | -0.015095 | 0 | 0.0062419 | 0.00012392 | 0.0053018 | 0 | 5.1699e-06 | 1.5481e-05 | 0.00029512 | 2.5849e-06 | 3.8049e-05 | 2.5819e-05 | 196830.8838 | 258233.7298 | 14098.9543 | 280289.0073 | 226489.4263 | 235016.4348 | 5122.2517 | 2.3929e-06 | 2.1378e-05 | 0.000182 | 0.00027398
3 | CAL BLANK |  | 06/08/2023 13:35:02 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CAL BLANK.358 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4.8356e-05 | 0 | 0.017354 | 4.7386e-05 | 0.0080957 | 0.045479 | 0.0012355 | 4.7231e-05 | 4.7386e-05 | 0.00047503 | 8.9328e-05 | 0.0029639 | 5.2394e-05 | 0.10991 | 0.00023887 | 0.00076629 | 0.00085881 | 0.00026796 | 0.0012977 | -0.015707 | 4.7386e-05 | 0.0053781 | 0.00010745 | 0.0056953 | 1.8338e-05 | 5.189e-06 | 1.0537e-05 | 0.00020821 | 0 | 2.1709e-05 | 4.4331e-05 | 180102.215 | 255065.2033 | 13990.1811 | 276357.489 | 220254.432 | 230485.1249 | 5076.2351 | 4.829e-06 | 2.8946e-05 | 0.000196 | 0.00031493
4 | LOW 0.01 |  | 06/08/2023 13:40:50 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LOW 0.01.359 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.00081924 | 0.0004832 | 0.23202 | 0.0016846 | 0.015805 | 0.00086927 | 0.00092751 | 0.00024121 | 0.00062795 | 0.0013032 | 0.00035141 | 0.0021788 | 0.0023566 | 0.0045865 | 0.0020768 | 0.0006839 | 0.0033747 | 0.0014182 | 0.0042953 | 0.00058488 | 9.7081e-05 | 0.020421 | 0.00040199 | 0.00013406 | -7.8476e-06 | 3.1611e-05 | 5.5174e-05 | 0.00013048 | 1.3092e-05 | 0.0014894 | 0.0015006 | 179124.7892 | 253794.7497 | 13848.7099 | 273488.6462 | 218102.3684 | 224211.1157 | 5017.5478 | 0.00032219 | 0.00083445 | 0.0020678 | 0.0035424
5 | LOW 0.05 |  | 06/08/2023 13:46:39 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LOW 0.05.360 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.0074869 | 0.0021671 | 0.27263 | 0.0031596 | 0.026814 | 0.0028702 | 0.0028565 | 0.0014792 | 0.0042491 | 0.0074039 | 0.0019311 | 0.0025263 | 0.0034571 | 0.011602 | 0.011407 | 0.0037178 | 0.010775 | 0.0046606 | 0.0035461 | 0.0024827 | 0.00059586 | 0.085586 | 0.0016891 | 0.00027647 | 3.2578e-05 | 0.00011248 | 0.00025487 | 0.0003067 | 4.2864e-05 | 0.0032758 | 0.0033439 | 174172.7239 | 248800.1366 | 13522.3981 | 265464.5214 | 219779.0539 | 224388.5039 | 4909.5102 | 0.0017438 | 0.0040819 | 0.0091888 | 0.015696
6 | LOW 0.1 |  | 06/08/2023 13:52:27 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LOW 0.1.361 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.01191 | 0.0057629 | 0.36011 | 0.0033363 | 0.031078 | 0.0025916 | 0.0064659 | 0.0029566 | 0.008006 | 0.012041 | 0.0036888 | 0.0092849 | 0.0076336 | 0.03176 | 0.026673 | 0.0066861 | 0.019441 | 0.0082308 | 0.009071 | 0.002988 | 0.001174 | 0.18132 | 0.0035802 | 0.00058516 | 8.5637e-05 | 0.00024871 | 0.00049354 | 0.00022757 | 0.00010147 | 0.0068791 | 0.0068072 | 176741.3101 | 236645.9464 | 13715.924 | 272424.4973 | 218327.9206 | 221955.0675 | 4684.1037 | 0.0034071 | 0.0077321 | 0.018836 | 0.032396
7 | LOW 0.5 |  | 06/08/2023 13:58:15 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LOW 0.5.362 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.067659 | 0.029335 | 0.54334 | 0.0042787 | 0.037774 | 0.012941 | 0.032215 | 0.013254 | 0.048565 | 0.06634 | 0.018279 | 0.02406 | 0.033858 | 0.11641 | 0.13826 | 0.036867 | 0.09753 | 0.039555 | 0.023211 | 0.017124 | 0.0046034 | 0.86037 | 0.017107 | 0.0021941 | 0.00033729 | 0.0012323 | 0.0025285 | 0.0018241 | 0.00046051 | 0.032769 | 0.032927 | 177339.2711 | 244575.8035 | 13070.6542 | 269232.914 | 215426.8278 | 216603.764 | 4858.8262 | 0.016462 | 0.038877 | 0.090908 | 0.16135
8 | STD 1 |  | 06/08/2023 14:04:03 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 1.363 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.13488 | 0.06095 | 0.82088 | 0.0063906 | 0.03598 | 0.019425 | 0.06226 | 0.022617 | 0.096201 | 0.13056 | 0.036357 | 0.043361 | 0.068312 | 0.20512 | 0.2658 | 0.075699 | 0.18989 | 0.076643 | 0.039931 | 0.037036 | 0.0094293 | 1.7672 | 0.033966 | 0.0038487 | 0.00072061 | 0.0023321 | 0.0047754 | 0.0037315 | 0.00080189 | 0.066797 | 0.066159 | 174759.7721 | 242942.2923 | 13358.9102 | 263003.7451 | 210478.1675 | 215783.7034 | 4674.7656 | 0.032535 | 0.077707 | 0.17807 | 0.32903
9 | STD 5 |  | 06/08/2023 14:09:52 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 5.364 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.65964 | 0.28232 | 2.6337 | 0.020356 | 0.099592 | 0.099651 | 0.30558 | 0.11072 | 0.433 | 0.62285 | 0.18079 | 0.1248 | 0.33281 | 0.56583 | 1.2815 | 0.36201 | 0.86246 | 0.35814 | 0.17535 | 0.19493 | 0.041825 | 8.8731 | 0.16741 | 0.020449 | 0.0036954 | 0.011471 | 0.024221 | 0.019051 | 0.0044141 | 0.32728 | 0.32498 | 157415.1079 | 233196.5948 | 13165.3972 | 250165.4314 | 206727.7047 | 210675.8643 | 4400.0108 | 0.1608 | 0.38178 | 0.93466 | 1.5291
10 | STD 10 |  | 06/08/2023 14:15:40 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 10.365 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.3162 | 0.57525 | 4.6405 | 0.035472 | 0.17184 | 0.19793 | 0.63169 | 0.23574 | 0.9177 | 1.2538 | 0.37046 | 0.24985 | 0.66096 | 1.1209 | 2.6884 | 0.72622 | 1.7161 | 0.71214 | 0.3439 | 0.36865 | 0.092094 | 17.2916 | 0.33206 | 0.037662 | 0.0072456 | 0.021641 | 0.047121 | 0.034112 | 0.0082357 | 0.66859 | 0.6391 | 165063.4829 | 232818.3318 | 12664.9452 | 247069.2922 | 201124.1935 | 205356.1414 | 4470.6996 | 0.32471 | 0.76959 | 1.7108 | 3.0813
11 | STD 25 |  | 06/08/2023 14:21:28 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 25.366 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.3594 | 1.4062 | 10.8929 | 0.08523 | 0.41718 | 0.53608 | 1.6256 | 0.61981 | 2.3818 | 3.2344 | 0.95305 | 0.63196 | 1.734 | 2.8653 | 6.7843 | 1.8337 | 4.3625 | 1.7822 | 0.8723 | 0.96704 | 0.22258 | 46.8386 | 0.8836 | 0.094454 | 0.01928 | 0.059989 | 0.12674 | 0.09081 | 0.022492 | 1.7316 | 1.7049 | 157929.5161 | 215280.0777 | 12048.412 | 230930.2945 | 204800.4543 | 189602.7374 | 4067.9133 | 0.83243 | 1.9485 | 4.1861 | 6.8361
12 | STD 50 |  | 06/08/2023 14:27:16 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 50.367 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 6.5352 | 2.7392 | 23.6723 | 0.16654 | 0.8018 | 1.0376 | 3.2588 | 1.2386 | 4.8476 | 6.5284 | 1.912 | 1.2539 | 3.5237 | 5.6697 | 13.4778 | 3.5977 | 8.4739 | 3.5463 | 1.7294 | 1.9329 | 0.44284 | 94.3201 | 1.808 | 0.18812 | 0.038671 | 0.12143 | 0.25799 | 0.18552 | 0.047149 | 3.5617 | 3.5507 | 153121.6643 | 203870.2062 | 11699.4572 | 222303.2713 | 177857.0724 | 183748.8782 | 3909.202 | 1.6349 | 3.834 | 8.0438 | 14.6703
13 | STD 100 |  | 06/08/2023 14:33:03 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 100.368 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 13.228 | 5.2221 | 47.2306 | 0.32026 | 1.5494 | 2.1044 | 6.6219 | 2.5503 | 9.9808 | 13.4976 | 3.8784 | 2.4749 | 7.0922 | 11.3638 | 27.7884 | 7.1902 | 16.8963 | 6.8676 | 3.3108 | 3.8881 | 0.8782 | 189.129 | 3.5374 | 0.37483 | 0.078845 | 0.24576 | 0.51676 | 0.37466 | 0.093812 | 7.3399 | 7.2194 | 146596.3508 | 191260.51 | 10986.8901 | 210352.8534 | 168709.7639 | 170137.9821 | 3577.7814 | 3.2213 | 7.6089 | 15.5312 | 30.0536
14 | Calibration Curves |  | 06/08/2023 14:33:03
15 | LBLANK |  | 06/08/2023 14:38:52 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.369 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.00023087 | 9.3358e-05 | 0.0079016 | 0.00013751 | 0.0030963 | 0.0032924 | -0.00013214 | -4.7652e-05 | 0.00032537 | 0.00076586 | 8.2807e-05 | -0.00057822 | 0.00013268 | -0.0033228 | 8.8658e-05 | 0.0015875 | 0.00073841 | 0.00023216 | 0.00016852 | -0.0013159 | 8.8714e-05 | 0.017565 | 0.0003492 | 0.00073164 | -1.5696e-05 | 7.7966e-06 | 3.6442e-05 | 0.00017443 | 7.8003e-06 | 0.0025455 | 0.00083706 | 185301.4246 | 255627.5175 | 14510.0361 | 278697.7125 | 203255.3257 | 226092.5268 | 5073.5683 | 4.7709e-05 | 0.00011701 | 8.8957e-06 | 6.5729e-05
16 | LBLANK |  | 06/08/2023 14:57:55 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.370 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.00047376 | 0 | 0.019873 | 7.624e-05 | -0.0015159 | 0.0014791 | -0.00015255 | -4.7652e-05 | 0.00021669 | 0.00093233 | 9.6585e-05 | -0.00055147 | 2.2303e-05 | 0.0014367 | -6.3316e-05 | 0.00027292 | 0.00016309 | 0.00015206 | 0.00042359 | -0.00061613 | 0.0002294 | 0.0021024 | 4.4761e-05 | 0.00043908 | -1.0016e-05 | 1.4138e-05 | 1.1623e-05 | 9.597e-06 | 5.513e-06 | 0.00042787 | 0.0003092 | 172111.9419 | 241266.9573 | 12453.4771 | 265282.5642 | 216199.9653 | 212227.489 | 4894.1716 | 5.2695e-06 | 2.8875e-05 | -7.5959e-05 | -9.9298e-05
17 | ICV PE |  | 06/08/2023 15:03:43 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\ICV PE.371 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.3115 | 0.55713 | 4.8582 | 0.033966 | 0.31718 | 0.19732 | 0.63971 | 0.24386 | 0.91524 | 1.2992 | 0.36318 | 0.25457 | 0.69377 | 1.1733 | 2.686 | 0.74678 | 1.7404 | 0.70606 | 0.33312 | 0.38806 | 0.079592 | 17.8084 | 0.34611 | 0.029906 | 0.0075818 | 0.02357 | 0.049457 | 0.037657 | 0.0090067 | 0.64783 | 0.65043 | 158845.0618 | 223452.7512 | 12724.3322 | 244373.328 | 210139.3361 | 200004.7924 | 4343.9939 | 0.32176 | 0.76622 | 1.7546 | 2.8839
18 | ICV PE |  | 06/08/2023 15:09:59 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\ICV PE.372 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.2972 | 0.57134 | 4.6203 | 0.034928 | 0.3239 | 0.20287 | 0.62869 | 0.24699 | 0.92528 | 1.3086 | 0.37153 | 0.26127 | 0.69805 | 1.2029 | 2.7093 | 0.75133 | 1.7527 | 0.72431 | 0.34154 | 0.37369 | 0.080714 | 17.0058 | 0.34477 | 0.028966 | 0.007422 | 0.022772 | 0.049179 | 0.034658 | 0.0085794 | 0.65198 | 0.65384 | 165620.5593 | 223078.3379 | 12597.5533 | 241561.9807 | 198734.2322 | 198046.47 | 4524.7169 | 0.3279 | 0.78351 | 1.6879 | 3.0269
19 | CCV 1ppb TRACE |  | 06/08/2023 15:15:47 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.373 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.12913 | 0.056279 | 0.78889 | 0.0058794 | 0.021063 | 0.026424 | 0.062208 | 0.023734 | 0.094816 | 0.13069 | 0.037198 | 0.026343 | 0.068328 | 0.12588 | 0.27317 | 0.076853 | 0.18794 | 0.07684 | 0.035027 | 0.03633 | 0.0077371 | 1.7131 | 0.034079 | 0.0043944 | 0.00071749 | 0.0023381 | 0.0050815 | 0.0037685 | 0.00075751 | 0.067889 | 0.067055 | 174225.4615 | 240593.6612 | 13270.1609 | 264432.4115 | 211804.6168 | 208274.8251 | 4784.1345 | 0.032167 | 0.077649 | 0.17389 | 0.31582
20 | CCV 1ppb TRACE |  | 06/08/2023 15:22:03 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.374 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.13611 | 0.060899 | 0.79049 | 0.0056921 | 0.014795 | 0.020774 | 0.062842 | 0.022362 | 0.090474 | 0.12913 | 0.037005 | 0.027934 | 0.069996 | 0.13122 | 0.27405 | 0.077572 | 0.19475 | 0.077509 | 0.03516 | 0.039856 | 0.0088157 | 1.7474 | 0.034735 | 0.0043682 | 0.00073621 | 0.0022315 | 0.0048481 | 0.0035989 | 0.00095755 | 0.067986 | 0.066674 | 169822.8241 | 229717.0459 | 13235.4618 | 253338.6182 | 208333.467 | 202986.7037 | 4569.3978 | 0.033168 | 0.078268 | 0.17322 | 0.30792
21 | LFB |  | 06/08/2023 15:27:51 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.375 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.2694 | 0.5417 | 4.3349 | 0.033097 | 0.17904 | 0.20809 | 0.61003 | 0.22493 | 0.90299 | 1.2454 | 0.37454 | 0.25369 | 0.68036 | 1.1223 | 2.6768 | 0.74097 | 1.7269 | 0.73264 | 0.35711 | 0.36545 | 0.086565 | 17.4571 | 0.34912 | 0.037291 | 0.0074165 | 0.022985 | 0.049194 | 0.03426 | 0.0086517 | 0.68736 | 0.66677 | 160831.1039 | 214239.1821 | 12368.6856 | 229659.7313 | 192086.1967 | 193641.9106 | 4284.6425 | 0.32812 | 0.7835 | 1.6059 | 2.935
22 | LLOQ 0.01 |  | 06/08/2023 15:33:39 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LLOQ 0.01.376 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.0015629 | 0.00070129 | 0.21779 | 0.0019286 | 0.013916 | 0.0059804 | 0.0008238 | 0.00030535 | 0.00095394 | 0.001594 | 0.00038833 | 0.0029363 | 0.0022116 | 0.019026 | 0.0022373 | 0.00084902 | 0.0030147 | 0.0015314 | 0.004372 | 0.0014895 | 5.0877e-05 | 0.016047 | 0.00032466 | 0.00015445 | 1.4398e-06 | 3.429e-05 | 4.8735e-05 | -0.00025099 | 1.6939e-05 | 0.0017313 | 0.0016173 | 178985.5082 | 236009.8031 | 13201.4379 | 259771.9315 | 214332.7099 | 209699.8883 | 4760.1264 | 0.00033916 | 0.00084361 | 0.0020803 | 0.0032654
23 | CCV 10 ppb MAJORS |  | 06/08/2023 15:39:28 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.377 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 8.594e-05 | 8.6264e-05 | -0.0062512 | -2.3859e-06 | -0.0062034 | -0.040683 | -0.00058601 | -4.5203e-06 | -4.5928e-06 | 0.00031451 | 1.3634e-05 | -0.0010771 | -1.4277e-05 | -0.068827 | 0.0004559 | -0.00010721 | -3.013e-05 | 5.771e-06 | -0.00052516 | 0.020835 | -4.7652e-05 | -0.0023435 | -4.8791e-05 | -0.0045647 | -1.3329e-05 | 2.2096e-06 | 9.5129e-06 | -8.608e-05 | 4.9667e-06 | 9.7886e-05 | -1.158e-05 | 192941.5958 | 264644.2699 | 15222.1099 | 311491.2091 | 245395.0214 | 226129.2022 | 5131.5911 | 8.0216e-06 | -3.4029e-06 | -0.00013735 | -0.00023411
24 | CCV 10 ppb MAJORS |  | 06/08/2023 15:45:42 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.378 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.3987 | 0.60501 | 4.184 | 0.035761 | 0.14843 | 0.2208 | 0.70028 | 0.25647 | 0.99622 | 1.3785 | 0.3696 | 0.25037 | 0.68908 | 1.2106 | 2.8598 | 0.78911 | 1.8902 | 0.71389 | 0.33827 | 0.38138 | 0.093148 | 17.3691 | 0.34278 | 0.03683 | 0.0073862 | 0.022766 | 0.047769 | 0.034565 | 0.0085021 | 0.67493 | 0.67529 | 159516.3546 | 213695.1862 | 11627.3988 | 235110.433 | 189586.4368 | 190975.2933 | 4217.956 | 0.32127 | 0.76097 | 1.6399 | 3.0233
25 | LLOQ 0.1 |  | 06/08/2023 15:51:31 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LLOQ 0.1.379 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.014977 | 0.0065856 | 0.35561 | 0.0025262 | 0.021048 | 0.0043363 | 0.0063103 | 0.0031748 | 0.0086259 | 0.014592 | 0.0038359 | 0.0095133 | 0.0076164 | 0.043913 | 0.025241 | 0.0086677 | 0.021932 | 0.0087806 | 0.0091552 | 0.0060431 | 0.00054814 | 0.16498 | 0.0033269 | 0.00066259 | 2.5177e-05 | 0.0002583 | 0.0005251 | 0.00044255 | 0.00011593 | 0.0070739 | 0.0069849 | 170695.8654 | 230021.1196 | 13451.6644 | 256347.5928 | 207106.0947 | 204106.3219 | 4639.4205 | 0.0031213 | 0.0076823 | 0.016961 | 0.030968
26 | STD 0.5 |  | 06/08/2023 15:57:19 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 0.5.380 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.068043 | 0.029526 | 0.53919 | 0.0052403 | 0.034325 | 0.01357 | 0.029866 | 0.012237 | 0.047283 | 0.071178 | 0.020883 | 0.028619 | 0.035834 | 0.11941 | 0.13614 | 0.038326 | 0.10005 | 0.045738 | 0.026666 | 0.02066 | 0.0041413 | 0.85993 | 0.017508 | 0.0024818 | 0.0003319 | 0.0012575 | 0.0025941 | 0.001642 | 0.00035936 | 0.0379 | 0.033634 | 165637.4035 | 224791.7929 | 12877.8031 | 227709.2961 | 204238.282 | 201097.0991 | 4578.7339 | 0.018429 | 0.043294 | 0.086966 | 0.15817
27 | 8_2_21_UF_1 |  | 06/08/2023 16:03:07 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_UF_1.381 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.72901 | 0.30876 | 1.01 | 0.0078113 | 0.11183 | 0.25412 | 0.077739 | 0.6811 | 0.036339 | 0.0081064 | 0.0021742 | 0.019691 | 1.0063 | 0.12989 | 0.0048641 | 0.017438 | 0.16819 | 0.067988 | 0.055276 | 0.16581 | 0.035746 | 7.9259 | 0.15025 | 0.0013853 | 2.4292e-05 | 0.00011581 | 0.00024459 | 0.0011295 | 3.2716e-05 | 0.33614 | 0.3387 | 150700.1075 | 203902.2025 | 11794.2009 | 224297.3137 | 184519.3858 | 184880.9711 | 3865.1895 | 0.00022431 | 0.00053877 | 0.0027333 | 0.0053388
28 | 8_2_21_UF_6 |  | 06/08/2023 16:08:55 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_UF_6.382 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.4124 | 0.39085 | 0.41595 | 0.0030935 | 0.077264 | 0.25071 | 0.13393 | 0.83846 | 0.093054 | 0.01302 | 0.0033198 | 0.023847 | 0.91275 | 0.13849 | 0.0052821 | 0.015221 | 0.38622 | 0.16242 | 0.083847 | 0.22218 | 0.046275 | 10.714 | 0.21047 | 0.0015158 | 6.561e-05 | 0.00025992 | 0.00059031 | 0.0014609 | 0.00010404 | 0.22353 | 0.22269 | 147345.3276 | 198637.4546 | 11471.2707 | 213822.6619 | 180092.8703 | 183839.3902 | 3903.867 | 0.00059966 | 0.0012707 | 0.006009 | 0.010946
29 | 8_2_21_UF_6_SP |  | 06/08/2023 16:14:43 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_UF_6_SP.383 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.1516 | 0.71624 | 2.8378 | 0.021432 | 0.19201 | 0.35472 | 0.49696 | 0.97776 | 0.59753 | 0.7018 | 0.20873 | 0.16912 | 1.3217 | 0.8001 | 1.4725 | 0.41434 | 1.3103 | 0.54328 | 0.2604 | 0.41998 | 0.098644 | 19.7203 | 0.39502 | 0.020964 | 0.0040927 | 0.012645 | 0.026734 | 0.020687 | 0.0048562 | 0.58444 | 0.59444 | 146550.5408 | 195912.2096 | 11266.4411 | 213861.1387 | 175224.3329 | 159294.8273 | 3925.8728 | 0.17511 | 0.40929 | 0.91113 | 1.7036
30 | STD 1 |  | 06/08/2023 16:20:31 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 1.384 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.14212 | 0.05917 | 0.85998 | 0.0063676 | 0.028715 | 0.029294 | 0.070068 | 0.026274 | 0.10133 | 0.13727 | 0.037221 | 0.049412 | 0.072421 | 0.20231 | 0.27867 | 0.074976 | 0.18907 | 0.077609 | 0.039734 | 0.038758 | 0.0090272 | 1.9032 | 0.037429 | 0.0046071 | 0.00076265 | 0.0023056 | 0.0049975 | 0.0037577 | 0.00081863 | 0.068162 | 0.070097 | 159581.9135 | 211702.4534 | 12275.2732 | 235254.7311 | 196503.6768 | 188082.9369 | 4170.6101 | 0.033582 | 0.078987 | 0.17875 | 0.32066
31 | 8_2_21_UF_6_SP_LD |  | 06/08/2023 16:26:19 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_UF_6_SP_LD.385 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.1589 | 0.72615 | 3.2567 | 0.022976 | 0.19234 | 0.36433 | 0.48585 | 1.0015 | 0.59564 | 0.71288 | 0.20956 | 0.15563 | 1.2861 | 0.7334 | 1.4308 | 0.413 | 1.3274 | 0.54473 | 0.2643 | 0.42505 | 0.096424 | 20.0305 | 0.3951 | 0.020805 | 0.0038452 | 0.012445 | 0.027197 | 0.021865 | 0.0048603 | 0.5919 | 0.58486 | 146648.2037 | 193026.7149 | 11320.4842 | 211799.357 | 175886.384 | 174758.1496 | 3808.5077 | 0.17677 | 0.41371 | 0.90549 | 1.682
32 | 8_2_21_UF_6_SP_LSP |  | 06/08/2023 16:32:06 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_UF_6_SP_LSP.386 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9.7743 | 3.7096 | 9.4397 | 0.073291 | 0.43202 | 1.506 | 4.1212 | 2.4649 | 1.9897 | 2.579 | 0.71152 | 0.51206 | 2.2388 | 2.4734 | 5.229 | 1.4144 | 3.573 | 1.426 | 0.69018 | 0.91543 | 0.21563 | 42.9361 | 0.84647 | 0.069974 | 0.013573 | 0.042973 | 0.092504 | 0.06938 | 0.016443 | 1.5142 | 1.478 | 136689.0046 | 178097.9886 | 9751.3272 | 191979.4359 | 161286.7315 | 158762.707 | 3511.0981 | 0.57389 | 1.3599 | 2.8254 | 5.289
33 | 8_2_21_UF_8 |  | 06/08/2023 16:37:54 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_UF_8.387 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.5234 | 0.3944 | 0.38797 | 0.002567 | 0.16775 | 0.21891 | 0.15121 | 0.78051 | 0.12601 | 0.022124 | 0.0057044 | 0.039138 | 0.91286 | 0.21165 | 0.0051023 | 0.013375 | 0.38137 | 0.15806 | 0.037674 | 0.2488 | 0.055812 | 11.6959 | 0.23459 | 0.0017569 | 4.4555e-05 | 0.00027592 | 0.00062796 | 0.001924 | 9.263e-05 | 0.18475 | 0.18597 | 149884.2088 | 201549.8253 | 11741.4903 | 219251.2124 | 180413.526 | 182606.8062 | 4045.2398 | 0.00043001 | 0.00093752 | 0.007277 | 0.013412
34 | 8_2_21_UF_9 |  | 06/08/2023 16:43:42 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_UF_9.388 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.1183 | 0.35657 | 0.3868 | 0.0033019 | 0.14004 | 0.1766 | 0.11739 | 0.67217 | 0.08861 | 0.011512 | 0.0027235 | 0.01522 | 0.44507 | 0.10263 | 0.0048683 | 0.012275 | 0.2638 | 0.11044 | 0.031959 | 0.17719 | 0.037322 | 8.34 | 0.16506 | 0.0014674 | 4.8881e-05 | 0.00020998 | 0.00042331 | 0.001203 | 9.4115e-05 | 0.128 | 0.12814 | 148400.844 | 198288.0317 | 11520.6455 | 217387.5642 | 181235.0806 | 182653.6809 | 3924.5391 | 0.00033253 | 0.0007594 | 0.0082086 | 0.014961
35 | LBLANK |  | 06/08/2023 16:49:31 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.389 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.00044952 | 0 | 0.015803 | 0.00011767 | -0.0033319 | 0.005221 | -0.00017392 | 6.4016e-06 | 0.00011257 | 0.00019168 | 9.3762e-05 | -0.00038431 | 3.3981e-05 | -0.005332 | -7.61e-05 | 0.00091857 | 0.00015057 | 0.00014625 | 1.8586e-05 | 0.00080312 | 0.00021242 | 0.057944 | 0.0011582 | 0.00049293 | -2.3714e-06 | 3.3113e-05 | 4.3743e-05 | 0.00011985 | 1.278e-05 | 0.00017889 | 9.6299e-05 | 155367.6581 | 209031.676 | 12534.8534 | 229849.8427 | 188542.031 | 184944.8598 | 4180.6116 | 3.9456e-06 | 3.4895e-05 | -9.3155e-05 | -4.6012e-05
36 | CCV 1ppb TRACE |  | 06/08/2023 16:55:19 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.390 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.12842 | 0.055379 | 0.82236 | 0.0061492 | 0.012217 | 0.023284 | 0.0612 | 0.025343 | 0.091544 | 0.12458 | 0.038566 | 0.028709 | 0.07419 | 0.11108 | 0.26391 | 0.072824 | 0.18068 | 0.079077 | 0.037243 | 0.037848 | 0.0083283 | 1.837 | 0.036303 | 0.0043178 | 0.00073144 | 0.0022288 | 0.0050452 | 0.0034382 | 0.00095423 | 0.070016 | 0.071898 | 154180.2794 | 204695.4891 | 12720.9987 | 226998.6517 | 186328.8333 | 180473.103 | 4047.2401 | 0.034861 | 0.08274 | 0.17752 | 0.33069
37 | CCV 1ppb TRACE |  | 06/08/2023 17:01:37 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.391 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.13554 | 0.062087 | 0.81971 | 0.0056907 | 0.012072 | 0.031159 | 0.066015 | 0.02695 | 0.10085 | 0.13005 | 0.038997 | 0.028453 | 0.072722 | 0.12694 | 0.27977 | 0.084782 | 0.19188 | 0.079731 | 0.036956 | 0.037241 | 0.0092826 | 1.8628 | 0.036212 | 0.0040969 | 0.00075671 | 0.0024768 | 0.0054414 | 0.0034955 | 0.0010154 | 0.070039 | 0.069944 | 154669.5814 | 208221.206 | 12073.7672 | 228916.9389 | 186871.6718 | 182822.6778 | 4047.907 | 0.034218 | 0.080815 | 0.17609 | 0.32943
38 | CCV 10 ppb MAJORS |  | 06/08/2023 17:07:25 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.392 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.3828 | 0.60385 | 4.3603 | 0.034088 | 0.15386 | 0.22187 | 0.6676 | 0.25675 | 0.98512 | 1.338 | 0.38864 | 0.2574 | 0.71586 | 1.2182 | 2.8518 | 0.78347 | 1.82 | 0.75072 | 0.35518 | 0.38332 | 0.093265 | 18.5441 | 0.36401 | 0.038807 | 0.0075415 | 0.023487 | 0.051089 | 0.037031 | 0.0090155 | 0.70252 | 0.69598 | 147179.8263 | 193137.6825 | 11233.7491 | 213718.8203 | 177136.8634 | 174505.559 | 3791.1697 | 0.33397 | 0.79311 | 1.6882 | 3.1618
39 | CCV 10 ppb MAJORS |  | 06/08/2023 17:13:42 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.393 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.3689 | 0.58554 | 4.3414 | 0.035424 | 0.15364 | 0.2241 | 0.6538 | 0.25559 | 0.96867 | 1.3482 | 0.38137 | 0.25647 | 0.71533 | 1.1982 | 2.8635 | 0.7861 | 1.8187 | 0.73914 | 0.34995 | 0.38206 | 0.090478 | 17.6572 | 0.35675 | 0.039469 | 0.0074164 | 0.023967 | 0.050433 | 0.037037 | 0.0090373 | 0.69471 | 0.69364 | 150429.3186 | 197867.6336 | 11463.9321 | 218072.5314 | 180179.1539 | 175562.3409 | 3999.2267 | 0.33018 | 0.78143 | 1.6859 | 3.0892
40 | 8_2_21_UF_10 |  | 06/08/2023 17:19:31 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_UF_10.394 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.2703 | 0.55156 | 0.41396 | 0.0033628 | 0.18647 | 0.34297 | 0.14987 | 0.9586 | 0.096358 | 0.0092151 | 0.0020396 | 0.016875 | 0.39097 | 0.12914 | 0.0061875 | 0.015453 | 0.2602 | 0.098857 | 0.038469 | 0.2113 | 0.050432 | 9.592 | 0.19296 | 0.0018765 | 4.7912e-05 | 0.00028417 | 0.00055787 | 0.001615 | 0.00010801 | 0.13214 | 0.12959 | 141714.8952 | 191441.6223 | 10568.5826 | 206090.6642 | 172985.7128 | 166552.644 | 3851.8527 | 0.0002604 | 0.0006859 | 0.011252 | 0.020342
41 | 8_2_21_UF_11 |  | 06/08/2023 17:25:19 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_UF_11.395 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.198 | 0.56625 | 0.37537 | 0.0030009 | 0.16581 | 0.35033 | 0.15258 | 1.0002 | 0.10717 | 0.11735 | 0.03185 | 0.18537 | 0.4343 | 0.96359 | 0.0081331 | 0.015471 | 0.22756 | 0.091446 | 0.032184 | 0.20883 | 0.053661 | 9.811 | 0.19646 | 0.0018518 | 4.1582e-05 | 0.00022448 | 0.00047012 | 0.0014794 | 7.0859e-05 | 0.12578 | 0.12614 | 138918.181 | 188629.9576 | 10415.2098 | 204933.4859 | 169878.2416 | 167941.6015 | 3778.4999 | 0.00020343 | 0.00046883 | 0.018798 | 0.034334
42 | 8_2_21_UF_11_DUP |  | 06/08/2023 17:31:09 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_UF_11_DUP.396 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.0909 | 0.52104 | 0.36162 | 0.0024789 | 0.16275 | 0.29912 | 0.13551 | 0.87058 | 0.090305 | 0.0075154 | 0.0019814 | 0.010803 | 0.43069 | 0.082277 | 0.0050469 | 0.010944 | 0.20754 | 0.091897 | 0.033365 | 0.20307 | 0.043108 | 9.7198 | 0.19052 | 0.0016498 | 6.3297e-05 | 0.00023574 | 0.00049366 | 0.0015187 | 0.00012766 | 0.12335 | 0.12204 | 141184.2324 | 187988.3203 | 11100.3119 | 204619.8945 | 172105.852 | 168189.8171 | 3686.4759 | 0.00016775 | 0.00044969 | 0.018311 | 0.033315
43 | 8_2_21_F_1 |  | 06/08/2023 17:36:57 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_F_1.397 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.63518 | 0.34094 | 1.5893 | 0.012159 | 0.086338 | 0.29512 | 0.078893 | 0.73811 | 0.035063 | 0.0067435 | 0.0017626 | 0.16988 | 1.1201 | 0.81808 | 0.0048548 | 0.02147 | 0.22484 | 0.093656 | 0.071097 | 0.17529 | 0.037926 | 8.782 | 0.16943 | 0.0012323 | 2.3532e-05 | 0.00015138 | 0.00028573 | 0.0012049 | 6.6066e-05 | 0.38631 | 0.38776 | 143002.9033 | 191407.137 | 11253.7654 | 209945.4742 | 173891.2714 | 171254.2122 | 3697.1454 | 0.00024596 | 0.00066279 | 0.0079964 | 0.014542
44 | 8_2_21_F_6 |  | 06/08/2023 17:44:00 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_F_6.398 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.3062 | 0.41787 | 1.2675 | 0.0090432 | 0.059592 | 0.28211 | 0.14253 | 0.90071 | 0.091122 | 0.053993 | 0.016192 | 0.26034 | 0.90958 | 1.2029 | 0.0073479 | 0.019623 | 0.721 | 0.30309 | 0.12515 | 0.23795 | 0.053044 | 10.2605 | 0.2043 | 0.0014167 | 8.3042e-05 | 0.00027733 | 0.00049507 | 0.0018691 | 0.00012723 | 0.2381 | 0.21821 | 138960.7794 | 206382.0338 | 11087.6348 | 204228.8262 | 170392.8388 | 165946.9844 | 4103.9333 | 0.00064813 | 0.0014365 | 0.017053 | 0.029751
45 | 8_2_21_F_8 |  | 06/08/2023 17:49:48 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_F_8.399 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.5215 | 0.41874 | 1.3596 | 0.01223 | 0.16028 | 0.24264 | 0.16878 | 0.86302 | 0.13047 | 0.016558 | 0.0048512 | 0.16029 | 0.98212 | 0.79552 | 0.0062437 | 0.015468 | 0.77598 | 0.31356 | 0.08134 | 0.26839 | 0.061104 | 12.1633 | 0.23942 | 0.0018158 | 4.6789e-05 | 0.0003016 | 0.0004816 | 0.0021573 | 8.738e-05 | 0.20776 | 0.1988 | 141115.7477 | 196680.3371 | 10802.0825 | 204996.4034 | 173271.152 | 167617.0527 | 3868.524 | 0.00038565 | 0.0010871 | 0.017609 | 0.031805
46 | 8_2_21_F_8_DUP |  | 06/08/2023 17:55:36 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_F_8_DUP.400 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.7963 | 0.4916 | 1.3549 | 0.011692 | 0.15523 | 0.29349 | 0.18916 | 1.0433 | 0.14959 | 0.0082035 | 0.0020033 | 0.15155 | 1.0042 | 0.88575 | 0.008613 | 0.017749 | 0.91957 | 0.32086 | 0.087959 | 0.26508 | 0.066688 | 12.6202 | 0.24795 | 0.0018937 | 9.0068e-05 | 0.00024664 | 0.00068656 | 0.002142 | 0.0001263 | 0.20854 | 0.20028 | 143629.8085 | 190364.7214 | 9710.0266 | 202265.8802 | 173683.4356 | 172098.6823 | 3740.4897 | 0.00056885 | 0.0012438 | 0.018052 | 0.033224
47 | 8_2_21_F_8_DUP_LD |  | 06/08/2023 18:01:24 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_F_8_DUP_LD.401 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.4568 | 0.42236 | 2.1719 | 0.017514 | 0.22421 | 0.23241 | 0.1584 | 0.85123 | 0.12946 | 0.010834 | 0.0028503 | 0.32858 | 1.0207 | 1.5631 | 0.0068494 | 0.016982 | 0.77132 | 0.30479 | 0.093369 | 0.2684 | 0.055031 | 12.3388 | 0.24361 | 0.0018641 | 7.9353e-05 | 0.00026633 | 0.00060357 | 0.0021685 | 0.00011193 | 0.19746 | 0.20555 | 139953.0035 | 191332.3436 | 11197.7203 | 214642.7606 | 173848.8998 | 147442.4127 | 3777.1661 | 0.00045347 | 0.0010656 | 0.01954 | 0.034607
48 | 8_2_21_F_8_DUP_LD |  | 06/08/2023 18:07:40 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_F_8_DUP_LD.402 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.5159 | 0.41645 | 2.2183 | 0.018228 | 0.18905 | 0.2414 | 0.16258 | 0.85441 | 0.13016 | 0.011683 | 0.002869 | 0.33751 | 1.0335 | 1.6443 | 0.0063491 | 0.016513 | 0.78842 | 0.31579 | 0.09649 | 0.26612 | 0.053553 | 12.9777 | 0.24914 | 0.0016129 | 9.0727e-05 | 0.0002589 | 0.00059132 | 0.0018334 | 7.038e-05 | 0.19983 | 0.20485 | 142691.7677 | 189445.2161 | 11046.2701 | 210031.9512 | 173362.9562 | 167332.3386 | 3637.1296 | 0.00046167 | 0.001174 | 0.019076 | 0.035546
49 | 8_2_21_F_8_DUP_LSP |  | 06/08/2023 18:13:28 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_F_8_DUP_LSP.403 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9.2954 | 3.4547 | 8.5401 | 0.065913 | 0.4406 | 1.4311 | 3.8732 | 2.2525 | 1.4707 | 1.8229 | 0.53312 | 0.53302 | 2.0309 | 2.4527 | 3.7151 | 1.0105 | 3.0536 | 1.2519 | 0.53579 | 0.78333 | 0.17978 | 36.2795 | 0.72636 | 0.051923 | 0.010425 | 0.032092 | 0.067554 | 0.051309 | 0.012024 | 1.153 | 1.1367 | 130290.8385 | 167895.5146 | 9833.3833 | 182819.0457 | 152266.5472 | 149065.8359 | 3362.3958 | 0.42463 | 0.99418 | 2.0665 | 3.8463
50 | LBLANK |  | 06/08/2023 18:19:16 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.404 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.00087572 | 0.00021706 | 0.0087953 | 0.000115 | 0.0013182 | 0.0099769 | 6.4702e-05 | -4.7652e-05 | 6.5159e-06 | 0.00044667 | 0.00012634 | -0.0005345 | 7.2162e-05 | -0.00020763 | -7.5457e-05 | 0.001246 | 0.00039068 | 0.00024205 | 1.3525e-05 | -0.00064382 | 6.0835e-05 | 0.01114 | 0.00023391 | 0.00080625 | -1.8296e-05 | 1.1517e-06 | 5.4903e-06 | 0.00012297 | 9.5678e-06 | 0.00040194 | 0.00019792 | 157424.6455 | 209045.8127 | 12277.9406 | 237920.4186 | 191342.6014 | 182190.4784 | 4319.3198 | 3.5248e-06 | 5.2051e-05 | -1.3394e-06 | -7.7222e-06
51 | CCV 1ppb TRACE |  | 06/08/2023 18:25:05 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.405 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.14994 | 0.063589 | 0.84606 | 0.0058492 | 0.01361 | 0.036915 | 0.074035 | 0.028611 | 0.10341 | 0.14433 | 0.0397 | 0.029466 | 0.077554 | 0.152 | 0.29771 | 0.087004 | 0.21279 | 0.080383 | 0.037087 | 0.041173 | 0.0097627 | 1.8356 | 0.037139 | 0.0044256 | 0.00084692 | 0.0024854 | 0.0055995 | 0.004112 | 0.0008946 | 0.072264 | 0.072444 | 155603.599 | 205001.1585 | 11812.885 | 230954.4828 | 189019.91 | 177227.3535 | 4156.6055 | 0.034001 | 0.080893 | 0.17554 | 0.32902
52 | CCV 1ppb TRACE |  | 06/08/2023 18:31:23 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.406 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.14372 | 0.059913 | 0.88281 | 0.0062352 | 0.012559 | 0.032279 | 0.067143 | 0.026547 | 0.1064 | 0.13873 | 0.039662 | 0.028864 | 0.075682 | 0.14667 | 0.29117 | 0.080906 | 0.20249 | 0.08157 | 0.036752 | 0.041055 | 0.0080465 | 1.8405 | 0.037576 | 0.0042891 | 0.00077929 | 0.0024269 | 0.0051872 | 0.0037446 | 0.00086137 | 0.069732 | 0.071146 | 155818.6164 | 205852.185 | 12191.2006 | 229669.5428 | 189213.0606 | 183852.91 | 4202.6181 | 0.033961 | 0.081147 | 0.17729 | 0.32433
53 | LFB |  | 06/08/2023 18:37:11 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.407 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.3447 | 0.58241 | 4.5616 | 0.033745 | 0.19581 | 0.23011 | 0.64284 | 0.25729 | 0.95747 | 1.3358 | 0.38646 | 0.26416 | 0.73158 | 1.1884 | 2.8475 | 0.7784 | 1.8235 | 0.7612 | 0.36718 | 0.38954 | 0.090149 | 18.4781 | 0.36874 | 0.040173 | 0.0077377 | 0.02388 | 0.052428 | 0.038543 | 0.0095424 | 0.69251 | 0.70693 | 148419.0478 | 192719.3654 | 11637.405 | 213285.1017 | 177100.4344 | 171743.0406 | 3847.8517 | 0.33206 | 0.78894 | 1.639 | 3.0505
54 | CCV 10 ppb MAJORS |  | 06/08/2023 18:42:59 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.408 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.384 | 0.58333 | 4.4119 | 0.032985 | 0.16207 | 0.22634 | 0.65331 | 0.25644 | 0.96419 | 1.3437 | 0.40159 | 0.27174 | 0.74539 | 1.2001 | 2.8298 | 0.79127 | 1.8595 | 0.78221 | 0.36772 | 0.38921 | 0.088705 | 18.2139 | 0.36582 | 0.039628 | 0.0079992 | 0.024391 | 0.051996 | 0.036512 | 0.0095953 | 0.72121 | 0.7121 | 150614.591 | 194488.9655 | 11758.1732 | 214766.9989 | 178196.5142 | 168840.6463 | 3908.5348 | 0.34062 | 0.80296 | 1.6554 | 3.1051
55 | CCV 10 ppb MAJORS |  | 06/08/2023 18:49:21 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.409 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.4854 | 0.62296 | 3.9874 | 0.036755 | 0.15954 | 0.24503 | 0.71474 | 0.27646 | 1.0506 | 1.4817 | 0.40047 | 0.26768 | 0.74817 | 1.3019 | 3.0589 | 0.86818 | 1.9848 | 0.77707 | 0.36343 | 0.35941 | 0.099048 | 18.4949 | 0.36697 | 0.035013 | 0.0079599 | 0.024066 | 0.051834 | 0.034064 | 0.0094636 | 0.71534 | 0.71033 | 164932.4617 | 194418.8386 | 11044.2771 | 216782.4593 | 176700.1186 | 170397.5888 | 3857.8543 | 0.33781 | 0.79649 | 1.5055 | 3.1232
56 | LBLANK |  | 06/08/2023 19:26:31 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.410 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.00012631 | 0 | 0.010621 | 0.0001255 | -0.0036093 | 0.00038491 | 3.2311e-05 | -4.7652e-05 | 9.7361e-06 | 0.00067436 | 0.00012856 | -0.00029338 | 6.2243e-05 | 0.0039468 | -0.00012188 | 0.00033426 | 0.00046894 | 0.00017076 | 0.00065033 | 0.0033139 | 9.7361e-06 | 0.0034871 | 8.5056e-05 | 0.00083691 | -1.4983e-05 | 1.1506e-05 | 1.9563e-05 | 0.00030638 | 0 | 0.00021616 | 0.00018381 | 153146.8073 | 198249.6115 | 11550.6679 | 226570.886 | 187776.0628 | 177993.3655 | 4279.3082 | 3.9174e-06 | 3.9034e-05 | -6.5739e-05 | -7.3081e-05
57 | PE ICV |  | 06/08/2023 19:32:19 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\PE ICV.411 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.4113 | 0.60419 | 4.5563 | 0.037607 | 0.33446 | 0.23414 | 0.67428 | 0.26208 | 0.99506 | 1.3844 | 0.3966 | 0.2744 | 0.7588 | 1.2678 | 2.9014 | 0.80362 | 1.8981 | 0.77446 | 0.3611 | 0.38076 | 0.084509 | 18.3036 | 0.36736 | 0.028006 | 0.0076169 | 0.024942 | 0.05275 | 0.036991 | 0.0094497 | 0.69202 | 0.68843 | 150832.7773 | 189196.6808 | 11243.7564 | 210836.4536 | 178062.3108 | 166278.8323 | 3805.1743 | 0.34047 | 0.80551 | 1.6453 | 3.0913
58 | 8_2_21_F_9 |  | 06/08/2023 19:38:07 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_F_9.412 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.0241 | 0.37722 | 1.1937 | 0.010472 | 0.13397 | 0.20131 | 0.12943 | 0.74734 | 0.095796 | 0.017771 | 0.0044023 | 0.14071 | 0.52047 | 0.6987 | 0.005224 | 0.01421 | 0.60364 | 0.23664 | 0.056938 | 0.19279 | 0.043224 | 8.8968 | 0.18133 | 0.0017636 | 2.7337e-05 | 0.00018474 | 0.00042986 | 0.0012889 | 7.7473e-05 | 0.14543 | 0.14571 | 146125.5808 | 189581.738 | 11090.9724 | 207808.5665 | 173119.4761 | 167317.4741 | 3868.5244 | 0.0003997 | 0.0010395 | 0.037569 | 0.068934
59 | 8_2_21_F_10 |  | 06/08/2023 19:43:55 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_F_10.413 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.0755 | 0.54894 | 1.935 | 0.015492 | 0.17498 | 0.35601 | 0.15266 | 0.99711 | 0.099129 | 0.0084994 | 0.0020483 | 0.11719 | 0.47814 | 0.63514 | 0.0085926 | 0.016743 | 0.47738 | 0.18508 | 0.088349 | 0.21636 | 0.053091 | 10.1603 | 0.20501 | 0.0022577 | 3.5396e-05 | 0.00019848 | 0.00055436 | 0.0017755 | 0.00011087 | 0.14168 | 0.14126 | 143542.9568 | 186466.3898 | 10458.5104 | 203880.5971 | 174174.7306 | 165691.4187 | 3763.1623 | 0.00031846 | 0.00089694 | 0.024223 | 0.044836
60 | 8_2_21_F_10_DUP |  | 06/08/2023 19:49:43 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_F_10_DUP.414 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.0585 | 0.52038 | 1.9345 | 0.014142 | 0.16942 | 0.34056 | 0.14457 | 0.9374 | 0.091029 | 0.034797 | 0.0094882 | 0.14586 | 0.4852 | 0.70883 | 0.0066232 | 0.016243 | 0.45176 | 0.18807 | 0.084896 | 0.22016 | 0.049538 | 9.8692 | 0.20381 | 0.0019005 | 5.3188e-05 | 0.00023755 | 0.00053795 | 0.0017878 | 0.0001224 | 0.14277 | 0.14285 | 141721.5329 | 186065.4945 | 10902.1588 | 204107.7743 | 173744.2505 | 167025.4711 | 3844.5197 | 0.00038028 | 0.00082696 | 0.024351 | 0.043557
61 | 8_2_21_F_11 |  | 06/08/2023 19:55:32 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_F_11.415 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.99858 | 0.53749 | 2.0871 | 0.016005 | 0.15598 | 0.33942 | 0.14365 | 0.94876 | 0.098747 | 0.010445 | 0.0023512 | 0.12604 | 0.53332 | 0.60928 | 0.0067158 | 0.016682 | 0.49145 | 0.20763 | 0.057174 | 0.21698 | 0.045301 | 10.0402 | 0.20243 | 0.0019237 | 9.3251e-05 | 0.00027518 | 0.00052505 | 0.0016573 | 8.2736e-05 | 0.13256 | 0.13193 | 142883.0785 | 185509.3924 | 10924.8426 | 202828.5259 | 157711.7236 | 164075.1074 | 3740.4897 | 0.00022548 | 0.00054989 | 0.069498 | 0.13832
62 | 8_2_21_F_11_SP |  | 06/08/2023 20:01:20 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_F_11_SP.416 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.8081 | 0.86563 | 4.6354 | 0.036861 | 0.24855 | 0.46066 | 0.52477 | 1.1013 | 0.64487 | 0.77535 | 0.22865 | 0.30715 | 0.94288 | 1.376 | 1.5248 | 0.42859 | 1.4305 | 0.5892 | 0.24349 | 0.43569 | 0.095379 | 19.8943 | 0.39646 | 0.022121 | 0.0042459 | 0.013221 | 0.027413 | 0.021958 | 0.004844 | 0.52064 | 0.51845 | 136851.2987 | 182981.896 | 10730.0293 | 199557.4819 | 167023.4015 | 164584.0651 | 3646.4654 | 0.18275 | 0.42794 | 1.0189 | 1.8375
63 | 8_2_21_W_1 |  | 06/08/2023 20:07:08 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_W_1.417 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.1289 | 0.068971 | 3.4086 | 0.030281 | 0.045848 | 0.068243 | 0.017526 | 0.14661 | 0.0076209 | 0.0047528 | 0.0011806 | 0.20452 | 0.34409 | 0.95639 | 0.0026665 | 0.0071395 | 0.081662 | 0.035428 | 0.010417 | 0.04006 | 0.0073883 | 1.8205 | 0.036182 | 0.00090155 | 6.1511e-06 | 3.6168e-05 | 8.9386e-05 | 0.00021188 | 1.3738e-05 | 0.076098 | 0.077444 | 150415.0533 | 193127.6652 | 11473.2722 | 218350.7842 | 184506.5458 | 173453.4059 | 3837.8488 | 8.0674e-05 | 0.00025825 | 0.0065989 | 0.011244
64 | 8_2_21_W_6 |  | 06/08/2023 20:12:56 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_W_6.418 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.26037 | 0.082462 | 6.06 | 0.050858 | 0.051378 | 0.063794 | 0.031684 | 0.18191 | 0.021542 | 0.0071526 | 0.0018573 | 0.38905 | 0.41537 | 1.8409 | 0.0035225 | 0.0034963 | 0.2855 | 0.11254 | 0.040139 | 0.051795 | 0.011993 | 2.3996 | 0.047367 | 0.0012004 | 8.9346e-06 | 5.542e-05 | 0.00012835 | 0.00044403 | 3.3665e-05 | 0.048003 | 0.047003 | 149818.8704 | 196984.4239 | 11312.4788 | 217093.8882 | 180223.854 | 172807.0388 | 3889.8629 | 0.00024409 | 0.0005551 | 0.022956 | 0.040699
65 | 8_2_21_W_6_DUP |  | 06/08/2023 20:18:44 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_W_6_DUP.419 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.25897 | 0.084227 | 5.3807 | 0.043643 | 0.050302 | 0.062562 | 0.023521 | 0.1753 | 0.019482 | 0.0057551 | 0.0014979 | 0.33795 | 0.39955 | 1.634 | 0.002318 | 0.0058228 | 0.26197 | 0.1062 | 0.034748 | 0.050225 | 0.009966 | 2.314 | 0.047003 | 0.00088487 | -1.1385e-05 | 8.46e-05 | 0.0001208 | 0.00035888 | 6.8992e-06 | 0.047196 | 0.04733 | 149390.6352 | 192939.5012 | 11236.4177 | 216034.316 | 179134.816 | 171321.8523 | 3919.8712 | 0.00024198 | 0.00050279 | 0.021135 | 0.036747
66 | LBLANK |  | 06/08/2023 20:24:33 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.420 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.0701e-05 | 0 | 0.013164 | 0.00018072 | -0.00090585 | 0.0074533 | 0.00024966 | -4.7652e-05 | 6.7941e-05 | 0.00061568 | 8.8829e-05 | -0.00046048 | 2.4476e-05 | 0.0028995 | -0.00012155 | -0.00019571 | -7.9354e-06 | 0.00019974 | 9.4164e-05 | 0.001887 | 1.0701e-05 | 0.001323 | 2.6376e-05 | 0.00069707 | -1.5001e-05 | 1.4587e-06 | 1.6183e-05 | 0.00014222 | 0 | 8.3472e-05 | -1.1892e-06 | 149603.9947 | 199196.0122 | 11633.4032 | 220990.5323 | 183355.1755 | 176524.0803 | 3990.5574 | 4.1228e-06 | 3.7165e-05 | -7.1363e-05 | -0.00010024
67 | CCV 1ppb TRACE |  | 06/08/2023 20:30:21 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.421 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.14756 | 0.059078 | 0.89747 | 0.0069923 | 0.012686 | 0.029041 | 0.069504 | 0.027501 | 0.10932 | 0.14033 | 0.040028 | 0.030152 | 0.07536 | 0.1418 | 0.29393 | 0.081587 | 0.19858 | 0.081948 | 0.037766 | 0.042884 | 0.0096214 | 1.8379 | 0.036618 | 0.0052735 | 0.00067722 | 0.0024558 | 0.0054139 | 0.0040691 | 0.00090628 | 0.073701 | 0.070861 | 145024.7405 | 197237.8309 | 11592.704 | 218079.963 | 183122.8984 | 174560.7156 | 3929.207 | 0.035507 | 0.083399 | 0.19105 | 0.33247
68 | CCV 1ppb TRACE |  | 06/08/2023 20:36:39 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.422 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.14132 | 0.059637 | 0.83296 | 0.0065874 | 0.011949 | 0.031197 | 0.064096 | 0.0262 | 0.10392 | 0.13745 | 0.040038 | 0.02951 | 0.076069 | 0.15004 | 0.30498 | 0.081804 | 0.20401 | 0.083007 | 0.037058 | 0.040523 | 0.0099911 | 1.8025 | 0.037296 | 0.0046778 | 0.00069517 | 0.0023517 | 0.0053853 | 0.0037081 | 0.0010452 | 0.071575 | 0.07211 | 152999.6571 | 197114.6703 | 11760.8397 | 222417.4376 | 181983.2261 | 173412.0472 | 4079.9163 | 0.035004 | 0.08292 | 0.17841 | 0.33373
69 | CCV 10 ppb MAJORS |  | 06/08/2023 20:42:27 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.423 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.3755 | 0.56473 | 4.3991 | 0.037657 | 0.15885 | 0.21902 | 0.63717 | 0.25018 | 0.95346 | 1.3137 | 0.40195 | 0.27292 | 0.7547 | 1.1523 | 2.7409 | 0.76256 | 1.7916 | 0.78348 | 0.37275 | 0.38307 | 0.0847 | 18.0218 | 0.36972 | 0.038411 | 0.0078894 | 0.024205 | 0.052178 | 0.036293 | 0.0095941 | 0.72541 | 0.72532 | 148737.1186 | 186170.5961 | 11652.7747 | 207116.6258 | 175274.824 | 165644.8926 | 3819.1771 | 0.3442 | 0.81422 | 1.7107 | 3.1397
70 | CCV 10 ppb MAJORS |  | 06/08/2023 20:48:44 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.424 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.4175 | 0.57935 | 4.3969 | 0.035 | 0.16179 | 0.22825 | 0.66763 | 0.26109 | 0.99016 | 1.3682 | 0.40219 | 0.26991 | 0.75355 | 1.2171 | 2.9266 | 0.80025 | 1.8741 | 0.77626 | 0.36795 | 0.39219 | 0.094355 | 18.4676 | 0.37146 | 0.040164 | 0.0078016 | 0.024962 | 0.053315 | 0.036928 | 0.0095279 | 0.71903 | 0.7306 | 148366.8781 | 188017.1987 | 11362.518 | 209434.1675 | 176612.7415 | 165476.9444 | 3781.8339 | 0.33885 | 0.80962 | 1.7051 | 3.0857
71 | 8_2_21_W_8 |  | 06/08/2023 20:54:33 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_W_8.425 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.28949 | 0.081429 | 3.5279 | 0.02762 | 0.07306 | 0.053755 | 0.028657 | 0.16885 | 0.025944 | 0.0098525 | 0.0028342 | 0.27771 | 0.51577 | 1.2247 | 0.0025445 | 0.0041144 | 0.23304 | 0.098695 | 0.025416 | 0.058432 | 0.011987 | 2.632 | 0.052665 | 0.0011282 | -1.3856e-06 | 8.6079e-05 | 0.00015899 | 0.00042071 | 2.7066e-05 | 0.041984 | 0.039797 | 151008.6145 | 196944.9727 | 11736.8207 | 209014.0186 | 179169.2427 | 178018.3857 | 3943.878 | 0.00015842 | 0.0004674 | 0.012851 | 0.023612
72 | 8_2_21_W_9 |  | 06/08/2023 21:00:21 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_W_9.426 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.19968 | 0.073564 | 3.1231 | 0.025308 | 0.067213 | 0.042959 | 0.022583 | 0.14841 | 0.017319 | 0.0034722 | 0.0010796 | 0.20431 | 0.32404 | 0.94848 | 0.0019389 | 0.003821 | 0.15132 | 0.065525 | 0.015906 | 0.040526 | 0.0084224 | 1.7741 | 0.036381 | 0.00066372 | 2.1049e-06 | 5.2727e-05 | 8.5025e-05 | 9.9835e-05 | 1.0268e-05 | 0.028727 | 0.028574 | 153587.9282 | 195787.9059 | 11634.7376 | 214507.6995 | 181533.241 | 171918.3455 | 4014.5642 | 0.00013192 | 0.00030728 | 0.0074436 | 0.014126
73 | 8_2_21_W_9_DUP |  | 06/08/2023 21:06:09 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_W_9_DUP.427 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.19715 | 0.072721 | 2.9568 | 0.027269 | 0.059283 | 0.041809 | 0.022946 | 0.13712 | 0.017927 | 0.0027726 | 0.00055982 | 0.183 | 0.25145 | 0.84476 | 0.0013583 | 0.0040832 | 0.1518 | 0.061834 | 0.016022 | 0.042133 | 0.0083982 | 1.8063 | 0.035803 | 0.0015506 | -1.1621e-05 | 5.8359e-05 | 9.3123e-05 | 0.00019734 | 2.9941e-05 | 0.028546 | 0.027764 | 151475.7488 | 199085.0722 | 11688.1129 | 218782.5761 | 181715.0784 | 172982.3857 | 3957.2158 | 0.00010186 | 0.00029101 | 0.0072799 | 0.013307
74 | 8_2_21_W_10 |  | 06/08/2023 21:11:57 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_W_10.428 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.20182 | 0.1017 | 10.2775 | 0.085604 | 0.081976 | 0.069131 | 0.02939 | 0.18161 | 0.01973 | 0.0050887 | 0.0014249 | 0.50445 | 0.46702 | 2.3766 | 0.0044441 | 0.0043342 | 0.19427 | 0.079072 | 0.042222 | 0.045536 | 0.0095973 | 2.097 | 0.042437 | 0.0012002 | 5.9834e-06 | 3.9503e-05 | 0.00015214 | 0.0004404 | 1.3843e-05 | 0.026979 | 0.027585 | 153897.9236 | 192903.149 | 11392.5414 | 214494.3991 | 172239.4765 | 171600.1991 | 3902.5331 | 0.00022204 | 0.00050576 | 0.0211 | 0.041774
75 | 8_2_21_W_11 |  | 06/08/2023 21:17:45 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_W_11.429 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.18701 | 0.10244 | 17.2867 | 0.12651 | 0.092247 | 0.070657 | 0.024986 | 0.17728 | 0.023984 | 0.0067056 | 0.0020313 | 0.78515 | 0.77703 | 3.5255 | 0.0054162 | 0.0044386 | 0.24259 | 0.10436 | 0.076565 | 0.049626 | 0.0090515 | 2.1411 | 0.043769 | 0.0014201 | -7.7216e-06 | 4.3792e-05 | 0.00014363 | 0.0003694 | 3.8176e-05 | 0.026306 | 0.026045 | 151160.2085 | 191623.62 | 11433.9086 | 212452.1725 | 179374.0728 | 170905.0172 | 3917.2039 | 0.00037831 | 0.00073571 | 0.033673 | 0.061408
76 | 5_10_23_DB |  | 06/08/2023 21:23:33 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_10_23_DB.430 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.00030007 | 0 | 0.2026 | 0.0016202 | 0.068421 | 0.0073901 | -0.00025928 | -4.7652e-05 | -4.7652e-05 | 0.0029122 | 0.00049228 | 0.0036459 | 0.00014701 | 0.016743 | -0.00012349 | 0.00044086 | 0.0079052 | 0.0030024 | 4.7049e-05 | 0.0016685 | 1.0437e-05 | 0.002596 | 5.7788e-05 | 0.0010349 | -1.4869e-05 | 1.6582e-06 | 1.0154e-05 | 0.00010199 | 1.0312e-05 | 2.7075e-05 | -1.3496e-05 | 148073.5955 | 193943.6244 | 11582.027 | 218987.3978 | 181805.5748 | 170013.9311 | 4001.8943 | -1.8385e-06 | 4.6588e-06 | 0.00042995 | 0.00084739
77 | 5_10_23_MFB |  | 06/08/2023 21:29:21 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_10_23_MFB.431 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.71093 | 0.30358 | 2.3811 | 0.019929 | 0.191 | 0.11561 | 0.33468 | 0.13002 | 0.50955 | 0.71974 | 0.19835 | 0.14201 | 0.37384 | 0.68248 | 1.4762 | 0.39934 | 0.98038 | 0.3868 | 0.18088 | 0.18678 | 0.044223 | 8.7709 | 0.17916 | 0.019732 | 0.0037794 | 0.012336 | 0.025876 | 0.01778 | 0.0046366 | 0.34369 | 0.3501 | 151965.338 | 187628.8113 | 10844.7824 | 211356.8588 | 177157.7026 | 165793.4124 | 3833.181 | 0.16884 | 0.39628 | 0.85344 | 1.6093
78 | 5_10_23_T-239 |  | 06/08/2023 21:35:10 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_10_23_T-239.432 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 6.8254 | 0.76635 | 53.7707 | 0.41402 | 0.24435 | 0.025912 | 0.2625 | 1.3161 | 0.19384 | 0.19594 | 0.054219 | 1.1026 | 0.2687 | 5.22 | 0.42435 | 0.13355 | 0.14798 | 0.063404 | 0.21598 | 0.093907 | 0.020879 | 4.2945 | 0.087311 | 0.019114 | 0.0015015 | 0.0047196 | 0.010211 | 0.008077 | 0.0018854 | 0.075269 | 0.076407 | 142701.2124 | 175083.7789 | 10433.1448 | 200398.5356 | 165380.551 | 153588.5626 | 3561.1106 | 0.11496 | 0.26269 | 0.21181 | 0.38974
79 | 5_10_23_DB |  | 06/08/2023 21:40:58 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_10_23_DB.433 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.00046993 | 0 | 0.17743 | 0.0018834 | 0.075267 | 0.0081915 | -0.00016212 | 0.00012254 | 0.00011883 | 0.00065222 | 5.3475e-05 | 0.0011573 | 5.1023e-05 | 0.0072972 | -0.00012717 | 0.0003698 | 0.0047063 | 0.0019152 | -0.00015693 | 0.0042733 | 6.7044e-05 | 0.0010201 | 2.8727e-05 | 0.00032449 | -1.8296e-05 | -1.9536e-06 | 2.191e-05 | 1.8556e-05 | 3.1539e-06 | 9.7601e-05 | 4.8151e-07 | 173509.0536 | 206492.0689 | 11761.5104 | 229050.3853 | 190236.6092 | 177987.5134 | 4372.0026 | 6.8092e-06 | 3.2139e-06 | 0.00016452 | 0.00049452
80 | 8_16_21_UF_1 |  | 06/08/2023 21:46:46 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_UF_1.434 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.3988 | 0.34696 | 0.91014 | 0.008125 | 0.40257 | 0.55667 | 0.14629 | 0.97705 | 0.045919 | 0.010672 | 0.0030017 | 0.016039 | 0.96376 | 0.12905 | 0.0083354 | 0.026528 | 0.13434 | 0.055732 | 0.033158 | 0.24419 | 0.055058 | 11.0932 | 0.23024 | 0.0024178 | 2.1541e-05 | 0.00014352 | 0.00034842 | 0.001517 | 6.1661e-05 | 1.2808 | 1.2919 | 146394.5047 | 183870.3976 | 10848.1206 | 204865.8915 | 170801.193 | 164289.3307 | 3816.5098 | 0.00027499 | 0.00067722 | 0.0027631 | 0.0046026
81 | LBLANK |  | 06/08/2023 21:52:35 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.435 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.00017742 | 0 | 0.01238 | 0.00017393 | 0.006737 | 0.0045604 | -0.0001212 | 6.4882e-05 | 6.4781e-05 | 0.0005326 | 0.00012027 | -0.00038316 | 1.8007e-05 | -0.00083455 | -0.00018284 | 0.00046965 | 9.3961e-05 | 0.00019642 | 0.00012362 | 0.0024076 | 7.7689e-06 | 0.020455 | 0.00042191 | 0.00075466 | -1.1584e-05 | 1.1445e-05 | 6.2257e-06 | 1.1691e-05 | 1.0034e-05 | 0.00019104 | 0.00017651 | 154284.8747 | 199157.9982 | 11917.6367 | 222626.8242 | 185305.4694 | 176873.9887 | 4075.9151 | -4.8247e-06 | 1.2857e-05 | -0.00010136 | -0.00010608
82 | CCV 1ppb TRACE |  | 06/08/2023 21:58:24 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.436 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.14502 | 0.06204 | 0.84378 | 0.0081015 | 0.015716 | 0.027893 | 0.068423 | 0.027035 | 0.10395 | 0.14359 | 0.040693 | 0.030692 | 0.079016 | 0.14462 | 0.29969 | 0.085047 | 0.20158 | 0.083558 | 0.037584 | 0.041074 | 0.0092429 | 1.8305 | 0.037669 | 0.0047702 | 0.00078316 | 0.0024777 | 0.0051952 | 0.0039357 | 0.0010404 | 0.072897 | 0.073198 | 153867.6887 | 195447.924 | 11630.7343 | 218924.7673 | 183883.909 | 174643.8236 | 4020.5657 | 0.035426 | 0.082563 | 0.18048 | 0.33358
83 | CCV 1ppb TRACE |  | 06/08/2023 22:04:41 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.437 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.14318 | 0.060168 | 0.83861 | 0.0064047 | 0.012402 | 0.028648 | 0.065489 | 0.026107 | 0.10379 | 0.14196 | 0.040529 | 0.029753 | 0.076884 | 0.14674 | 0.29895 | 0.07826 | 0.21226 | 0.083523 | 0.038294 | 0.03925 | 0.0092148 | 1.8431 | 0.03747 | 0.0048389 | 0.00082141 | 0.0023984 | 0.0053879 | 0.00378 | 0.00092073 | 0.071768 | 0.073925 | 157807.5934 | 196153.6249 | 11787.5288 | 220287.0751 | 175559.214 | 186761.9388 | 3989.8908 | 0.034796 | 0.08122 | 0.18108 | 0.35595
84 | LFB |  | 06/08/2023 22:10:29 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.438 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.3948 | 0.57944 | 4.4763 | 0.038203 | 0.19889 | 0.22849 | 0.62715 | 0.25572 | 0.97156 | 1.3357 | 0.39327 | 0.26875 | 0.73922 | 1.1823 | 2.8551 | 0.77807 | 1.8292 | 0.76868 | 0.36946 | 0.38866 | 0.092565 | 18.1757 | 0.36778 | 0.039558 | 0.0078105 | 0.025053 | 0.053134 | 0.037996 | 0.0096994 | 0.71266 | 0.70967 | 148815.4723 | 186452.3577 | 11102.3126 | 210449.5088 | 173192.189 | 165984.0832 | 3772.4981 | 0.33555 | 0.79332 | 1.6708 | 3.0904
85 | CCV 10 ppb MAJORS |  | 06/08/2023 22:16:17 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.439 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.4092 | 0.5928 | 4.3336 | 0.036052 | 0.1673 | 0.2259 | 0.66736 | 0.26729 | 0.98912 | 1.4032 | 0.41061 | 0.27591 | 0.75932 | 1.215 | 2.9339 | 0.80303 | 1.8625 | 0.78711 | 0.36984 | 0.38657 | 0.090292 | 18.0986 | 0.36745 | 0.039976 | 0.0076827 | 0.023927 | 0.052547 | 0.035891 | 0.0095446 | 0.72826 | 0.72592 | 150395.0798 | 188900.064 | 11344.5042 | 208679.1408 | 174541.9607 | 167381.5579 | 3835.8485 | 0.34006 | 0.80538 | 1.6865 | 3.1532
86 | CCV 10 ppb MAJORS |  | 06/08/2023 22:22:34 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.440 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.3763 | 0.58305 | 4.3393 | 0.034116 | 0.16195 | 0.22143 | 0.6558 | 0.25961 | 0.96246 | 1.3573 | 0.41022 | 0.27358 | 0.75991 | 1.1875 | 2.8727 | 0.77059 | 1.8227 | 0.78575 | 0.37088 | 0.392 | 0.091364 | 18.0229 | 0.37269 | 0.040675 | 0.0080092 | 0.025282 | 0.052592 | 0.037234 | 0.0093055 | 0.72637 | 0.73493 | 149959.6381 | 187757.9117 | 11565.3474 | 211660.6552 | 174267.939 | 166663.2808 | 3884.5283 | 0.33994 | 0.80007 | 1.6928 | 3.1556
87 | 8_16_21_UF_6 |  | 06/08/2023 22:28:23 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_UF_6.441 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.466 | 0.37152 | 0.51692 | 0.0051794 | 0.22758 | 0.37987 | 0.13002 | 0.89112 | 0.079366 | 0.0079305 | 0.0020297 | 0.011697 | 0.84499 | 0.094669 | 0.005945 | 0.05462 | 0.2928 | 0.12493 | 0.078915 | 0.22003 | 0.048299 | 10.1429 | 0.20688 | 0.0026811 | 6.536e-05 | 0.00019064 | 0.00053683 | 0.0019584 | 6.8734e-05 | 0.63854 | 0.64043 | 145853.4517 | 183913.6891 | 11098.3094 | 203017.8234 | 168273.9935 | 166532.4473 | 3751.8262 | 0.00050366 | 0.00113 | 0.018495 | 0.034857
88 | 8_16_21_UF_8 |  | 06/08/2023 22:34:11 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_UF_8.442 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.4892 | 0.40373 | 0.51278 | 0.0038086 | 0.30099 | 0.31896 | 0.14227 | 0.85139 | 0.099512 | 0.012997 | 0.0031405 | 0.017803 | 0.86266 | 0.11701 | 0.0052843 | 0.03167 | 0.3248 | 0.13861 | 0.049302 | 0.22114 | 0.052385 | 10.1555 | 0.20522 | 0.0023183 | 4.3552e-05 | 0.00016553 | 0.00041571 | 0.0017808 | 9.4402e-05 | 0.50899 | 0.50783 | 144637.8543 | 183308.6429 | 10733.3644 | 201475.4643 | 166790.8451 | 166001.7068 | 3704.4804 | 0.00040206 | 0.00087091 | 0.0082907 | 0.015843
89 | 8_16_21_UF_9 |  | 06/08/2023 22:39:59 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_UF_9.443 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.2472 | 0.37779 | 0.69629 | 0.0058202 | 0.30006 | 0.27208 | 0.11098 | 0.75347 | 0.082716 | 0.008192 | 0.0023014 | 0.0098404 | 0.4738 | 0.072949 | 0.0042872 | 0.058106 | 0.23729 | 0.10259 | 0.026058 | 0.16983 | 0.031826 | 7.9723 | 0.16259 | 0.0022093 | 3.6372e-05 | 0.00019841 | 0.00034583 | 0.0012641 | 4.3601e-05 | 0.33316 | 0.34218 | 146112.1135 | 183269.5479 | 10950.2069 | 202531.2955 | 166667.9997 | 164515.2169 | 3737.8223 | 0.00022176 | 0.00057762 | 0.0050056 | 0.0090118
90 | 8_16_21_UF_10 |  | 06/08/2023 22:45:47 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_UF_10.444 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.2367 | 0.52503 | 0.93902 | 0.0072045 | 0.32952 | 0.32621 | 0.12983 | 0.93872 | 0.085936 | 0.012621 | 0.0036682 | 0.023737 | 0.27276 | 0.14566 | 0.0040665 | 0.022539 | 0.21898 | 0.093105 | 0.03937 | 0.18276 | 0.037997 | 8.4436 | 0.17142 | 0.0023532 | 7.9315e-05 | 0.00023674 | 0.00045204 | 0.0014353 | 0.00010117 | 0.16386 | 0.16296 | 145006.2357 | 184453.1832 | 10659.9811 | 200308.5813 | 164348.2471 | 164848.3349 | 3745.8245 | 0.00026821 | 0.00066014 | 0.0089749 | 0.017407
91 | 8_16_21_UF_11 |  | 06/08/2023 22:51:35 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_UF_11.445 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.0867 | 0.50419 | 0.43588 | 0.0041331 | 0.29868 | 0.31895 | 0.12828 | 0.83942 | 0.081413 | 0.0063215 | 0.0015569 | 0.0064689 | 0.18643 | 0.064854 | 0.0028194 | 0.010401 | 0.1716 | 0.075116 | 0.035432 | 0.17941 | 0.03781 | 8.0932 | 0.16535 | 0.0023875 | 5.5038e-05 | 0.00021807 | 0.00042446 | 0.0011971 | 8.7872e-05 | 0.13252 | 0.13297 | 141989.4967 | 182475.4969 | 10968.2289 | 199627.873 | 164480.1972 | 162984.1069 | 3730.4874 | 9.8614e-05 | 0.00030135 | 0.01018 | 0.018453
92 | 8_16_21_F_1 |  | 06/08/2023 22:57:23 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_F_1.446 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.2379 | 0.35329 | 7.5518 | 0.056757 | 0.25121 | 0.53282 | 0.14246 | 0.94697 | 0.040604 | 0.009015 | 0.0024098 | 0.14823 | 1.1224 | 0.67818 | 0.009362 | 0.026302 | 0.18163 | 0.081998 | 0.041079 | 0.25616 | 0.056534 | 12.3978 | 0.25271 | 0.0025603 | 3.1353e-05 | 0.00016218 | 0.00038881 | 0.0016277 | 7.2738e-05 | 1.3762 | 1.4265 | 143605.6724 | 174878.3571 | 10888.1484 | 198571.1648 | 166624.8557 | 154279.6997 | 3571.781 | 0.00026755 | 0.00056508 | 0.0046847 | 0.0090402
93 | 8_16_21_F_1_DUP |  | 06/08/2023 23:03:11 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_F_1_DUP.447 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.2559 | 0.36528 | 1.0958 | 0.0092349 | 0.22628 | 0.55538 | 0.14439 | 0.98546 | 0.040841 | 0.0095603 | 0.0026221 | 0.14586 | 1.0631 | 0.68241 | 0.0089168 | 0.025461 | 0.20779 | 0.089206 | 0.040582 | 0.25072 | 0.055816 | 11.9458 | 0.23741 | 0.0018966 | 2.9074e-05 | 0.0001696 | 0.0003828 | 0.0013678 | 5.8198e-05 | 1.3555 | 1.3358 | 143505.8625 | 183033.7676 | 10559.2349 | 198436.695 | 166562.1039 | 165975.1206 | 3639.1303 | 0.00024043 | 0.00058932 | 0.005819 | 0.011192
94 | 8_16_21_F_1_DUP_LD |  | 06/08/2023 23:08:58 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_F_1_DUP_LD.448 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.1827 | 0.3345 | 1.6958 | 0.011285 | 0.25326 | 0.52905 | 0.137 | 0.93098 | 0.037517 | 0.0098158 | 0.00286 | 0.14917 | 1.0609 | 0.64951 | 0.01035 | 0.026452 | 0.20106 | 0.091691 | 0.043349 | 0.26051 | 0.053147 | 11.7838 | 0.24073 | 0.0019509 | 2.2259e-05 | 0.00012391 | 0.00031007 | 0.0015821 | 7.3754e-05 | 1.3346 | 1.3604 | 142068.9374 | 180872.5584 | 11028.2656 | 199426.4782 | 165730.5151 | 164186.149 | 3695.1445 | 0.00025264 | 0.00066637 | 0.005904 | 0.011646
95 | 8_16_21_F_1_DUP_LD |  | 06/08/2023 23:15:14 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_F_1_DUP_LD.449 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.2399 | 0.35778 | 1.6988 | 0.012454 | 0.25696 | 0.54177 | 0.14497 | 0.97469 | 0.042337 | 0.010903 | 0.002876 | 0.14973 | 1.0606 | 0.72488 | 0.0095852 | 0.030304 | 0.21264 | 0.091067 | 0.042686 | 0.25411 | 0.056819 | 11.3579 | 0.2366 | 0.0023859 | 2.1863e-05 | 0.00016188 | 0.00031005 | 0.0017803 | 9.0807e-05 | 1.3398 | 1.3473 | 142759.0929 | 183040.6045 | 10669.3186 | 198842.1051 | 168291.5005 | 162047.9195 | 3812.5087 | 0.00032087 | 0.00066946 | 0.0065051 | 0.011643
96 | 8_16_21_F_1_DUP_LSP |  | 06/08/2023 23:21:03 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_F_1_DUP_LSP.450 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9.0498 | 3.4418 | 8.2943 | 0.062953 | 0.52991 | 1.739 | 3.8307 | 2.4379 | 1.4393 | 1.8678 | 0.55855 | 0.55072 | 2.0987 | 2.4181 | 3.8784 | 1.0489 | 2.5709 | 1.0614 | 0.50338 | 0.77933 | 0.17867 | 36.8137 | 0.73534 | 0.053499 | 0.010577 | 0.032692 | 0.069923 | 0.052251 | 0.012621 | 2.3928 | 2.3417 | 130991.1669 | 162453.1648 | 9491.1523 | 177873.9432 | 163572.6962 | 143774.5962 | 3245.0352 | 0.42722 | 1.0129 | 2.0465 | 3.4976
97 | LBLANK |  | 06/08/2023 23:26:52 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.451 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.00046927 | 5.8018e-05 | 0.010315 | 6.5853e-05 | 0.0013675 | 0.013046 | 0.00013032 | 8.9173e-06 | 9.1004e-06 | 0.00026794 | 9.5886e-05 | -0.00034186 | 4.1304e-05 | 0.0083418 | -1.0168e-05 | 0.00083675 | 0.00067853 | 0.00012864 | 6.2647e-05 | 5.3152e-05 | 9.1004e-06 | 0.012074 | 0.00026459 | 0.0020437 | -1.1822e-05 | 7.6622e-06 | 2.8498e-05 | 0.00045389 | 9.8236e-06 | 0.0007656 | 0.00040898 | 151805.4569 | 206019.8779 | 11674.1019 | 233794.9761 | 190370.0657 | 182678.618 | 4385.3399 | 8.5024e-07 | 5.9331e-05 | -2.2144e-05 | -0.00011176
98 | CCV 1ppb TRACE |  | 06/08/2023 23:32:40 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.452 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.14395 | 0.061571 | 0.86331 | 0.0070934 | 0.013563 | 0.030655 | 0.068885 | 0.027568 | 0.10332 | 0.14006 | 0.040991 | 0.031505 | 0.078056 | 0.14338 | 0.29228 | 0.083171 | 0.20609 | 0.084775 | 0.038469 | 0.041231 | 0.00916 | 1.8147 | 0.0376 | 0.0055054 | 0.00078495 | 0.0025253 | 0.0051103 | 0.0038647 | 0.0010167 | 0.075 | 0.073113 | 158698.786 | 202687.2153 | 11868.9288 | 225014.6365 | 187276.2877 | 177635.8258 | 4200.6177 | 0.033966 | 0.082911 | 0.17774 | 0.32678
99 | CCV 1ppb TRACE |  | 06/08/2023 23:38:58 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.453 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.14389 | 0.061001 | 0.83436 | 0.0067302 | 0.012419 | 0.034195 | 0.067299 | 0.027315 | 0.10362 | 0.13757 | 0.040393 | 0.029895 | 0.076384 | 0.15055 | 0.29492 | 0.082268 | 0.20536 | 0.082615 | 0.037602 | 0.039997 | 0.0097835 | 1.7842 | 0.036966 | 0.0053283 | 0.00068174 | 0.0024272 | 0.005589 | 0.0038471 | 0.00099501 | 0.072832 | 0.071722 | 158892.8906 | 203710.2316 | 11976.3523 | 225889.6732 | 188049.6282 | 178297.8512 | 4220.6235 | 0.034603 | 0.081324 | 0.17686 | 0.32618
100 | CCV 10 ppb MAJORS |  | 06/08/2023 23:44:46 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.454 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.3924 | 0.57196 | 4.33 | 0.03525 | 0.16431 | 0.22643 | 0.631 | 0.2582 | 0.98633 | 1.4044 | 0.40809 | 0.27407 | 0.75862 | 1.2089 | 2.9381 | 0.80469 | 1.8845 | 0.7863 | 0.37155 | 0.3858 | 0.089986 | 17.8337 | 0.36481 | 0.040446 | 0.0077058 | 0.024413 | 0.052569 | 0.037229 | 0.0092474 | 0.72602 | 0.71941 | 152904.0275 | 192373.9894 | 11333.1613 | 211060.7211 | 174236.1495 | 167534.7402 | 3937.2095 | 0.33829 | 0.79855 | 1.6834 | 3.1771
101 | CCV 10 ppb MAJORS |  | 06/08/2023 23:51:04 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.455 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.3492 | 0.57597 | 4.3072 | 0.034088 | 0.16351 | 0.21923 | 0.64052 | 0.25998 | 0.98225 | 1.3475 | 0.4103 | 0.27644 | 0.76502 | 1.1896 | 2.9194 | 0.78844 | 1.8366 | 0.78431 | 0.37024 | 0.39239 | 0.094292 | 17.9697 | 0.36655 | 0.040833 | 0.0078238 | 0.024359 | 0.051858 | 0.037318 | 0.0097049 | 0.72328 | 0.72278 | 152966.1041 | 190688.2939 | 11556.6785 | 212473.2897 | 176740.6762 | 170801.7246 | 3889.196 | 0.33596 | 0.80388 | 1.7008 | 3.1236
102 | 8_16_21_F_6 |  | 06/08/2023 23:56:52 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_F_6.456 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.3295 | 0.38903 | 3.1741 | 0.02072 | 0.20019 | 0.41927 | 0.14092 | 0.92116 | 0.07608 | 0.010971 | 0.0027769 | 0.20382 | 0.93723 | 0.92939 | 0.0081689 | 0.070908 | 0.61897 | 0.25917 | 0.19053 | 0.25128 | 0.050444 | 10.8132 | 0.22452 | 0.0031639 | 5.7937e-05 | 0.00025587 | 0.00049052 | 0.0022984 | 6.5377e-05 | 0.71201 | 0.7095 | 140643.0302 | 183690.8585 | 10948.1937 | 201553.521 | 168587.6045 | 166661.9183 | 3816.5101 | 0.00053798 | 0.0014739 | 1.2987 | 2.2903
103 | 8_16_21_F_6_SP |  | 06/09/2023 00:02:41 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_F_6_SP.457 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.1522 | 0.71797 | 5.0439 | 0.039527 | 0.32349 | 0.53477 | 0.48567 | 1.0932 | 0.6297 | 0.75032 | 0.21976 | 0.3374 | 1.299 | 1.551 | 1.5798 | 0.48897 | 1.5786 | 0.65449 | 0.3829 | 0.434 | 0.1004 | 19.4479 | 0.40387 | 0.022159 | 0.0041769 | 0.012803 | 0.026704 | 0.020685 | 0.0051201 | 1.0745 | 1.056 | 146933.8133 | 181844.2791 | 10369.7656 | 196539.514 | 165737.2397 | 160337.5133 | 3776.4992 | 0.17844 | 0.42191 | 0.95951 | 1.8083
104 | 8_16_21_F_8 |  | 06/09/2023 00:08:29 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_F_8.458 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.4352 | 0.43448 | 1.488 | 0.013274 | 0.2962 | 0.37625 | 0.14677 | 0.93834 | 0.11367 | 0.0097096 | 0.0026476 | 0.18164 | 0.96992 | 0.87579 | 0.0069919 | 0.041389 | 0.72931 | 0.28901 | 0.082823 | 0.23844 | 0.054632 | 10.6932 | 0.22439 | 0.0022329 | 2.545e-05 | 0.00028464 | 0.00052642 | 0.0019661 | 6.8854e-05 | 0.55285 | 0.56113 | 148334.8289 | 183929.9477 | 10610.6121 | 206282.1662 | 169214.4402 | 166853.4334 | 3859.1879 | 0.00053861 | 0.00098289 | 0.02605 | 0.04828
105 | 8_16_21_F_8_DUP |  | 06/09/2023 00:14:18 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_F_8_DUP.459 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.4038 | 0.40196 | 1.515 | 0.012689 | 0.2895 | 0.35583 | 0.141 | 0.86176 | 0.097616 | 0.014271 | 0.0036592 | 0.188 | 0.93689 | 0.87444 | 0.0075394 | 0.04074 | 0.68485 | 0.28271 | 0.089451 | 0.23571 | 0.049644 | 10.6836 | 0.22223 | 0.0023617 | 6.6292e-05 | 0.00023323 | 0.00045946 | 0.0015969 | 0.00010677 | 0.55516 | 0.55514 | 146150.4905 | 181524.745 | 10626.6181 | 200667.6459 | 171527.392 | 162111.3451 | 3776.4992 | 0.00046044 | 0.0010346 | 0.026969 | 0.048259
106 | 8_16_21_F_9 |  | 06/09/2023 00:20:06 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_F_9.460 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.1801 | 0.37648 | 2.5208 | 0.02259 | 0.30991 | 0.32199 | 0.1168 | 0.80428 | 0.087643 | 0.033295 | 0.0085387 | 0.16638 | 0.52986 | 0.76797 | 0.0064722 | 0.094395 | 0.53888 | 0.22263 | 0.052828 | 0.19246 | 0.041677 | 8.4471 | 0.17728 | 0.0024069 | 7.2029e-05 | 0.00020075 | 0.0004848 | 0.0013423 | 9.0209e-05 | 0.37598 | 0.37681 | 146340.505 | 184519.367 | 10719.3537 | 201528.1908 | 169107.6903 | 158704.1271 | 3874.5257 | 0.00027644 | 0.00077143 | 0.018141 | 0.033784
107 | 8_16_21_F_10 |  | 06/09/2023 00:25:54 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_F_10.461 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.0704 | 0.53187 | 1.7031 | 0.013185 | 0.29049 | 0.36423 | 0.14054 | 0.97685 | 0.090023 | 0.0067304 | 0.0022669 | 0.1081 | 0.32318 | 0.53618 | 0.0054765 | 0.012925 | 0.41799 | 0.17229 | 0.057857 | 0.19643 | 0.041204 | 9.1766 | 0.18912 | 0.0024538 | 3.3785e-05 | 0.00021056 | 0.00062923 | 0.0016647 | 8.5162e-05 | 0.17654 | 0.18183 | 146472.4056 | 179483.427 | 10625.9509 | 201351.284 | 166230.9905 | 161828.1226 | 3701.1465 | 0.00038939 | 0.00079034 | 0.01802 | 0.033471
108 | 8_16_21_F_10_DUP |  | 06/09/2023 00:31:42 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_F_10_DUP.462 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.02 | 0.508 | 1.7285 | 0.013648 | 0.27719 | 0.34485 | 0.13257 | 0.9228 | 0.083586 | 0.0093177 | 0.0026693 | 0.099605 | 0.29706 | 0.50147 | 0.0048346 | 0.036565 | 0.42241 | 0.16778 | 0.051412 | 0.19465 | 0.039758 | 8.3844 | 0.17329 | 0.0025365 | 7.3021e-05 | 0.00027102 | 0.00041139 | 0.0016299 | 7.0927e-05 | 0.16686 | 0.16649 | 145883.1163 | 188496.433 | 10777.3976 | 206909.6969 | 171108.0873 | 167377.3039 | 3895.8652 | 0.00031763 | 0.00078938 | 0.018056 | 0.03226
109 | 8_16_21_F_11 |  | 06/09/2023 00:37:30 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_F_11.463 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.0248 | 0.52663 | 2.1334 | 0.019034 | 0.26687 | 0.36782 | 0.13376 | 0.92188 | 0.089958 | 0.0067944 | 0.0022948 | 0.11105 | 0.23728 | 0.5343 | 0.0058586 | 0.014848 | 0.40255 | 0.16894 | 0.060055 | 0.19744 | 0.040211 | 8.7478 | 0.18427 | 0.0026287 | 7.7873e-05 | 0.00026053 | 0.00058025 | 0.0017079 | 0.00010322 | 0.15086 | 0.14934 | 144666.9391 | 180493.8691 | 10726.6984 | 198145.6165 | 167471.2832 | 165066.9248 | 3801.8392 | 0.00014997 | 0.00042461 | 0.0217 | 0.038009
110 | 8_16_21_W_1 |  | 06/09/2023 00:43:18 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_W_1.464 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.23506 | 0.066536 | 1.8331 | 0.014705 | 0.12154 | 0.12253 | 0.025208 | 0.18863 | 0.0074681 | 0.0021174 | 0.00068568 | 0.15079 | 0.3243 | 0.59707 | 0.0028473 | 0.0057104 | 0.054457 | 0.026732 | 0.0087202 | 0.053535 | 0.011313 | 2.4348 | 0.049254 | 0.0014069 | -1.1368e-05 | 5.0145e-05 | 4.838e-05 | 0.00043088 | 2.0755e-05 | 0.29942 | 0.26949 | 151712.7536 | 192632.0577 | 11662.1026 | 195194.3093 | 176202.049 | 164380.6292 | 3897.1982 | 0.00010036 | 0.00019238 | 0.003771 | 0.0070336
111 | 8_16_21_W_6 |  | 06/09/2023 00:49:05 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_W_6.465 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.27023 | 0.080021 | 4.1532 | 0.035618 | 0.05594 | 0.089542 | 0.027169 | 0.19462 | 0.017059 | 0.0038492 | 0.00090611 | 0.28221 | 0.33221 | 1.2937 | 0.0019765 | 0.0047052 | 0.21968 | 0.089343 | 0.025728 | 0.049715 | 0.011606 | 2.2591 | 0.046961 | 0.001202 | 9.2691e-06 | 7.4323e-05 | 0.00010012 | 0.0006729 | 2.4209e-05 | 0.14024 | 0.13792 | 159036.6443 | 193000.3809 | 11276.5336 | 214954.5927 | 178160.0998 | 171570.2899 | 4012.5635 | 0.00015975 | 0.00037986 | 0.017892 | 0.032315
112 | LBLANK |  | 06/09/2023 00:54:53 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.466 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.00011425 | 0 | 0.0065569 | 0.00017256 | 0.0021652 | 0.0065314 | -0.00030986 | -4.7652e-05 | -4.7652e-05 | 0.00018281 | 8.425e-05 | -0.00035323 | 3.2378e-05 | -0.0021897 | -0.0001828 | -4.4789e-05 | 0.00018383 | 0.00019495 | 5.0709e-05 | 0.0016368 | -4.7652e-05 | -0.0014418 | -2.6101e-05 | 0.00088314 | -1.8296e-05 | 1.5712e-06 | 6.4199e-06 | -0.00012081 | 1.0181e-05 | 0.00010004 | 8.7253e-05 | 154609.6876 | 197141.0332 | 12136.4873 | 219060.2398 | 177084.2991 | 175147.5835 | 4040.5715 | -1.7937e-06 | 2.8941e-05 | -4.0541e-05 | -7.3547e-05
113 | CCV 1ppb TRACE |  | 06/09/2023 01:00:42 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.467 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.14174 | 0.059233 | 0.81971 | 0.0068174 | 0.01597 | 0.032058 | 0.067248 | 0.025872 | 0.10187 | 0.13778 | 0.040111 | 0.030222 | 0.077168 | 0.14699 | 0.29577 | 0.082251 | 0.20258 | 0.082382 | 0.037348 | 0.038326 | 0.0091747 | 1.8265 | 0.036583 | 0.0048954 | 0.00085507 | 0.0025161 | 0.0054843 | 0.0035022 | 0.00090805 | 0.071825 | 0.073212 | 159935.097 | 190152.5741 | 11445.2525 | 218982.0206 | 181401.1956 | 167418.0196 | 3811.8422 | 0.034118 | 0.081733 | 0.17992 | 0.33656
114 | CCV 1ppb TRACE |  | 06/09/2023 01:06:59 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.468 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.14205 | 0.060655 | 0.80296 | 0.0068527 | 0.013599 | 0.029552 | 0.063322 | 0.026627 | 0.098884 | 0.14052 | 0.040026 | 0.030681 | 0.075844 | 0.13368 | 0.29363 | 0.081477 | 0.19983 | 0.081035 | 0.036702 | 0.038833 | 0.009092 | 1.7749 | 0.036298 | 0.0048535 | 0.00075473 | 0.0023781 | 0.0052932 | 0.0038083 | 0.0010332 | 0.071001 | 0.073217 | 155123.7462 | 193065.505 | 11474.609 | 220523.6063 | 173170.0192 | 169710.2093 | 3948.546 | 0.034022 | 0.081663 | 0.18238 | 0.35433
115 | LFB |  | 06/09/2023 01:12:47 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.469 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.3777 | 0.56768 | 4.2129 | 0.037883 | 0.20061 | 0.23184 | 0.64253 | 0.26272 | 0.96971 | 1.3659 | 0.39486 | 0.26349 | 0.74573 | 1.2025 | 2.8937 | 0.77441 | 1.8534 | 0.76736 | 0.36339 | 0.37569 | 0.093527 | 17.3475 | 0.36256 | 0.038265 | 0.0078425 | 0.02508 | 0.05203 | 0.036195 | 0.0096376 | 0.70276 | 0.7067 | 152840.9525 | 180707.3417 | 10693.3349 | 203713.53 | 167346.5458 | 161909.3781 | 3777.833 | 0.33718 | 0.79318 | 1.6541 | 3.1583
116 | CCV 10 ppb MAJORS |  | 06/09/2023 01:18:35 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.470 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.3902 | 0.56476 | 4.1693 | 0.035223 | 0.16485 | 0.22347 | 0.6534 | 0.26202 | 1.0101 | 1.3737 | 0.40316 | 0.2709 | 0.76032 | 1.2196 | 2.932 | 0.79848 | 1.8933 | 0.77888 | 0.36494 | 0.38394 | 0.096787 | 17.514 | 0.36724 | 0.040245 | 0.0079646 | 0.024488 | 0.051673 | 0.035942 | 0.0094278 | 0.72407 | 0.72213 | 152522.6461 | 184431.6044 | 10843.4475 | 206672.4951 | 170578.3672 | 166636.2405 | 3867.857 | 0.33931 | 0.79788 | 1.7124 | 3.2338
117 | CCV 10 ppb MAJORS |  | 06/09/2023 01:24:53 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.471 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.4141 | 0.57696 | 4.216 | 0.033144 | 0.17035 | 0.22868 | 0.65042 | 0.25612 | 0.99422 | 1.3887 | 0.40753 | 0.27788 | 0.76619 | 1.19 | 2.9441 | 0.80904 | 1.8909 | 0.78823 | 0.36855 | 0.38877 | 0.088143 | 17.2204 | 0.3688 | 0.040232 | 0.0081276 | 0.024999 | 0.053611 | 0.037084 | 0.0097065 | 0.72915 | 0.73657 | 151829.2381 | 182936.5618 | 11046.9374 | 207304.6763 | 170545.4041 | 161884.7959 | 3917.8705 | 0.34161 | 0.80704 | 1.7201 | 3.2005
118 | 8_16_21_W_6_DUP |  | 06/09/2023 01:30:41 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_W_6_DUP.472 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.2488 | 0.074143 | 4.1972 | 0.037567 | 0.046444 | 0.080883 | 0.022651 | 0.17863 | 0.01574 | 0.0091299 | 0.0025774 | 0.29221 | 0.32876 | 1.3424 | 0.0027998 | 0.0040099 | 0.20942 | 0.087037 | 0.027818 | 0.047343 | 0.010799 | 2.2416 | 0.045776 | 0.0013331 | -1.4781e-05 | 6.0567e-05 | 0.00012241 | 0.00030819 | 2.1029e-05 | 0.13335 | 0.13608 | 157186.0673 | 191150.6317 | 11651.4181 | 215234.9897 | 178146.2796 | 168181.2275 | 3904.534 | 0.00015343 | 0.00044812 | 0.015961 | 0.028888
119 | 8_16_21_W_8 |  | 06/09/2023 01:36:29 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_W_8.473 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.25009 | 0.072739 | 4.2898 | 0.034468 | 0.078351 | 0.063215 | 0.024155 | 0.16228 | 0.019969 | 0.0025951 | 0.00099149 | 0.28281 | 0.43675 | 1.2274 | 0.0027788 | 0.0034472 | 0.20005 | 0.082795 | 0.025446 | 0.049363 | 0.0081966 | 2.2084 | 0.04614 | 0.0013621 | 1.7297e-05 | 8.0262e-05 | 0.00012526 | 0.00064356 | 1.7781e-05 | 0.10944 | 0.10976 | 153664.0594 | 186870.3252 | 11763.5155 | 210412.3552 | 174385.4129 | 167083.5198 | 3907.2013 | 0.00016614 | 0.00040837 | 0.014382 | 0.025916
120 | 8_16_21_W_9 |  | 06/09/2023 01:42:18 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_W_9.474 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.22215 | 0.074361 | 4.6273 | 0.037242 | 0.064329 | 0.067487 | 0.021912 | 0.16192 | 0.018691 | 0.0050325 | 0.0015282 | 0.34283 | 0.35711 | 1.3559 | 0.0022165 | 0.0037685 | 0.16228 | 0.080003 | 0.021852 | 0.040907 | 0.0090534 | 1.7565 | 0.036915 | 0.0011967 | 1.3116e-05 | 4.3961e-05 | 0.0001125 | 0.00050091 | 1.7648e-05 | 0.087723 | 0.074861 | 152249.6904 | 189506.9578 | 11139.6759 | 185701.4303 | 174261.0732 | 166642.354 | 3981.2218 | 0.00017737 | 0.00031426 | 0.0122 | 0.022446
121 | 8_16_21_W_9_DUP |  | 06/09/2023 01:48:06 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_W_9_DUP.475 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.20732 | 0.069652 | 5.0786 | 0.042917 | 0.062037 | 0.060495 | 0.020958 | 0.14748 | 0.017302 | 0.0041668 | 0.0009717 | 0.32148 | 0.39179 | 1.4524 | 0.0031677 | 0.0038407 | 0.16634 | 0.06889 | 0.020064 | 0.040265 | 0.0078757 | 1.794 | 0.036409 | 0.0011092 | -1.8296e-05 | 7.3675e-05 | 0.00010409 | 8.0585e-05 | 2.51e-05 | 0.072905 | 0.072663 | 151239.4751 | 186157.7555 | 11336.4975 | 211220.3146 | 171844.9131 | 165230.0716 | 3779.1669 | 0.00010271 | 0.00029913 | 0.013493 | 0.024676
122 | 8_16_21_W_10 |  | 06/09/2023 01:53:55 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_W_10.476 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.20039 | 0.102 | 5.6175 | 0.047955 | 0.059271 | 0.070807 | 0.025373 | 0.18438 | 0.018123 | 0.013676 | 0.0039986 | 0.29833 | 0.25412 | 1.4027 | 0.0016459 | 0.0034886 | 0.12417 | 0.054005 | 0.019191 | 0.041697 | 0.0074859 | 1.7279 | 0.037405 | 0.0017445 | 1.0532e-05 | 2.7101e-05 | 0.00016923 | 0.00050529 | 2.1651e-05 | 0.034581 | 0.034848 | 151226.2238 | 185767.4561 | 10973.5464 | 209357.7602 | 172670.002 | 165164.7253 | 4021.2327 | 0.00012567 | 0.00023219 | 0.012981 | 0.023812
123 | 8_16_21_W_11 |  | 06/09/2023 01:59:43 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_W_11.477 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.19106 | 0.10315 | 2.8675 | 0.024662 | 0.045541 | 0.074899 | 0.023664 | 0.1844 | 0.016285 | 0.0018283 | 0.00061213 | 0.13952 | 0.12636 | 0.6402 | 0.00085168 | 0.0028144 | 0.10627 | 0.044518 | 0.013352 | 0.041266 | 0.0072453 | 1.7646 | 0.03696 | 0.0012211 | -1.7065e-07 | 3.4713e-05 | 0.00014958 | 0.00059931 | 2.5468e-05 | 0.029966 | 0.030013 | 151125.9096 | 183538.3254 | 10980.2183 | 209996.9283 | 164320.9463 | 165821.031 | 3845.1846 | 6.1857e-05 | 0.00015826 | 0.0075906 | 0.014973
124 | 5_10_23_DB |  | 06/09/2023 02:05:32 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_10_23_DB.478 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | -4.7652e-05 | 0 | 0.16493 | 0.0014294 | 0.069251 | 0.0047218 | -0.00037201 | -4.7652e-05 | -4.7652e-05 | 0.00086052 | 0.00015446 | 0.0017924 | 8.1296e-05 | 0.0060798 | -0.00012429 | 0.00041978 | 0.0040056 | 0.00195 | -0.00017076 | 0.0010502 | 8.9269e-06 | -0.0034042 | -6.5e-05 | 0.0010085 | -1.4795e-05 | -1.704e-06 | 3.6101e-06 | -0.00051639 | 7.0212e-06 | 3.3734e-05 | -2.2739e-06 | 153193.6997 | 189671.5077 | 11423.9047 | 216446.2047 | 173479.6869 | 165689.4295 | 4045.9064 | -4.8247e-06 | 1.9633e-06 | 0.00025147 | 0.00041542
125 | 5_10_23_MFB |  | 06/09/2023 02:11:19 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_10_23_MFB.479 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.65456 | 0.27497 | 2.2657 | 0.018874 | 0.17306 | 0.11052 | 0.30014 | 0.12411 | 0.47151 | 0.66275 | 0.2035 | 0.1469 | 0.37652 | 0.60962 | 1.3889 | 0.37632 | 0.90566 | 0.3982 | 0.18584 | 0.18732 | 0.042217 | 8.6216 | 0.17884 | 0.020451 | 0.0034665 | 0.012118 | 0.025656 | 0.017618 | 0.0046445 | 0.3572 | 0.34893 | 151540.2265 | 182216.8451 | 11169.0312 | 201993.8553 | 169098.5126 | 163205.0453 | 3781.8342 | 0.1737 | 0.4118 | 0.87874 | 1.6547
126 | 5_10_23_T-239 |  | 06/09/2023 02:17:07 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_10_23_T-239.480 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 6.4439 | 0.70262 | 56.5296 | 0.37289 | 0.22441 | 0.022894 | 0.2397 | 1.2978 | 0.18678 | 0.1886 | 0.055501 | 1.0786 | 0.27166 | 5.0648 | 0.40572 | 0.13208 | 0.13679 | 0.061686 | 0.21241 | 0.10237 | 0.019589 | 3.9861 | 0.086351 | 0.021419 | 0.0015896 | 0.0044987 | 0.010431 | 0.0089192 | 0.0019566 | 0.075276 | 0.07651 | 134950.3434 | 166261.9233 | 10125.5876 | 192095.5733 | 157767.1281 | 148297.0806 | 3601.7874 | 0.11358 | 0.26832 | 0.23821 | 0.39875
127 | 5_10_23_DB |  | 06/09/2023 02:22:55 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_10_23_DB.481 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.00035216 | 5.981e-05 | 0.30747 | 0.0023039 | 0.082712 | 0.0084461 | 0.00014837 | 6.8063e-05 | 0.00018054 | 0.0023184 | 0.00063802 | 0.0054241 | 0.00011156 | 0.031768 | -0.00017845 | -0.00024949 | 0.0062554 | 0.0029513 | -0.00013605 | 0.0011213 | -4.7652e-05 | 0.00013388 | 8.3259e-06 | 0.0010898 | -1.4939e-05 | -5.2274e-06 | 3.2252e-06 | -0.00011796 | 6.7988e-06 | 0.00017192 | 3.2398e-06 | 165083.9809 | 195350.1009 | 11618.7269 | 227238.3974 | 180133.9764 | 169696.6886 | 4140.6005 | 1.0691e-06 | 4.0316e-07 | 0.00035693 | 0.00085836
128 | LBLANK |  | 06/09/2023 02:28:44 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.482 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 6.8569e-05 | 0 | 0.0058526 | 1.128e-05 | 0.01258 | 0.0079936 | -0.00013733 | 6.7656e-05 | 1.0366e-05 | 0.00068185 | 6.2301e-05 | -0.00059003 | 1.7171e-05 | 0.0029445 | -0.00023826 | -6.4565e-05 | 6.8561e-05 | 0.00011742 | 2.826e-05 | 0.0015314 | -4.7652e-05 | -0.002098 | -3.755e-05 | 0.0011976 | -1.1373e-05 | 1.7356e-06 | 2.0816e-05 | -4.3698e-05 | 0 | 0.00010675 | 4.6007e-05 | 159508.4482 | 191692.2178 | 11479.9447 | 233592.4566 | 181720.6605 | 170107.4668 | 4047.2406 | -4.8247e-06 | 5.4204e-06 | -5.4455e-05 | -0.000109
129 | CCV 1ppb TRACE |  | 06/09/2023 02:34:32 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.483 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.14089 | 0.05736 | 0.82474 | 0.007413 | 0.01551 | 0.029691 | 0.057979 | 0.023189 | 0.10276 | 0.14463 | 0.040464 | 0.029594 | 0.074246 | 0.12828 | 0.29053 | 0.081694 | 0.1973 | 0.08334 | 0.037128 | 0.040906 | 0.0092166 | 1.6664 | 0.036688 | 0.0052785 | 0.00067507 | 0.0024279 | 0.0051914 | 0.0041086 | 0.00099444 | 0.072317 | 0.070035 | 151904.7743 | 192980.2487 | 11530.6529 | 216206.6737 | 177311.6285 | 166390.0633 | 4249.9657 | 0.034455 | 0.082583 | 0.1907 | 0.33851
130 | CCV 1ppb TRACE |  | 06/09/2023 02:40:49 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.484 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.13936 | 0.055554 | 0.77346 | 0.0064328 | 0.013155 | 0.026902 | 0.061822 | 0.024515 | 0.10195 | 0.13645 | 0.040146 | 0.029986 | 0.076971 | 0.13441 | 0.29337 | 0.080965 | 0.1994 | 0.082547 | 0.037723 | 0.03918 | 0.0078677 | 1.7185 | 0.037295 | 0.005406 | 0.00084093 | 0.0023923 | 0.0055177 | 0.0036943 | 0.00098248 | 0.072098 | 0.074403 | 157277.7462 | 188288.4277 | 11631.4006 | 218625.1381 | 176456.4382 | 168787.9704 | 4091.2535 | 0.034652 | 0.082063 | 0.17971 | 0.3345
131 | CCV 10 ppb MAJORS |  | 06/09/2023 02:46:38 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.485 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.3705 | 0.55695 | 4.0028 | 0.034399 | 0.16408 | 0.22637 | 0.63662 | 0.25685 | 1.0035 | 1.3749 | 0.3979 | 0.26708 | 0.76516 | 1.1837 | 2.9154 | 0.80165 | 1.8803 | 0.76603 | 0.36439 | 0.38101 | 0.093884 | 17.3361 | 0.37024 | 0.040552 | 0.0080332 | 0.02467 | 0.053596 | 0.035814 | 0.0096142 | 0.72146 | 0.73463 | 152080.6684 | 176782.485 | 10836.7791 | 204328.9365 | 165990.1326 | 160080.0665 | 3778.5001 | 0.33664 | 0.80536 | 1.7207 | 3.2548
132 | CCV 10 ppb MAJORS |  | 06/09/2023 02:52:55 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.486 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.3592 | 0.55435 | 4.146 | 0.033999 | 0.17146 | 0.22475 | 0.6278 | 0.26508 | 0.98409 | 1.3592 | 0.40529 | 0.26992 | 0.74835 | 1.2041 | 2.9008 | 0.78455 | 1.8773 | 0.78391 | 0.36508 | 0.38703 | 0.096078 | 17.6144 | 0.36402 | 0.04135 | 0.0078471 | 0.023889 | 0.05081 | 0.03763 | 0.0091459 | 0.71916 | 0.72571 | 151154.3482 | 182573.949 | 10828.7695 | 204679.2284 | 169119.1861 | 154172.1581 | 3773.8319 | 0.33773 | 0.79561 | 1.7297 | 3.1961
133 | MT 34 D |  | 06/09/2023 02:58:43 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\MT 34 D.487 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.00035165 | 0 | 0.43438 | 0.0041576 | 0.10261 | 0.00982 | -0.00023276 | 0.00034035 | 0.00023519 | 0.0017108 | 0.00052861 | 0.0048786 | 0.0001752 | 0.035045 | 0.00020897 | 0.0024218 | 0.0030312 | 0.0014507 | 0.0022378 | 0.00073221 | -4.7652e-05 | 0.00055553 | 2.3345e-05 | 0.0012528 | -1.8296e-05 | 1.4627e-06 | -5.0147e-07 | 0.0001252 | 0 | 0.001733 | 0.0015743 | 161493.0446 | 199449.4328 | 11861.591 | 230691.2597 | 187731.6967 | 174468.5825 | 4384.6729 | -1.899e-06 | 8.1139e-05 | 0.0047812 | 0.0087542
134 | MT 35 D |  | 06/09/2023 03:04:31 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\MT 35 D.488 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 6.2987e-05 | 0 | 0.52211 | 0.0040818 | 0.11605 | 0.0039063 | -0.00046618 | 0.0001717 | 8.1496e-06 | 0.0011178 | 0.00033926 | 0.0020507 | 0.00016291 | 0.012137 | -0.00018367 | 0.0011136 | 0.0029405 | 0.0014581 | 0.0015556 | 0.00024993 | 8.1496e-06 | -0.00074683 | -8.1752e-06 | 0.00082185 | -1.4887e-05 | -5.2274e-06 | 3.2062e-06 | -0.00010817 | 1.371e-05 | 0.00032428 | 5.1093e-05 | 161450.6199 | 195089.748 | 12105.127 | 224045.4971 | 182896.0398 | 169988.8256 | 4179.9452 | -1.798e-06 | 6.9131e-05 | 0.0049051 | 0.0088351
135 | MT 36 D |  | 06/09/2023 03:10:18 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\MT 36 D.489 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.00023492 | 0 | 0.41421 | 0.0025045 | 0.11671 | 0.0093331 | -4.4604e-05 | 0.00018088 | -4.7652e-05 | 0.0025301 | 0.00054254 | 0.0017894 | 0.00019279 | 0.010679 | -0.00023826 | 0.00031407 | 0.0033413 | 0.0016012 | 0.0023546 | 0.0015907 | -4.7652e-05 | 5.1723e-05 | 5.1484e-06 | 0.00064807 | -1.4891e-05 | -1.8287e-06 | 9.9619e-06 | 0.00051825 | 3.3987e-06 | 0.00034708 | 0.00017684 | 161483.6097 | 195860.1693 | 11742.1582 | 222677.5172 | 183453.1857 | 169181.3664 | 4063.245 | 7.128e-06 | 4.3028e-05 | 0.0046381 | 0.0087048
136 | MT 37 D |  | 06/09/2023 03:16:06 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\MT 37 D.490 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.00057056 | 5.7928e-05 | 0.49488 | 0.0035638 | 0.12183 | 0.0050461 | -0.00034308 | 0.00018107 | -4.7652e-05 | 0.0012653 | 0.00031687 | 0.00088094 | 0.00015483 | 0.0070054 | -1.056e-05 | 0.00036815 | 0.002537 | 0.0012842 | 0.0021575 | 0.0016912 | -4.7652e-05 | -0.0022188 | -3.8038e-05 | 0.00098543 | -1.8296e-05 | 1.6767e-06 | 1.7333e-05 | -8.8846e-05 | 0 | 0.0003066 | 0.00012466 | 159608.0099 | 193151.7987 | 11854.9192 | 222976.4641 | 186045.171 | 168698.5418 | 4221.9573 | 7.0842e-06 | 3.9986e-05 | 0.0047287 | 0.0087715
137 | MT 38 D |  | 06/09/2023 03:21:55 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\MT 38 D.491 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.002175 | 5.7221e-05 | 0.65777 | 0.0067162 | 0.13231 | 0.0041429 | 0.00039162 | 0.00030436 | 6.6789e-05 | 0.0016176 | 0.00047848 | 0.099003 | 0.0016696 | 0.48048 | 0.0001679 | 0.0013936 | 0.0184 | 0.008182 | 0.015319 | 0.0011815 | 0.00012729 | -0.00061703 | -6.7842e-06 | 0.0010814 | -1.8296e-05 | 1.5626e-05 | 1.0297e-05 | 0.00029755 | 0 | 0.00030399 | 0.00019397 | 159955.2136 | 192834.127 | 11434.5754 | 220888.7015 | 180703.1438 | 161098.7126 | 4053.9085 | 7.2503e-06 | 4.6539e-05 | 0.0048374 | 0.0090261
138 | MT 39 D |  | 06/09/2023 03:27:43 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\MT 39 D.492 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.00023026 | 0 | 0.45143 | 0.0035774 | 0.12326 | 0.0055206 | -0.00018149 | 0.00011958 | 0.00017531 | 0.0019771 | 0.00053298 | 0.027409 | 0.00051539 | 0.10113 | -0.00018171 | 0.00018997 | 0.0044426 | 0.0020734 | 0.0020893 | 0.00031442 | -4.7652e-05 | -0.0023025 | -4.0878e-05 | 0.0012743 | -1.8296e-05 | 1.632e-06 | -3.5954e-06 | -3.6915e-05 | 0 | 0.0010352 | 0.00089996 | 157944.3707 | 190602.0175 | 11961.0059 | 211904.2564 | 181300.54 | 166649.7679 | 4106.5902 | 2.1676e-05 | 8.7238e-05 | 0.0048017 | 0.0089782
139 | MT 40 D |  | 06/09/2023 03:33:31 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\MT 40 D.493 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.00012361 | 0 | 0.47387 | 0.0039569 | 0.12989 | 0.0079156 | -0.00033386 | 0.00012374 | -4.7652e-05 | 0.0011508 | 0.00033049 | 0.0022148 | 0.00029137 | 0.016256 | 0.00026429 | 0.00058253 | 0.0038921 | 0.0016729 | 0.0020632 | 8.5876e-06 | -4.7652e-05 | -0.0014975 | -2.4869e-05 | 0.00091796 | -1.8296e-05 | -1.7613e-06 | 6.7761e-06 | -2.7843e-05 | 0 | 0.00018785 | 0.00017908 | 161051.3949 | 194074.1994 | 11839.5726 | 222378.1882 | 174125.4025 | 169613.7934 | 4106.5905 | 1.3186e-05 | 6.699e-05 | 0.0051353 | 0.0093858
140 | MT 32 D |  | 06/09/2023 03:39:20 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\MT 32 D.494 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.00024797 | 0 | 0.62673 | 0.0055131 | 0.11974 | 0.0069734 | -0.00029567 | 0.00024707 | 6.879e-05 | 0.0015328 | 0.00036164 | 0.0028378 | 0.00062936 | 0.01868 | -0.00023826 | 0.00031535 | 0.0095991 | 0.003928 | 0.0023288 | 0.001945 | 1.0569e-05 | 0.0005962 | 2.1294e-05 | 0.00071006 | -1.1304e-05 | 1.765e-06 | 7.1466e-06 | -4.0105e-05 | 0 | 0.00018714 | 0.00017434 | 160923.418 | 192090.4047 | 11185.7198 | 223452.8091 | 183556.4562 | 170866.1154 | 4131.931 | 1.1133e-06 | 5.4653e-05 | 0.0067963 | 0.012964
141 | MT 32 D LD |  | 06/09/2023 03:45:07 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\MT 32 D LD.495 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.00024223 | 0 | 0.45827 | 0.0036046 | 0.10523 | 0.0068872 | 3.6591e-05 | 6.8384e-05 | -4.7652e-05 | 0.0021906 | 0.00046954 | 0.0016403 | 0.00019106 | 0.0091845 | -0.00018024 | 0.00033927 | 0.0076646 | 0.0031806 | 0.0030898 | 0.00010064 | -4.7652e-05 | -0.0011388 | -1.5555e-05 | 0.00094499 | -1.4692e-05 | -1.7319e-06 | 3.6357e-06 | -4.9754e-05 | 0 | 0.00022864 | 0.00019565 | 157008.2093 | 188898.9592 | 11499.9602 | 226528.6723 | 170615.265 | 192084.2645 | 4091.2525 | -4.8247e-06 | 8.602e-05 | 0.0065553 | 0.013013
142 | MT 32 D LD |  | 06/09/2023 03:51:24 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\MT 32 D LD.496 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.00018574 | 5.8786e-05 | 0.44771 | 0.0039252 | 0.12538 | 0.010396 | 4.4327e-05 | 6.8658e-05 | 1.0457e-05 | 0.0018573 | 0.00045555 | 0.0014841 | 0.00017173 | 0.0078302 | -0.00023826 | 0.00040359 | 0.007014 | 0.003134 | 0.0030945 | 0.0016068 | -4.7652e-05 | -0.00065215 | -6.168e-06 | 0.00073399 | -1.1237e-05 | -1.7378e-06 | 1.5167e-05 | 0.00011403 | 1.3998e-05 | 0.00021259 | 0.00018834 | 156307.1775 | 185459.4917 | 11422.5649 | 219152.9643 | 176112.24 | 162841.207 | 3962.5518 | 1.3529e-05 | 5.6003e-05 | 0.0066501 | 0.012447
143 | LBLANK |  | 06/09/2023 03:57:12 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.497 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.00012052 | 0 | 0.0046488 | 0.00012105 | 0.012533 | 0.0050578 | 0.00011389 | 8.583e-06 | 8.583e-06 | 0.00031061 | 8.5685e-05 | -0.00045767 | 2.3672e-05 | -0.00046762 | -0.00023826 | 0.00013429 | 0.0011092 | 0.0001696 | -1.2235e-06 | 0.0021867 | -4.7652e-05 | -0.0020489 | -3.7291e-05 | 0.00082267 | -1.1114e-05 | 5.2816e-06 | 5.4187e-06 | -4.4855e-05 | 0 | 5.9937e-05 | -2.4036e-06 | 157064.7519 | 173960.1467 | 11847.5779 | 227799.8033 | 180086.8477 | 168707.9526 | 3641.1344 | 3.4796e-06 | 2.8293e-06 | -5.598e-05 | -9.6131e-05
144 | CCV 1ppb TRACE |  | 06/09/2023 04:03:01 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.498 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.13878 | 0.056049 | 0.75808 | 0.0056043 | 0.016828 | 0.028495 | 0.059755 | 0.024457 | 0.1026 | 0.13581 | 0.039745 | 0.029673 | 0.075189 | 0.13105 | 0.29166 | 0.084535 | 0.1929 | 0.083299 | 0.037322 | 0.039359 | 0.008621 | 1.6622 | 0.036762 | 0.0051149 | 0.00089622 | 0.0024726 | 0.0052291 | 0.003528 | 0.0010234 | 0.073098 | 0.072342 | 154572.681 | 185131.2128 | 11311.8107 | 212380.4019 | 172943.2455 | 164252.2592 | 4095.2537 | 0.034205 | 0.082886 | 0.1828 | 0.34264
145 | CCV 1ppb TRACE |  | 06/09/2023 04:09:18 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.499 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.12766 | 0.053503 | 0.73463 | 0.0073688 | 0.013719 | 0.022789 | 0.056998 | 0.023095 | 0.098204 | 0.13215 | 0.040306 | 0.030373 | 0.075507 | 0.1252 | 0.27893 | 0.076931 | 0.19194 | 0.082242 | 0.037059 | 0.038431 | 0.0081283 | 1.6534 | 0.036568 | 0.0047063 | 0.00083275 | 0.0025655 | 0.0054514 | 0.0038086 | 0.0008363 | 0.071573 | 0.073568 | 159567.9868 | 184953.5998 | 11684.111 | 211545.504 | 174184.1932 | 162126.9152 | 4091.2528 | 0.035183 | 0.081736 | 0.1757 | 0.3427
146 | LFB |  | 06/09/2023 04:15:06 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.500 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.2778 | 0.52583 | 3.9478 | 0.034323 | 0.19856 | 0.20757 | 0.58348 | 0.24649 | 0.92207 | 1.3042 | 0.39427 | 0.26394 | 0.74285 | 1.1352 | 2.7436 | 0.7623 | 1.8267 | 0.75862 | 0.36188 | 0.3764 | 0.08757 | 17.4388 | 0.36885 | 0.038296 | 0.0080693 | 0.025003 | 0.053766 | 0.035749 | 0.0093444 | 0.70161 | 0.7195 | 149287.4518 | 171337.0217 | 10671.3178 | 198549.629 | 162675.7784 | 153381.075 | 3623.793 | 0.33627 | 0.7911 | 1.6767 | 3.162
147 | LFB |  | 06/09/2023 04:21:22 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.501 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.278 | 0.5175 | 4.1229 | 0.035649 | 0.19648 | 0.20592 | 0.57508 | 0.24528 | 0.94091 | 1.2913 | 0.3921 | 0.2623 | 0.74196 | 1.1486 | 2.7825 | 0.75945 | 1.79 | 0.75847 | 0.36519 | 0.38602 | 0.091686 | 16.9469 | 0.36376 | 0.040343 | 0.0080114 | 0.0249 | 0.053213 | 0.037128 | 0.009614 | 0.70284 | 0.70917 | 149086.14 | 176177.9428 | 10772.0605 | 201436.2583 | 163549.5724 | 159836.5468 | 3783.1679 | 0.33422 | 0.79536 | 1.6879 | 3.1505
148 | CCV 10 ppb MAJORS |  | 06/09/2023 04:27:10 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.502 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.3226 | 0.5366 | 3.9043 | 0.033791 | 0.16214 | 0.21604 | 0.60735 | 0.2513 | 0.97705 | 1.3835 | 0.41493 | 0.28035 | 0.75995 | 1.1727 | 2.8491 | 0.79053 | 1.8253 | 0.7957 | 0.3708 | 0.37592 | 0.091262 | 17.2114 | 0.36617 | 0.040006 | 0.0078179 | 0.025022 | 0.052443 | 0.034603 | 0.0095189 | 0.74298 | 0.73351 | 154462.7777 | 176586.1935 | 10881.4764 | 200509.839 | 164486.2834 | 158628.5185 | 3757.8278 | 0.34198 | 0.81106 | 1.692 | 3.2251
149 | CCV 10 ppb MAJORS |  | 06/09/2023 04:33:27 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.503 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.3471 | 0.5577 | 4.0454 | 0.034928 | 0.16303 | 0.22064 | 0.61827 | 0.26221 | 1.0174 | 1.3767 | 0.37857 | 0.25316 | 0.75709 | 1.2188 | 2.9089 | 0.79092 | 1.8606 | 0.73439 | 0.34231 | 0.38409 | 0.09017 | 17.3725 | 0.36385 | 0.041466 | 0.0076442 | 0.024497 | 0.05175 | 0.035769 | 0.0094701 | 0.68563 | 0.72655 | 152288.8117 | 178980.9075 | 10780.7346 | 217471.477 | 169664.2779 | 158755.9251 | 3749.8256 | 0.31515 | 0.75695 | 1.718 | 3.1767
150 | MT 33 D |  | 06/09/2023 04:39:15 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\MT 33 D.504 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.00053379 | 0 | 0.46976 | 0.0041468 | 0.1047 | 0.0064885 | 8.7224e-05 | 1.0034e-05 | 0.00012214 | 0.0014241 | 0.00035133 | 0.0015712 | 0.00026525 | 0.011797 | -1.078e-05 | 0.00034058 | 0.0086673 | 0.0033766 | 0.0015154 | 0.0014853 | -4.7652e-05 | 0.0021663 | 5.4943e-05 | 0.0018717 | -1.8296e-05 | -1.8096e-06 | 3.4008e-06 | 0.00027525 | 6.9275e-06 | 0.0012913 | 0.00028329 | 157169.3217 | 193333.4825 | 11542.6627 | 223903.7031 | 179278.6031 | 169033.6076 | 4152.6036 | 2.1945e-05 | 9.3126e-05 | 0.0055009 | 0.0093575
151 | MT 33 D LSP |  | 06/09/2023 04:45:03 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\MT 33 D LSP.505 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.4437 | 1.3545 | 11.3402 | 0.093732 | 0.51042 | 0.53416 | 1.6077 | 0.65559 | 2.5374 | 3.5014 | 1.0449 | 0.70577 | 1.9544 | 2.977 | 7.2476 | 1.917 | 4.555 | 1.9312 | 0.88239 | 0.98753 | 0.20915 | 42.584 | 0.89534 | 0.09847 | 0.01894 | 0.059925 | 0.12795 | 0.091917 | 0.022643 | 1.8524 | 1.8768 | 140417.4447 | 169745.2168 | 10445.1523 | 193889.6775 | 160903.1263 | 160461.6438 | 3572.447 | 0.82374 | 1.938 | 4.4098 | 7.7519
152 | cal blank |  | 06/09/2023 04:50:51 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\cal blank.506 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.00024563 | 0 | -0.0028266 | -4.7652e-05 | 0.0082711 | 0.0066371 | -0.00036701 | -4.7652e-05 | 6.7422e-05 | -7.5986e-06 | -1.2813e-05 | -0.00066978 | 2.2965e-05 | 0.0049808 | -0.0001211 | 0.00040483 | -0.00045033 | 0.00011049 | -0.00018979 | 0.00024155 | 1.1976e-05 | 0.0048043 | 0.00010493 | 0.0013431 | -1.1386e-05 | 1.8701e-05 | -3.6482e-06 | 0.00015587 | 6.8591e-06 | 0.0024762 | 0.00026342 | 156570.9012 | 194908.0099 | 11419.8971 | 218492.9724 | 175194.5306 | 167602.1307 | 4056.576 | 1.0375e-05 | 1.6772e-05 | 4.3101e-06 | 3.0872e-05
153 | STD 0.01 |  | 06/09/2023 04:56:40 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 0.01.507 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.0016656 | 0.0010057 | 0.22942 | 0.0022596 | 0.016678 | 0.0043879 | 0.00023898 | 0.00024829 | 0.00072233 | 0.0024735 | 0.00040699 | 0.0031717 | 0.0027661 | 0.021879 | 0.0036046 | 0.00071375 | 0.0038668 | 0.001523 | 0.0050274 | 0.0013094 | -4.7652e-05 | 0.021098 | 0.00046677 | 0.0015994 | -1.1309e-05 | 2.6227e-05 | 4.2207e-05 | 0.0001818 | 1.0437e-05 | 0.0028815 | 0.0017864 | 157612.2063 | 190424.8422 | 11286.4569 | 217771.4967 | 177936.0379 | 163157.2009 | 4127.93 | 0.00027361 | 0.00076613 | 0.0019779 | 0.0037241
154 | std 0.05 |  | 06/09/2023 05:02:29 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\std 0.05.508 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.0073787 | 0.0037709 | 0.25562 | 0.0028374 | 0.027433 | 0.0051515 | 0.0032433 | 0.0017814 | 0.0050884 | 0.0083675 | 0.0021607 | 0.0037328 | 0.003911 | 0.017329 | 0.017087 | 0.0047741 | 0.012758 | 0.0049926 | 0.0037503 | 0.0037389 | 0.00059929 | 0.08864 | 0.0018667 | 0.0012782 | 2.0272e-05 | 0.00014593 | 0.00032982 | 0.00037614 | 4.5887e-05 | 0.0044111 | 0.0036665 | 156044.4225 | 190232.9325 | 11310.4761 | 215163.7124 | 173056.9367 | 167774.3412 | 3991.2249 | 0.0018265 | 0.0040655 | 0.0090133 | 0.017727
155 | std 0.1 |  | 06/09/2023 05:08:17 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\std 0.1.509 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.013749 | 0.0053965 | 0.12392 | 0.00081139 | 0.01282 | 0.0096339 | 0.005493 | 0.0027692 | 0.010641 | 0.013511 | 0.0039284 | 0.0078639 | 0.007888 | 0.045043 | 0.030126 | 0.0083845 | 0.023996 | 0.0089336 | 0.0067369 | 0.005452 | 0.0010437 | 0.16323 | 0.0035837 | 0.0020049 | 7.5784e-05 | 0.00022783 | 0.00055274 | 0.00042616 | 9.3845e-05 | 0.0076561 | 0.0072785 | 157956.4651 | 191635.5126 | 11592.7024 | 219999.8112 | 178343.3103 | 168154.1191 | 4199.284 | 0.0033673 | 0.0079928 | 0.017819 | 0.032897
156 | std 0.5 |  | 06/09/2023 05:14:05 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\std 0.5.510 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.068619 | 0.029504 | 0.51797 | 0.0040841 | 0.033776 | 0.01888 | 0.03192 | 0.016156 | 0.048521 | 0.073126 | 0.020314 | 0.027301 | 0.038524 | 0.12947 | 0.15185 | 0.044862 | 0.10922 | 0.042949 | 0.024514 | 0.020197 | 0.0040731 | 0.85627 | 0.01841 | 0.0029888 | 0.00040364 | 0.0012758 | 0.0026007 | 0.0018603 | 0.00046405 | 0.036281 | 0.036448 | 157815.5821 | 191228.2808 | 11349.1761 | 220630.1822 | 176173.6669 | 167704.3018 | 4109.9245 | 0.017086 | 0.039588 | 0.090817 | 0.16541
157 | std 1 |  | 06/09/2023 05:19:54 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\std 1.511 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.13311 | 0.055043 | 0.79014 | 0.0065144 | 0.030014 | 0.023763 | 0.058732 | 0.026748 | 0.10157 | 0.13752 | 0.042259 | 0.04056 | 0.076752 | 0.18098 | 0.29419 | 0.085656 | 0.19457 | 0.085559 | 0.04268 | 0.037215 | 0.0091347 | 1.6821 | 0.03614 | 0.0046097 | 0.00075096 | 0.0023943 | 0.005447 | 0.0032806 | 0.00092272 | 0.075487 | 0.072383 | 156218.8549 | 187083.8969 | 11381.2006 | 206843.3169 | 181069.7969 | 165431.8387 | 4018.5652 | 0.035037 | 0.083321 | 0.17735 | 0.313
158 | std 5 |  | 06/09/2023 05:25:42 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\std 5.512 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.6518 | 0.2701 | 2.1918 | 0.018673 | 0.11772 | 0.1115 | 0.30909 | 0.12292 | 0.48651 | 0.67892 | 0.2145 | 0.14951 | 0.37417 | 0.61126 | 1.4121 | 0.39323 | 0.91409 | 0.42083 | 0.20054 | 0.18364 | 0.043308 | 8.6924 | 0.18153 | 0.019935 | 0.0038119 | 0.012319 | 0.026104 | 0.017289 | 0.0046474 | 0.38228 | 0.35349 | 157039.8582 | 181181.7057 | 10974.2139 | 192757.2274 | 167886.087 | 148275.1615 | 3783.8345 | 0.1813 | 0.43089 | 0.82879 | 1.6001
159 | std 10 |  | 06/09/2023 05:31:30 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\std 10.513 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.3554 | 0.55107 | 4.1593 | 0.033776 | 0.18359 | 0.21738 | 0.60224 | 0.25278 | 0.98405 | 1.3689 | 0.40382 | 0.27225 | 0.75045 | 1.2064 | 2.9112 | 0.79936 | 1.8614 | 0.78556 | 0.36532 | 0.38405 | 0.091241 | 16.7249 | 0.35993 | 0.041321 | 0.0077353 | 0.023718 | 0.051245 | 0.036236 | 0.0091691 | 0.72474 | 0.71275 | 148900.0915 | 177732.8159 | 10631.2883 | 201299.9614 | 165117.4928 | 154503.0031 | 3825.8458 | 0.33747 | 0.79844 | 1.7107 | 3.1578
160 | LBLANK |  | 06/09/2023 05:37:19 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.514 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 6.8792e-05 | 5.8797e-05 | 0.0060998 | 0.00012759 | 0.0014421 | 0.0077148 | -0.00041533 | -4.7652e-05 | 1.1552e-05 | 0.00045807 | 0.00012941 | -0.00025826 | 3.1838e-05 | 0.0049526 | 5.1868e-05 | 0.00016986 | 0.00019289 | 0.00016808 | -4.028e-05 | 0.0018354 | 7.0756e-05 | -2.9859e-06 | 8.6792e-06 | 0.001486 | -1.8296e-05 | -1.7438e-06 | 1.1058e-08 | 8.7261e-05 | 0 | 0.0011906 | 0.00017686 | 154786.778 | 190070.8361 | 11415.2266 | 219988.3727 | 169769.3334 | 165354.9203 | 4104.5897 | 1.2665e-06 | -1.683e-06 | -3.6545e-05 | -8.2576e-05
161 | CCV 1ppb TRACE |  | 06/09/2023 05:43:07 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.515 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.14652 | 0.060826 | 0.7752 | 0.0076739 | 0.014125 | 0.031665 | 0.06688 | 0.026955 | 0.11117 | 0.14994 | 0.040825 | 0.030429 | 0.075851 | 0.16786 | 0.32253 | 0.091144 | 0.21431 | 0.083281 | 0.036796 | 0.039978 | 0.01152 | 1.7069 | 0.037225 | 0.0050922 | 0.00073027 | 0.0025954 | 0.0051847 | 0.0036381 | 0.00094024 | 0.072825 | 0.07309 | 159196.888 | 187135.0609 | 10774.0736 | 216810.1785 | 174450.1546 | 167448.3915 | 4079.9159 | 0.034433 | 0.081549 | 0.17606 | 0.33289
162 | CCV 1ppb TRACE |  | 06/09/2023 05:49:25 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.516 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.13712 | 0.054099 | 0.77511 | 0.0070503 | 0.01243 | 0.027465 | 0.059781 | 0.023306 | 0.10473 | 0.14064 | 0.042129 | 0.031049 | 0.081126 | 0.14484 | 0.29085 | 0.076796 | 0.19424 | 0.085564 | 0.037908 | 0.03964 | 0.0086964 | 1.8319 | 0.038558 | 0.0050453 | 0.00081879 | 0.0026308 | 0.0055069 | 0.003925 | 0.00095042 | 0.074344 | 0.078502 | 155607.6932 | 180676.6122 | 11362.5172 | 208313.3726 | 172798.3227 | 165632.8745 | 3799.1723 | 0.034361 | 0.082997 | 0.17928 | 0.33621
163 | CCV 10 ppb MAJORS |  | 06/09/2023 05:55:13 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.517 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.3337 | 0.54248 | 3.9993 | 0.032943 | 0.16324 | 0.2192 | 0.61696 | 0.24972 | 0.99379 | 1.3685 | 0.40191 | 0.26713 | 0.75803 | 1.1928 | 2.86 | 0.76294 | 1.8388 | 0.77661 | 0.35832 | 0.38575 | 0.094559 | 17.3917 | 0.36576 | 0.041278 | 0.0072885 | 0.024411 | 0.051652 | 0.036112 | 0.0092004 | 0.71306 | 0.72645 | 147337.2497 | 175272.6682 | 10731.3629 | 200944.846 | 134163.6594 | 157002.2312 | 3690.4773 | 0.33348 | 0.79188 | 1.725 | 4.1907
164 | CCV 10 ppb MAJORS |  | 06/09/2023 06:01:30 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.518 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.3402 | 0.53904 | 4.1121 | 0.033009 | 0.16444 | 0.21321 | 0.60933 | 0.25157 | 0.96624 | 1.3447 | 0.40579 | 0.2682 | 0.76913 | 1.1518 | 2.83 | 0.77541 | 1.8118 | 0.78152 | 0.36477 | 0.38951 | 0.094872 | 17.3806 | 0.36913 | 0.040981 | 0.0078536 | 0.02508 | 0.052647 | 0.036087 | 0.0096532 | 0.72293 | 0.73685 | 147490.2418 | 174108.6037 | 10716.6913 | 199590.6324 | 164111.9679 | 155449.3013 | 3698.4788 | 0.33484 | 0.79082 | 1.713 | 3.1661
165 | std 25 |  | 06/09/2023 06:07:19 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\std 25.519 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.3583 | 1.315 | 10.3533 | 0.082332 | 0.45284 | 0.54158 | 1.5751 | 0.63908 | 2.4945 | 3.4091 | 1.0506 | 0.70675 | 1.9364 | 2.8872 | 7.2235 | 1.9116 | 4.492 | 1.9724 | 0.9404 | 1.0075 | 0.22926 | 45.3521 | 0.958 | 0.10187 | 0.021272 | 0.066323 | 0.13727 | 0.09586 | 0.024671 | 1.8784 | 1.8882 | 141982.8347 | 165340.0478 | 10245.0083 | 188267.1813 | 151611.6471 | 149413.9832 | 3492.4268 | 0.84501 | 2.0133 | 4.1882 | 7.7833
166 | std 50 |  | 06/09/2023 06:13:06 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\std 50.520 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 6.8279 | 2.6948 | 25.7645 | 0.17266 | 0.93515 | 1.1096 | 3.2089 | 1.3424 | 5.1882 | 7.044 | 2.1678 | 1.3352 | 4.2071 | 5.9216 | 14.5975 | 3.8545 | 9.013 | 3.8945 | 1.8699 | 2.0323 | 0.46173 | 98.8449 | 2.0618 | 0.20508 | 0.044393 | 0.14108 | 0.2963 | 0.19217 | 0.053577 | 3.836 | 4.0721 | 135239.8081 | 148310.2448 | 9762.6687 | 176054.3368 | 144712.6205 | 139015.1891 | 3093.6685 | 1.6876 | 3.9709 | 8.1348 | 15.0734
167 | std 50 |  | 06/09/2023 06:19:23 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\std 50.521 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 7.6028 | 3.0076 | 26.5071 | 0.19237 | 0.96665 | 1.2573 | 3.54 | 1.4711 | 5.7076 | 7.7796 | 2.1753 | 1.3335 | 3.9899 | 6.5145 | 15.838 | 4.2089 | 9.9208 | 3.9056 | 1.8731 | 2.0501 | 0.50412 | 91.9438 | 1.9321 | 0.20531 | 0.041646 | 0.13331 | 0.28208 | 0.19631 | 0.050838 | 3.8128 | 3.8716 | 134595.8068 | 156061.2002 | 9015.5131 | 176618.8651 | 143217.7041 | 138680.5929 | 3281.7104 | 1.6629 | 3.9531 | 8.1074 | 15.022
168 | std 100 |  | 06/09/2023 06:25:11 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\std 100.522 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 14.9185 | 5.8041 | 54.8178 | 0.36455 | 1.8804 | 2.4861 | 7.1932 | 2.9268 | 11.4587 | 15.3298 | 4.5555 | 2.7109 | 8.3147 | 12.4739 | 31.1385 | 8.1193 | 18.963 | 7.6956 | 3.7033 | 4.178 | 0.98367 | 187.6849 | 3.8928 | 0.41278 | 0.087979 | 0.27055 | 0.57475 | 0.39728 | 0.10353 | 7.9528 | 7.981 | 120715.3029 | 140686.0347 | 8557.8968 | 157305.5139 | 130202.6722 | 125507.0919 | 2918.2981 | 3.3412 | 7.8683 | 15.5713 | 29.1331
169 | std 100 |  | 06/09/2023 06:31:28 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\std 100.523 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 15.1445 | 5.9977 | 54.5914 | 0.3695 | 1.8951 | 2.569 | 7.3004 | 3.0272 | 11.9141 | 15.852 | 4.5539 | 2.7543 | 8.2669 | 12.9353 | 31.7749 | 8.1983 | 19.6306 | 7.766 | 3.6816 | 4.1204 | 1.0093 | 184.1457 | 3.9057 | 0.40669 | 0.086818 | 0.27045 | 0.57392 | 0.39511 | 0.10348 | 7.9937 | 7.9742 | 123629.387 | 142489.1334 | 8438.4922 | 158408.0955 | 130628.5452 | 125956.8827 | 3022.3197 | 3.3103 | 7.8057 | 15.4722 | 32.867
170 | LBLANK |  | 06/09/2023 06:37:16 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.524 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.00065304 | 0.00011963 | 0.0077622 | 0.00025106 | 0.0034886 | 0.01167 | 0.0002271 | -4.7652e-05 | 0.00031333 | 0.00059511 | 0.0001253 | -0.00018876 | 0.00010792 | 0.013546 | 0.0010768 | 0.0010212 | 0.00085293 | 0.0003004 | -1.749e-05 | 0.00028207 | 0.0004988 | 0.048678 | 0.0010642 | 0.0014256 | -7.6919e-06 | 2.3187e-05 | 5.0117e-05 | 2.8382e-05 | 0 | 0.0073997 | 0.0013601 | 157106.9357 | 187116.3508 | 11189.7223 | 214600.6064 | 174616.9113 | 166481.1169 | 4063.2446 | 4.4688e-05 | 0.00010161 | -3.4407e-05 | 0.00011033
171 | PE ICV |  | 06/09/2023 06:43:04 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\PE ICV.525 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.427 | 0.6027 | 4.2441 | 0.036896 | 0.34935 | 0.23369 | 0.66132 | 0.272 | 1.0257 | 1.4136 | 0.41574 | 0.28016 | 0.77857 | 1.2485 | 2.9644 | 0.82567 | 1.9338 | 0.77818 | 0.3568 | 0.38512 | 0.083938 | 16.9564 | 0.36502 | 0.030259 | 0.0078014 | 0.025041 | 0.052904 | 0.036735 | 0.0095 | 0.69846 | 0.69403 | 149682.131 | 175567.142 | 10213.6574 | 198430.7182 | 164337.1699 | 156597.1076 | 3779.1665 | 0.33162 | 0.78757 | 1.6575 | 3.0911
172 | 5ppb CCV |  | 06/09/2023 06:48:52 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5ppb CCV.526 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.67567 | 0.28347 | 2.4061 | 0.020822 | 0.1094 | 0.11789 | 0.30701 | 0.12677 | 0.49385 | 0.68662 | 0.20538 | 0.14548 | 0.37982 | 0.6121 | 1.4399 | 0.39054 | 0.94342 | 0.39975 | 0.18815 | 0.20328 | 0.044735 | 8.324 | 0.18315 | 0.022583 | 0.0040327 | 0.012778 | 0.026815 | 0.019112 | 0.0046036 | 0.36095 | 0.36139 | 146905.2432 | 179876.3598 | 10960.203 | 203162.4332 | 164918.5365 | 159012.9743 | 3958.5487 | 0.16511 | 0.39365 | 0.8704 | 1.5608
173 | 25ppb CCV |  | 06/09/2023 06:54:40 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\25ppb CCV.527 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.577 | 1.4322 | 10.49 | 0.09201 | 0.47797 | 0.58462 | 1.6507 | 0.6867 | 2.6744 | 3.6404 | 1.0759 | 0.71608 | 1.9552 | 3.0473 | 7.5581 | 2.0332 | 4.8415 | 1.9891 | 0.94969 | 1.0299 | 0.24801 | 45.5962 | 0.9651 | 0.10477 | 0.020702 | 0.065437 | 0.13889 | 0.099406 | 0.02544 | 1.8989 | 1.8782 | 141419.2986 | 163972.4334 | 9870.0757 | 184495.8479 | 150914.9618 | 145452.1598 | 3470.4215 | 0.85514 | 2.0193 | 4.1587 | 7.7513
174 | 8_30_21_W_1 |  | 06/09/2023 07:00:27 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_W_1.528 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.19278 | 0.063923 | 9.5961 | 0.078207 | 0.099473 | 0.097653 | 0.022638 | 0.16963 | 0.0071638 | 0.0046821 | 0.0012622 | 0.17491 | 0.23731 | 0.78606 | 0.0025085 | 0.0060476 | 0.062623 | 0.025528 | 0.015334 | 0.038302 | 0.0086939 | 1.7195 | 0.036373 | 0.0016984 | -6.9524e-06 | 3.2307e-05 | 4.9678e-05 | 0.00016363 | 1.4873e-05 | 0.20867 | 0.21376 | 154875.1973 | 177708.5803 | 10856.1252 | 207322.7759 | 168202.6371 | 155482.2986 | 3757.8276 | 6.9126e-05 | 0.00017336 | 0.0071078 | 0.013093
175 | LBLANK |  | 06/09/2023 07:06:16 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.529 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 7.0358e-05 | 5.8953e-05 | 0.0076724 | 0.00019103 | 0.0068967 | 0.0088255 | 0.00030794 | -4.7652e-05 | 7.3069e-05 | 0.00053412 | 5.7708e-05 | -0.00026165 | 4.4075e-05 | 0.0077383 | -0.00017926 | 0.00068033 | 0.00015577 | 0.00015659 | 0.00010973 | -0.00079036 | 1.1353e-05 | 0.0042466 | 0.0001005 | 0.0017265 | -1.4633e-05 | -1.4792e-06 | 2.284e-05 | -0.00016922 | 3.7482e-06 | 0.0029157 | 0.00027116 | 145640.501 | 180032.0656 | 11136.3409 | 208867.196 | 169715.6274 | 159602.0507 | 3883.1945 | 7.9469e-06 | 2.2147e-05 | 2.6307e-06 | -6.3481e-05
176 | CCV 1ppb TRACE |  | 06/09/2023 07:12:04 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.530 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.1407 | 0.059037 | 0.77077 | 0.0069227 | 0.015559 | 0.0257 | 0.063619 | 0.026743 | 0.10657 | 0.13991 | 0.042543 | 0.032009 | 0.073457 | 0.14258 | 0.29146 | 0.082642 | 0.20039 | 0.086388 | 0.038375 | 0.03732 | 0.0079531 | 1.623 | 0.035243 | 0.0052323 | 0.0008065 | 0.0023487 | 0.0052105 | 0.0038586 | 0.00082615 | 0.076618 | 0.069389 | 155464.839 | 187777.7061 | 10998.8992 | 199842.7774 | 165632.8403 | 152632.8781 | 4063.9127 | 0.034534 | 0.084175 | 0.17172 | 0.33511
177 | CCV 1ppb TRACE |  | 06/09/2023 07:18:22 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.531 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.00016686 | 0 | -0.0065092 | -4.7652e-05 | -0.006557 | -0.039581 | -0.00075905 | -4.7652e-05 | -4.7652e-05 | 0.00011783 | 2.8672e-05 | -0.00097526 | -1.7042e-05 | -0.070257 | 0.00023317 | -0.00034019 | -0.00027356 | 7.4199e-06 | -0.00048602 | 0.021382 | 7.4571e-06 | -0.0047335 | -9.3083e-05 | -0.0046361 | -1.4744e-05 | 1.7477e-06 | -6.9029e-06 | -0.00027122 | 0 | 0.0023759 | -2.6911e-05 | 165357.9153 | 190031.871 | 12569.5375 | 232457.9999 | 188271.8174 | 156045.3393 | 4049.9081 | 6.6907e-06 | -9.0504e-06 | -0.00012035 | -0.00021188
178 | LFB |  | 06/09/2023 07:24:08 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.532 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.357 | 0.54549 | 3.9103 | 0.036095 | 0.20535 | 0.2142 | 0.60605 | 0.24328 | 0.98074 | 1.3697 | 0.40393 | 0.27565 | 0.76123 | 1.1569 | 2.8251 | 0.76407 | 1.8296 | 0.7749 | 0.3683 | 0.37011 | 0.088282 | 17.2418 | 0.37031 | 0.038107 | 0.0079814 | 0.0253 | 0.053944 | 0.035556 | 0.0091243 | 0.70878 | 0.71737 | 151337.2326 | 167137.4519 | 10294.3758 | 191588.2543 | 156344.8833 | 148907.4403 | 3592.4522 | 0.33239 | 0.78583 | 1.5709 | 3.0863
179 | LFB |  | 06/09/2023 07:30:24 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.533 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.3257 | 0.55098 | 4.1205 | 0.035617 | 0.20614 | 0.21474 | 0.60466 | 0.25205 | 0.97918 | 1.3476 | 0.41034 | 0.27415 | 0.76301 | 1.1845 | 2.8405 | 0.76669 | 1.842 | 0.77466 | 0.36912 | 0.38854 | 0.089347 | 17.2755 | 0.37257 | 0.041282 | 0.0084955 | 0.024996 | 0.05285 | 0.036089 | 0.0097752 | 0.71756 | 0.72274 | 145004.9207 | 167536.209 | 10230.9967 | 190355.9021 | 157376.7803 | 144830.4943 | 3612.4567 | 0.33037 | 0.78094 | 1.6446 | 3.0633
180 | CCV 10 ppb MAJORS |  | 06/09/2023 07:36:12 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.534 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.4267 | 0.57664 | 3.9516 | 0.036329 | 0.17118 | 0.24356 | 0.65334 | 0.27831 | 1.0714 | 1.4711 | 0.41339 | 0.27683 | 0.7855 | 1.272 | 3.0862 | 0.839 | 1.9757 | 0.78721 | 0.36786 | 0.37704 | 0.097818 | 17.4239 | 0.37372 | 0.039677 | 0.0080148 | 0.025175 | 0.052542 | 0.035811 | 0.0098528 | 0.72729 | 0.74548 | 149186.4271 | 168394.7296 | 9946.1314 | 193726.1363 | 157914.8244 | 148606.2723 | 3613.1238 | 0.33822 | 0.80297 | 1.6543 | 3.1479
181 | CCV 10 ppb MAJORS |  | 06/09/2023 07:42:30 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.535 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.4329 | 0.57934 | 4.0209 | 0.036777 | 0.16763 | 0.2308 | 0.64626 | 0.25558 | 1.0224 | 1.423 | 0.41585 | 0.28026 | 0.78127 | 1.255 | 3.0061 | 0.81184 | 1.9348 | 0.79663 | 0.37102 | 0.38319 | 0.091736 | 17.5684 | 0.37053 | 0.041347 | 0.0076304 | 0.024426 | 0.052606 | 0.035694 | 0.0092451 | 0.74073 | 0.73937 | 146961.4027 | 168933.8475 | 10128.9273 | 192408.6597 | 149013.4214 | 150194.8965 | 3563.1111 | 0.3411 | 0.79945 | 1.686 | 3.3786
182 | 8_30_21_W_6 |  | 06/09/2023 07:48:18 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_W_6.536 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.22821 | 0.079307 | 4.2258 | 0.037154 | 0.059131 | 0.082927 | 0.023591 | 0.17897 | 0.015077 | 0.004406 | 0.0012663 | 0.23004 | 0.24264 | 1.0496 | 0.0020135 | 0.0074432 | 0.1747 | 0.070813 | 0.024518 | 0.0373 | 0.0069901 | 1.7428 | 0.03758 | 0.0016198 | 1.6157e-05 | 6.7629e-05 | 9.2929e-05 | 0.00016482 | 4.9857e-05 | 0.11839 | 0.11848 | 149406.6056 | 173925.0338 | 10468.5261 | 198941.1726 | 155244.7533 | 153147.2831 | 3750.4925 | 0.00013759 | 0.000361 | 0.013797 | 0.027493
183 | 8_30_21_W_6_LD |  | 06/09/2023 07:54:05 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_W_6_LD.537 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.21719 | 0.073988 | 5.5297 | 0.046799 | 0.074613 | 0.079886 | 0.022744 | 0.17103 | 0.014396 | 0.0032056 | 0.00096113 | 0.22711 | 0.26686 | 1.0106 | 0.0017946 | 0.0081353 | 0.16554 | 0.069818 | 0.024435 | 0.037976 | 0.0064948 | 1.9531 | 0.040898 | 0.0019591 | -1.315e-06 | 7.4843e-05 | 0.00012015 | 0.00076591 | 2.4813e-05 | 0.11652 | 0.13063 | 150621.2853 | 158629.5763 | 10493.8533 | 198241.8724 | 167095.3 | 152701.2267 | 3326.3886 | 0.00024372 | 0.00051001 | 0.014629 | 0.026086
184 | 8_30_21_W_6_LD |  | 06/09/2023 08:00:20 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_W_6_LD.538 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.23043 | 0.080607 | 5.3431 | 0.049171 | 0.079867 | 0.088433 | 0.024812 | 0.19166 | 0.015455 | 0.0036295 | 0.0009437 | 0.22827 | 0.24408 | 1.1179 | 0.0019842 | 0.007921 | 0.18156 | 0.070291 | 0.024838 | 0.037502 | 0.0080684 | 1.7372 | 0.037751 | 0.0015296 | 5.2633e-06 | 5.7579e-05 | 0.00011907 | 0.00044898 | 3.532e-05 | 0.11836 | 0.11905 | 145182.7198 | 169908.4411 | 9872.0843 | 192877.3723 | 140220.5532 | 149360.8168 | 3691.8103 | 0.00018186 | 0.00036879 | 0.015088 | 0.031636
185 | 8_30_21_W_6_DUP |  | 06/09/2023 08:06:08 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_W_6_DUP.539 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.21907 | 0.074133 | 5.2045 | 0.045647 | 0.080673 | 0.081259 | 0.022882 | 0.17135 | 0.012162 | 0.0030239 | 0.00093907 | 0.24632 | 0.24602 | 1.1102 | 0.0018597 | 0.0070023 | 0.18344 | 0.077732 | 0.025069 | 0.036814 | 0.0071451 | 1.6932 | 0.036486 | 0.0014203 | -6.6613e-06 | 2.2274e-05 | 0.00012349 | 0.00054186 | 1.5824e-05 | 0.11822 | 0.11971 | 147359.4521 | 169787.1242 | 10476.5069 | 193953.283 | 163512.6069 | 150696.0416 | 3660.4694 | 0.00014331 | 0.00038726 | 0.017162 | 0.029265
186 | 8_30_21_W_6 |  | 06/09/2023 08:11:56 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_W_6.540 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.22009 | 0.078354 | 4.0327 | 0.03616 | 0.072695 | 0.085643 | 0.023348 | 0.18383 | 0.01317 | 0.003994 | 0.0012482 | 0.22757 | 0.24168 | 1.0543 | 0.0016519 | 0.0082585 | 0.17073 | 0.06943 | 0.024226 | 0.035958 | 0.0096038 | 1.7583 | 0.037616 | 0.0011424 | -1.4308e-05 | 2.6323e-05 | 8.4349e-05 | 0.00020455 | 2.3663e-05 | 0.11862 | 0.11945 | 150790.3772 | 169304.0782 | 9877.415 | 190506.5364 | 157168.124 | 157471.6131 | 3622.4596 | 0.00017705 | 0.00045099 | 0.014131 | 0.027731
187 | 8_30_21_W_6_DUP_LSP |  | 06/09/2023 08:17:43 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_W_6_DUP_LSP.541 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.2278 | 1.2696 | 13.6339 | 0.12384 | 0.43434 | 0.59004 | 1.3679 | 0.78035 | 2.1945 | 2.9759 | 0.83625 | 0.82495 | 1.7942 | 3.8358 | 6.1485 | 1.7065 | 4.1058 | 1.623 | 0.75341 | 0.8401 | 0.20213 | 35.7783 | 0.77979 | 0.082648 | 0.016756 | 0.050607 | 0.10642 | 0.075011 | 0.019631 | 1.6098 | 1.612 | 138130.4884 | 152982.1015 | 8740.6751 | 174316.4669 | 142274.3473 | 135158.6776 | 3334.3891 | 0.68057 | 1.6017 | 3.4372 | 6.3076
188 | 8_30_21_W_6_DUP_LSP |  | 06/09/2023 08:24:01 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_W_6_DUP_LSP.542 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.0807 | 1.229 | 13.5603 | 0.11919 | 0.4388 | 0.56675 | 1.3273 | 0.74815 | 2.1307 | 2.9063 | 0.83619 | 0.82055 | 1.8268 | 3.6916 | 6.0007 | 1.6464 | 3.9777 | 1.636 | 0.76145 | 0.83349 | 0.20315 | 37.3557 | 0.79213 | 0.081692 | 0.016292 | 0.05142 | 0.1085 | 0.074474 | 0.01977 | 1.6076 | 1.636 | 136944.5899 | 154600.0637 | 9270.3431 | 177227.2315 | 144773.2118 | 138068.7103 | 3279.043 | 0.67209 | 1.5922 | 3.313 | 6.2272
189 | 8_30_21_W_8 |  | 06/09/2023 08:29:49 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_W_8.543 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.25376 | 0.082241 | 4.3104 | 0.040999 | 0.10812 | 0.088803 | 0.024766 | 0.1822 | 0.016009 | 0.0062182 | 0.0014625 | 0.20317 | 0.27002 | 0.97746 | 0.0024621 | 0.004768 | 0.18523 | 0.073661 | 0.017965 | 0.040909 | 0.0091325 | 1.7835 | 0.03921 | 0.0014288 | -1.0375e-05 | 5.4372e-05 | 0.00012479 | 0.00040911 | 2.3865e-05 | 0.10749 | 0.10424 | 146493.2843 | 167701.2629 | 9810.7183 | 185225.8458 | 157585.0562 | 121553.3444 | 3689.8102 | 0.00015071 | 0.00035327 | 0.011171 | 0.020703
190 | 8_30_21_W_8 |  | 06/09/2023 08:36:07 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_W_8.544 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.23078 | 0.071078 | 4.3706 | 0.037252 | 0.081126 | 0.078694 | 0.022667 | 0.17427 | 0.018265 | 0.0051999 | 0.0014072 | 0.19557 | 0.27251 | 0.87515 | 0.0014489 | 0.0033506 | 0.17421 | 0.071048 | 0.018226 | 0.039998 | 0.0085599 | 1.8667 | 0.040309 | 0.0015348 | 1.0514e-05 | 5.2148e-05 | 9.2099e-05 | 0.00038825 | 8.1993e-06 | 0.1036 | 0.10635 | 142407.6454 | 162685.7771 | 10230.3291 | 187488.5767 | 154796.3362 | 143734.2788 | 3515.0993 | 0.00012669 | 0.00040181 | 0.011095 | 0.021155
191 | 8_30_21_W_8_SP |  | 06/09/2023 08:41:55 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_W_8_SP.545 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.38319 | 0.13519 | 4.7941 | 0.044367 | 0.10323 | 0.10114 | 0.091412 | 0.20977 | 0.12507 | 0.15417 | 0.043408 | 0.21004 | 0.3457 | 0.98192 | 0.31282 | 0.093136 | 0.36803 | 0.14688 | 0.054715 | 0.081312 | 0.019339 | 3.7817 | 0.080676 | 0.0054803 | 0.00084362 | 0.0027792 | 0.0058523 | 0.0038074 | 0.0011128 | 0.17665 | 0.18361 | 139602.8258 | 159451.6435 | 9715.3099 | 188522.4853 | 150926.4931 | 141145.9553 | 3401.7384 | 0.035739 | 0.084199 | 0.19821 | 0.37931
192 | LBLANK |  | 06/09/2023 08:47:44 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.546 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.2875e-05 | 0 | 0.006082 | 8.3708e-05 | 0.0087324 | 0.0064157 | -0.00034563 | -4.7652e-05 | -4.7652e-05 | 0.00042155 | 0.00010175 | -0.00027679 | 5.341e-05 | -0.0047342 | -3.9959e-05 | 0.0014072 | -0.00010218 | 0.00012545 | 5.8813e-05 | -0.0010302 | 1.7607e-05 | 0.00015574 | 1.3695e-05 | 0.0011936 | -1.4161e-05 | 1.9336e-06 | 1.6752e-05 | 0.00016659 | 0 | 0.00029384 | 0.00011677 | 146077.7297 | 170535.4398 | 10438.4848 | 206665.3545 | 160056.34 | 140850.6715 | 3745.8264 | 4.776e-06 | 9.9955e-06 | -4.9994e-05 | -0.00011476
193 | CCV 1ppb TRACE |  | 06/09/2023 08:53:33 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.547 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.13815 | 0.055332 | 0.81163 | 0.0070784 | 0.015758 | 0.032207 | 0.065859 | 0.025998 | 0.1029 | 0.14627 | 0.041568 | 0.030349 | 0.078394 | 0.13058 | 0.29307 | 0.080039 | 0.1939 | 0.084873 | 0.037356 | 0.038596 | 0.0093023 | 1.7148 | 0.037409 | 0.0053216 | 0.00079782 | 0.0026038 | 0.0056238 | 0.0038653 | 0.00091222 | 0.073713 | 0.074273 | 139102.7437 | 165912.5776 | 10287.7137 | 190525.4522 | 154163.044 | 144623.2522 | 3619.1251 | 0.034966 | 0.081256 | 0.18887 | 0.34034
194 | CCV 1ppb TRACE |  | 06/09/2023 08:59:50 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.548 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.14448 | 0.057039 | 0.77025 | 0.0063322 | 0.013862 | 0.035535 | 0.062947 | 0.026668 | 0.10449 | 0.14014 | 0.04161 | 0.031272 | 0.078307 | 0.14369 | 0.28798 | 0.081228 | 0.20523 | 0.0846 | 0.038113 | 0.038859 | 0.0086555 | 1.7052 | 0.036576 | 0.0049256 | 0.00075752 | 0.0027671 | 0.0055013 | 0.003868 | 0.0010609 | 0.074401 | 0.075697 | 143992.0534 | 163308.1811 | 10189.6379 | 187778.1753 | 154068.3852 | 144204.8541 | 3506.4305 | 0.033686 | 0.082384 | 0.17903 | 0.33868
195 | CCV 10 ppb MAJORS |  | 06/09/2023 09:05:39 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.549 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.4546 | 0.58817 | 3.9445 | 0.039512 | 0.17014 | 0.24705 | 0.65694 | 0.27636 | 1.0565 | 1.4616 | 0.41299 | 0.2782 | 0.77969 | 1.2798 | 3.0469 | 0.83208 | 1.9693 | 0.78961 | 0.37034 | 0.38696 | 0.096826 | 16.8741 | 0.37078 | 0.042004 | 0.0082605 | 0.024935 | 0.052706 | 0.036507 | 0.0096236 | 0.73773 | 0.7428 | 137774.8852 | 155351.8355 | 9238.9936 | 180607.9567 | 146073.6886 | 136056.8645 | 3413.0745 | 0.34016 | 0.81288 | 1.6998 | 3.2098
196 | CCV 10 ppb MAJORS |  | 06/09/2023 09:11:56 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.550 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.3749 | 0.55606 | 3.91 | 0.035921 | 0.16801 | 0.23712 | 0.62691 | 0.26198 | 1.0049 | 1.3867 | 0.41584 | 0.28114 | 0.78346 | 1.2229 | 2.889 | 0.79318 | 1.8831 | 0.79941 | 0.37232 | 0.39019 | 0.092237 | 17.7874 | 0.37716 | 0.040947 | 0.0078483 | 0.024939 | 0.053277 | 0.037228 | 0.0095391 | 0.74528 | 0.74982 | 140225.5466 | 157235.3302 | 9821.3751 | 180583.1551 | 148733.0281 | 138385.754 | 3339.0574 | 0.34259 | 0.81342 | 1.6727 | 3.1833
197 | 8_30_21_W_9 |  | 06/09/2023 09:17:44 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_W_9.551 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.20262 | 0.072611 | 6.6824 | 0.059595 | 0.074958 | 0.063437 | 0.021665 | 0.1678 | 0.016896 | 0.0045669 | 0.0012617 | 0.38624 | 0.3689 | 1.7791 | 0.0032382 | 0.003174 | 0.1877 | 0.075947 | 0.026529 | 0.033235 | 0.0078538 | 1.5674 | 0.033596 | 0.0013227 | -1.8806e-06 | 3.5446e-05 | 0.00014097 | 0.00032976 | 2.0573e-05 | 0.074898 | 0.075284 | 142005.8827 | 162761.3844 | 10196.3048 | 188945.3119 | 157727.3767 | 146899.4537 | 3488.4259 | 0.00010815 | 0.00036646 | 0.016343 | 0.029276
198 | 8_30_21_W_10 |  | 06/09/2023 09:23:32 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_W_10.552 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.20119 | 0.098372 | 5.9951 | 0.052514 | 0.091042 | 0.081892 | 0.024525 | 0.19836 | 0.017828 | 0.0034952 | 0.0011401 | 0.29647 | 0.23901 | 1.4036 | 0.0027205 | 0.0032884 | 0.14612 | 0.05708 | 0.025009 | 0.035036 | 0.0081964 | 1.6341 | 0.034417 | 0.00085832 | 6.8066e-06 | 7.0091e-05 | 0.00012396 | 0.00043811 | 2.0972e-05 | 0.047605 | 0.049443 | 143255.8303 | 159111.9429 | 9711.3019 | 185788.0793 | 150884.6424 | 143142.0652 | 3350.3929 | 0.00014489 | 0.00037998 | 0.014229 | 0.02599
199 | 8_30_21_W_11 |  | 06/09/2023 09:29:20 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_W_11.553 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.19742 | 0.09014 | 7.0158 | 0.059455 | 0.094048 | 0.079062 | 0.022687 | 0.17387 | 0.019766 | 0.003504 | 0.001018 | 0.34361 | 0.27268 | 1.5046 | 0.0020909 | 0.0040243 | 0.14067 | 0.061551 | 0.02766 | 0.036083 | 0.0068597 | 1.6019 | 0.034827 | 0.00095249 | -9.9217e-06 | 4.1134e-05 | 0.00012891 | 0.0003972 | 2.1114e-05 | 0.045512 | 0.045437 | 141494.6974 | 157955.833 | 10029.5196 | 182388.4181 | 145814.266 | 138908.8901 | 3433.0792 | 0.00014821 | 0.00051854 | 0.043508 | 0.085717
200 | 8_30_21_F_1 |  | 06/09/2023 09:35:08 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_F_1.554 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.99917 | 0.32976 | 2.2002 | 0.020872 | 0.15377 | 0.48107 | 0.12262 | 0.93446 | 0.03365 | 0.016959 | 0.0051913 | 0.10468 | 0.85215 | 0.52503 | 0.008765 | 0.034147 | 0.16548 | 0.069154 | 0.19901 | 0.18129 | 0.039536 | 7.6459 | 0.17398 | 0.0023955 | 3.0206e-05 | 0.00010529 | 0.00025902 | 0.0014716 | 2.6469e-05 | 1.0685 | 1.0993 | 135488.9286 | 150937.8691 | 9196.9612 | 174879.9116 | 143267.5115 | 138261.8851 | 3436.4135 | 0.00025045 | 0.00058878 | 0.016949 | 0.032253
201 | 8_30_21_F_1 |  | 06/09/2023 09:41:25 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_F_1.555 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.95844 | 0.29944 | 2.2828 | 0.019289 | 0.14392 | 0.46546 | 0.11472 | 0.86626 | 0.033178 | 0.017446 | 0.005444 | 0.10787 | 0.84787 | 0.48674 | 0.0084503 | 0.031341 | 0.1559 | 0.070664 | 0.20071 | 0.18421 | 0.039818 | 8.0328 | 0.17354 | 0.0020696 | 3.7785e-05 | 7.2514e-05 | 0.00022671 | 0.0012336 | 8.6338e-05 | 1.0855 | 1.0931 | 134448.4145 | 154494.249 | 9762.6797 | 173234.7122 | 142973.9136 | 138336.2719 | 3337.7233 | 0.00018758 | 0.00064078 | 0.01744 | 0.032172
202 | 8_30_21_F_1_DUP |  | 06/09/2023 09:47:13 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_F_1_DUP.556 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.0263 | 0.3225 | 2.5581 | 0.020711 | 0.14452 | 0.48192 | 0.1147 | 0.90249 | 0.035016 | 0.017121 | 0.004569 | 0.10574 | 0.83161 | 0.51952 | 0.0084333 | 0.03444 | 0.16508 | 0.070731 | 0.19914 | 0.19717 | 0.039159 | 8.0242 | 0.17153 | 0.0029732 | 2.4831e-05 | 0.00013277 | 0.00021781 | 0.0016512 | 4.7278e-05 | 1.0712 | 1.088 | 125248.1746 | 154788.079 | 9396.4262 | 174900.8757 | 145856.2936 | 136968.2361 | 3309.0499 | 0.00024698 | 0.00058034 | 0.018772 | 0.032661
203 | 8_30_21_F_1_DUP |  | 06/09/2023 09:53:30 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_F_1_DUP.557 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.0128 | 0.31787 | 2.3127 | 0.022037 | 0.14224 | 0.47263 | 0.11774 | 0.88093 | 0.034368 | 0.016002 | 0.0047215 | 0.10482 | 0.84762 | 0.49304 | 0.0077896 | 0.032869 | 0.1585 | 0.070679 | 0.20016 | 0.17875 | 0.038975 | 8.0548 | 0.17345 | 0.0020559 | 1.2092e-05 | 9.9047e-05 | 0.0002419 | 0.0014288 | 6.0916e-05 | 1.0665 | 1.0898 | 134013.1678 | 153482.832 | 9447.1242 | 173912.8774 | 144006.8963 | 136154.5398 | 3305.049 | 0.00026346 | 0.00052673 | 0.017168 | 0.033146
204 | 8_30_21_F_6 |  | 06/09/2023 09:59:18 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_F_6.558 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.1169 | 0.37225 | 1.4973 | 0.011966 | 0.043139 | 0.3932 | 0.11697 | 0.86837 | 0.066334 | 0.0064046 | 0.0023346 | 0.16548 | 0.82228 | 0.72292 | 0.0070601 | 0.018371 | 0.49255 | 0.21904 | 0.12171 | 0.18724 | 0.042188 | 8.4345 | 0.17985 | 0.0021319 | 3.8227e-05 | 0.00023673 | 0.00041857 | 0.0013278 | 0.00011705 | 0.61042 | 0.61041 | 132673.2622 | 154082.5815 | 9600.5602 | 172166.0171 | 121183.5689 | 136890.1629 | 3291.0464 | 0.00066831 | 0.0013262 | 0.022491 | 0.051117
205 | 8_30_21_F_6 |  | 06/09/2023 10:05:38 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_F_6.559 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.1071 | 0.38349 | 1.4739 | 0.011982 | 0.041365 | 0.40829 | 0.11404 | 0.91321 | 0.06579 | 0.0082472 | 0.0023781 | 0.16641 | 0.84733 | 0.74506 | 0.0066691 | 0.017983 | 0.52189 | 0.2216 | 0.12326 | 0.18588 | 0.040125 | 8.4108 | 0.18161 | 0.0024555 | 5.5821e-05 | 0.00025149 | 0.00046402 | 0.0013808 | 9.1399e-05 | 0.61819 | 0.6195 | 134181.2189 | 153182.2632 | 9563.2318 | 169663.5949 | 143883.7791 | 139349.6789 | 3313.7184 | 0.00068714 | 0.0013894 | 0.021878 | 0.04186
206 | 8_30_21_F_8 |  | 06/09/2023 10:11:26 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_F_8.560 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.1926 | 0.36917 | 1.0138 | 0.0089077 | 0.066578 | 0.37446 | 0.12292 | 0.87773 | 0.08348 | 0.0089547 | 0.0023149 | 0.17712 | 1.0015 | 0.80953 | 0.0065412 | 0.018041 | 0.56689 | 0.2475 | 0.063288 | 0.2022 | 0.042466 | 9.7373 | 0.21251 | 0.0022586 | 2.9155e-05 | 0.00023206 | 0.00045056 | 0.0015317 | 8.1776e-05 | 0.54887 | 0.60596 | 132809.8683 | 141995.4328 | 9601.2285 | 170832.7465 | 141558.0105 | 136273.5505 | 3096.3362 | 0.00057651 | 0.001138 | 0.018407 | 0.034803
207 | 8_30_21_F_8 |  | 06/09/2023 10:17:43 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_F_8.561 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.2015 | 0.38478 | 1.0146 | 0.0068724 | 0.064387 | 0.38399 | 0.12367 | 0.91874 | 0.083144 | 0.0087419 | 0.0023441 | 0.17657 | 0.88919 | 0.83673 | 0.0070364 | 0.016606 | 0.58618 | 0.24157 | 0.06311 | 0.20338 | 0.047154 | 8.8039 | 0.18722 | 0.0024633 | 6.3504e-05 | 0.00022678 | 0.00046327 | 0.0016424 | 7.7297e-05 | 0.54679 | 0.53664 | 132822.6086 | 155467.2235 | 9254.9971 | 172077.5975 | 144428.5365 | 134589.083 | 3303.7156 | 0.00050572 | 0.0012067 | 0.019064 | 0.03455
208 | LBLANK |  | 06/09/2023 10:23:31 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.562 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 8.6683e-05 | 0 | 0.0072873 | 0.00028695 | -0.0048369 | 0.011058 | -0.00023487 | -4.7652e-05 | 1.9651e-05 | 0.00019242 | 0.00011658 | -0.00017457 | 5.3944e-05 | 0.0064504 | -0.00023826 | 0.00057681 | -0.00018798 | 0.00022111 | 0.00012859 | -0.0010596 | -4.7652e-05 | 0.0033246 | 8.4652e-05 | 0.0010712 | -1.8296e-05 | 7.0275e-06 | 2.6215e-05 | 4.9185e-05 | 8.0842e-06 | 0.00028038 | 0.00023479 | 140983.1136 | 163242.9673 | 9959.4706 | 177274.6418 | 150627.5382 | 141895.9901 | 3597.1195 | -1.4171e-06 | 2.0773e-05 | -0.00010582 | -7.1273e-05
209 | CCV 1ppb TRACE |  | 06/09/2023 10:29:20 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.563 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.14517 | 0.057544 | 0.77421 | 0.0067512 | 0.012977 | 0.037895 | 0.062223 | 0.030126 | 0.10681 | 0.14428 | 0.042384 | 0.031646 | 0.078113 | 0.15045 | 0.30518 | 0.085595 | 0.20896 | 0.084228 | 0.03903 | 0.037967 | 0.0095275 | 1.7363 | 0.03751 | 0.004742 | 0.00085989 | 0.0025009 | 0.0054919 | 0.0038071 | 0.00089849 | 0.074679 | 0.075469 | 141737.7045 | 161751.9551 | 9828.7139 | 185708.0486 | 152713.9277 | 139299.2505 | 3493.0937 | 0.035259 | 0.083237 | 0.17455 | 0.33428
210 | CCV 1ppb TRACE |  | 06/09/2023 10:35:37 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.564 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.14601 | 0.062209 | 0.77211 | 0.0073114 | 0.011799 | 0.037289 | 0.061576 | 0.028622 | 0.10897 | 0.1486 | 0.042729 | 0.031696 | 0.07706 | 0.15585 | 0.30636 | 0.085671 | 0.20958 | 0.084528 | 0.038411 | 0.036622 | 0.0094819 | 1.697 | 0.037295 | 0.0050317 | 0.00077399 | 0.002617 | 0.0054211 | 0.0039291 | 0.00088241 | 0.074535 | 0.075938 | 143472.211 | 163173.4062 | 9865.413 | 185901.1078 | 153517.8255 | 142429.9098 | 3587.1172 | 0.035255 | 0.082724 | 0.17395 | 0.33633
211 | LFB |  | 06/09/2023 10:41:25 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.565 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.3013 | 0.5308 | 3.9137 | 0.037224 | 0.21131 | 0.22559 | 0.59414 | 0.25688 | 0.96212 | 1.3339 | 0.40441 | 0.2684 | 0.75636 | 1.1474 | 2.7612 | 0.76434 | 1.8071 | 0.76142 | 0.36929 | 0.37847 | 0.086919 | 16.6404 | 0.37348 | 0.038589 | 0.0080625 | 0.025613 | 0.05348 | 0.036985 | 0.0098363 | 0.71333 | 0.73406 | 135856.9594 | 149746.7804 | 9551.8589 | 172889.2251 | 141707.3804 | 133988.7977 | 3361.0621 | 0.33253 | 0.79554 | 1.5965 | 3.0847
212 | LFB |  | 06/09/2023 10:47:41 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.566 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.3152 | 0.51642 | 4.0161 | 0.03596 | 0.21962 | 0.2251 | 0.59878 | 0.24975 | 0.95731 | 1.3292 | 0.40317 | 0.26846 | 0.75966 | 1.1401 | 2.7981 | 0.76356 | 1.8124 | 0.76717 | 0.36956 | 0.39063 | 0.091029 | 17.343 | 0.37102 | 0.040291 | 0.0080619 | 0.025378 | 0.054419 | 0.03745 | 0.0097637 | 0.72113 | 0.73111 | 135624.1461 | 152638.5361 | 9719.3057 | 175064.9123 | 143577.9427 | 131328.9584 | 3269.041 | 0.33462 | 0.79659 | 1.6035 | 3.0982
213 | CCV 10 ppb MAJORS |  | 06/09/2023 10:53:35 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.567 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.3992 | 0.56331 | 3.9476 | 0.035307 | 0.17251 | 0.24129 | 0.63511 | 0.2697 | 1.0266 | 1.4103 | 0.41905 | 0.27889 | 0.79959 | 1.2501 | 2.9445 | 0.79553 | 1.882 | 0.78681 | 0.3689 | 0.38943 | 0.09282 | 17.2213 | 0.38191 | 0.041492 | 0.0083258 | 0.025797 | 0.054961 | 0.0374 | 0.0095776 | 0.74266 | 0.76198 | 140834.169 | 152930.3529 | 9772.6763 | 180007.8255 | 147427.5373 | 137080.5732 | 3391.0691 | 0.3375 | 0.80202 | 1.6432 | 3.1549
214 | LBLANK |  | 06/09/2023 11:28:06 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.568 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 5.2996e-05 | 0 | 0.10443 | 0.0007393 | 0.00486 | 0.049709 | 0.00058032 | 0 | 5.3191e-05 | 0.0011069 | 0.000171 | 0.0025961 | 5.735e-05 | 0.10432 | 5.2996e-05 | 0.00089544 | 0.00089468 | 0.00039492 | 0.0046962 | -0.015859 | 0 | 0.0021255 | 4.7189e-05 | 0.006255 | 0 | 1.0082e-05 | 0 | 0.00030591 | 0 | 8.5457e-05 | 8.0899e-05 | 176517.9465 | 197865.5394 | 12640.9241 | 234349.9989 | 192913.21 | 172238.9985 | 4402.012 | 2.8716e-06 | 7.3881e-05 | 0.0001021 | 0.00020417
215 | Cal blank |  | 06/09/2023 11:33:55 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\Cal blank.569 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0 | 0 | 0.08785 | 0.00052549 | 0.0035722 | 0.058554 | 0.00068378 | 0 | 5.3533e-05 | 0.00078853 | 9.33e-05 | 0.0025753 | 5.3687e-05 | 0.1039 | 0.00026419 | 0.0014775 | 0.00089386 | 0.00025732 | 0.0033841 | -0.016 | 5.1219e-05 | 0.0019805 | 4.3515e-05 | 0.0060054 | 0 | 3.3825e-06 | 1.0061e-05 | 0.00037263 | 6.7772e-06 | 6.1931e-05 | 3.3565e-05 | 173145.1592 | 198833.7946 | 12686.2991 | 235708.7554 | 194173.4048 | 174881.419 | 4366.6677 | 1.1345e-05 | 3.1063e-05 | 0.00017799 | 0.00031946
216 | CAL BLANK |  | 06/09/2023 11:39:43 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CAL BLANK.570 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 5.331e-05 | 5.3327e-05 | 0.070929 | 0.00069693 | 0.0031799 | 0.056763 | 0.00080208 | 0 | 0 | 0.00058956 | 9.2823e-05 | 0.0028162 | 7.7512e-05 | 0.10002 | 0.0001607 | 0.00091026 | 0.00053546 | 0.00038005 | 0.0096058 | -0.017095 | 0 | 0.00077751 | 1.6838e-05 | 0.0060881 | 0 | 6.7348e-06 | 1.3499e-05 | 0.00033903 | 0 | 3.7693e-05 | 2.0204e-05 | 175923.9588 | 197782.4116 | 12444.7517 | 230195.4537 | 188175.9642 | 172199.7869 | 4290.6443 | 0 | 3.1753e-05 | 0.00011748 | 0.00030113
217 | LOW 0.01 |  | 06/09/2023 11:45:32 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LOW 0.01.571 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.0010534 | 0.00058073 | 0.20288 | 0.0016296 | 0.017348 | -0.00071729 | 0.00062216 | 0.00015847 | 0.00074076 | 0.0016825 | 0.00044429 | 0.0033462 | 0.0024008 | 0.014738 | 0.0023213 | 0.00056481 | 0.0035297 | 0.0012329 | 0.0039537 | -9.8126e-05 | 0 | 0.014565 | 0.0003153 | 3.5508e-05 | 1.6998e-05 | 5.0852e-05 | 4.7216e-05 | -9.4442e-05 | 1.0016e-05 | 0.0013853 | 0.001657 | 175443.415 | 197349.4216 | 12621.5735 | 232419.4697 | 188365.2744 | 173930.4519 | 4250.6324 | 0.0002725 | 0.00079445 | 0.0019799 | 0.0033125
218 | LOW 0.05 |  | 06/09/2023 11:51:20 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LOW 0.05.572 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.0065026 | 0.0022956 | 0.18768 | 0.0018167 | 0.02927 | 0.0014173 | 0.002747 | 0.0015306 | 0.0047 | 0.0061929 | 0.0019112 | 0.0038405 | 0.0033623 | 0.020461 | 0.013123 | 0.0039474 | 0.01034 | 0.0040891 | 0.0010156 | 0.0024345 | 0.00032843 | 0.073199 | 0.0016145 | 0.00028323 | 2.7406e-05 | 0.00010902 | 0.00021796 | 6.2855e-06 | 8.1967e-05 | 0.0032307 | 0.0031844 | 176694.8327 | 195555.4608 | 12197.2049 | 230936.3626 | 188325.8823 | 173952.1167 | 4305.9824 | 0.0014691 | 0.0037003 | 0.0078521 | 0.015534
219 | LOW 0.1 |  | 06/09/2023 11:57:09 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LOW 0.1.573 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.011933 | 0.004359 | 0.077226 | 0.00017909 | 0.016815 | 0.00060277 | 0.0067219 | 0.0026141 | 0.0071925 | 0.013857 | 0.0036953 | 0.0075477 | 0.0065846 | 0.031696 | 0.025994 | 0.0067195 | 0.018219 | 0.0074875 | -0.00011766 | 0.0037661 | 0.00048939 | 0.13293 | 0.0029767 | 0.0007653 | 0.00010741 | 0.00021547 | 0.00044332 | 0.00021708 | 9.6819e-05 | 0.006251 | 0.0066462 | 169485.0049 | 192628.1385 | 12231.9015 | 226869.9285 | 185109.8056 | 166298.4401 | 4314.6516 | 0.0030648 | 0.007171 | 0.016127 | 0.028983
220 | LOW 0.5 |  | 06/09/2023 12:02:57 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LOW 0.5.574 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.057662 | 0.023293 | 0.39798 | 0.0034904 | 0.036425 | 0.0063078 | 0.028033 | 0.0109 | 0.040851 | 0.058043 | 0.01715 | 0.023637 | 0.033014 | 0.10746 | 0.12441 | 0.034311 | 0.087301 | 0.037592 | 0.013858 | 0.015791 | 0.0040803 | 0.68358 | 0.015222 | 0.0017174 | 0.00036051 | 0.0011469 | 0.0024104 | 0.0012961 | 0.00043963 | 0.031729 | 0.032046 | 171573.4124 | 194178.1122 | 12417.3945 | 226097.4724 | 158207.1445 | 167798.4167 | 4323.3211 | 0.015007 | 0.035227 | 0.077348 | 0.17701
221 | STD 1 |  | 06/09/2023 12:08:45 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 1.575 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.12386 | 0.048444 | 0.60264 | 0.0055044 | 0.030595 | 0.016997 | 0.052074 | 0.022452 | 0.086064 | 0.12185 | 0.035264 | 0.036679 | 0.067084 | 0.1587 | 0.2554 | 0.065135 | 0.1678 | 0.071227 | 0.028445 | 0.032642 | 0.0079092 | 1.4445 | 0.03162 | 0.002975 | 0.00067309 | 0.0022317 | 0.0047054 | 0.002864 | 0.00089169 | 0.063711 | 0.066596 | 174104.5338 | 189099.3622 | 12252.5882 | 223026.1748 | 182543.7305 | 164892.5193 | 4141.2671 | 0.03054 | 0.070097 | 0.14975 | 0.29787
222 | STD 5 |  | 06/09/2023 12:14:33 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 5.576 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.57098 | 0.23104 | 1.8251 | 0.01869 | 0.10213 | 0.090226 | 0.26405 | 0.11633 | 0.41803 | 0.57854 | 0.17649 | 0.12155 | 0.33556 | 0.53342 | 1.2118 | 0.32624 | 0.80859 | 0.33827 | 0.15308 | 0.16388 | 0.036897 | 7.3489 | 0.15913 | 0.017013 | 0.0033801 | 0.010902 | 0.023655 | 0.015869 | 0.0041653 | 0.31399 | 0.32373 | 168752.8865 | 182406.8136 | 11901.6228 | 213341.9135 | 173902.2226 | 162384.5163 | 3950.5464 | 0.14911 | 0.35315 | 0.72992 | 1.433
223 | STD 10 |  | 06/09/2023 12:20:21 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 10.577 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.1984 | 0.47789 | 3.3977 | 0.033295 | 0.1689 | 0.19714 | 0.56426 | 0.24425 | 0.87383 | 1.2063 | 0.36068 | 0.24379 | 0.68194 | 1.0419 | 2.5486 | 0.70171 | 1.6282 | 0.68506 | 0.3132 | 0.33259 | 0.076492 | 14.7666 | 0.32216 | 0.033296 | 0.0072002 | 0.021582 | 0.046392 | 0.031346 | 0.0083849 | 0.65048 | 0.66389 | 163515.8753 | 178238.4195 | 11359.8517 | 208310.5196 | 169219.0075 | 158155.3068 | 3888.5293 | 0.29633 | 0.70789 | 1.4414 | 2.8455
224 | STD 25 |  | 06/09/2023 12:26:09 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 25.578 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.9794 | 1.1834 | 9.5805 | 0.077814 | 0.41772 | 0.51753 | 1.4032 | 0.59907 | 2.2305 | 3.0396 | 0.92443 | 0.62351 | 1.7243 | 2.6267 | 6.4126 | 1.7059 | 4.0369 | 1.7179 | 0.81542 | 0.97422 | 0.20646 | 38.2578 | 0.84642 | 0.098232 | 0.018107 | 0.058184 | 0.1222 | 0.092412 | 0.022589 | 1.6682 | 1.7093 | 145258.2501 | 170786.9197 | 11163.6982 | 198765.9467 | 162498.379 | 151561.2151 | 3779.1666 | 0.75799 | 1.792 | 3.9164 | 6.8343
225 | STD 50 |  | 06/09/2023 12:31:57 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 50.579 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 6.1906 | 2.389 | 21.924 | 0.15883 | 0.83759 | 1.0522 | 2.9061 | 1.2354 | 4.6463 | 6.2458 | 1.8873 | 1.1428 | 3.4944 | 5.4121 | 13.0659 | 3.4932 | 8.1819 | 3.3595 | 1.6029 | 1.7251 | 0.41373 | 81.0733 | 1.6887 | 0.17024 | 0.03738 | 0.11698 | 0.24696 | 0.16497 | 0.044347 | 3.3534 | 3.4585 | 153680.8801 | 165955.1023 | 10521.207 | 190482.8915 | 143988.2514 | 144207.632 | 3458.4188 | 1.4685 | 3.4616 | 6.6241 | 14.0806
226 | STD 100 |  | 06/09/2023 12:37:44 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 100.580 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 13.3334 | 5.0674 | 47.5866 | 0.33106 | 1.6557 | 2.278 | 6.381 | 2.6459 | 10.1091 | 13.6098 | 4.0695 | 2.3945 | 7.3527 | 11.2535 | 27.807 | 7.2458 | 16.8856 | 6.8437 | 3.238 | 3.5737 | 0.863 | 160.8094 | 3.435 | 0.35562 | 0.077066 | 0.24224 | 0.51088 | 0.34652 | 0.091454 | 7.1294 | 7.1855 | 140896.1143 | 151051.0732 | 9675.2863 | 168924.2849 | 142293.1523 | 135144.2426 | 3227.0312 | 2.9535 | 6.9896 | 12.8219 | 24.9735
227 | Calibration Curves |  | 06/09/2023 12:37:44
228 | LBLANK |  | 06/09/2023 12:43:32 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.581 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.00031865 | -3.7312e-08 | -0.010163 | -0.00032331 | 0.005963 | 0.0054983 | 0.00020669 | 0 | 0.00042273 | 0.00047626 | 7.623e-05 | -8.6853e-05 | 6.3866e-05 | 0.015567 | 0.00074258 | 0.00046861 | 0.0011129 | 7.9415e-05 | -0.0075469 | -0.0025149 | 0.00010619 | 0.0085683 | 0.00019516 | 0.0011923 | 0 | 1.3415e-05 | 1.6773e-05 | -1.1849e-05 | 1.01e-05 | 0.0062062 | 0.00061618 | 164451.7702 | 198035.9042 | 12542.1702 | 236784.2628 | 192278.7257 | 181885.2717 | 4476.0347 | 1.9728e-05 | 6.9446e-05 | 6.9584e-05 | 4.8062e-07
229 | LBLANK |  | 06/09/2023 13:13:17 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.583 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.00092879 | -5.357e-05 | -0.014943 | 7.044e-05 | 0.0042248 | -0.010679 | 0.00029496 | 0 | 0.00022538 | 0.00035194 | 6.6093e-05 | 0.00013404 | 1.4077e-05 | -4.9461e-06 | 0.00028436 | 0.00019063 | 0.00077112 | 8.1383e-05 | -0.0022697 | -0.00012569 | 0.00011053 | 0.0023818 | 5.3499e-05 | 0.00052717 | 7.0664e-06 | 1.4351e-05 | -9.9496e-06 | -4.8383e-05 | 0 | 0.0031421 | 0.00033161 | 167029.6465 | 189368.1968 | 12176.526 | 218481.4505 | 183016.9291 | 163111.2583 | 4249.966 | 6.0699e-06 | 2.0039e-05 | 6.7077e-06 | -6.4163e-05
230 | LBLANK |  | 06/09/2023 13:19:06 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.584 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.00022094 | -5.357e-05 | -0.047015 | -0.00047538 | 0.0015515 | 0.0045238 | 0.00030198 | 0 | 5.5783e-05 | -0.00025878 | 2.6102e-05 | -0.00023658 | -7.0185e-08 | 0.0031662 | 0.00016728 | 0.00018274 | 0.00040209 | 1.0628e-05 | -0.0078546 | 0.00059006 | 0 | 0.0016168 | 3.5922e-05 | 3.294e-05 | 0 | 1.4412e-05 | 4.1198e-06 | -0.00016888 | 0 | 0.0028377 | 8.8944e-05 | 175238.3405 | 189243.8936 | 12108.4649 | 223811.9604 | 186084.1879 | 145761.3934 | 4190.6146 | 2.9456e-06 | 1.2645e-05 | 1.2046e-05 | -0.00011101
231 | ICV PE |  | 06/09/2023 13:24:53 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\ICV PE.585 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.2754 | 0.51879 | 3.3782 | 0.033148 | 0.30974 | 0.21195 | 0.58328 | 0.25207 | 0.92571 | 1.2763 | 0.35878 | 0.24871 | 0.68227 | 1.1464 | 2.7211 | 0.74105 | 1.7289 | 0.67511 | 0.30306 | 0.32169 | 0.074049 | 14.764 | 0.32375 | 0.023576 | 0.0070692 | 0.022352 | 0.046584 | 0.031154 | 0.0085105 | 0.61632 | 0.63386 | 169053.537 | 180471.6125 | 10930.1849 | 209168.8781 | 171884.5702 | 157104.4866 | 3958.5486 | 0.2896 | 0.68862 | 1.365 | 2.7104
232 | ICV PE |  | 06/09/2023 13:31:10 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\ICV PE.586 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.241 | 0.52194 | 3.4768 | 0.032372 | 0.30919 | 0.19772 | 0.56883 | 0.25322 | 0.89775 | 1.2521 | 0.41545 | 0.28698 | 0.68494 | 1.1235 | 2.6275 | 0.73476 | 1.7204 | 0.78386 | 0.35696 | 0.32654 | 0.072569 | 14.8522 | 0.3258 | 0.024249 | 0.0070471 | 0.021885 | 0.047061 | 0.031205 | 0.0084854 | 0.7096 | 0.6357 | 167031.7105 | 179179.4538 | 11273.1152 | 190352.0582 | 178124.3157 | 158277.3523 | 3931.8745 | 0.33534 | 0.79538 | 1.3917 | 2.6595
233 | CCV 1ppb TRACE |  | 06/09/2023 13:36:58 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.587 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.12116 | 0.049028 | 0.62218 | 0.0056806 | 0.018742 | 0.018229 | 0.058024 | 0.023246 | 0.088823 | 0.12003 | 0.036999 | 0.027208 | 0.068291 | 0.11965 | 0.25034 | 0.068078 | 0.17416 | 0.074258 | 0.025492 | 0.031708 | 0.0081164 | 1.4604 | 0.032499 | 0.0038062 | 0.00071601 | 0.0020754 | 0.0047253 | 0.0027887 | 0.00094199 | 0.06647 | 0.066235 | 174867.7828 | 191773.9823 | 12717.0281 | 221887.507 | 182124.9223 | 165839.0968 | 4267.9709 | 0.030348 | 0.071637 | 0.14662 | 0.29033
234 | CCV 1ppb TRACE |  | 06/09/2023 13:43:15 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.588 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.12506 | 0.050575 | 0.60259 | 0.0058411 | 0.016358 | 0.021181 | 0.059427 | 0.024438 | 0.092961 | 0.12575 | 0.037886 | 0.027997 | 0.070186 | 0.13616 | 0.26007 | 0.069987 | 0.18108 | 0.073634 | 0.025402 | 0.038078 | 0.0085444 | 1.4564 | 0.032275 | 0.0036376 | 0.00071144 | 0.002324 | 0.0049545 | 0.0033206 | 0.00083597 | 0.067548 | 0.067988 | 172274.1229 | 187383.6332 | 12250.5853 | 216997.7681 | 179712.6665 | 164046.5937 | 4152.6035 | 0.031011 | 0.072655 | 0.14371 | 0.28646
235 | LFB |  | 06/09/2023 13:49:03 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.589 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.1677 | 0.46751 | 3.2963 | 0.029895 | 0.18656 | 0.19809 | 0.53406 | 0.22818 | 0.83559 | 1.1805 | 0.35257 | 0.23674 | 0.6639 | 1.0229 | 2.4481 | 0.67364 | 1.6213 | 0.67997 | 0.31333 | 0.3287 | 0.081461 | 14.3672 | 0.32204 | 0.033663 | 0.0073541 | 0.022078 | 0.047342 | 0.031806 | 0.0083482 | 0.63371 | 0.64231 | 165891.7449 | 176375.5944 | 11571.3517 | 207108.501 | 170242.4573 | 159151.6726 | 3954.5475 | 0.29054 | 0.69035 | 1.3256 | 2.6081
236 | LLOQ 0.01 |  | 06/09/2023 13:54:51 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LLOQ 0.01.590 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.0010371 | 0.00051638 | 0.14982 | 0.0015415 | 0.020615 | -0.0052623 | 0.0011153 | 0.00025833 | 0.0017138 | 0.0012332 | 0.00037419 | 0.0036712 | 0.002445 | 0.012289 | 0.0021774 | 0.00070404 | 0.0031555 | 0.0012558 | -0.0024881 | 0.0027503 | 5.2297e-05 | 0.017083 | 0.00037279 | 0.00029334 | 1.0565e-05 | 2.8342e-05 | 2.1942e-05 | -0.00013384 | 1.0607e-05 | 0.0027325 | 0.0014246 | 172294.9636 | 188598.3873 | 12837.0989 | 224144.3613 | 182511.9433 | 165423.6253 | 4112.5926 | 0.00033938 | 0.00065861 | 0.0017399 | 0.0032395
237 | CCV 10 ppb MAJORS |  | 06/09/2023 14:00:40 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.591 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.269 | 0.49981 | 3.2522 | 0.028618 | 0.15873 | 0.21558 | 0.58231 | 0.2434 | 0.91437 | 1.2711 | 0.34615 | 0.23175 | 0.69271 | 1.114 | 2.6816 | 0.73198 | 1.7245 | 0.65273 | 0.29483 | 0.33914 | 0.088175 | 15.2218 | 0.32842 | 0.033948 | 0.0070332 | 0.021721 | 0.047009 | 0.031789 | 0.0082202 | 0.6154 | 0.67076 | 164650.1675 | 174070.1064 | 11066.2875 | 216230.2278 | 167364.7994 | 155332.7551 | 3756.494 | 0.28049 | 0.66969 | 1.3925 | 2.7309
238 | CCV 10 ppb MAJORS |  | 06/09/2023 14:06:57 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.592 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.2133 | 0.47827 | 3.3032 | 0.029346 | 0.15492 | 0.19804 | 0.56672 | 0.23439 | 0.88149 | 1.2107 | 0.37281 | 0.25202 | 0.69913 | 1.0754 | 2.6081 | 0.70604 | 1.6645 | 0.69931 | 0.31301 | 0.3395 | 0.08186 | 14.934 | 0.32842 | 0.035271 | 0.0070162 | 0.022687 | 0.046912 | 0.031451 | 0.0083934 | 0.66239 | 0.67721 | 164944.1004 | 174527.6692 | 11487.2838 | 204894.916 | 169757.3701 | 156259.9837 | 3838.5157 | 0.30033 | 0.70846 | 1.3976 | 2.7084
239 | LLOQ 0.1 |  | 06/09/2023 14:12:45 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LLOQ 0.1.593 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.013126 | 0.0056345 | 0.062777 | 0.00083461 | 0.0091107 | -8.5655e-05 | 0.0062449 | 0.0019639 | 0.008858 | 0.012973 | 0.0037576 | 0.0081938 | 0.0072866 | 0.049907 | 0.02937 | 0.0069525 | 0.021343 | 0.0079437 | 0.011432 | 0.0074686 | 0.0010959 | 0.1518 | 0.0033343 | 0.00060724 | 0.00010203 | 0.00020087 | 0.00048644 | 0.00020698 | 9.513e-05 | 0.0076752 | 0.0066168 | 172099.3916 | 189418.2249 | 12191.2001 | 220946.8479 | 183963.0479 | 167602.804 | 4159.939 | 0.0029509 | 0.0069931 | 0.014674 | 0.027801
240 | 8_30_21_F_9 |  | 06/09/2023 14:19:27 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_F_9.594 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.90304 | 0.32309 | 0.62148 | 0.0069776 | 0.070972 | 0.27173 | 0.092466 | 0.7407 | 0.062519 | 0.0074325 | 0.0021983 | 0.10623 | 0.44095 | 0.50076 | 0.0041133 | 0.011043 | 0.43684 | 0.18577 | 0.024151 | 0.14611 | 0.03006 | 6.4345 | 0.13796 | 0.001214 | 3.5106e-05 | 0.00016085 | 0.00037645 | 0.001274 | 8.1874e-05 | 0.33433 | 0.34423 | 157654.5672 | 170918.3974 | 11379.8642 | 198497.5764 | 165564.5667 | 153600.8079 | 3666.4706 | 0.00033594 | 0.00080448 | 0.013728 | 0.02659
241 | 8_30_21_F_10 |  | 06/09/2023 14:25:14 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_F_10.595 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.93163 | 0.41156 | 1.1211 | 0.010811 | 0.10675 | 0.32147 | 0.11448 | 0.87096 | 0.070877 | 0.0091682 | 0.0026687 | 0.10249 | 0.3247 | 0.5013 | 0.0040625 | 0.011629 | 0.36613 | 0.14761 | 0.069086 | 0.15542 | 0.03259 | 6.7633 | 0.14776 | 0.00099297 | 7.0307e-05 | 0.00018881 | 0.00042407 | 0.0014916 | 6.2571e-05 | 0.2141 | 0.22731 | 154888.495 | 170596.5753 | 10997.5684 | 199749.3069 | 159227.9494 | 154807.6712 | 3727.8197 | 0.00029885 | 0.0007118 | 0.015244 | 0.0307
242 | 8_30_21_F_11 |  | 06/09/2023 14:31:03 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_F_11.596 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.92222 | 0.44185 | 1.1474 | 0.010171 | 0.075931 | 0.31908 | 0.11123 | 0.83728 | 0.07798 | 0.0093222 | 0.0025819 | 0.089297 | 0.2559 | 0.45299 | 0.0044644 | 0.012398 | 0.36149 | 0.1478 | 0.042384 | 0.15939 | 0.034313 | 6.7156 | 0.14959 | 0.0017698 | 8.9145e-05 | 0.00018705 | 0.00052884 | 0.0015204 | 9.3119e-05 | 0.20688 | 0.21371 | 156392.0727 | 172010.0595 | 10974.2145 | 196433.4078 | 174833.348 | 150443.4887 | 3832.5142 | 0.00025457 | 0.00059899 | 0.015763 | 0.028075
243 | 8_30_21_UF_1 |  | 06/09/2023 14:36:51 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_UF_1.597 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.0195 | 0.28008 | 0.91988 | 0.0084882 | 0.33475 | 0.39946 | 0.10861 | 0.81485 | 0.033322 | 0.0075118 | 0.0023018 | 0.01029 | 0.70915 | 0.089005 | 0.0063742 | 0.025931 | 0.10714 | 0.044748 | 0.074322 | 0.15685 | 0.034529 | 6.7774 | 0.14483 | 0.0019243 | 2.7193e-05 | 0.00010988 | 0.00016518 | 0.00099699 | 3.1005e-05 | 0.88141 | 0.91689 | 154904.7603 | 171594.4283 | 10861.4699 | 198745.7775 | 169308.8116 | 156555.9593 | 3670.472 | 0.00021828 | 0.00047473 | 0.011875 | 0.021135
244 | 8_30_21_UF_6 |  | 06/09/2023 14:42:39 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_UF_6.598 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.1673 | 0.33323 | 0.53163 | 0.0038942 | 0.26336 | 0.34108 | 0.10785 | 0.83805 | 0.056229 | 0.0074899 | 0.0021227 | 0.017364 | 0.73298 | 0.12268 | 0.0058996 | 0.017822 | 0.29459 | 0.11494 | 0.041227 | 0.16241 | 0.033695 | 7.3694 | 0.1587 | 0.0018381 | 6.5425e-05 | 0.00012587 | 0.00048 | 0.0012647 | 5.5728e-05 | 0.50561 | 0.55075 | 157758.9775 | 169231.0863 | 10898.1556 | 204934.1458 | 162283.2728 | 147752.8932 | 3660.4744 | 0.00041949 | 0.0010955 | 0.012492 | 0.025463
245 | 8_30_21_UF_8 |  | 06/09/2023 14:48:27 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_UF_8.599 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.1687 | 0.34133 | 0.73133 | 0.0070463 | 0.28875 | 0.32087 | 0.11372 | 0.80246 | 0.071918 | 0.0092367 | 0.0019882 | 0.009247 | 0.76339 | 0.079541 | 0.0049511 | 0.013503 | 0.31143 | 0.12771 | 0.024185 | 0.16757 | 0.037581 | 7.4285 | 0.1609 | 0.0018688 | 9.3464e-05 | 0.00021927 | 0.00035257 | 0.0013816 | 6.2294e-05 | 0.44864 | 0.46156 | 157932.9208 | 171141.7007 | 11002.236 | 196669.8639 | 165470.8352 | 153854.8571 | 3707.1476 | 0.00042728 | 0.00089661 | 0.0094724 | 0.017627
246 | 8_30_21_UF_9 |  | 06/09/2023 14:54:15 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_UF_9.600 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.0299 | 0.31345 | 0.30616 | 0.0027338 | 0.26652 | 0.25647 | 0.092712 | 0.72441 | 0.061245 | 0.0074432 | 0.0019585 | 0.0082265 | 0.41673 | 0.080307 | 0.0047332 | 0.013013 | 0.2834 | 0.11564 | 0.099171 | 0.13466 | 0.034525 | 5.9488 | 0.13151 | 0.00095787 | 4.7236e-05 | 0.00015373 | 0.00032415 | 0.0011872 | 5.1015e-05 | 0.3103 | 0.32038 | 159191.3271 | 169927.184 | 10757.3901 | 195032.3293 | 162931.215 | 159564.9216 | 3755.8271 | 0.00029396 | 0.00061799 | 0.009321 | 0.018641
247 | 8_30_21_UF_9_LD |  | 06/09/2023 15:00:03 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_UF_9_LD.601 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.98398 | 0.31301 | 0.5927 | 0.0062611 | 0.32281 | 0.25731 | 0.091407 | 0.71985 | 0.060678 | 0.006309 | 0.0020183 | 0.01068 | 0.40481 | 0.075193 | 0.0041926 | 0.011167 | 0.26476 | 0.12086 | 0.046392 | 0.13329 | 0.027507 | 5.9833 | 0.12836 | 0.00093429 | 5.4373e-05 | 0.00014849 | 0.00042465 | 0.00089245 | 5.0489e-05 | 0.32582 | 0.31159 | 158190.0522 | 171767.1448 | 10917.5051 | 185790.9352 | 161741.1304 | 144922.1089 | 3685.142 | 0.00031293 | 0.0006528 | 0.010189 | 0.019761
248 | 8_30_21_UF_9_LD |  | 06/09/2023 15:06:19 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_UF_9_LD.602 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.97584 | 0.31453 | 0.59621 | 0.0059377 | 0.29819 | 0.25239 | 0.092389 | 0.70465 | 0.06267 | 0.0075257 | 0.0021155 | 0.0097234 | 0.409 | 0.074216 | 0.0027324 | 0.012768 | 0.27003 | 0.11403 | 0.026404 | 0.13137 | 0.030675 | 5.8495 | 0.1308 | 0.0013124 | 6.6764e-05 | 0.00017393 | 0.000348 | 0.00079612 | 9.8194e-05 | 0.30554 | 0.31623 | 157726.0398 | 169691.8891 | 10998.236 | 194799.3252 | 160909.8404 | 150391.517 | 3794.5039 | 0.00026353 | 0.00072529 | 0.0098835 | 0.019241
249 | 8_30_21_UF_10 |  | 06/09/2023 15:12:07 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_UF_10.603 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.0382 | 0.40779 | 0.91894 | 0.0089695 | 0.30569 | 0.29971 | 0.10996 | 0.85526 | 0.070028 | 0.0087013 | 0.002396 | 0.031568 | 0.29093 | 0.17382 | 0.0051539 | 0.01082 | 0.23098 | 0.098109 | 0.040069 | 0.14755 | 0.03046 | 6.4908 | 0.14149 | 0.0012816 | 5.5429e-05 | 0.00014396 | 0.00047372 | 0.0013921 | 9.1036e-05 | 0.20878 | 0.21631 | 159257.4515 | 168409.5991 | 10912.8338 | 193455.1561 | 162175.4354 | 150978.279 | 3672.4722 | 0.00031016 | 0.00067473 | 0.012333 | 0.023693
250 | LBLANK |  | 06/09/2023 15:17:56 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.604 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | -2.586e-06 | 4.7523e-07 | -0.039135 | -0.00048436 | 0.0091634 | -0.0015784 | -0.00015623 | 0.00010702 | 0 | -9.995e-05 | 6.0321e-05 | -4.1842e-05 | 2.7256e-06 | -0.0012226 | -5.369e-05 | 0.00012185 | 0.0001545 | -2.0219e-05 | -0.0081619 | 0.0036763 | 0.0002171 | 0.02389 | 0.00053411 | 0.00069806 | 7.2887e-06 | -3.1086e-06 | 4.7274e-06 | -5.6269e-05 | 3.6443e-06 | 0.00011853 | 0.00011102 | 159256.6936 | 182698.1831 | 12436.0856 | 213509.0207 | 176272.9041 | 158008.4244 | 4083.9173 | 3.1322e-06 | -3.7671e-06 | 1.6974e-05 | -5.1384e-05
251 | CCV 1ppb TRACE |  | 06/09/2023 15:23:44 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.605 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.12341 | 0.052686 | 0.61841 | 0.005737 | 0.018347 | 0.02194 | 0.057826 | 0.025372 | 0.088165 | 0.12976 | 0.03736 | 0.028478 | 0.068512 | 0.12376 | 0.25848 | 0.074456 | 0.17901 | 0.073506 | 0.02475 | 0.037112 | 0.0090791 | 1.4565 | 0.032857 | 0.0038039 | 0.0007068 | 0.0020765 | 0.0046284 | 0.0034573 | 0.00083902 | 0.065761 | 0.067467 | 167872.7012 | 181138.514 | 11538.6623 | 207255.5188 | 160040.6706 | 157844.5907 | 4085.9176 | 0.030459 | 0.072661 | 0.15125 | 0.32423
252 | CCV 1ppb TRACE |  | 06/09/2023 15:30:02 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.606 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.12569 | 0.054589 | 0.6174 | 0.0042837 | 0.016621 | 0.019975 | 0.058353 | 0.025573 | 0.090744 | 0.1224 | 0.037577 | 0.029537 | 0.069611 | 0.12237 | 0.2607 | 0.071595 | 0.17639 | 0.075429 | 0.025712 | 0.036637 | 0.0082174 | 1.5372 | 0.033058 | 0.0034952 | 0.00081602 | 0.0023583 | 0.0049356 | 0.002894 | 0.00072541 | 0.066702 | 0.068724 | 165207.267 | 176467.9117 | 11519.3092 | 202915.2345 | 169456.4926 | 155932.5772 | 3795.8378 | 0.031561 | 0.075386 | 0.15204 | 0.29553
253 | CCV 10 ppb MAJORS |  | 06/09/2023 15:35:51 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.607 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.2357 | 0.49503 | 3.3303 | 0.033077 | 0.15208 | 0.20462 | 0.56947 | 0.24267 | 0.90427 | 1.2208 | 0.36964 | 0.24834 | 0.79194 | 1.053 | 2.5389 | 0.69595 | 1.6601 | 0.69608 | 0.31393 | 0.33847 | 0.081962 | 16.9617 | 0.3766 | 0.034341 | 0.0080052 | 0.025301 | 0.054408 | 0.031787 | 0.0098989 | 0.66245 | 0.77411 | 160752.0881 | 151558.9081 | 10878.8077 | 195473.6895 | 163351.9898 | 148964.2488 | 3369.7335 | 0.30474 | 0.72602 | 1.449 | 2.8403
254 | CCV 10 ppb MAJORS |  | 06/09/2023 15:42:08 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.608 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.2334 | 0.49178 | 3.2388 | 0.032867 | 0.11223 | 0.21058 | 0.56313 | 0.24926 | 0.89419 | 1.238 | 0.36417 | 0.24802 | 0.70054 | 1.0872 | 2.5764 | 0.7179 | 1.6772 | 0.68529 | 0.3078 | 0.32661 | 0.081843 | 15.1304 | 0.3276 | 0.034045 | 0.0071666 | 0.022283 | 0.04771 | 0.030397 | 0.0083416 | 0.65221 | 0.67763 | 166570.5814 | 171940.575 | 11047.6038 | 203038.0168 | 165543.1178 | 195893.0789 | 3723.1518 | 0.29934 | 0.71045 | 1.4164 | 2.8259
255 | 8_30_21_UF_10_LSP |  | 06/09/2023 15:47:56 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_UF_10_LSP.609 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.1313 | 0.84574 | 4.3165 | 0.038911 | 0.39603 | 0.47966 | 0.67367 | 1.0079 | 0.94137 | 1.2054 | 0.3687 | 0.27372 | 0.89062 | 1.2324 | 2.5065 | 0.68835 | 1.7818 | 0.75384 | 0.34729 | 0.46286 | 0.11129 | 19.5116 | 0.42437 | 0.035 | 0.0066165 | 0.021259 | 0.044839 | 0.032697 | 0.0081602 | 0.84003 | 0.81349 | 159485.9182 | 174468.6768 | 10783.4027 | 190077.5517 | 162659.7979 | 147925.2571 | 3797.1717 | 0.29517 | 0.69403 | 1.3848 | 2.6835
256 | 8_30_21_UF_10_DUP_SP |  | 06/09/2023 15:53:45 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_UF_10_DUP_SP.610 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.731 | 0.67738 | 2.1476 | 0.019445 | 0.41262 | 0.42864 | 0.43223 | 1.0392 | 0.55578 | 0.67076 | 0.20999 | 0.1463 | 0.67222 | 0.65186 | 1.3851 | 0.39117 | 1.1074 | 0.47525 | 0.21326 | 0.32863 | 0.07994 | 14.5278 | 0.31798 | 0.020126 | 0.003857 | 0.012111 | 0.025634 | 0.0179 | 0.0047768 | 0.58327 | 0.57651 | 154990.2497 | 164720.891 | 10662.6449 | 183944.5001 | 155834.0925 | 147640.3538 | 3605.7884 | 0.16359 | 0.39276 | 0.76275 | 1.5263
257 | 8_30_21_UF_11 |  | 06/09/2023 15:59:33 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_UF_11.611 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.9882 | 0.40877 | 0.44141 | 0.004078 | 0.2933 | 0.30674 | 0.1122 | 0.80754 | 0.071777 | 0.008585 | 0.002154 | 0.0097129 | 0.22932 | 0.082619 | 0.0036885 | 0.013225 | 0.23881 | 0.099476 | 0.023519 | 0.1526 | 0.035183 | 6.6982 | 0.14567 | 0.0013906 | 6.8382e-05 | 0.00018629 | 0.00046524 | 0.00090589 | 8.4502e-05 | 0.19808 | 0.20621 | 154873.2797 | 165714.27 | 10896.8211 | 190776.7468 | 159561.5185 | 147525.8433 | 3605.7885 | 0.00027253 | 0.00043993 | 0.0092755 | 0.018422
258 | 8_30_21_BB |  | 06/09/2023 16:05:21 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_BB.612 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.00035252 | 6.2299e-05 | 1.3334 | 0.011444 | 0.015216 | 0.0025628 | 0.00099594 | 0.00017413 | 0.00034772 | 0.003472 | 0.0008451 | 0.029388 | 0.00076902 | 0.15358 | 7.245e-05 | 0.00036903 | 0.019152 | 0.0070883 | -0.0031147 | 0.0031635 | 0.00017424 | 0.017053 | 0.00036799 | 0.00017988 | 0 | 3.0241e-05 | 5.0775e-06 | -0.00010098 | 0 | 0.00021556 | 0.00017573 | 164925.9305 | 180325.0025 | 11478.6099 | 218546.5256 | 172215.8312 | 158356.8941 | 3891.8634 | 6.4732e-06 | 3.0403e-05 | 0.0055131 | 0.010713
259 | 8_30_21_FB |  | 06/09/2023 16:11:09 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_FB.613 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.0035772 | 0.00053196 | 0.73924 | 0.0082073 | 0.010586 | 0.0020206 | 0.00048547 | 0.00093736 | 0.00023456 | 0.019795 | 0.0055519 | 0.016391 | 0.0091555 | 0.073082 | 0.0015962 | 0.6569 | 0.02272 | 0.0087174 | 0.17458 | 0.003061 | 0 | 0.014593 | 0.00032267 | 0.00037615 | 7.5346e-06 | 1.206e-05 | 1.6741e-05 | -0.00017901 | 3.8027e-06 | 0.00023176 | 0.00019091 | 162856.5311 | 176856.062 | 11381.1985 | 207952.0766 | 168893.8634 | 151726.0966 | 3905.8674 | 4.1716e-05 | 8.0398e-05 | 0.0099005 | 0.019241
260 | 8_30_21_LAB_A |  | 06/09/2023 16:16:56 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_LAB_A.614 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.00012211 | -5.357e-05 | 0.12185 | 0.0015786 | 0.004443 | 0.00097632 | 0.00019154 | 0 | 0.00029138 | 0.0090171 | 0.0029694 | 0.014349 | 0.00052711 | 0.065655 | 1.3124e-05 | 0.019673 | 0.018277 | 0.0072249 | -0.0028893 | 0.0030952 | 0 | 0.0090601 | 0.00019716 | 0.00057805 | 0 | 9.6792e-07 | -9.6387e-06 | -0.00014459 | 0 | 0.0014965 | 0.0015369 | 159584.5806 | 174301.6427 | 11404.551 | 207289.8331 | 169534.3961 | 152347.6691 | 3800.5055 | 6.4337e-06 | 6.6833e-06 | 0.0053894 | 0.0098969
261 | 8_30_21_LAB_B |  | 06/09/2023 16:22:44 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_LAB_B.615 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.00023939 | -5.357e-05 | 0.13386 | 0.00093686 | 0.0047776 | -0.0012067 | 0.00042156 | 0.00017635 | 0 | 0.0030816 | 0.0011179 | 0.01005 | 0.00027739 | 0.037685 | -4.5506e-05 | 0.018381 | 0.014347 | 0.0069634 | -0.003982 | 0.0026133 | 5.932e-05 | 0.0068024 | 0.00014535 | 0.00039289 | 0 | 4.7029e-06 | -5.9561e-06 | -0.00017916 | 3.8205e-06 | 5.4515e-05 | 5.118e-05 | 162564.5501 | 176420.8721 | 11461.9305 | 193529.9419 | 168388.6056 | 152089.3451 | 3783.1681 | 0 | -1.0839e-06 | 0.0053408 | 0.0095849
262 | 9_13_21_W_1 |  | 06/09/2023 16:28:32 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_W_1.616 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.17267 | 0.056037 | 2.5939 | 0.027984 | 0.057744 | 0.08288 | 0.024248 | 0.1656 | 0.0049597 | 0.0039401 | 0.0013543 | 0.18194 | 0.29717 | 0.79905 | 0.0021878 | 0.0046567 | 0.057131 | 0.023496 | -0.0007399 | 0.030111 | 0.0057388 | 1.2093 | 0.026738 | 0.00036987 | 1.9456e-05 | 2.0514e-05 | 4.1051e-05 | 4.3653e-05 | 1.17e-05 | 0.20609 | 0.20808 | 163432.384 | 171261.7893 | 11281.7916 | 197333.7035 | 169715.5045 | 147401.2447 | 3789.8364 | 7.1233e-05 | 0.00015435 | 0.0056442 | 0.010258
263 | 9_13_21_W_6 |  | 06/09/2023 16:34:20 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_W_6.617 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.17896 | 0.063033 | 2.3602 | 0.022263 | 0.035128 | 0.059717 | 0.017485 | 0.15134 | 0.010506 | 0.0043679 | 0.0011595 | 0.19358 | 0.21462 | 0.8603 | 0.0013141 | 0.0038697 | 0.14211 | 0.057951 | 0.012546 | 0.030948 | 0.0069079 | 1.302 | 0.028172 | 0.00064511 | 2.3771e-05 | 5.2337e-05 | 6.1381e-05 | 0.0003457 | 2.7508e-05 | 0.075119 | 0.077118 | 159159.7309 | 169279.1047 | 11295.7974 | 196473.1867 | 162303.4709 | 148564.7522 | 3664.4702 | 0.0001525 | 0.00029376 | 0.0094325 | 0.018428
264 | 9_13_21_W_8 |  | 06/09/2023 16:40:07 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_W_8.618 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.1968 | 0.065485 | 2.4029 | 0.024969 | 0.042224 | 0.052446 | 0.021236 | 0.15778 | 0.014655 | 0.0078146 | 0.002827 | 0.23608 | 0.19189 | 0.8963 | 0.0016911 | 0.0076811 | 0.11602 | 0.058803 | 0.0073736 | 0.03487 | 0.0060426 | 1.3094 | 0.029115 | 0.0003303 | 1.5688e-05 | 5.2307e-05 | 9.312e-05 | 0.000182 | 3.5299e-05 | 0.075857 | 0.065563 | 158989.513 | 169779.954 | 10800.7485 | 176350.876 | 165772.9917 | 145308.7216 | 3776.4996 | 0.00012725 | 0.00028424 | 0.0077411 | 0.01496
265 | LBLANK |  | 06/09/2023 16:45:55 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.619 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.00017619 | -5.357e-05 | -0.053701 | -0.00045902 | 0.0029932 | -0.0019283 | 0.00013075 | 0 | 6.0156e-05 | 0.00012158 | 5.3543e-05 | -0.00073581 | -1.5406e-05 | -0.0021757 | -0.00010364 | -0.00021055 | 0.00016884 | -5.7859e-05 | -0.0084145 | 0.00403 | 0 | 0.0023487 | 5.2832e-05 | 0.00030662 | 0 | -2.8897e-06 | -1.7897e-06 | -0.00026619 | 0 | 7.2523e-05 | 7.297e-05 | 159532.534 | 171701.7426 | 11383.8694 | 205271.1165 | 161143.1292 | 150040.478 | 3829.8468 | 9.6542e-06 | 7.0519e-06 | -4.2348e-05 | -0.00016063
266 | CCV 1ppb TRACE |  | 06/09/2023 16:51:44 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.620 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.11885 | 0.049489 | 0.39554 | 0.0037067 | 0.014868 | 0.015679 | 0.055555 | 0.024651 | 0.083866 | 0.11734 | 0.035969 | 0.030271 | 0.064918 | 0.13646 | 0.24812 | 0.066728 | 0.16296 | 0.069193 | 0.023915 | 0.033906 | 0.008259 | 1.4252 | 0.032001 | 0.0032697 | 0.00074256 | 0.0020759 | 0.0048258 | 0.0025535 | 0.00082113 | 0.063448 | 0.065627 | 159453.0097 | 169633.8955 | 11212.4 | 199390.5577 | 162733.0951 | 148207.5225 | 3809.1745 | 0.029443 | 0.070221 | 0.14748 | 0.29337
267 | CCV 1ppb TRACE |  | 06/09/2023 16:58:02 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.621 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.11561 | 0.04612 | 0.38701 | 0.0040438 | 0.014126 | 0.017743 | 0.052712 | 0.025698 | 0.08763 | 0.11983 | 0.035827 | 0.030639 | 0.06389 | 0.13271 | 0.24533 | 0.065957 | 0.17153 | 0.068843 | 0.024111 | 0.035621 | 0.0067971 | 1.4133 | 0.031284 | 0.0034136 | 0.00065694 | 0.0020298 | 0.004889 | 0.0028939 | 0.00084774 | 0.063302 | 0.065606 | 161537.5976 | 171529.695 | 11388.538 | 198984.9287 | 165344.7391 | 146943.2919 | 3797.1713 | 0.0302 | 0.070776 | 0.14704 | 0.29279
268 | LFB |  | 06/09/2023 17:03:50 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.622 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.1353 | 0.45532 | 3.1909 | 0.031141 | 0.18911 | 0.18886 | 0.52713 | 0.23397 | 0.82474 | 1.1401 | 0.34532 | 0.23356 | 0.65491 | 1.0029 | 2.3884 | 0.64635 | 1.5537 | 0.65848 | 0.3034 | 0.32106 | 0.080803 | 14.4341 | 0.31748 | 0.03341 | 0.0071333 | 0.022164 | 0.04694 | 0.031282 | 0.0085804 | 0.62618 | 0.65199 | 160063.3364 | 162170.6476 | 10765.3885 | 192455.2681 | 160074.6067 | 141135.7754 | 3567.112 | 0.29421 | 0.69532 | 1.3981 | 2.7531
269 | CCV 10 ppb MAJORS |  | 06/09/2023 17:09:38 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.623 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.1863 | 0.47934 | 3.204 | 0.029809 | 0.14862 | 0.21268 | 0.56083 | 0.24293 | 0.888 | 1.1909 | 0.38863 | 0.26263 | 0.67493 | 1.055 | 2.5281 | 0.69813 | 1.6375 | 0.72835 | 0.34378 | 0.33222 | 0.083292 | 14.2478 | 0.31369 | 0.035487 | 0.0070628 | 0.022135 | 0.045856 | 0.030988 | 0.0083236 | 0.68984 | 0.65829 | 160314.4304 | 167684.7637 | 10935.5191 | 185986.6369 | 161683.2303 | 145919.4565 | 3688.4763 | 0.32015 | 0.75927 | 1.4411 | 2.8204
270 | CCV 10 ppb MAJORS |  | 06/09/2023 17:15:55 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.624 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.182 | 0.46593 | 3.1248 | 0.030398 | 0.14978 | 0.19011 | 0.55248 | 0.24792 | 0.84506 | 1.1545 | 0.36032 | 0.24434 | 0.69121 | 1.0413 | 2.479 | 0.68617 | 1.6152 | 0.67899 | 0.32139 | 0.32538 | 0.080925 | 14.4253 | 0.32263 | 0.033254 | 0.0072064 | 0.021837 | 0.046519 | 0.030427 | 0.0084512 | 0.64665 | 0.66925 | 161575.319 | 166616.7156 | 11254.4316 | 196405.5933 | 162412.0151 | 148356.5541 | 3726.486 | 0.29552 | 0.70391 | 1.4028 | 2.7531
271 | 9_13_21_W_9 |  | 06/09/2023 17:21:43 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_W_9.625 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.17531 | 0.064632 | 1.8813 | 0.016965 | 0.035453 | 0.050905 | 0.017289 | 0.14298 | 0.012028 | 0.0027218 | 0.00079745 | 0.1183 | 0.11737 | 0.5279 | 0.001238 | 0.0023466 | 0.098814 | 0.040843 | -0.00074962 | 0.030837 | 0.0061005 | 1.1475 | 0.025351 | 0.0011955 | 1.5646e-05 | 3.2352e-05 | 0.00010406 | 6.1e-05 | 2.7426e-05 | 0.049203 | 0.050596 | 149125.9397 | 170271.4113 | 11461.9301 | 199969.2262 | 165669.8014 | 150075.5308 | 3761.8287 | 6.6708e-05 | 0.0001714 | 0.0055208 | 0.010078
272 | 9_13_21_W_10 |  | 06/09/2023 17:27:31 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_W_10.626 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.16316 | 0.080771 | 4.3223 | 0.041125 | 0.047149 | 0.052159 | 0.019027 | 0.15787 | 0.014362 | 0.0020552 | 0.00073203 | 0.25122 | 0.19374 | 1.0885 | 0.0018222 | 0.0029952 | 0.10192 | 0.043984 | 0.0091852 | 0.0317 | 0.0057686 | 1.3355 | 0.029564 | 0.00060732 | 4.0632e-06 | 6.2234e-05 | 0.00010003 | 0.00018366 | 8.1265e-06 | 0.031963 | 0.032416 | 159383.5299 | 164357.4043 | 11094.3063 | 195769.6515 | 162580.6617 | 148448.6508 | 3639.7971 | 8.8552e-05 | 0.00023728 | 0.010566 | 0.019216
273 | 9_13_21_W_10_DUP |  | 06/09/2023 17:33:20 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_W_10_DUP.627 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.15721 | 0.082281 | 3.7289 | 0.033372 | 0.04816 | 0.058375 | 0.020667 | 0.1556 | 0.013586 | 0.0029636 | 0.0010855 | 0.21193 | 0.16972 | 0.9581 | 0.0019389 | 0.0023351 | 0.085569 | 0.039031 | 0.0075271 | 0.031469 | 0.005784 | 1.2864 | 0.02858 | 0.00061661 | 2.0104e-05 | 4.5636e-05 | 8.7604e-05 | 0.00032558 | 2.0214e-05 | 0.031278 | 0.032474 | 159025.5691 | 165309.0619 | 11086.9676 | 195164.7613 | 161241.1285 | 145331.5677 | 3673.1389 | 9.9075e-05 | 0.00021759 | 0.012275 | 0.023746
274 | 9_13_21_W_11 |  | 06/09/2023 17:39:08 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_W_11.628 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.15192 | 0.078706 | 8.5646 | 0.085907 | 0.048671 | 0.055145 | 0.018394 | 0.15431 | 0.013919 | 0.004684 | 0.0015429 | 0.43559 | 0.2431 | 1.9322 | 0.0021709 | 0.0032563 | 0.14102 | 0.057794 | 0.022613 | 0.029651 | 0.005953 | 1.2624 | 0.027597 | 0.00070554 | 1.6324e-05 | 5.0349e-05 | 9.6223e-05 | 2.0406e-05 | 1.6343e-05 | 0.028788 | 0.029458 | 159151.4813 | 163437.0557 | 10870.1342 | 191954.3842 | 160543.2201 | 143286.3335 | 3571.113 | 0.00014238 | 0.00032591 | 0.014135 | 0.026546
275 | 9_13_21_F_1 |  | 06/09/2023 17:44:56 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_F_1.629 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.88989 | 0.27571 | 1.4447 | 0.013007 | 0.14216 | 0.42242 | 0.10675 | 0.84796 | 0.027424 | 0.021292 | 0.0065765 | 0.11899 | 0.90009 | 0.55181 | 0.0084426 | 0.033594 | 0.13715 | 0.062534 | 0.047771 | 0.14251 | 0.028222 | 5.8703 | 0.12861 | 0.0020534 | 3.3042e-05 | 0.00011714 | 0.00020164 | 0.00092339 | 4.9532e-05 | 1.0313 | 1.074 | 145172.7875 | 161248.3629 | 10704.0089 | 190337.6361 | 158473.5079 | 144698.5206 | 3533.7706 | 0.00026619 | 0.00055287 | 0.018336 | 0.03348
276 | 9_13_21_F_6 |  | 06/09/2023 17:50:44 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_F_6.630 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.99119 | 0.35748 | 1.1323 | 0.010004 | 0.03866 | 0.32397 | 0.10635 | 0.88877 | 0.059766 | 0.01317 | 0.0032903 | 0.11349 | 0.82102 | 0.58078 | 0.0070688 | 0.028094 | 0.43116 | 0.17318 | 0.083022 | 0.1381 | 0.030521 | 5.8056 | 0.13222 | 0.0015895 | 6.2057e-05 | 0.00017818 | 0.00044742 | 0.0013843 | 7.3787e-05 | 0.38916 | 0.40594 | 157033.0306 | 162034.4927 | 10235.6745 | 191922.9625 | 157589.9569 | 145931.6365 | 3693.8113 | 0.00045415 | 0.0012076 | 0.020763 | 0.041704
277 | 9_13_21_F_8 |  | 06/09/2023 17:56:32 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_F_8.631 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.0662 | 0.33911 | 1.2546 | 0.011684 | 0.059149 | 0.30458 | 0.10584 | 0.84705 | 0.076546 | 0.0080858 | 0.0025393 | 0.13398 | 0.73521 | 0.6669 | 0.0058268 | 0.018556 | 0.46167 | 0.18719 | 0.062114 | 0.1521 | 0.035953 | 6.4803 | 0.14401 | 0.0018527 | 4.5034e-05 | 0.00016118 | 0.00036328 | 0.0014843 | 6.1472e-05 | 0.33128 | 0.33953 | 154232.8321 | 162778.1336 | 10241.009 | 188700.1472 | 156249.8939 | 145106.0021 | 3617.7915 | 0.00037105 | 0.0008867 | 0.028728 | 0.055687
278 | 9_13_21_F_9 |  | 06/09/2023 18:02:20 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_F_9.632 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.94255 | 0.33956 | 2.7248 | 0.025169 | 0.053578 | 0.2595 | 0.097775 | 0.76139 | 0.061978 | 0.0086205 | 0.0025233 | 0.13964 | 0.26681 | 0.66167 | 0.0044918 | 0.013055 | 0.38944 | 0.15815 | 0.16321 | 0.12793 | 0.026149 | 5.2835 | 0.11842 | 0.0016296 | 5.7789e-05 | 0.00016124 | 0.00041134 | 0.0011044 | 6.5873e-05 | 0.24333 | 0.25033 | 154150.5613 | 163531.4867 | 10322.4035 | 187335.2108 | 154546.8433 | 147849.22 | 3673.1401 | 0.00023552 | 0.00050577 | 0.016832 | 0.034113
279 | 9_13_21_F_10 |  | 06/09/2023 18:08:08 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_F_10.633 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.82285 | 0.40691 | 1.3398 | 0.013352 | 0.069107 | 0.28681 | 0.10374 | 0.82697 | 0.065919 | 0.0080701 | 0.0024491 | 0.10037 | 0.3136 | 0.45475 | 0.0041432 | 0.012605 | 0.29544 | 0.13202 | 0.054803 | 0.14242 | 0.031765 | 6.2091 | 0.13703 | 0.001432 | 5.4777e-05 | 0.00020435 | 0.00037391 | 0.0010702 | 6.7421e-05 | 0.16576 | 0.16465 | 150921.6929 | 158165.4578 | 10461.8318 | 179040.4429 | 152831.855 | 140991.68 | 3490.4264 | 0.00026418 | 0.00062568 | 0.017183 | 0.034532
280 | 9_13_21_F_11 |  | 06/09/2023 18:13:56 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_F_11.634 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.80142 | 0.40798 | 3.794 | 0.03674 | 0.053301 | 0.29882 | 0.10268 | 0.8239 | 0.066853 | 0.0082517 | 0.002426 | 0.11351 | 0.33517 | 0.49835 | 0.0036642 | 0.012252 | 0.35304 | 0.15773 | 0.11938 | 0.13783 | 0.029723 | 6.0733 | 0.13377 | 0.0014537 | 5.0642e-05 | 0.00018306 | 0.00037471 | 0.0013096 | 8.4339e-05 | 0.15592 | 0.1527 | 152608.998 | 158049.4801 | 10329.0684 | 175760.739 | 151378.2202 | 137645.4125 | 3483.7584 | 0.00030476 | 0.0006554 | 0.016271 | 0.032206
281 | 9_13_21_F_11_LD |  | 06/09/2023 18:19:44 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_F_11_LD.635 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.80884 | 0.4158 | 3.3918 | 0.030654 | 0.070172 | 0.29502 | 0.10468 | 0.8359 | 0.062135 | 0.0077525 | 0.0022262 | 0.10625 | 0.34114 | 0.49035 | 0.0045714 | 0.012014 | 0.35882 | 0.15352 | 0.11185 | 0.14159 | 0.028531 | 6.0972 | 0.13493 | 0.0015813 | 3.4107e-05 | 0.00019495 | 0.00040665 | 0.00094667 | 5.9842e-05 | 0.14958 | 0.15401 | 148609.2518 | 155475.038 | 10319.0652 | 181123.636 | 153391.8415 | 138382.4885 | 3439.7474 | 0.00027606 | 0.00072269 | 0.016527 | 0.031788
282 | 9_13_21_F_11_LD |  | 06/09/2023 18:26:00 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_F_11_LD.636 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.78591 | 0.40413 | 3.312 | 0.031983 | 0.071314 | 0.28105 | 0.10322 | 0.80406 | 0.061797 | 0.0082662 | 0.002077 | 0.10391 | 0.33858 | 0.47437 | 0.0042436 | 0.013356 | 0.34819 | 0.14966 | 0.10956 | 0.13801 | 0.029408 | 6.1184 | 0.13238 | 0.0014679 | 7.1839e-05 | 0.00020024 | 0.00040485 | 0.0012444 | 6.3361e-05 | 0.14582 | 0.15266 | 151939.8729 | 157798.0554 | 10608.6044 | 182867.7624 | 152827.19 | 141087.3334 | 3419.7432 | 0.00031723 | 0.00056613 | 0.015716 | 0.031319
283 | LBLANK |  | 06/09/2023 18:31:48 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.637 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 6.9113e-05 | -5.357e-05 | -0.052495 | -0.00039447 | 0.0024317 | -0.003348 | 0.00059059 | 0 | 6.055e-05 | -0.00022601 | 4.031e-05 | -0.00055403 | -1.8973e-05 | -0.001877 | -0.00010016 | 0.00023808 | 0.00025378 | -4.0888e-07 | -0.008552 | 0.0022019 | 0 | 0.0010211 | 2.2313e-05 | 0.00048018 | 1.5824e-05 | 5.0353e-06 | -5.6304e-06 | 3.9247e-05 | 3.8963e-06 | 4.5457e-05 | 7.3543e-05 | 158194.0697 | 170708.6534 | 11012.2433 | 200547.8689 | 162999.4166 | 146209.1395 | 3677.807 | 3.3258e-06 | 4.7148e-06 | -1.6922e-05 | -0.00016224
284 | CCV 1ppb TRACE |  | 06/09/2023 18:37:37 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.638 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.11852 | 0.047704 | 0.40527 | 0.0038882 | 0.014795 | 0.018394 | 0.057006 | 0.022993 | 0.087284 | 0.12137 | 0.035907 | 0.030309 | 0.066169 | 0.13485 | 0.25056 | 0.069219 | 0.16134 | 0.068952 | 0.023044 | 0.033725 | 0.0076091 | 1.4779 | 0.031588 | 0.0038586 | 0.00071144 | 0.0022393 | 0.0045745 | 0.0029586 | 0.00086653 | 0.063305 | 0.066152 | 159893.1791 | 167710.0397 | 10877.4739 | 197601.9376 | 143975.8015 | 147160.7663 | 3587.7841 | 0.029244 | 0.069037 | 0.14996 | 0.33745
285 | CCV 1ppb TRACE |  | 06/09/2023 18:43:54 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.639 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.11705 | 0.046749 | 0.40313 | 0.0033026 | 0.013968 | 0.02369 | 0.059484 | 0.024111 | 0.088897 | 0.12324 | 0.03492 | 0.030762 | 0.066099 | 0.14543 | 0.25062 | 0.068591 | 0.17986 | 0.069098 | 0.023818 | 0.035187 | 0.0080953 | 1.3939 | 0.031595 | 0.0038382 | 0.00068277 | 0.0023126 | 0.0047649 | 0.0029065 | 0.00091057 | 0.062493 | 0.065005 | 159704.5068 | 166951.9846 | 10670.6531 | 198876.1248 | 146275.5679 | 146509.4461 | 3786.502 | 0.029525 | 0.069873 | 0.15038 | 0.34054
286 | CCV 10 ppb MAJORS |  | 06/09/2023 18:49:43 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.640 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.147 | 0.47761 | 3.1562 | 0.026287 | 0.14701 | 0.19717 | 0.54822 | 0.23882 | 0.85613 | 1.1679 | 0.3537 | 0.24124 | 0.73385 | 1.0169 | 2.4554 | 0.65327 | 1.5907 | 0.66734 | 0.31252 | 0.32635 | 0.078598 | 16.1096 | 0.35093 | 0.03363 | 0.0073736 | 0.023915 | 0.051256 | 0.030892 | 0.0090393 | 0.63332 | 0.72306 | 156223.1342 | 146720.6129 | 10525.2093 | 187645.0088 | 154943.0842 | 142979.2812 | 3202.3608 | 0.29299 | 0.70131 | 1.4485 | 2.8664
287 | CCV 10 ppb MAJORS |  | 06/09/2023 18:56:01 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.641 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.1472 | 0.46911 | 3.2202 | 0.029879 | 0.14548 | 0.19724 | 0.53298 | 0.23864 | 0.84191 | 1.1611 | 0.36636 | 0.24845 | 0.68989 | 1.0223 | 2.479 | 0.65305 | 1.5823 | 0.69187 | 0.32495 | 0.33664 | 0.075087 | 14.8187 | 0.32529 | 0.034263 | 0.0068501 | 0.021662 | 0.047801 | 0.031157 | 0.008317 | 0.65861 | 0.66907 | 155119.0767 | 159108.515 | 10870.1345 | 186237.0452 | 168712.2642 | 147164.8231 | 3492.4268 | 0.30029 | 0.72042 | 1.4513 | 2.5888
288 | 9_13_21_F_11_DUP |  | 06/09/2023 19:01:49 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_F_11_DUP.642 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.77359 | 0.39415 | 2.8048 | 0.027377 | 0.053679 | 0.28827 | 0.099023 | 0.82018 | 0.060378 | 0.01185 | 0.0040664 | 0.11262 | 0.33433 | 0.5347 | 0.0047302 | 0.011836 | 0.35704 | 0.15292 | 0.10489 | 0.1361 | 0.028746 | 6.0322 | 0.13199 | 0.0013257 | 4.721e-05 | 0.00016924 | 0.00041176 | 0.001389 | 9.8716e-05 | 0.14578 | 0.15154 | 153665.3086 | 155265.8744 | 10353.7508 | 180977.1401 | 150526.2305 | 138162.5291 | 3398.4044 | 0.00030943 | 0.00069024 | 0.017188 | 0.035208
289 | 9_13_21_F_11_DUP_LSP |  | 06/09/2023 19:07:37 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_F_11_DUP_LSP.643 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.8754 | 0.8255 | 6.0325 | 0.058725 | 0.21413 | 0.45216 | 0.63585 | 0.98104 | 0.92554 | 1.195 | 0.36242 | 0.3479 | 1.0545 | 1.5366 | 2.4897 | 0.674 | 1.8829 | 0.78303 | 0.39869 | 0.44001 | 0.1075 | 20.9554 | 0.47383 | 0.0338 | 0.0073383 | 0.023393 | 0.049628 | 0.030864 | 0.0084926 | 0.77332 | 0.86661 | 150813.302 | 143898.8762 | 10018.846 | 178887.201 | 145983.0033 | 133959.1477 | 3247.0358 | 0.28778 | 0.68834 | 1.4032 | 2.8322
290 | 9_13_21_UF_1 |  | 06/09/2023 19:13:24 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_UF_1.644 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.0301 | 0.27744 | 0.69091 | 0.0054519 | 0.12131 | 0.42709 | 0.108 | 0.81737 | 0.02664 | 0.0068653 | 0.0022536 | 0.022683 | 0.87565 | 0.13204 | 0.0076643 | 0.033826 | 0.094635 | 0.043784 | 0.035215 | 0.14518 | 0.027894 | 5.8173 | 0.12688 | 0.0022841 | 2.9984e-05 | 8.7372e-05 | 0.00020506 | 0.0011015 | 3.8208e-05 | 1.003 | 1.0348 | 136520.8102 | 155629.9396 | 10236.3341 | 180972.4119 | 150809.5251 | 148241.1838 | 3395.0702 | 0.00025418 | 0.00057623 | 0.0081293 | 0.014795
291 | 9_13_21_UF_6 |  | 06/09/2023 19:19:12 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_UF_6.645 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.0527 | 0.32084 | 0.23961 | 0.0017585 | 0.050749 | 0.29944 | 0.096312 | 0.81448 | 0.05648 | 0.0085749 | 0.0022089 | 0.014919 | 0.77427 | 0.10122 | 0.0046783 | 0.015483 | 0.27572 | 0.12234 | 0.050811 | 0.13643 | 0.029901 | 5.845 | 0.12697 | 0.0016724 | 5.787e-05 | 0.00018718 | 0.0003312 | 0.0010198 | 6.2859e-05 | 0.37891 | 0.37647 | 149410.6322 | 161346.5643 | 10325.0634 | 182078.9603 | 151179.1856 | 140666.5734 | 3509.0983 | 0.00047292 | 0.0011106 | 0.0099446 | 0.01998
292 | 9_13_21_UF_6_DUP |  | 06/09/2023 19:25:00 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_UF_6_DUP.646 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.0696 | 0.32828 | 0.22968 | 0.0018178 | 0.050995 | 0.30127 | 0.10018 | 0.80037 | 0.052701 | 0.011162 | 0.0030149 | 0.01727 | 0.91367 | 0.1152 | 0.0053234 | 0.016266 | 0.28352 | 0.11982 | 0.052614 | 0.13477 | 0.027483 | 6.8954 | 0.15093 | 0.001876 | 7.5797e-05 | 0.00023983 | 0.00044866 | 0.0010979 | 6.8385e-05 | 0.37165 | 0.4502 | 149244.2293 | 138862.6069 | 10328.3993 | 183446.3607 | 150067.5837 | 140708.2695 | 3029.6549 | 0.00047233 | 0.00099594 | 0.010073 | 0.020338
293 | 9_13_21_UF_8 |  | 06/09/2023 19:30:49 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_UF_8.647 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.126 | 0.3438 | 0.22048 | 0.0033333 | 0.058411 | 0.27745 | 0.10597 | 0.80352 | 0.070924 | 0.0069356 | 0.0018766 | 0.012595 | 0.70098 | 0.1049 | 0.0051512 | 0.015123 | 0.26391 | 0.1096 | 0.027527 | 0.14164 | 0.033035 | 6.5049 | 0.14113 | 0.0014226 | 5.958e-05 | 0.0001591 | 0.00040667 | 0.00099498 | 4.2459e-05 | 0.31969 | 0.3282 | 154403.409 | 157092.5138 | 9906.7677 | 180468.1254 | 151445.203 | 139541.7489 | 3407.7398 | 0.00031802 | 0.00065184 | 0.0079126 | 0.01671

## Net Intensities RSDs
| Sample Id | R | Acquisition Time | QC Status | Dataset File | Method File | Na KED 23/23 Helium KED_Low | Mg KED 24/24 Helium KED_Low | Al 27/27 | Al KED 27/27 Helium KED_Low | P Oshift 31/47 Oxygen DRC | S (34) -1 34/34 Helium KED_Low | K KED 39/39 Helium KED_Low | Ca KED 43/43 Helium KED_Low | V KED 51/51 Helium KED_Low | Cr (52) 52/52 Helium KED_Low | Cr (53) NH3 53/53 Ammonia DRC | Fe (54) NH3 54/54 Ammonia DRC | Mn Oshift 55/71 Oxygen DRC | Fe (56) KED 56/56 Helium KED_Low | Co KED 59/59 Helium KED_Low | Ni KED 60/60 Helium KED_Low | Cu (63) KED 63/63 Helium KED_Low | Cu (65) NH3 65/65 Ammonia DRC | Zn NH3 66/66 Ammonia DRC | As 75/75 | As KED 75/75 Helium KED_Low | As RH Oshift 75/91 Oxygen DRC | As Oshift 75/91 Oxygen DRC | Se (77) 77/77 | Se (77) Oshift 77/93 Oxygen DRC | Se (78) Oshift 78/94 Oxygen DRC | Se (80) Oshift 80/96 Oxygen DRC | Se (82) 82/82 | Se (82) Oshift 82/98 Oxygen DRC | Mo (95) NH3 95/95 Ammonia DRC | Mo (95) OShift 95/127 Oxygen DRC | Rh (IS1) 103/103 (IS) | Rh (IS2) 103/103 (IS) Oxygen DRC | Rh (IS3) 103/103 (IS) Helium KED_Low | Rh (IS4) 103/103 (IS) Ammonia DRC | Rh (IS5) 103/103 (IS) Ammonia DRC | Rh (IS6) 103/103 (IS) Oxygen DRC | Rh OShift (IS1) 103/119 (IS) Oxygen DRC | Cd (111) NH3 111/111 Ammonia DRC | Cd (114) NH3 114/114 Ammonia DRC | Pb 208/208 | Pb NH3 208/208 Ammonia DRC
1 | Sac Blank |  | 6/8/2023 1:23:24 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\Sac Blank.356 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth |  |  | 2.382% | 113.7% | 2.176% | 6.414% | 42.74% |  | 86.62% | 51.43% | 19.39% | 6.626% | 6.693% | 5.133% | 98.13% | 32.18% | 25.92% | 20.35% | 4.915% | 4.667% |  | 9.741% | 9.015% | 5.445% | 86.61% | 0.6749% | 62.44% | 33.14% | 0.6749% | 42.84% | 53.63% | 1.627% | 0.6725% | 2.877% | 1.585% | 1.201% | 0.8656% | 1.379% | 100.2% | 26.03% | 16.07% | 16.9%
2 | sac blank cal |  | 6/8/2023 1:29:13 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\sac blank cal.357 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 86.6% |  | 24.61% | 170.3% | 4.795% | 10.59% | 74.78% |  | 173.2% | 32.26% | 31.06% | 4.516% | 12.29% | 0.9044% | 173.2% | 45.61% | 49.43% | 12.98% | 1.543% | 4.664% |  | 11.23% | 12.53% | 8.783% |  | 173.2% | 49.83% | 46.13% | 173.2% | 60.14% | 17.41% | 4.13% | 0.1976% | 0.9564% | 0.5423% | 1.078% | 1.771% | 1.743% | 173.2% | 88.12% | 21.11% | 12.71%
3 | CAL BLANK |  | 6/8/2023 1:35:02 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CAL BLANK.358 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 173.2% |  | 1.232% | 173.2% | 3.828% | 6.429% | 39.78% | 173.2% | 173.2% | 44.91% | 51.62% | 1.93% | 24.22% | 3.552% | 91.18% | 61.64% | 17.76% | 28.51% | 8.859% | 9.642% | 173.2% | 26.9% | 29.18% | 7.659% | 26.18% | 86.61% | 115.6% | 25.07% |  | 33.22% | 35.54% | 0.7181% | 1.453% | 1.271% | 0.3834% | 1.229% | 0.3919% | 1.067% | 86.6% | 24.91% | 25.53% | 14.04%
4 | LOW 0.01 |  | 6/8/2023 1:40:50 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LOW 0.01.359 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 17.88% | 47.04% | 0.9628% | 25.26% | 2.183% | 294.7% | 2.804% | 1.438% | 36.27% | 15.45% | 20% | 7.517% | 5.871% | 76.25% | 26.53% | 76.9% | 28.16% | 9.661% | 2.451% | 170.9% | 148.3% | 22.57% | 20.6% | 117.6% | 152.7% | 39.04% | 49.55% | 140.2% | 68.52% | 4.085% | 4.264% | 1.278% | 1.13% | 1.195% | 1.267% | 0.8522% | 0.2996% | 1.297% | 13.87% | 6.69% | 7.211% | 3.988%
5 | LOW 0.05 |  | 6/8/2023 1:46:39 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LOW 0.05.360 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10.15% | 12.86% | 0.1548% | 30.38% | 3.22% | 50.35% | 21.87% | 14% | 16.67% | 9.933% | 15.87% | 4.46% | 5.82% | 85.43% | 8.323% | 37.27% | 2.864% | 5.398% | 4.959% | 19.36% | 39.82% | 3.57% | 5.708% | 96.61% | 36.6% | 20.53% | 31.85% | 39.89% | 10.41% | 1.283% | 4.409% | 0.9218% | 1.429% | 1.689% | 0.6189% | 0.4923% | 2.372% | 0.5563% | 10.43% | 3.678% | 1.219% | 2.012%
6 | LOW 0.1 |  | 6/8/2023 1:52:27 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LOW 0.1.361 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.2408% | 17.63% | 2.793% | 29.47% | 5.624% | 28.56% | 15.07% | 14.9% | 16.1% | 9.203% | 2.203% | 11.9% | 14.15% | 20.56% | 11.35% | 6.111% | 3.233% | 7.196% | 2.106% | 22.69% | 28.89% | 13.06% | 11.3% | 48.03% | 73.97% | 32.51% | 19.19% | 77.04% | 7.635% | 4.307% | 8.568% | 0.8594% | 5.115% | 4.619% | 1% | 1.438% | 1.021% | 7.014% | 2.273% | 2.575% | 5.531% | 2.95%
7 | LOW 0.5 |  | 6/8/2023 1:58:15 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LOW 0.5.362 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9.706% | 2.07% | 0.8426% | 32.62% | 1.974% | 37.54% | 13.39% | 22.06% | 7.337% | 5.34% | 2.879% | 3.268% | 2.919% | 10.71% | 5.495% | 7.076% | 8.048% | 2.173% | 2.332% | 7.94% | 9.113% | 2.639% | 4.897% | 12.32% | 17.82% | 9.107% | 7.774% | 8.4% | 5.789% | 2.447% | 4.822% | 0.2665% | 3.091% | 5.196% | 1.212% | 0.3302% | 0.4057% | 0.8037% | 0.7364% | 0.2065% | 0.5932% | 0.9194%
8 | STD 1 |  | 6/8/2023 2:04:03 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 1.363 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 5.691% | 9.404% | 2.044% | 21.43% | 0.6152% | 13.87% | 6.273% | 6.207% | 5.934% | 3.947% | 2.453% | 2.569% | 1.053% | 2.359% | 1.699% | 2.093% | 1.805% | 1.539% | 0.8274% | 6.202% | 18.69% | 5.183% | 2.626% | 16.1% | 8.927% | 6.428% | 7.856% | 8.273% | 8.951% | 1.782% | 1.27% | 0.2087% | 2.187% | 0.8447% | 1.061% | 1.73% | 1.465% | 4.04% | 1.403% | 1.626% | 0.9196% | 2.375%
9 | STD 5 |  | 6/8/2023 2:09:52 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 5.364 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4.016% | 3.062% | 3.835% | 4.29% | 2.262% | 7.388% | 3.325% | 3.935% | 2.979% | 1.644% | 0.1979% | 0.6126% | 1.049% | 2.168% | 2.132% | 3.092% | 1.816% | 1.038% | 0.5154% | 4.196% | 4.293% | 0.3878% | 0.05706% | 5.39% | 2.227% | 0.8362% | 2.124% | 7.23% | 6.132% | 1.1% | 1.054% | 4.676% | 0.3397% | 0.7155% | 0.5782% | 0.3337% | 2.109% | 0.3813% | 0.2322% | 1.502% | 3.571% | 1.611%
10 | STD 10 |  | 6/8/2023 2:15:40 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 10.365 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.1569% | 1.203% | 1.091% | 5.751% | 0.7952% | 1.883% | 0.8502% | 0.8055% | 0.5885% | 1.355% | 1.682% | 1.905% | 1.717% | 1.373% | 1.961% | 0.9786% | 1.499% | 1.082% | 1.882% | 0.4181% | 3.99% | 1.274% | 1.931% | 1.459% | 1.85% | 5.599% | 4.142% | 3.014% | 3.042% | 2.355% | 1.404% | 0.5486% | 2.039% | 0.9279% | 2.81% | 1.823% | 0.4268% | 1.998% | 2.321% | 2.548% | 0.4661% | 0.9703%
11 | STD 25 |  | 6/8/2023 2:21:28 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 25.366 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.766% | 0.6069% | 1.45% | 2.71% | 1.485% | 2.634% | 0.6523% | 1.455% | 2.148% | 0.5357% | 1.085% | 1.079% | 1.29% | 0.8556% | 2.498% | 2.54% | 1.355% | 0.5545% | 0.9911% | 0.3396% | 1.867% | 5.416% | 0.36% | 1.452% | 2.322% | 1.182% | 0.8204% | 1.559% | 0.7888% | 1.567% | 0.6343% | 0.6048% | 1.045% | 0.5095% | 0.8364% | 4.03% | 1.321% | 4.63% | 0.4667% | 1.32% | 0.8075% | 1.922%
12 | STD 50 |  | 6/8/2023 2:27:16 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 50.367 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.252% | 5.37% | 1.033% | 5.335% | 0.6558% | 4.448% | 3.985% | 1.758% | 1.901% | 3.96% | 1.588% | 1.219% | 3.807% | 3.95% | 3.9% | 4.638% | 3.012% | 2.068% | 2.13% | 1.538% | 2.855% | 4.033% | 3.375% | 2.368% | 5.031% | 4.343% | 4.151% | 2.907% | 3.541% | 1.557% | 2.657% | 1.27% | 3.464% | 2.381% | 1.638% | 1.147% | 0.6748% | 3.989% | 1.12% | 1.662% | 1.068% | 0.713%
13 | STD 100 |  | 6/8/2023 2:33:03 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 100.368 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.076% | 0.8128% | 0.5699% | 1.016% | 1.03% | 1.801% | 1.881% | 1.895% | 1.198% | 1.972% | 1.474% | 2.085% | 1.69% | 1.022% | 1.352% | 1.36% | 1.241% | 0.4601% | 0.8971% | 0.7968% | 1.512% | 1.655% | 0.4559% | 0.4891% | 1.176% | 0.7752% | 0.7085% | 1.16% | 1.199% | 1.454% | 1.804% | 0.6864% | 0.6608% | 0.6473% | 1.208% | 1.065% | 1.602% | 1.482% | 0.7816% | 1.224% | 1.38% | 1.052%
15 | LBLANK |  | 6/8/2023 2:38:52 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.369 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 63.15% | 86.62% | 1.22% | 61.96% | 7.914% | 86.94% | 382.6% | 0% | 92.02% | 20.78% | 64.72% | 25.98% | 59.97% | 123% | 373% | 21.68% | 95.68% | 7.3% | 38.26% | 82.74% | 150.8% | 11.1% | 15.73% | 50.52% | 28.69% | 57.23% | 42.42% | 10.7% | 100% | 5.16% | 8.374% | 0.826% | 0.5315% | 2.954% | 0.7769% | 1.662% | 0.9978% | 3.796% | 42.84% | 10.09% | 19.08% | 9.787%
16 | LBLANK |  | 6/8/2023 2:57:55 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.370 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 67.04% |  | 92.73% | 281.5% | 12.96% | 678.5% | 136.1% | 0% | 78.49% | 95.24% | 50.25% | 32.52% | 73.46% | 776.7% | 295% | 84.46% | 137% | 38.37% | 17.31% | 100.3% | 106.3% | 96% | 92.7% | 61.83% | 82.73% | 90% | 147.9% | 608.2% | 173.2% | 18.06% | 10.57% | 0.6473% | 0.2683% | 11.92% | 0.8077% | 1.582% | 0.5199% | 1.146% | 165.9% | 16.46% | 62.35% | 34.36%
17 | ICV PE |  | 6/8/2023 3:03:43 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\ICV PE.371 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.22% | 0.8248% | 2.266% | 6.7% | 0.867% | 4.862% | 2.314% | 3.414% | 1.681% | 1.886% | 1.548% | 0.1117% | 0.6089% | 1.204% | 1.646% | 3.179% | 3.406% | 1.285% | 1.228% | 3.056% | 5.002% | 2.175% | 1.177% | 1.785% | 2.545% | 2.087% | 0.355% | 2.408% | 5.166% | 0.6722% | 0.9597% | 1.695% | 0.7044% | 1.85% | 1.344% | 6.61% | 2.238% | 2.261% | 0.7057% | 1.047% | 1.497% | 6.893%
18 | ICV PE |  | 6/8/2023 3:09:59 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\ICV PE.372 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.152% | 2.752% | 0.9446% | 12.9% | 1.444% | 2.909% | 3.433% | 4.696% | 5.658% | 3.959% | 1.046% | 3.395% | 0.5026% | 4.587% | 2.038% | 2.63% | 2.988% | 1.793% | 0.7269% | 1.845% | 4.95% | 2.725% | 1.04% | 1.695% | 3.345% | 0.5396% | 2.422% | 0.7369% | 4.982% | 1.246% | 0.9418% | 0.7618% | 0.9079% | 1.909% | 0.9095% | 1.115% | 2.288% | 2.925% | 2.22% | 1.873% | 0.02433% | 0.7606%
19 | CCV 1ppb TRACE |  | 6/8/2023 3:15:47 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.373 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.146% | 4.79% | 1.355% | 10.32% | 4.322% | 6.652% | 2.177% | 12.88% | 1.733% | 4.503% | 1.898% | 2.961% | 0.6762% | 4.194% | 3.016% | 2.486% | 2.404% | 1.364% | 1.701% | 3.438% | 12.08% | 0.5862% | 3.408% | 5.604% | 15.02% | 8.324% | 3.768% | 8.266% | 9.856% | 3.354% | 1.205% | 1.514% | 2.122% | 0.9313% | 1.015% | 0.5324% | 1.54% | 1.954% | 3.669% | 2.096% | 1.248% | 0.6185%
20 | CCV 1ppb TRACE |  | 6/8/2023 3:22:03 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.374 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.7328% | 2.502% | 1.749% | 39.49% | 2.806% | 17.08% | 6.032% | 10.38% | 6.039% | 0.4855% | 3.597% | 1.215% | 2.522% | 7.259% | 1.708% | 7.103% | 3.176% | 2.383% | 2.925% | 4.034% | 15.74% | 3.732% | 1.593% | 7.796% | 8.342% | 6.163% | 3.611% | 1.984% | 13.43% | 3.194% | 1.354% | 0.9699% | 1.165% | 0.4048% | 2.209% | 0.5599% | 0.9466% | 3.065% | 1.588% | 3.924% | 1.889% | 1.751%
21 | LFB |  | 6/8/2023 3:27:51 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.375 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.4293% | 4.148% | 1.378% | 2.942% | 2.17% | 4.11% | 0.3902% | 2.878% | 2.35% | 2.248% | 3.542% | 4.659% | 1.177% | 2.465% | 2.204% | 3.981% | 0.5024% | 4.348% | 5.069% | 1.551% | 3.605% | 2.208% | 2.16% | 4.446% | 1.343% | 1.226% | 2.018% | 2.878% | 2.604% | 5.176% | 1.782% | 1.704% | 0.5031% | 0.6252% | 3.826% | 1.025% | 2.003% | 0.9338% | 5.601% | 4.301% | 1.217% | 1.099%
22 | LLOQ 0.01 |  | 6/8/2023 3:33:39 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LLOQ 0.01.376 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 43.4% | 29.37% | 2.702% | 16.44% | 1.068% | 47.81% | 32.42% | 73.52% | 57.25% | 0.3449% | 30.82% | 10.17% | 6.85% | 19.07% | 23.05% | 19.75% | 29.76% | 2.373% | 4.494% | 50.19% | 167.7% | 21.66% | 20.24% | 304.1% | 885.8% | 27.87% | 28.65% | 160.5% | 49.45% | 6.032% | 5.157% | 2.784% | 1.412% | 4.349% | 0.4254% | 1.262% | 0.3261% | 1.502% | 14.13% | 1.143% | 5.832% | 1.255%
23 | CCV 10 ppb MAJORS |  | 6/8/2023 3:39:28 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.377 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 158.1% | 173.2% | 2.944% | 3286% | 1.687% | 1.682% | 36.29% | 1653% | 1624% | 75.87% | 177.9% | 8.954% | 62.05% | 2.266% | 139.1% | 202.8% | 1042% | 464.7% | 6.963% | 0.5532% | 0% | 57.3% | 48.58% | 2.951% | 32.3% | 328.8% | 82.59% | 101.2% | 86.69% | 17.14% | 46.14% | 2.953% | 3.68% | 2.815% | 2.368% | 3.804% | 4.976% | 6.926% | 3.8% | 308.3% | 5.105% | 14.47%
24 | CCV 10 ppb MAJORS |  | 6/8/2023 3:45:42 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.378 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.919% | 3.811% | 1.545% | 6.698% | 0.6869% | 5.103% | 4.182% | 7.146% | 5.052% | 3.04% | 0.3672% | 1.975% | 1.535% | 5.026% | 2.081% | 3.351% | 3.393% | 0.5696% | 1.133% | 1.956% | 5.211% | 1.748% | 0.5819% | 2.732% | 2.749% | 3.096% | 2.426% | 2.531% | 5.962% | 1.11% | 0.6576% | 0.8399% | 0.918% | 2.645% | 1.254% | 0.8326% | 2.677% | 1.684% | 2.287% | 2.446% | 1.225% | 0.7776%
25 | LLOQ 0.1 |  | 6/8/2023 3:51:31 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LLOQ 0.1.379 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 8.812% | 12.04% | 0.6505% | 17.69% | 4.757% | 70.92% | 27.28% | 3.854% | 6.898% | 2.596% | 4.701% | 4.832% | 1.826% | 26.18% | 7.103% | 19.78% | 8.99% | 0.992% | 0.4282% | 24.98% | 28.4% | 5.786% | 5.205% | 23.52% | 34.31% | 23.35% | 24.63% | 19.09% | 26.23% | 3.711% | 2.782% | 1.363% | 0.7574% | 1.414% | 0.5907% | 1.786% | 2.48% | 3.45% | 5.923% | 4.285% | 2.723% | 1.285%
26 | STD 0.5 |  | 6/8/2023 3:57:19 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 0.5.380 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.702% | 4.546% | 2.371% | 11.95% | 3.825% | 41.58% | 3.693% | 15.46% | 10.66% | 2.395% | 7.097% | 6.012% | 3.059% | 3.122% | 5.626% | 11.1% | 6.716% | 8.876% | 6.218% | 4.962% | 11.99% | 4.707% | 2.515% | 15.97% | 9.698% | 7.617% | 7.226% | 21.91% | 13.86% | 6.352% | 2.597% | 0.5691% | 1.706% | 1.928% | 5.45% | 0.3688% | 1.667% | 2.241% | 7.131% | 7.904% | 0.2489% | 2.505%
27 | 8_2_21_UF_1 |  | 6/8/2023 4:03:07 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_UF_1.381 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.7567% | 3.735% | 1.343% | 8.539% | 0.9086% | 1.417% | 1.202% | 1.431% | 8.1% | 9.737% | 12.09% | 1.763% | 1.386% | 10.27% | 29.47% | 10.26% | 4.144% | 0.7506% | 1.04% | 2.54% | 4.38% | 2.012% | 2.274% | 39.61% | 123.9% | 21.74% | 4.291% | 13.9% | 45.94% | 0.6115% | 1.043% | 1.21% | 0.4188% | 1.712% | 1.335% | 0.7152% | 2.049% | 0.07906% | 27.75% | 4.639% | 4.375% | 4.054%
28 | 8_2_21_UF_6 |  | 6/8/2023 4:08:55 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_UF_6.382 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.87% | 1.627% | 3.714% | 15.87% | 2.419% | 7.267% | 7.869% | 3.154% | 3.361% | 17.97% | 6.581% | 5.246% | 1.64% | 4.262% | 3.096% | 1.233% | 4.843% | 3.077% | 1.774% | 0.9874% | 11.1% | 3.128% | 2.105% | 19.62% | 8.888% | 11.85% | 21.94% | 12.28% | 31.1% | 2.819% | 0.8963% | 0.6371% | 0.02132% | 0.9061% | 1.816% | 1.39% | 0.8845% | 2.885% | 12.09% | 8.319% | 10.09% | 3.795%
29 | 8_2_21_UF_6_SP |  | 6/8/2023 4:14:43 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_UF_6_SP.383 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.485% | 4.295% | 1.317% | 2.029% | 5.489% | 3.118% | 3.781% | 1.606% | 1.625% | 3.01% | 2.568% | 3.219% | 0.8674% | 5.82% | 1.231% | 0.4726% | 1.188% | 1.747% | 1.378% | 1.12% | 2.508% | 2.991% | 0.9963% | 4.015% | 4.91% | 2.908% | 2.915% | 5.203% | 4.046% | 1.551% | 1.864% | 0.5132% | 0.1651% | 0.8316% | 1.552% | 1.917% | 5.294% | 1.965% | 2.354% | 1.769% | 0.5366% | 1.229%
30 | STD 1 |  | 6/8/2023 4:20:31 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 1.384 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.62% | 4.033% | 1.372% | 10.91% | 2.478% | 7.773% | 3.896% | 5.326% | 3.3% | 7.781% | 1.363% | 15.68% | 1.507% | 5.181% | 3.48% | 2.894% | 3.526% | 1.557% | 2.632% | 4.265% | 4.504% | 4.268% | 1.048% | 7.15% | 7.449% | 10.91% | 6.735% | 14.11% | 7.025% | 2.304% | 1.161% | 1.884% | 0.5413% | 2.148% | 1.09% | 1.041% | 1.707% | 5.861% | 1.916% | 1.927% | 3.638% | 1.466%
31 | 8_2_21_UF_6_SP_LD |  | 6/8/2023 4:26:19 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_UF_6_SP_LD.385 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4.412% | 3.769% | 1.569% | 9.785% | 0.7721% | 0.7884% | 4.988% | 2.344% | 1.938% | 1.765% | 0.458% | 1.192% | 2.457% | 1.748% | 3.899% | 0.4768% | 2.434% | 0.7915% | 1.63% | 1.364% | 4.316% | 2.399% | 0.5881% | 0.2507% | 6.391% | 3.5% | 1.132% | 1.981% | 4.4% | 1.151% | 1.939% | 0.5446% | 0.8162% | 1.402% | 1.325% | 1.871% | 2.392% | 2.015% | 1.524% | 0.9369% | 1.306% | 1.488%
32 | 8_2_21_UF_6_SP_LSP |  | 6/8/2023 4:32:06 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_UF_6_SP_LSP.386 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.9673% | 1.471% | 3.095% | 10.29% | 1.79% | 0.1854% | 1.02% | 1.739% | 1.412% | 0.7039% | 0.6872% | 1.821% | 1.075% | 1.669% | 1.125% | 0.6035% | 1.193% | 1.289% | 0.6705% | 3.492% | 2.683% | 0.5773% | 0.7237% | 4.442% | 3.422% | 1.165% | 2.079% | 2.935% | 4.056% | 0.5874% | 0.521% | 2.059% | 0.722% | 0.9421% | 0.9866% | 0.8148% | 0.9554% | 0.5294% | 1.146% | 0.6828% | 2.761% | 0.7277%
33 | 8_2_21_UF_8 |  | 6/8/2023 4:37:54 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_UF_8.387 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.87% | 2.442% | 1.029% | 14.85% | 0.8439% | 5.755% | 1.051% | 4.949% | 2.698% | 7.485% | 2.502% | 2.484% | 0.8291% | 2.22% | 14.22% | 1.668% | 1.529% | 1.082% | 0.4922% | 1.484% | 10.63% | 3.042% | 0.8156% | 10.42% | 13.02% | 13.58% | 13.49% | 5.292% | 12.54% | 1.956% | 1.175% | 0.7286% | 0.1808% | 1.148% | 1.363% | 1.293% | 2.475% | 3.389% | 7.943% | 6.047% | 1.942% | 2.047%
34 | 8_2_21_UF_9 |  | 6/8/2023 4:43:42 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_UF_9.388 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.161% | 0.9612% | 1.906% | 33.19% | 0.339% | 2.131% | 3.906% | 3.779% | 4.232% | 12.35% | 7.123% | 1.451% | 1.251% | 12.46% | 28.57% | 11.59% | 2.611% | 0.9629% | 0.4782% | 2.082% | 7.033% | 1.186% | 1.531% | 14.01% | 62.33% | 29.33% | 21.84% | 5.838% | 24.55% | 1.583% | 0.271% | 0.7503% | 0.7306% | 2.566% | 0.6562% | 0.7723% | 1.037% | 1.547% | 7.01% | 7.916% | 2.147% | 4.555%
35 | LBLANK |  | 6/8/2023 4:49:31 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.389 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 79.15% |  | 8.483% | 137.9% | 14.62% | 46.05% | 443.2% | 1463% | 11.11% | 231.9% | 48.48% | 34.16% | 51.53% | 167.5% | 369.1% | 66.21% | 207.1% | 62.48% | 1346% | 149.8% | 115.3% | 10.52% | 8.634% | 70.37% | 229.7% | 50.91% | 33.35% | 117.2% | 43.96% | 31.94% | 56.12% | 0.6767% | 1.232% | 7.964% | 1.232% | 0.2253% | 0.9105% | 0.5329% | 223.7% | 29.56% | 14.37% | 71.65%
36 | CCV 1ppb TRACE |  | 6/8/2023 4:55:19 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.390 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.446% | 7.902% | 3.09% | 13.83% | 1.796% | 23.87% | 3.261% | 18.36% | 2.916% | 6.403% | 3.066% | 0.5704% | 1.9% | 9.285% | 4.173% | 6.901% | 4.452% | 0.8311% | 1.339% | 3.043% | 10.9% | 3.29% | 1.744% | 6.283% | 21.01% | 1.345% | 3.267% | 10.19% | 18.27% | 1.297% | 0.9088% | 1.163% | 0.909% | 3.291% | 0.5874% | 3.356% | 1.97% | 2.573% | 2.765% | 1.841% | 1.956% | 1.375%
37 | CCV 1ppb TRACE |  | 6/8/2023 5:01:37 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.391 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.622% | 5.692% | 1.559% | 12.77% | 7.054% | 12.65% | 9.944% | 4.419% | 4.443% | 5.221% | 0.9798% | 4.732% | 0.651% | 4.314% | 1.65% | 3.512% | 1.741% | 2.077% | 1.107% | 2.758% | 8.868% | 1.243% | 2.708% | 2.737% | 5.184% | 6.986% | 3.564% | 11.12% | 12.9% | 2.155% | 1.301% | 1.534% | 0.7503% | 1.252% | 0.8184% | 0.7491% | 1.883% | 2.468% | 1.777% | 1.289% | 0.8885% | 0.3345%
38 | CCV 10 ppb MAJORS |  | 6/8/2023 5:07:25 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.392 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.396% | 2.292% | 1.977% | 3.877% | 1.746% | 5.353% | 3.829% | 4.536% | 1.319% | 2.001% | 0.6264% | 2.244% | 2.786% | 1.784% | 1.781% | 2.393% | 0.7656% | 1.57% | 0.698% | 2.015% | 6.086% | 1.97% | 1.771% | 5.44% | 6.63% | 2.415% | 2.156% | 2.221% | 2.961% | 1.781% | 1.948% | 2.064% | 1.555% | 1.419% | 0.6118% | 1.09% | 1.704% | 1.342% | 0.8862% | 1.062% | 2.642% | 0.7132%
39 | CCV 10 ppb MAJORS |  | 6/8/2023 5:13:42 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.393 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.16% | 3.505% | 2.308% | 1.763% | 2.866% | 3.069% | 3.209% | 3.498% | 2.41% | 1.684% | 1.536% | 2.427% | 1.105% | 1.429% | 1.972% | 0.9589% | 1.029% | 0.3912% | 0.63% | 2.387% | 2.378% | 2.941% | 2.14% | 3.456% | 3.701% | 1.002% | 4.182% | 3.17% | 0.7329% | 0.5242% | 1.058% | 1.514% | 0.9778% | 1.687% | 0.4846% | 0.4652% | 1.479% | 2.845% | 1.114% | 0.7688% | 1.816% | 2.849%
40 | 8_2_21_UF_10 |  | 6/8/2023 5:19:31 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_UF_10.394 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 5.503% | 3.409% | 2.592% | 25.18% | 0.5564% | 7.152% | 6.149% | 5.002% | 4.333% | 8.217% | 8.616% | 1.009% | 0.3477% | 10.67% | 30.99% | 13.93% | 3.307% | 0.1797% | 0.6805% | 0.7256% | 4.582% | 1.598% | 0.8547% | 12.32% | 14.01% | 37.88% | 20.53% | 22.74% | 11.76% | 2.041% | 0.4095% | 1.627% | 0.9795% | 5.576% | 0.8848% | 0.3215% | 0.8424% | 1.95% | 5.333% | 13.75% | 6.086% | 1.862%
41 | 8_2_21_UF_11 |  | 6/8/2023 5:25:19 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_UF_11.395 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 19.18% | 22.07% | 1.928% | 12.9% | 1.905% | 22.28% | 15.9% | 22.88% | 21.63% | 20.44% | 2.333% | 2.472% | 0.6556% | 22.12% | 6.031% | 20.33% | 20.84% | 1.062% | 2.194% | 0.5781% | 19.83% | 2.254% | 0.8809% | 16.31% | 62.58% | 9.625% | 7.508% | 20.78% | 35.38% | 1.977% | 1.667% | 0.4756% | 1.187% | 17.91% | 0.5905% | 1.555% | 1.831% | 2.621% | 7.858% | 5.808% | 1.792% | 3.044%
42 | 8_2_21_UF_11_DUP |  | 6/8/2023 5:31:09 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_UF_11_DUP.396 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.483% | 1.62% | 0.8809% | 45.67% | 2.132% | 2.531% | 4.451% | 1.702% | 3.744% | 5.224% | 10.86% | 6.769% | 2.248% | 2.945% | 13.37% | 3.978% | 1.794% | 2.86% | 0.8724% | 1.079% | 7.864% | 3.001% | 1.374% | 12.72% | 51.1% | 10.23% | 13.18% | 5.788% | 21.82% | 1.034% | 1.894% | 0.4259% | 1.515% | 1.876% | 1.526% | 0.4075% | 1.775% | 2.95% | 13.82% | 12.79% | 2.489% | 1.161%
43 | 8_2_21_F_1 |  | 6/8/2023 5:36:57 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_F_1.397 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.503% | 0.6914% | 2.05% | 2.361% | 2.291% | 2.224% | 4.083% | 0.8618% | 6.609% | 8.611% | 9.879% | 1.537% | 1.31% | 5.132% | 4.175% | 7.284% | 5.044% | 0.1512% | 0.3851% | 1.467% | 11.59% | 4.793% | 0.6282% | 42.06% | 118% | 17.45% | 36.89% | 9.631% | 48.01% | 0.05269% | 1.021% | 1.102% | 0.7922% | 2.018% | 0.5353% | 0.2488% | 2.767% | 3.616% | 18.83% | 21.61% | 6.444% | 1.876%
44 | 8_2_21_F_6 |  | 6/8/2023 5:44:00 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_F_6.398 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.8101% | 3.354% | 2.975% | 27.11% | 2.208% | 1.987% | 2.054% | 3.507% | 2.368% | 4.024% | 0.4678% | 0.7283% | 18.54% | 1.138% | 21.22% | 1.951% | 0.6547% | 0.2064% | 1.256% | 1.846% | 6.22% | 18.66% | 19.2% | 4.327% | 50.78% | 6.552% | 14.9% | 5.725% | 32.36% | 1.082% | 18.02% | 2.107% | 16.96% | 1.116% | 0.5041% | 0.4939% | 0.7912% | 16.41% | 4.274% | 9.841% | 2.725% | 0.437%
45 | 8_2_21_F_8 |  | 6/8/2023 5:49:48 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_F_8.399 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.832% | 0.6194% | 2.624% | 15.25% | 1.018% | 3.478% | 3.846% | 0.509% | 5.92% | 0.4865% | 4.806% | 0.4874% | 6.445% | 4.481% | 17.98% | 9.502% | 1.745% | 0.1123% | 1.435% | 0.7862% | 7.538% | 3.85% | 5.087% | 13.05% | 59.33% | 31.17% | 19.94% | 10.01% | 37.11% | 2.076% | 6.581% | 2.065% | 4.363% | 0.4728% | 0.7762% | 0.5153% | 1.324% | 2.711% | 14.62% | 11.99% | 3.794% | 3.143%
46 | 8_2_21_F_8_DUP |  | 6/8/2023 5:55:36 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_F_8_DUP.400 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 20.88% | 22.31% | 2.775% | 22.31% | 1.144% | 29.05% | 18.72% | 23.31% | 21.25% | 18.45% | 5.109% | 3.764% | 1.885% | 24.45% | 8.328% | 9.118% | 20.32% | 1.484% | 1.072% | 0.7361% | 27.9% | 1.522% | 1.28% | 18.41% | 39.76% | 18.59% | 4.003% | 11.82% | 33.93% | 2.721% | 2.328% | 1.061% | 1.065% | 16.67% | 0.6848% | 0.3242% | 1.959% | 1.914% | 6.876% | 10.62% | 2.903% | 0.4356%
47 | 8_2_21_F_8_DUP_LD |  | 6/8/2023 6:01:24 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_F_8_DUP_LD.401 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.794% | 1.711% | 1.278% | 14.73% | 2.114% | 2.782% | 2.328% | 2.711% | 5.679% | 7.585% | 6.116% | 7.61% | 1.245% | 1.333% | 18.8% | 6.813% | 3.854% | 7.861% | 8.353% | 1.115% | 4.693% | 1.441% | 2.776% | 11.9% | 20.88% | 19.05% | 19.85% | 3.916% | 37.16% | 8.354% | 3.106% | 1.757% | 1.732% | 0.4066% | 7.31% | 1.44% | 1.493% | 2.337% | 19.84% | 20.69% | 0.06361% | 2.821%
48 | 8_2_21_F_8_DUP_LD |  | 6/8/2023 6:07:40 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_F_8_DUP_LD.402 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.507% | 3.767% | 1.834% | 5.168% | 1.94% | 5.585% | 4.779% | 5.018% | 1.254% | 22.52% | 7.032% | 2.865% | 1.313% | 6.509% | 11.24% | 6.666% | 1.272% | 2.779% | 4.109% | 0.9771% | 4.98% | 0.8117% | 0.5945% | 46.35% | 43.76% | 33.7% | 8.241% | 5.494% | 22.96% | 2.175% | 0.9531% | 0.4498% | 0.5478% | 1.953% | 2.404% | 0.8697% | 0.7976% | 0.9844% | 17.64% | 4.893% | 2.382% | 2.512%
49 | 8_2_21_F_8_DUP_LSP |  | 6/8/2023 6:13:28 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_F_8_DUP_LSP.403 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.7741% | 0.3391% | 0.9871% | 6.109% | 0.6243% | 1.592% | 0.7787% | 1.466% | 2.044% | 0.8003% | 0.7972% | 0.9045% | 1.026% | 2.881% | 1.528% | 1.723% | 1.174% | 0.7268% | 0.4105% | 1.182% | 4.356% | 1.95% | 0.4972% | 1.078% | 3.974% | 1.533% | 2.228% | 1.823% | 0.8756% | 0.4895% | 0.5238% | 1.162% | 0.03388% | 0.7448% | 0.4913% | 0.8345% | 1.718% | 2.177% | 2.057% | 1.221% | 0.8289% | 1.051%
50 | LBLANK |  | 6/8/2023 6:19:16 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.404 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 38.91% | 43.07% | 2.899% | 141.3% | 34.35% | 24.94% | 1155% | 0% | 1440% | 21.46% | 53.55% | 21.1% | 48.09% | 2275% | 216% | 49.31% | 95.83% | 19.26% | 529.1% | 154.5% | 154.4% | 18.18% | 16.17% | 53.05% | 0% | 479.7% | 201.2% | 78.09% | 0.9153% | 17.84% | 9.854% | 0.4256% | 0.915% | 0.2225% | 1.344% | 2.11% | 0.9712% | 2.467% | 234.9% | 58.69% | 1153% | 426.5%
51 | CCV 1ppb TRACE |  | 6/8/2023 6:25:05 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.405 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.303% | 8.367% | 1.045% | 23.52% | 3.439% | 6.581% | 6.988% | 10.33% | 9.627% | 2.519% | 1.773% | 1.403% | 3.282% | 2.382% | 5.026% | 7.02% | 2.892% | 1.601% | 3.131% | 4.263% | 24% | 7.134% | 2.486% | 8.037% | 5.324% | 3.007% | 1.335% | 14.14% | 5.325% | 1.414% | 0.5261% | 1.242% | 0.9058% | 3.012% | 2.135% | 1.906% | 1.345% | 4.724% | 1.55% | 3.181% | 0.8987% | 2.082%
52 | CCV 1ppb TRACE |  | 6/8/2023 6:31:23 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.406 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.417% | 5.498% | 9.766% | 12.18% | 6.92% | 18.82% | 5.271% | 13.28% | 4.203% | 3.895% | 2.845% | 1.81% | 1.563% | 4.569% | 3.742% | 3.79% | 1.862% | 1.77% | 2.801% | 11.72% | 5.952% | 2.301% | 0.939% | 0.261% | 9.332% | 0.2285% | 9.381% | 12.98% | 1.19% | 1.673% | 0.7915% | 6.156% | 1.851% | 1.638% | 0.8231% | 1.744% | 0.9822% | 0.5946% | 1.274% | 2.244% | 6.368% | 2.389%
53 | LFB |  | 6/8/2023 6:37:11 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.407 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.85% | 0.4251% | 1.53% | 4.791% | 0.7876% | 4.255% | 3.693% | 3.667% | 1.672% | 2.105% | 0.9435% | 1.669% | 2.916% | 1.35% | 0.8573% | 1.152% | 2.266% | 1.076% | 1.714% | 1.478% | 5.035% | 3.349% | 0.8653% | 1.928% | 5.544% | 4.132% | 4.198% | 0.6814% | 2.537% | 0.6443% | 1.833% | 0.9397% | 1.322% | 0.8294% | 0.7835% | 1.203% | 1.693% | 2.628% | 0.3681% | 0.9522% | 1.54% | 1.922%
54 | CCV 10 ppb MAJORS |  | 6/8/2023 6:42:59 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.408 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.108% | 4.886% | 1.734% | 9.845% | 1.184% | 9.494% | 3.581% | 4.761% | 4.359% | 3.016% | 1.021% | 1.419% | 0.6991% | 5.295% | 3.581% | 4.051% | 3.979% | 1.077% | 0.2044% | 1.702% | 7.327% | 3.429% | 0.8732% | 3.116% | 4.141% | 0.3015% | 1.247% | 2.739% | 1.72% | 1.022% | 1.686% | 1.526% | 0.3053% | 3.025% | 0.2478% | 0.4219% | 1.176% | 2.545% | 0.8978% | 0.5974% | 0.72% | 1.52%
55 | CCV 10 ppb MAJORS |  | 6/8/2023 6:49:21 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.409 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 5.813% | 4.893% | 16.77% | 12.09% | 1.036% | 6.458% | 4.535% | 2.786% | 3.299% | 5.644% | 3.362% | 1.237% | 1.614% | 6.414% | 6.422% | 5.501% | 5.534% | 1.492% | 0.9395% | 15.32% | 7.288% | 2.032% | 1.809% | 16.63% | 3.088% | 1.778% | 1.037% | 16.2% | 1.469% | 0.7357% | 1.875% | 12.12% | 1.358% | 5.813% | 0.5542% | 0.6777% | 0.8079% | 1.634% | 0.612% | 0.7002% | 14.85% | 2.218%
56 | LBLANK |  | 6/8/2023 7:26:31 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.410 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 138.2% |  | 1.168% | 1.003% | 2.893% | 993.7% | 1138% | 0% | 1021% | 106% | 15.2% | 52.55% | 69.11% | 98.4% | 165.4% | 78.97% | 75.58% | 39.56% | 55.41% | 33.42% | 1021% | 84.18% | 80.4% | 46.36% | 38.3% | 132.5% | 152.4% | 91.43% |  | 19.64% | 28.68% | 2.344% | 1.316% | 0.7243% | 1.651% | 0.5655% | 0.742% | 4.12% | 221.2% | 81.38% | 15.43% | 107.6%
57 | PE ICV |  | 6/8/2023 7:32:19 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\PE ICV.411 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.885% | 2.977% | 2.587% | 5.495% | 1.537% | 1.739% | 0.3546% | 5.453% | 1.412% | 0.8395% | 0.8167% | 2.825% | 1.846% | 1.49% | 0.6632% | 3.709% | 1.306% | 0.6351% | 1.189% | 2.827% | 2.919% | 6.022% | 0.5977% | 3.486% | 5.173% | 1.742% | 1.145% | 1.798% | 2.606% | 0.5333% | 0.854% | 0.6745% | 1.012% | 0.3535% | 0.5433% | 1.359% | 0.9893% | 5.248% | 1.501% | 1.209% | 1.556% | 1.506%
58 | 8_2_21_F_9 |  | 6/8/2023 7:38:07 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_F_9.412 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.337% | 2.693% | 1.189% | 5.598% | 0.9763% | 1.606% | 2.78% | 1.471% | 5.218% | 10.91% | 1.883% | 2.406% | 1.701% | 3.417% | 12.04% | 2.148% | 3.523% | 1.918% | 0.8399% | 1.874% | 6.073% | 4.338% | 0.7181% | 21.64% | 87.99% | 10.43% | 29.22% | 31.65% | 21.88% | 0.6651% | 1.541% | 0.8986% | 1.027% | 2.668% | 0.5173% | 1.548% | 0.4929% | 4.268% | 24.97% | 10.44% | 1.337% | 0.3846%
59 | 8_2_21_F_10 |  | 6/8/2023 7:43:55 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_F_10.413 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10.92% | 7.059% | 1.057% | 2.758% | 1.87% | 12.79% | 6.36% | 9.999% | 10.52% | 5.873% | 11.33% | 2.468% | 0.7843% | 10.83% | 9.796% | 6.514% | 7.702% | 1.298% | 1.162% | 0.8427% | 4.371% | 2.02% | 0.5671% | 4.023% | 91.39% | 21.32% | 14.63% | 16.85% | 24.61% | 0.7856% | 1.709% | 1.2% | 0.2992% | 8.127% | 0.7232% | 0.7729% | 1.136% | 1.7% | 28.78% | 12.14% | 3.658% | 1.764%
60 | 8_2_21_F_10_DUP |  | 6/8/2023 7:49:43 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_F_10_DUP.414 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.176% | 2.21% | 2.302% | 2.451% | 1.406% | 2.083% | 3.355% | 2.284% | 4.367% | 4.084% | 1.668% | 2.331% | 9.444% | 1.481% | 24.88% | 4.464% | 1.374% | 1.225% | 1.044% | 0.7969% | 3.686% | 8.374% | 7.915% | 12.27% | 28% | 29.22% | 9.486% | 26.03% | 44.75% | 1.453% | 7.909% | 0.7766% | 7.909% | 1.151% | 0.4265% | 1.137% | 1.019% | 8.433% | 32.09% | 6.301% | 0.1872% | 2.237%
61 | 8_2_21_F_11 |  | 6/8/2023 7:55:32 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_F_11.415 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4.12% | 2.983% | 0.9285% | 9.556% | 2.258% | 1.002% | 8.36% | 0.3749% | 4.5% | 4.79% | 8.754% | 1.837% | 0.8339% | 0.4221% | 21.98% | 11.08% | 1.129% | 1.638% | 2.003% | 2.237% | 4.21% | 1.804% | 1.199% | 18.14% | 52.65% | 22.12% | 10.65% | 19.31% | 33.2% | 0.8118% | 0.9886% | 1.377% | 0.4379% | 0.7896% | 1.174% | 3.282% | 2.071% | 1.158% | 36.05% | 13.48% | 0.9719% | 3.996%
62 | 8_2_21_F_11_SP |  | 6/8/2023 8:01:20 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_F_11_SP.416 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.292% | 1.489% | 2.782% | 2.956% | 1.466% | 1.402% | 3.292% | 0.4989% | 0.6856% | 1.955% | 1.212% | 1.852% | 1.705% | 2.733% | 1.987% | 2.014% | 3.733% | 1.5% | 1.271% | 3.144% | 3.905% | 1.216% | 1.158% | 2.405% | 3.92% | 3.885% | 2.951% | 3.87% | 4.556% | 1.467% | 1.666% | 2.49% | 1.337% | 2.096% | 0.2242% | 0.5709% | 1.169% | 1.045% | 2.448% | 1.073% | 2.466% | 1.561%
63 | 8_2_21_W_1 |  | 6/8/2023 8:07:08 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_W_1.417 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.868% | 1.923% | 0.6004% | 5.994% | 2.539% | 6.208% | 6.111% | 4.209% | 5.538% | 7.165% | 4.326% | 0.7109% | 2.283% | 0.5307% | 30.05% | 22.6% | 4.503% | 1.493% | 2.857% | 5.8% | 24.58% | 1.838% | 2.153% | 54.38% | 355.4% | 28.37% | 39.13% | 43.51% | 41.21% | 1.516% | 0.8951% | 0.2326% | 1.946% | 0.5275% | 0.8372% | 0.6291% | 1.08% | 1.133% | 28.75% | 16.7% | 4.692% | 1.884%
64 | 8_2_21_W_6 |  | 6/8/2023 8:12:56 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_W_6.418 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9.319% | 6.314% | 1.671% | 3.582% | 1.628% | 5.945% | 4.194% | 4.943% | 6.482% | 27% | 4.039% | 0.5259% | 1.898% | 0.7304% | 28.34% | 30.54% | 3.781% | 2.238% | 0.1417% | 4.015% | 15.59% | 3.819% | 1.524% | 16.2% | 176.4% | 46.25% | 12.55% | 83.25% | 44% | 0.8938% | 1.959% | 0.8327% | 1.807% | 2.276% | 1.053% | 1.456% | 1.445% | 1.505% | 29% | 15.59% | 3.277% | 0.6477%
65 | 8_2_21_W_6_DUP |  | 6/8/2023 8:18:44 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_W_6_DUP.419 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.6155% | 5.154% | 0.3485% | 8.093% | 2.028% | 7.882% | 10.55% | 2.188% | 6.22% | 4.311% | 5.586% | 0.7943% | 1.332% | 4.098% | 44.51% | 2.178% | 2.808% | 1.377% | 3.389% | 2.995% | 22.36% | 1.463% | 0.5499% | 87.51% | 105.1% | 18.58% | 29.98% | 64.19% | 86.6% | 1.602% | 2.094% | 2.209% | 0.3325% | 1.169% | 2.007% | 0.5503% | 0.5122% | 2.329% | 4.637% | 16.62% | 1.619% | 2.523%
66 | LBLANK |  | 6/8/2023 8:24:33 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.420 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 944.5% |  | 20.96% | 109.5% | 17.35% | 39.67% | 156.2% | 0% | 294.7% | 35.85% | 9.34% | 52.97% | 98.92% | 271.7% | 166.3% | 213% | 8397% | 24.22% | 114.7% | 16.61% | 944.5% | 181.3% | 171.3% | 18.93% | 38.04% | 397.1% | 92.71% | 149.7% |  | 52.36% | 2646% | 1.089% | 1.463% | 2.338% | 1.806% | 0.3488% | 1.454% | 1.954% | 215.2% | 57.84% | 75.96% | 13.28%
67 | CCV 1ppb TRACE |  | 6/8/2023 8:30:21 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.421 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4.726% | 2.034% | 5.325% | 29.31% | 1.835% | 6.883% | 4.47% | 9.457% | 6.665% | 3.352% | 0.7194% | 1.466% | 3.127% | 5.248% | 1.859% | 0.2963% | 6.208% | 1.07% | 3.051% | 4.919% | 4.506% | 0.6103% | 1.328% | 16.06% | 11.65% | 7.888% | 8.868% | 7.099% | 4.431% | 1.591% | 1.743% | 4.546% | 1.629% | 2.687% | 0.3267% | 0.4631% | 0.595% | 0.7419% | 4.182% | 0.8873% | 6.788% | 0.4374%
68 | CCV 1ppb TRACE |  | 6/8/2023 8:36:39 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.422 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.236% | 1.11% | 1.01% | 9.965% | 4.562% | 17.98% | 5.549% | 6.227% | 1.062% | 4.256% | 3.772% | 1.166% | 1.54% | 4.625% | 4.606% | 4.593% | 0.5538% | 1.7% | 0.387% | 4.366% | 12.41% | 2.269% | 1.349% | 11.81% | 8.03% | 3.387% | 3.041% | 6.001% | 2.945% | 2.392% | 1.214% | 0.8958% | 1.013% | 1.382% | 0.6356% | 1.164% | 2.832% | 3.492% | 2.501% | 1.355% | 1.964% | 0.6176%
69 | CCV 10 ppb MAJORS |  | 6/8/2023 8:42:27 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.423 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9.64% | 10.18% | 1.049% | 19.44% | 1.556% | 10.58% | 8.224% | 8.135% | 8.661% | 11.09% | 0.6992% | 2.574% | 1.548% | 12.06% | 12.59% | 9.953% | 10.68% | 0.6186% | 1.018% | 3.567% | 9.275% | 1.646% | 2.043% | 1.657% | 3.645% | 3.145% | 1.369% | 2.269% | 3.361% | 0.9501% | 1.659% | 1.809% | 0.6531% | 8.73% | 0.5137% | 1.389% | 1.196% | 1.013% | 1.47% | 1.216% | 2.053% | 1.159%
70 | CCV 10 ppb MAJORS |  | 6/8/2023 8:48:44 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.424 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.528% | 1.318% | 2.178% | 4.878% | 0.3747% | 1.426% | 2.842% | 3.089% | 3.617% | 3.01% | 1.803% | 2.481% | 0.6641% | 4.839% | 1.286% | 1.731% | 2.154% | 1.506% | 1.45% | 2.333% | 6.449% | 0.4759% | 0.5821% | 3.057% | 4.767% | 1.313% | 2.077% | 5.475% | 1.61% | 1.621% | 0.7442% | 2.857% | 0.5105% | 1.842% | 1.205% | 6.545% | 1.067% | 0.9986% | 1.333% | 0.7407% | 2.958% | 10.63%
71 | 8_2_21_W_8 |  | 6/8/2023 8:54:33 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_W_8.425 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.479% | 2.778% | 1.809% | 3.602% | 1.189% | 5.639% | 4.785% | 2.658% | 10.56% | 9.866% | 5.834% | 4.728% | 1.529% | 2.337% | 1.973% | 23.16% | 2.515% | 5.602% | 5.683% | 1.654% | 8.672% | 3.711% | 0.2752% | 25.79% | 419.4% | 40.55% | 27.2% | 51.24% | 21.33% | 4.684% | 1.214% | 0.5854% | 0.6998% | 2.047% | 4.571% | 0.5382% | 1.08% | 3.225% | 41.74% | 23.09% | 3.693% | 2.232%
72 | 8_2_21_W_9 |  | 6/8/2023 9:00:21 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_W_9.426 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.637% | 2.294% | 1.436% | 10.59% | 2.372% | 10.12% | 13.26% | 3.096% | 3.267% | 10.93% | 1.645% | 3.725% | 2.59% | 2.497% | 21.84% | 3.7% | 5.727% | 2.702% | 3.786% | 4.799% | 7.996% | 1.157% | 2.417% | 57.28% | 479.3% | 59.4% | 66.86% | 230.2% | 100.2% | 2.297% | 1.244% | 0.9821% | 0.7714% | 2.336% | 2.308% | 2.456% | 1.057% | 2.274% | 21.2% | 22.3% | 2.425% | 1.932%
73 | 8_2_21_W_9_DUP |  | 6/8/2023 9:06:09 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_W_9_DUP.427 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.514% | 2.514% | 2.458% | 7.718% | 1.381% | 18.59% | 3.515% | 3.975% | 16.31% | 24.16% | 25.55% | 0.1574% | 0.6608% | 1.16% | 14.1% | 22.62% | 2.58% | 1.515% | 1.532% | 5.979% | 17.55% | 8.31% | 3.81% | 21.85% | 49.75% | 8.64% | 36.6% | 122.7% | 99.35% | 0.771% | 2.749% | 2.159% | 1.14% | 0.4551% | 0.1802% | 0.19% | 0.8813% | 5.538% | 29.01% | 24.95% | 8.669% | 2.659%
74 | 8_2_21_W_10 |  | 6/8/2023 9:11:57 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_W_10.428 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 5.993% | 2.14% | 1.547% | 7.441% | 1.939% | 8.761% | 9.119% | 1.084% | 4.865% | 31.85% | 5.002% | 1.71% | 1.498% | 3.144% | 5.059% | 32.76% | 6.198% | 0.9044% | 1.38% | 2.35% | 12.44% | 1.412% | 3.531% | 24.12% | 267.3% | 58.73% | 15.35% | 53.15% | 43.79% | 0.79% | 1.343% | 1.147% | 1.297% | 1.335% | 1.083% | 0.7649% | 2.394% | 2.133% | 16.75% | 8.604% | 3.275% | 4.279%
75 | 8_2_21_W_11 |  | 6/8/2023 9:17:45 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_W_11.429 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4.348% | 3.872% | 0.7744% | 4.32% | 2.541% | 17.54% | 7.674% | 5.255% | 4.357% | 11.66% | 10.89% | 1.098% | 2.898% | 2.375% | 20.48% | 20.13% | 1.695% | 2.555% | 1.134% | 0.2818% | 21.45% | 1.175% | 1.728% | 13.77% | 137.5% | 71.08% | 53.38% | 103.6% | 40.63% | 0.9915% | 1.658% | 0.6368% | 1.979% | 2.123% | 0.8546% | 2.034% | 1.398% | 2.777% | 15.91% | 24.46% | 0.4642% | 1.02%
76 | 5_10_23_DB |  | 6/8/2023 9:23:33 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_10_23_DB.430 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 153.2% |  | 2.423% | 15.35% | 0.9973% | 93.01% | 80.31% | 0% | 0% | 34.97% | 15.14% | 15.43% | 35.61% | 6.073% | 80.51% | 152.9% | 29.11% | 9.106% | 131% | 124.7% | 964% | 43.38% | 46.41% | 24.26% | 39.92% | 719.2% | 175.8% | 371% | 0.2639% | 52.55% | 0.605% | 1.723% | 0.2643% | 1.344% | 1.887% | 0.5783% | 1.941% | 3.629% | 281.3% | 240% | 7.025% | 10.19%
77 | 5_10_23_MFB |  | 6/8/2023 9:29:21 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_10_23_MFB.431 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.529% | 2.101% | 1.487% | 1.009% | 1.108% | 5.343% | 1.381% | 1.052% | 3.346% | 2.62% | 2.093% | 2.379% | 0.6781% | 2.434% | 4.293% | 3.287% | 1.636% | 1.671% | 1.835% | 2.522% | 3.678% | 2.044% | 1.672% | 3.81% | 6.15% | 3.088% | 2.853% | 2.252% | 2.244% | 2.57% | 1.574% | 0.4281% | 0.5394% | 1.838% | 1.264% | 0.1646% | 1.932% | 1.84% | 1.252% | 1.058% | 0.7097% | 2.883%
78 | 5_10_23_T-239 |  | 6/8/2023 9:35:10 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_10_23_T-239.432 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4.464% | 5.337% | 1.665% | 2.116% | 0.3797% | 15.21% | 4.437% | 3.322% | 5.35% | 6.645% | 2.836% | 0.722% | 1.029% | 6.045% | 4.63% | 3.943% | 5.77% | 2.696% | 1.046% | 2.016% | 19.8% | 3.551% | 1.993% | 2.974% | 6.237% | 5.093% | 2.05% | 6.289% | 6.921% | 0.9972% | 2.686% | 0.7626% | 1.058% | 3.463% | 0.8974% | 1.372% | 0.4926% | 2.339% | 0.3804% | 2.28% | 1.385% | 0.9152%
79 | 5_10_23_DB |  | 6/8/2023 9:40:58 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_10_23_DB.433 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 76.65% |  | 5.439% | 19.79% | 1.235% | 76.74% | 155% | 5.024% | 242.7% | 34.68% | 90.8% | 26.81% | 9.292% | 82.75% | 75.65% | 16.03% | 13.55% | 10.86% | 44.34% | 34.41% | 148.4% | 116.1% | 96.64% | 99.31% | 0% | 290.3% | 113.3% | 550% | 173.2% | 4.211% | 6360% | 3.917% | 2.06% | 3.543% | 0.8157% | 2.021% | 1.893% | 2.819% | 73.48% | 840.7% | 22.71% | 11.34%
80 | 8_16_21_UF_1 |  | 6/8/2023 9:46:46 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_UF_1.434 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 7.488% | 4.122% | 0.7001% | 8.404% | 1.331% | 5.941% | 5.77% | 4.695% | 6.488% | 17.56% | 10.37% | 2.491% | 0.6691% | 4.03% | 16.74% | 13.4% | 4.193% | 0.9116% | 2.383% | 2.57% | 6.378% | 0.9007% | 0.8375% | 13.73% | 116.3% | 35.35% | 24.23% | 3.935% | 20.68% | 2.086% | 1.147% | 2.166% | 0.2769% | 3.495% | 0.8223% | 1.747% | 0.8733% | 1.835% | 3.26% | 6.506% | 2.309% | 2.192%
81 | LBLANK |  | 6/8/2023 9:52:35 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.435 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 109.9% |  | 11.01% | 146% | 3.152% | 60.04% | 150.5% | 150.3% | 150.4% | 84.03% | 40.26% | 9.989% | 110.7% | 293.1% | 52.5% | 25.56% | 112.3% | 25.91% | 108% | 7.323% | 1236% | 7.458% | 9.852% | 30.41% | 50.18% | 133.4% | 245.7% | 3356% | 99.59% | 8.552% | 5.882% | 1.986% | 0.6458% | 1.782% | 1.358% | 0.7688% | 2.472% | 3.367% | 0% | 102.9% | 17.58% | 16.24%
82 | CCV 1ppb TRACE |  | 6/8/2023 9:58:24 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.436 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 5.932% | 6.812% | 1.012% | 13.61% | 7.376% | 13.32% | 8.663% | 7.778% | 4.342% | 4.2% | 0.9483% | 1.323% | 0.2715% | 2.787% | 0.6% | 4.913% | 7.541% | 1.693% | 2.299% | 5.757% | 12.77% | 2.34% | 3.654% | 6.929% | 12.7% | 6.975% | 4.551% | 16.27% | 4.767% | 0.3379% | 3.736% | 1.714% | 0.365% | 2.303% | 0.6137% | 0.3899% | 1.344% | 1.048% | 2.849% | 0.9931% | 1.548% | 0.5658%
83 | CCV 1ppb TRACE |  | 6/8/2023 10:04:41 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.437 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4.531% | 1.451% | 3.043% | 12.8% | 5.475% | 24.36% | 15.62% | 8.639% | 0.469% | 7.779% | 1.599% | 2.105% | 1.113% | 7.402% | 1.897% | 8.559% | 1.794% | 2.057% | 1.36% | 2.01% | 12.2% | 4.209% | 2.629% | 12.03% | 4.602% | 15.27% | 4.957% | 7.419% | 21.47% | 0.9313% | 1.867% | 1.331% | 1.256% | 1.775% | 1.681% | 10.01% | 3.655% | 3.007% | 1.935% | 1.948% | 1.191% | 10.73%
84 | LFB |  | 6/8/2023 10:10:29 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.438 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.419% | 1.351% | 2.807% | 3.5% | 1.87% | 4.304% | 4.191% | 1.709% | 0.8207% | 0.9849% | 0.2087% | 1.939% | 1.601% | 1.323% | 2.036% | 1.402% | 0.922% | 1.133% | 0.194% | 1.734% | 4.899% | 1.616% | 1.929% | 5.705% | 2.766% | 3.397% | 1.75% | 2.747% | 3.042% | 1.698% | 2.256% | 2.722% | 1.402% | 0.7825% | 0.78% | 0.2444% | 0.838% | 0.7367% | 0.2889% | 1.944% | 1.932% | 1.276%
85 | CCV 10 ppb MAJORS |  | 6/8/2023 10:16:17 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.439 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 6.297% | 2.122% | 3.307% | 4.731% | 1.117% | 3.638% | 4.271% | 5.629% | 2.25% | 2.919% | 0.8148% | 1.81% | 2.439% | 1.965% | 3.013% | 4.065% | 4.562% | 1.907% | 1.303% | 2.203% | 9.698% | 1.515% | 1.926% | 0.9092% | 1.369% | 1.105% | 2.113% | 3.765% | 3.277% | 1.15% | 2.098% | 2.14% | 0.5004% | 2.31% | 0.6927% | 1.531% | 0.5834% | 2.649% | 0.8158% | 0.4725% | 0.7264% | 0.8946%
86 | CCV 10 ppb MAJORS |  | 6/8/2023 10:22:34 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.440 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.33% | 2.463% | 1.576% | 8.954% | 1.729% | 3.701% | 3.354% | 4.032% | 1.159% | 3.066% | 0.1845% | 2.248% | 0.6753% | 4.101% | 3.149% | 4.923% | 4.064% | 1.078% | 1.477% | 1.003% | 4.055% | 3.312% | 1.421% | 2.675% | 5.02% | 2.494% | 1.449% | 2.95% | 4.033% | 1.523% | 1.448% | 0.2995% | 1.173% | 1.881% | 1.318% | 1.333% | 1.045% | 2.697% | 1.205% | 0.8558% | 1.151% | 1.287%
87 | 8_16_21_UF_6 |  | 6/8/2023 10:28:23 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_UF_6.441 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.317% | 1.538% | 1.706% | 12.76% | 0.7002% | 4% | 2.552% | 1.894% | 0.08498% | 18.04% | 9.239% | 2.734% | 0.4687% | 9.953% | 34.89% | 4.506% | 4.661% | 0.2465% | 1.248% | 1.753% | 10.37% | 1.747% | 2.266% | 11.67% | 96.03% | 10.81% | 4.491% | 10.06% | 32.08% | 1.788% | 1.236% | 0.2729% | 1.009% | 0.316% | 1.029% | 0.999% | 1.135% | 3.033% | 29.28% | 4.281% | 3.055% | 1.076%
88 | 8_16_21_UF_8 |  | 6/8/2023 10:34:11 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_UF_8.442 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.688% | 3.442% | 1.151% | 20.9% | 1.683% | 3.989% | 4.904% | 4.198% | 3.118% | 26.85% | 2.786% | 3.03% | 1.905% | 6.372% | 21.53% | 4.271% | 4.507% | 2.677% | 1.284% | 0.668% | 4.251% | 1.171% | 1.337% | 17.02% | 29.29% | 18% | 34.95% | 10.94% | 28.21% | 0.9417% | 1.818% | 0.3269% | 0.9418% | 1.159% | 0.3786% | 1.574% | 1.443% | 1.964% | 26.6% | 9.635% | 6.618% | 2.96%
89 | 8_16_21_UF_9 |  | 6/8/2023 10:39:59 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_UF_9.443 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 6.623% | 6.811% | 1.315% | 9.291% | 1.782% | 8.32% | 1.774% | 8.683% | 2.341% | 9.717% | 3.961% | 8.185% | 1.778% | 10.45% | 33.69% | 7.073% | 9.2% | 2.896% | 2.322% | 1.268% | 20.99% | 1.526% | 1.224% | 33.08% | 52.93% | 27.23% | 11.02% | 12.19% | 89.46% | 1.606% | 2.129% | 0.9286% | 1.333% | 6.494% | 2.079% | 1.027% | 0.6065% | 1.406% | 21.53% | 11.43% | 1.442% | 4.986%
90 | 8_16_21_UF_10 |  | 6/8/2023 10:45:47 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_UF_10.444 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 7.885% | 5.539% | 2.391% | 13.09% | 0.7453% | 5.333% | 7.61% | 7.98% | 3.204% | 13.19% | 1.959% | 2.642% | 0.9986% | 7.819% | 9.421% | 6.256% | 6.14% | 2.597% | 2.788% | 0.8645% | 14.44% | 2.269% | 0.473% | 18.78% | 14.05% | 27.15% | 10.89% | 37.56% | 12.06% | 1.284% | 0.8726% | 0.53% | 0.4721% | 4.474% | 0.6153% | 2.17% | 0.0812% | 2.17% | 11.33% | 9.57% | 3.789% | 3.523%
91 | 8_16_21_UF_11 |  | 6/8/2023 10:51:35 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_UF_11.445 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 7.465% | 10.23% | 2.279% | 23.99% | 1.347% | 13.96% | 10.87% | 12.78% | 11.04% | 18.66% | 2.945% | 5.969% | 1.005% | 17.2% | 13.62% | 7.418% | 8.435% | 1.996% | 2.269% | 1.486% | 15.53% | 3.319% | 1.242% | 2.993% | 62.34% | 23.88% | 17.56% | 6.288% | 22.83% | 1.182% | 1.372% | 0.6087% | 1.42% | 8.435% | 0.9888% | 1.059% | 2.268% | 3.491% | 14.49% | 23.22% | 1.452% | 2.101%
92 | 8_16_21_F_1 |  | 6/8/2023 10:57:23 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_F_1.446 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.931% | 2.33% | 1.642% | 3.238% | 3.69% | 1.294% | 4.762% | 3.639% | 9.724% | 10.39% | 5.806% | 2.646% | 2.659% | 4.6% | 13.16% | 4.142% | 4.757% | 2.191% | 4.446% | 1.465% | 2.536% | 5.811% | 2.217% | 8.116% | 23.91% | 11.85% | 14.98% | 15.66% | 38.24% | 1.506% | 3.315% | 1.826% | 2.627% | 1.547% | 2.628% | 1.181% | 4.479% | 6.253% | 12.89% | 4.364% | 3.431% | 0.7289%
93 | 8_16_21_F_1_DUP |  | 6/8/2023 11:03:11 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_F_1_DUP.447 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.063% | 2.117% | 0.8875% | 8.43% | 2.692% | 3.13% | 4.737% | 2.064% | 10.18% | 16.08% | 7.06% | 1.858% | 1.503% | 1.907% | 2.183% | 13.62% | 1.263% | 1.007% | 1.435% | 0.7415% | 5.46% | 3.216% | 1.63% | 25.44% | 22.5% | 16.98% | 22.45% | 11.31% | 46.82% | 0.2388% | 0.117% | 0.4696% | 0.4723% | 1.549% | 0.1372% | 0.9902% | 2.864% | 2.161% | 2.561% | 23.22% | 8.244% | 6.372%
94 | 8_16_21_F_1_DUP_LD |  | 6/8/2023 11:08:58 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_F_1_DUP_LD.448 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 7.22% | 7.09% | 1.733% | 7.228% | 2.136% | 5.931% | 6.286% | 5.924% | 4.488% | 19.18% | 6.199% | 4.522% | 0.7321% | 5.665% | 31.3% | 8.529% | 4.304% | 0.7864% | 1.321% | 1.021% | 8.662% | 1.568% | 1.039% | 18.35% | 29.06% | 19.56% | 8.531% | 20.51% | 34.8% | 0.9258% | 0.8879% | 0.3947% | 0.9927% | 6.076% | 1.261% | 0.9871% | 0.8203% | 1.144% | 20.25% | 15.11% | 2.636% | 5.371%
95 | 8_16_21_F_1_DUP_LD |  | 6/8/2023 11:15:14 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_F_1_DUP_LD.449 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4.359% | 4.735% | 1.269% | 5.458% | 0.4498% | 4.84% | 3.812% | 3.229% | 9.483% | 31.65% | 6.32% | 2.299% | 2.645% | 4.766% | 26.64% | 4.071% | 2.129% | 3.009% | 5.388% | 0.367% | 5.757% | 0.454% | 1.464% | 22.25% | 60.07% | 36.24% | 7.976% | 14.62% | 35.74% | 2.2% | 1.393% | 0.09173% | 1.145% | 3.099% | 2.351% | 0.7248% | 0.5667% | 0.773% | 13.32% | 11.85% | 2.557% | 2.934%
96 | 8_16_21_F_1_DUP_LSP |  | 6/8/2023 11:21:03 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_F_1_DUP_LSP.450 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.815% | 0.5156% | 1.164% | 1.857% | 1.298% | 2.03% | 2.336% | 2.567% | 0.1337% | 2.218% | 1.907% | 3.218% | 0.7808% | 0.8306% | 2.039% | 2.035% | 3.29% | 1.133% | 1.373% | 0.6029% | 1.181% | 1.181% | 0.7572% | 1.868% | 5.692% | 1.999% | 1.748% | 2.769% | 7.838% | 0.6824% | 1.597% | 0.7187% | 0.7526% | 1.516% | 1.123% | 1.947% | 1.145% | 0.8549% | 1.722% | 1.773% | 0.8638% | 3.789%
97 | LBLANK |  | 6/8/2023 11:26:52 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.451 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 74.65% | 173.2% | 36.24% | 298.5% | 52.7% | 25.14% | 118.7% | 1099% | 1080% | 135.6% | 44.54% | 14.96% | 63.84% | 43.52% | 957.7% | 51.21% | 98.45% | 14.79% | 129.4% | 3037% | 1080% | 14.64% | 15.01% | 36.06% | 47.44% | 69.97% | 36.47% | 71.42% | 100.6% | 14.95% | 11.01% | 9.862% | 1.758% | 1.371% | 0.7345% | 0.6461% | 0.6941% | 2.442% | 1156% | 40.71% | 190.9% | 34.99%
98 | CCV 1ppb TRACE |  | 6/8/2023 11:32:40 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.452 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4.243% | 4.696% | 0.251% | 29.43% | 1.277% | 8.193% | 4.839% | 9.908% | 1.519% | 6.872% | 2.263% | 2.322% | 1.631% | 14.83% | 3.575% | 8.216% | 1.161% | 3.434% | 2.666% | 4.269% | 11.7% | 2.658% | 0.5207% | 6.016% | 24.62% | 8.063% | 7.696% | 10.8% | 8.001% | 2.188% | 0.8637% | 1.712% | 1.207% | 1.071% | 1.81% | 1.001% | 1.833% | 2.31% | 2.794% | 2.239% | 0.686% | 0.3067%
99 | CCV 1ppb TRACE |  | 6/8/2023 11:38:58 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.453 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.539% | 1.799% | 4.197% | 17.73% | 4.145% | 15.74% | 1.651% | 11.08% | 2.768% | 1.443% | 2.358% | 2.02% | 1.84% | 2.386% | 2.855% | 2.703% | 1.988% | 1.204% | 2.14% | 5.099% | 19.79% | 2.851% | 2.264% | 9.708% | 8.994% | 5.183% | 7.117% | 2.602% | 2.639% | 2.38% | 0.7494% | 1.523% | 0.9683% | 1.687% | 1.931% | 0.5873% | 1.156% | 1.603% | 1.193% | 1.474% | 1.201% | 1.447%
100 | CCV 10 ppb MAJORS |  | 6/8/2023 11:44:46 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.454 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.557% | 2.313% | 0.4251% | 0.5362% | 0.7993% | 0.7861% | 3.267% | 1.224% | 1.906% | 3.073% | 1.074% | 3.222% | 1.35% | 0.945% | 1.185% | 2.515% | 1.949% | 1.407% | 1.4% | 1.693% | 8.073% | 2.69% | 1.566% | 3.519% | 3.818% | 0.9127% | 3.742% | 3.142% | 2.019% | 1.442% | 1.014% | 1.337% | 0.7502% | 1.837% | 1.87% | 0.8617% | 1.188% | 3.266% | 2.369% | 1.789% | 0.8114% | 0.6659%
101 | CCV 10 ppb MAJORS |  | 6/8/2023 11:51:04 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.455 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.096% | 4.949% | 1.691% | 6.368% | 1.831% | 8.224% | 6.594% | 7.006% | 4.52% | 5.354% | 1.677% | 2.107% | 0.9472% | 4.795% | 5.133% | 6.708% | 5.049% | 1.647% | 0.7924% | 1.87% | 7.225% | 0.8273% | 2.029% | 1.193% | 4.678% | 2.953% | 1.326% | 1.36% | 4.715% | 0.7298% | 0.4553% | 1.83% | 0.8537% | 4.35% | 1.445% | 1.142% | 0.9676% | 0.6513% | 0.88% | 0.6275% | 1.501% | 1.009%
102 | 8_16_21_F_6 |  | 6/8/2023 11:56:52 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_F_6.456 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.249% | 0.914% | 4.454% | 7.702% | 1.123% | 2.543% | 5.786% | 2.55% | 4.972% | 7.606% | 8.583% | 0.6486% | 1.677% | 4.034% | 17.35% | 3.816% | 1.787% | 2.567% | 0.967% | 4.088% | 5.446% | 3.366% | 0.6581% | 20.66% | 32.7% | 44.82% | 7.144% | 13.07% | 17.38% | 0.7187% | 1.712% | 4.12% | 1.066% | 0.5566% | 1.839% | 0.4059% | 0.9384% | 3.185% | 5.611% | 6.952% | 5.313% | 1.064%
103 | 8_16_21_F_6_SP |  | 6/9/2023 12:02:41 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_F_6_SP.457 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.596% | 2.171% | 2.727% | 10.21% | 1.623% | 5.434% | 4.565% | 2.918% | 5.644% | 3.611% | 1.537% | 1.211% | 2.505% | 2.417% | 4.062% | 3.471% | 3.672% | 1.68% | 1.183% | 0.1546% | 3.341% | 2.096% | 1.55% | 0.6819% | 9.481% | 6.101% | 1.331% | 3.628% | 5.645% | 1.153% | 2.253% | 0.9558% | 1.787% | 3.645% | 1.124% | 0.8169% | 1.191% | 1.719% | 1.383% | 2.12% | 0.6514% | 0.7937%
104 | 8_16_21_F_8 |  | 6/9/2023 12:08:29 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_F_8.458 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 5.914% | 7.889% | 1.928% | 1.846% | 2.471% | 4.528% | 10.08% | 5.186% | 3.66% | 23.56% | 3.517% | 1.73% | 1.126% | 7.492% | 29.29% | 9.885% | 4.234% | 2.079% | 1.389% | 1.888% | 3.42% | 0.672% | 1.872% | 20.61% | 171.4% | 4.655% | 11.33% | 8.752% | 18.07% | 1.571% | 1.434% | 0.7318% | 1.297% | 4.947% | 1.349% | 1.596% | 0.583% | 1.253% | 14.03% | 7.158% | 1.49% | 3.244%
105 | 8_16_21_F_8_DUP |  | 6/9/2023 12:14:18 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_F_8_DUP.459 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.9234% | 4.921% | 1.472% | 3.842% | 2.041% | 3.072% | 4.474% | 1.968% | 4.698% | 11.08% | 4.704% | 1.743% | 1.067% | 3.367% | 2.443% | 7.589% | 2.55% | 1.42% | 1.346% | 1.028% | 7.128% | 2.677% | 2.127% | 8.038% | 42.74% | 18.26% | 8.336% | 13.02% | 40.19% | 0.8832% | 0.3155% | 0.7788% | 0.937% | 1.43% | 1.107% | 2.573% | 1.037% | 2.029% | 5.398% | 5.854% | 2.967% | 1.744%
106 | 8_16_21_F_9 |  | 6/9/2023 12:20:06 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_F_9.460 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.5444% | 1.614% | 1.786% | 2.605% | 2.558% | 4.242% | 1.301% | 3.29% | 1.169% | 8.207% | 1.555% | 1.404% | 0.294% | 1.974% | 27.7% | 1.666% | 1.111% | 0.4237% | 1.93% | 2.287% | 9.429% | 2.838% | 1.177% | 1.859% | 8.675% | 11.02% | 16.25% | 7.279% | 45.19% | 1.203% | 1.049% | 1.521% | 0.5233% | 0.8233% | 0.204% | 0.2324% | 1.431% | 3.191% | 19.98% | 18.08% | 4.409% | 1.789%
107 | 8_16_21_F_10 |  | 6/9/2023 12:25:54 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_F_10.461 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.913% | 0.4372% | 0.9428% | 4.437% | 1.274% | 2.994% | 3.133% | 1.806% | 4.732% | 14.96% | 5.398% | 1.345% | 1.182% | 3.325% | 25.16% | 9.536% | 1.982% | 1.171% | 1.591% | 0.5987% | 9.733% | 3.556% | 1.8% | 23.58% | 51.78% | 14.12% | 34.65% | 17.26% | 28.85% | 1.191% | 1.268% | 0.358% | 2.012% | 1.344% | 1.583% | 1.291% | 0.6617% | 3.59% | 10.94% | 17.8% | 1.254% | 2.74%
108 | 8_16_21_F_10_DUP |  | 6/9/2023 12:31:42 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_F_10_DUP.462 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4.477% | 2.132% | 0.3774% | 20.41% | 2.986% | 1.732% | 5.503% | 2.985% | 6.769% | 25.34% | 5.668% | 5.3% | 5.376% | 4.321% | 4.565% | 10.73% | 0.6818% | 3.665% | 4.839% | 0.7739% | 8.691% | 5.934% | 5.972% | 6.589% | 63.43% | 12.36% | 8.214% | 15.81% | 23.66% | 5.694% | 5.316% | 0.8373% | 4.53% | 1.341% | 3.78% | 1.365% | 2.745% | 4.607% | 4.769% | 7.759% | 1.337% | 1.066%
109 | 8_16_21_F_11 |  | 6/9/2023 12:37:30 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_F_11.463 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 6.922% | 7.144% | 0.6181% | 1.912% | 1.847% | 4.522% | 4.62% | 4.046% | 0.4546% | 8.056% | 8.257% | 0.8437% | 1.91% | 5.892% | 9.451% | 15.68% | 6.711% | 0.4502% | 1.279% | 0.6046% | 4.058% | 1.055% | 0.9353% | 9.138% | 66.33% | 38.27% | 5.389% | 6.388% | 23.34% | 2.527% | 1.852% | 1.65% | 1.315% | 4.767% | 1.686% | 0.7693% | 1.186% | 0.6053% | 10.28% | 23.45% | 2.481% | 0.9269%
110 | 8_16_21_W_1 |  | 6/9/2023 12:43:18 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_W_1.464 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10.18% | 9.068% | 1.379% | 11.5% | 2.333% | 10.14% | 10.13% | 4.424% | 15.1% | 39.76% | 21.16% | 11.65% | 0.7786% | 10.44% | 49.9% | 7.546% | 7.895% | 11.78% | 16.74% | 4.108% | 11.24% | 2.385% | 1.925% | 20.21% | 105.6% | 31.66% | 32.84% | 43.32% | 86.6% | 12.41% | 0.8025% | 1.295% | 0.2771% | 5.899% | 10.25% | 0.9086% | 2.217% | 1.242% | 39.92% | 13.98% | 7.318% | 7.522%
111 | 8_16_21_W_6 |  | 6/9/2023 12:49:05 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_W_6.465 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 17.41% | 18.73% | 3.036% | 18.15% | 1.835% | 26.72% | 28.05% | 20.03% | 26.02% | 17.15% | 14.25% | 1.617% | 0.614% | 22.79% | 39.17% | 16.9% | 18.36% | 0.3806% | 0.6414% | 2.707% | 32.76% | 2.558% | 0.3124% | 18.78% | 125.8% | 22.24% | 6.929% | 43% | 25.53% | 1.214% | 2.266% | 2.28% | 0.9384% | 16.91% | 1.167% | 0.8165% | 1.573% | 1.469% | 29.8% | 19.63% | 7.6% | 3.121%
112 | LBLANK |  | 6/9/2023 12:54:53 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.466 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 245.4% |  | 10.88% | 56.3% | 8.668% | 57.55% | 147.6% | 0% | 0% | 91.34% | 22.07% | 66.3% | 50.61% | 148.7% | 52.55% | 784.1% | 256% | 12.61% | 355.9% | 55.37% | 0% | 98.66% | 114.5% | 79.75% | 0% | 374.7% | 329.2% | 148.8% | 99.66% | 19.09% | 39.15% | 1.964% | 1.029% | 1.548% | 1.078% | 1.524% | 2.507% | 2.19% | 292.7% | 66.3% | 121.1% | 50.37%
113 | CCV 1ppb TRACE |  | 6/9/2023 1:00:42 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.467 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4.745% | 13.12% | 1.835% | 20.19% | 2.668% | 26.61% | 3.545% | 10.91% | 4.67% | 3.536% | 1.946% | 2.612% | 2.448% | 6.434% | 2.295% | 7.411% | 3.27% | 0.5652% | 1.944% | 2.832% | 10.5% | 4.24% | 1.012% | 7.464% | 7.374% | 6.525% | 5.874% | 13.35% | 0.9593% | 1.214% | 0.5415% | 1.548% | 0.9687% | 3.05% | 1.976% | 0.9733% | 0.9663% | 3.391% | 3.054% | 0.8172% | 1.031% | 0.9329%
114 | CCV 1ppb TRACE |  | 6/9/2023 1:06:59 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.468 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 8.022% | 10.18% | 0.255% | 11.11% | 1.645% | 12.82% | 10.09% | 12.94% | 5.901% | 6.25% | 3.888% | 5.269% | 1.562% | 9.014% | 1.891% | 1.817% | 4.413% | 3.612% | 2.401% | 4.975% | 10.49% | 1.449% | 2.074% | 4.103% | 6.13% | 3.187% | 4.759% | 8.628% | 6.314% | 3.72% | 3.585% | 1.173% | 1.962% | 2.863% | 3.339% | 1.21% | 1.275% | 3.203% | 2.485% | 4.301% | 0.8842% | 0.1669%
115 | LFB |  | 6/9/2023 1:12:47 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.469 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.899% | 3.383% | 1.907% | 11.16% | 0.9598% | 3.03% | 1.446% | 4.258% | 3.191% | 2.563% | 0.8826% | 1.372% | 0.754% | 3.193% | 3.377% | 2.75% | 3.371% | 0.6162% | 1.136% | 0.686% | 6.076% | 2.859% | 0.8803% | 2.941% | 3.061% | 5.303% | 1.386% | 1.989% | 2.697% | 1.37% | 1.327% | 2.58% | 1.871% | 1.793% | 0.279% | 1.137% | 2.477% | 2.268% | 0.6576% | 0.9301% | 0.8984% | 1.474%
116 | CCV 10 ppb MAJORS |  | 6/9/2023 1:18:35 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.470 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.9854% | 4.547% | 1.567% | 15.22% | 3.553% | 3.854% | 3.385% | 3.518% | 2.79% | 1.129% | 0.3295% | 0.6844% | 1.649% | 1.813% | 1.302% | 1.61% | 1.898% | 1.524% | 1.936% | 1.879% | 3.906% | 1.16% | 0.3743% | 2.1% | 2.9% | 2.222% | 2.557% | 2.679% | 4.618% | 1.624% | 0.8322% | 1.314% | 0.7835% | 1.261% | 1.627% | 0.504% | 0.3669% | 2.207% | 2.466% | 2.695% | 0.9916% | 0.4527%
117 | CCV 10 ppb MAJORS |  | 6/9/2023 1:24:53 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.471 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.732% | 3.171% | 1.835% | 0.2003% | 2.696% | 3.724% | 4.619% | 2.422% | 1.452% | 0.57% | 0.7688% | 1.249% | 0.5528% | 0.1745% | 0.4319% | 2.315% | 2.805% | 1.18% | 0.3052% | 2.288% | 2.124% | 0.9044% | 0.3816% | 1.314% | 2.564% | 1.257% | 3.439% | 2.711% | 3.773% | 0.2429% | 0.9804% | 0.4798% | 0.7623% | 1.984% | 0.6755% | 1.288% | 1.023% | 0.9139% | 1.279% | 0.5764% | 0.587% | 1.856%
118 | 8_16_21_W_6_DUP |  | 6/9/2023 1:30:41 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_W_6_DUP.472 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.5458% | 4.276% | 1.606% | 3.713% | 7.748% | 4.4% | 8.223% | 3.617% | 10.55% | 4.148% | 2.029% | 2.783% | 1.332% | 2.444% | 11.92% | 39.47% | 2.693% | 3.125% | 0.9322% | 2.501% | 19.92% | 1.401% | 0.9074% | 24.42% | 41.18% | 69.4% | 23.3% | 32.26% | 50.85% | 1.959% | 1.236% | 0.8148% | 1.833% | 2.474% | 1.861% | 2.343% | 5.142% | 3.532% | 33.17% | 22.22% | 5.213% | 0.8614%
119 | 8_16_21_W_8 |  | 6/9/2023 1:36:29 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_W_8.473 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.878% | 10.32% | 2.893% | 5.285% | 0.5869% | 12.66% | 1.524% | 10.05% | 2.034% | 25.69% | 4.526% | 1.571% | 1.823% | 7.497% | 18.76% | 22.64% | 8.664% | 0.8849% | 1.295% | 2.241% | 17.34% | 4.119% | 1.095% | 10% | 92.79% | 66.01% | 25.51% | 24.72% | 91.83% | 0.8536% | 0.8626% | 0.8363% | 0.7829% | 4.818% | 1.071% | 1.594% | 1.305% | 3.35% | 14.38% | 10.42% | 1.13% | 0.6807%
120 | 8_16_21_W_9 |  | 6/9/2023 1:42:18 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_W_9.474 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.801% | 3.827% | 1.153% | 1.041% | 1.257% | 3.448% | 4.028% | 1.322% | 5.587% | 12.21% | 12.24% | 8.822% | 3.011% | 3.13% | 19.03% | 46.42% | 7.486% | 10.33% | 7.215% | 5.409% | 11.67% | 2.263% | 4.314% | 35.98% | 157% | 11.98% | 12.36% | 35.78% | 35.64% | 8.973% | 3.466% | 0.6867% | 1.875% | 1.902% | 8.411% | 0.327% | 0.898% | 3.381% | 27.11% | 35.14% | 1.094% | 1.206%
121 | 8_16_21_W_9_DUP |  | 6/9/2023 1:48:06 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_W_9_DUP.475 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.975% | 3.595% | 1.614% | 6.946% | 1.248% | 9.516% | 9.411% | 1.964% | 0.5556% | 19.02% | 5.202% | 0.663% | 0.7941% | 1.611% | 11.84% | 36.39% | 2.025% | 1.334% | 2.734% | 1.873% | 15.97% | 2.975% | 2.981% | 29.52% | 0% | 51.83% | 15.28% | 576.7% | 49.7% | 2.482% | 1.502% | 2.945% | 0.568% | 2.032% | 1.667% | 1.656% | 0.7095% | 3.362% | 20.75% | 11.77% | 2.721% | 0.7271%
122 | 8_16_21_W_10 |  | 6/9/2023 1:53:55 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_W_10.476 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.526% | 5.887% | 1.792% | 7.749% | 1.14% | 7.176% | 6.988% | 5.221% | 10.56% | 8.213% | 3.134% | 0.5876% | 0.7086% | 1.493% | 49.97% | 16.25% | 2.08% | 2.541% | 1.883% | 1.892% | 2.94% | 1.887% | 0.8372% | 32.33% | 158.4% | 80.18% | 21.4% | 54.67% | 99.87% | 2.814% | 0.7979% | 0.4395% | 1.074% | 0.4996% | 1.338% | 1.091% | 1.271% | 2.239% | 7.63% | 15.57% | 3.636% | 0.6883%
123 | 8_16_21_W_11 |  | 6/9/2023 1:59:43 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_W_11.477 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.058% | 1.323% | 0.6305% | 7.988% | 1.514% | 14.12% | 13.19% | 3.057% | 5.517% | 31.3% | 9.45% | 1.696% | 0.8448% | 1.769% | 72.88% | 40.04% | 5.578% | 0.7852% | 4.863% | 3.064% | 19.7% | 1.664% | 2.162% | 19.72% | 3611% | 17.69% | 22.16% | 45.58% | 25.78% | 0.2656% | 2.934% | 1.254% | 1.455% | 0.7518% | 0.8231% | 13.11% | 1.793% | 3.945% | 15.68% | 29.33% | 5.632% | 13.1%
124 | 5_10_23_DB |  | 6/9/2023 2:05:32 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_10_23_DB.478 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0% |  | 2.645% | 44.74% | 0.7458% | 150% | 85.53% | 0% | 0% | 25.99% | 41.29% | 10.35% | 39.73% | 18.58% | 79.41% | 141.2% | 24.52% | 4.261% | 89.99% | 56.99% | 1098% | 25.4% | 28.03% | 15.8% | 40.99% | 358.1% | 169.3% | 33.61% | 86.6% | 81.62% | 458.1% | 0.6514% | 0.3468% | 4.246% | 1.196% | 0.8099% | 1.351% | 2.522% | 0% | 992.9% | 66.63% | 32.37%
125 | 5_10_23_MFB |  | 6/9/2023 2:11:19 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_10_23_MFB.479 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.631% | 3.009% | 0.6148% | 4.71% | 1.508% | 1.107% | 4.23% | 5.917% | 0.9735% | 2.478% | 2.408% | 1.47% | 1.514% | 3.557% | 3.019% | 2.918% | 2.338% | 3.409% | 1.612% | 0.8835% | 4.956% | 3.175% | 1.269% | 4.06% | 6.805% | 2.137% | 0.4931% | 2.109% | 4.287% | 2.748% | 0.7478% | 0.9238% | 1.177% | 0.4866% | 2.302% | 2.018% | 1.174% | 3.21% | 3.288% | 2.639% | 1.523% | 1.538%
126 | 5_10_23_T-239 |  | 6/9/2023 2:17:07 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_10_23_T-239.480 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.58% | 4.2% | 3.763% | 1.93% | 1.093% | 10.88% | 0.3535% | 1.741% | 2.253% | 1.41% | 0.5083% | 0.5513% | 0.9226% | 2.211% | 3.613% | 5.186% | 3.229% | 0.3065% | 0.3065% | 5.227% | 15.4% | 2.001% | 1.455% | 4.974% | 5.134% | 6.749% | 2.814% | 9.254% | 5.266% | 3.374% | 0.7845% | 3.786% | 0.4834% | 1.372% | 0.2909% | 1.177% | 0.1404% | 1.484% | 0.3196% | 0.615% | 3.394% | 0.7923%
127 | 5_10_23_DB |  | 6/9/2023 2:22:55 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_10_23_DB.481 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 71.86% | 173.2% | 6.395% | 6.755% | 1.756% | 42.41% | 442.1% | 147.5% | 51.51% | 30.69% | 24.01% | 7.584% | 2.09% | 13.36% | 58.05% | 62.54% | 2.579% | 7.344% | 74.68% | 151.5% | 0% | 2339% | 728% | 28.57% | 38.92% | 0% | 187% | 323.1% | 86.62% | 27.06% | 466.8% | 5.084% | 1.435% | 3.572% | 0.4082% | 0.9682% | 1.095% | 3.749% | 954.8% | 1291% | 5.141% | 8.206%
128 | LBLANK |  | 6/9/2023 2:28:44 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.482 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 146.8% |  | 0.2561% | 904.9% | 2.874% | 51.42% | 316% | 147.6% | 969.4% | 50.72% | 64.64% | 22.29% | 227.8% | 74.19% | 0% | 280.5% | 875.8% | 44.61% | 869.1% | 57.12% | 0% | 45.58% | 58.59% | 49.98% | 105.4% | 694.9% | 49.46% | 643.3% |  | 21.34% | 13.87% | 1.952% | 0.4323% | 1.415% | 1.112% | 1.548% | 0.4756% | 4.293% | 0% | 420.3% | 52.7% | 32.14%
129 | CCV 1ppb TRACE |  | 6/9/2023 2:34:32 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.483 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 6.325% | 3.578% | 10.4% | 4.594% | 2.191% | 11.46% | 3.852% | 7.552% | 6.779% | 2.08% | 1.39% | 3.766% | 2.115% | 3.684% | 1.086% | 0.6216% | 2.332% | 2.764% | 3.081% | 11.47% | 10.84% | 4.025% | 1.175% | 18.68% | 14.2% | 4.03% | 5.135% | 25.49% | 7.468% | 1.623% | 1.848% | 8.779% | 1.578% | 2.068% | 1.398% | 0.705% | 0.3807% | 2.378% | 1.806% | 3.351% | 10.87% | 1.049%
130 | CCV 1ppb TRACE |  | 6/9/2023 2:40:49 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.484 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.5563% | 3.893% | 0.8642% | 7.813% | 2.302% | 20.45% | 4.493% | 13.72% | 9.111% | 1.687% | 3.438% | 2.089% | 3.453% | 4.56% | 0.09559% | 9.101% | 2.01% | 2.801% | 2.099% | 1.577% | 8.729% | 4.71% | 1.334% | 4.59% | 20.24% | 12.52% | 7.124% | 7.039% | 12.68% | 2.57% | 2.009% | 0.9302% | 2.163% | 1.547% | 1.608% | 0.2209% | 0.6941% | 5.292% | 1.852% | 0.9516% | 1.038% | 1.211%
131 | CCV 10 ppb MAJORS |  | 6/9/2023 2:46:38 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.485 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.237% | 3.893% | 0.4106% | 9.06% | 1.038% | 2.002% | 4.331% | 9.171% | 3.941% | 2.605% | 2.506% | 1.708% | 0.6716% | 1.699% | 4.483% | 3.456% | 2.421% | 1.087% | 0.9885% | 1.592% | 7.147% | 3.568% | 0.7176% | 0.3764% | 1.179% | 1.739% | 1.423% | 3.424% | 2.419% | 2.418% | 0.5999% | 1.519% | 0.8867% | 3.707% | 0.735% | 0.6177% | 1.684% | 3.675% | 1.27% | 2.01% | 1.785% | 1.372%
132 | CCV 10 ppb MAJORS |  | 6/9/2023 2:52:55 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.486 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.7909% | 3.809% | 7.661% | 3.211% | 1.741% | 4.462% | 1.268% | 3.38% | 1.482% | 1.028% | 1.719% | 0.2523% | 1.2% | 0.2808% | 0.9852% | 1.616% | 1.572% | 0.2983% | 0.9751% | 8.433% | 3.851% | 1.287% | 0.5622% | 9.005% | 3.16% | 3.874% | 0.713% | 8.967% | 4.597% | 1.192% | 0.7655% | 7.42% | 0.5892% | 0.8202% | 0.7148% | 0.689% | 0.8504% | 2.359% | 0.8013% | 0.7988% | 8.416% | 1.884%
133 | MT 34 D |  | 6/9/2023 2:58:43 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\MT 34 D.487 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 114.7% |  | 0.6219% | 19.82% | 2.885% | 55.84% | 185.9% | 101.3% | 111.7% | 16.64% | 27.96% | 2.936% | 26.36% | 38.31% | 87.89% | 49.61% | 22.93% | 14.68% | 9.116% | 28.84% | 0% | 229.1% | 117.4% | 42.13% | 0% | 396.2% | 1979% | 27.61% |  | 6.008% | 9.382% | 0.8334% | 1.127% | 2.42% | 1.282% | 1.715% | 2.209% | 1.242% | 266.8% | 39.2% | 5.354% | 5.429%
134 | MT 35 D |  | 6/9/2023 3:04:31 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\MT 35 D.488 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 152.1% |  | 1.029% | 6.27% | 1.28% | 71.25% | 55.8% | 221.3% | 1186% | 32.74% | 2.593% | 12.89% | 37.92% | 25.02% | 51.48% | 47.02% | 25.06% | 7.27% | 19.84% | 352.5% | 1186% | 332.9% | 627.6% | 55.07% | 39.66% | 0% | 183.5% | 439.5% | 114.8% | 23.84% | 120.4% | 1.242% | 0.4161% | 1.153% | 1.601% | 0.6902% | 1.27% | 3.417% | 291.6% | 45.91% | 6.544% | 3.91%
135 | MT 36 D |  | 6/9/2023 3:10:18 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\MT 36 D.489 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 81.98% |  | 0.1147% | 22.64% | 0.3555% | 41.41% | 427.6% | 109.5% | 0% | 8.216% | 2.852% | 7.212% | 5.183% | 57.15% | 0% | 56.83% | 8.974% | 10.01% | 3.228% | 77.29% | 0% | 3531% | 713.6% | 25.6% | 39.6% | 321.9% | 102.3% | 49.05% | 173.2% | 23.25% | 18.72% | 0.8744% | 0.1385% | 1.874% | 1.032% | 1.176% | 1.504% | 3.88% | 71.32% | 43.48% | 4.832% | 5.505%
136 | MT 37 D |  | 6/9/2023 3:16:06 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\MT 37 D.490 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 33.57% | 173.2% | 2.068% | 22.38% | 0.4407% | 54.54% | 68.36% | 147.2% | 0% | 18.87% | 11.9% | 30.92% | 46.34% | 41.64% | 1868% | 104% | 60.32% | 5.986% | 6.734% | 24.69% | 0% | 27.63% | 33.62% | 22.71% | 0% | 356.7% | 126.8% | 357.9% |  | 24.23% | 12.27% | 0.5003% | 1.251% | 2.686% | 2.011% | 1.692% | 2.203% | 1.899% | 145.7% | 48.77% | 2.157% | 8.849%
137 | MT 38 D |  | 6/9/2023 3:21:55 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\MT 38 D.491 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 34.56% | 173.2% | 0.8437% | 5.404% | 0.5572% | 25.56% | 62.5% | 59.54% | 296.8% | 26.62% | 11.53% | 3.277% | 5.044% | 3.756% | 117.7% | 52.08% | 8.124% | 3.212% | 1.128% | 118.1% | 2.503% | 301.9% | 587.5% | 52.35% | 0% | 176.7% | 267.2% | 112.8% |  | 31.08% | 18.53% | 0.7602% | 0.7168% | 1.825% | 0.895% | 1.248% | 1.445% | 0.5458% | 190% | 12.72% | 9.594% | 1.286%
138 | MT 39 D |  | 6/9/2023 3:27:43 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\MT 39 D.492 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 40.73% |  | 1.024% | 12.51% | 1.271% | 44.67% | 229.9% | 2.002% | 55.05% | 27.63% | 24.72% | 9.206% | 20.81% | 13.62% | 53.9% | 288.1% | 27.01% | 11.8% | 16.19% | 507% | 0% | 31.62% | 36.15% | 17.72% | 0% | 728% | 330.4% | 659.2% |  | 16.21% | 20.73% | 0.8961% | 1.76% | 1.431% | 6.407% | 0.4438% | 0.8516% | 0.5133% | 144.4% | 40.99% | 1.843% | 2.398%
139 | MT 40 D |  | 6/9/2023 3:33:31 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\MT 40 D.493 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 140.5% |  | 0.6638% | 15.28% | 0.7104% | 57.12% | 84.25% | 140.4% | 0% | 30.98% | 13.59% | 4.34% | 8.977% | 31.75% | 166.1% | 72.36% | 37.5% | 19.68% | 4.382% | 6018% | 0% | 50.68% | 66.77% | 23.6% | 0% | 340.9% | 90.54% | 45.12% |  | 31.69% | 15.9% | 0.1994% | 1.359% | 2.365% | 0.7701% | 4.341% | 1.556% | 2.984% | 68.46% | 31.05% | 5.345% | 5.134%
140 | MT 32 D |  | 6/9/2023 3:39:20 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\MT 32 D.494 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 79.85% |  | 0.8422% | 23.44% | 1.148% | 93.65% | 141.6% | 36.42% | 293.2% | 64.48% | 17.03% | 2.141% | 19.63% | 36.28% | 0% | 202.9% | 7.72% | 3.184% | 8.864% | 92.49% | 954.2% | 143.2% | 75.87% | 43.36% | 53.58% | 343.2% | 309.6% | 521.7% |  | 22.11% | 7.637% | 1.434% | 1.689% | 5.538% | 0.9985% | 0.9056% | 2.31% | 2.124% | 461.9% | 26.16% | 11.78% | 1.813%
141 | MT 32 D LD |  | 6/9/2023 3:45:07 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\MT 32 D LD.495 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 41.49% |  | 4.219% | 12.78% | 3.277% | 30.49% | 1454% | 293.9% | 0% | 60.7% | 15.49% | 11.01% | 17.95% | 68.99% | 55.75% | 207.4% | 18.22% | 4.973% | 5.479% | 1248% | 0% | 128% | 190.7% | 17.37% | 42.49% | 349.6% | 439.6% | 222.6% |  | 14.59% | 4.702% | 1.454% | 1.796% | 0.07034% | 0.9252% | 11.27% | 1.47% | 1.56% | 0% | 36.96% | 7.047% | 11.72%
142 | MT 32 D LD |  | 6/9/2023 3:51:24 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\MT 32 D LD.496 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 143.5% | 173.2% | 0.5277% | 25.28% | 1.279% | 8.61% | 591.9% | 146.7% | 962.5% | 13.9% | 10.57% | 10.43% | 24.84% | 28.64% | 0% | 123.9% | 25.45% | 5.758% | 2.916% | 10.3% | 0% | 111.5% | 249.6% | 14.91% | 108.8% | 347.8% | 117.6% | 225.2% | 114.1% | 17.53% | 27.76% | 0.9039% | 4.251% | 0.6272% | 1.013% | 0.4387% | 2.442% | 7.927% | 118.8% | 66.31% | 3.392% | 1.873%
143 | LBLANK |  | 6/9/2023 3:57:12 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.497 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 138.4% |  | 9.024% | 241.4% | 3.356% | 44.92% | 272% | 1135% | 1135% | 28.75% | 26.79% | 51.13% | 37.41% | 1738% | 0% | 357.7% | 44.8% | 20.52% | 823.7% | 49.52% | 0% | 35.18% | 40.38% | 47.61% | 56.01% | 344.6% | 258.9% | 416.6% |  | 32.06% | 88.3% | 0.8985% | 12.04% | 1.277% | 5.761% | 1.625% | 2.086% | 11% | 413.4% | 582.3% | 47.53% | 56.46%
144 | CCV 1ppb TRACE |  | 6/9/2023 4:03:01 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.498 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.967% | 2.799% | 1.848% | 13.26% | 4.968% | 5.318% | 2.372% | 5.749% | 6.03% | 3.911% | 1.053% | 3.957% | 1.279% | 6.698% | 2.295% | 1.847% | 2.365% | 0.7403% | 1.803% | 4.425% | 13.1% | 4.523% | 2.864% | 10.4% | 5.154% | 3.463% | 1.912% | 19.49% | 17.32% | 1.236% | 1.445% | 2.215% | 0.591% | 1.411% | 0.5086% | 0.8483% | 1.025% | 1.69% | 1.115% | 1.36% | 1.294% | 0.5138%
145 | CCV 1ppb TRACE |  | 6/9/2023 4:09:18 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.499 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.3858% | 3.137% | 1.026% | 19.63% | 2.071% | 15.02% | 4.796% | 13.26% | 3.254% | 2.024% | 2.864% | 2.177% | 1.581% | 2.064% | 2.521% | 3.255% | 2.66% | 1.942% | 2.99% | 3.684% | 14.26% | 2.535% | 1.333% | 9.848% | 13.54% | 6.886% | 10.72% | 22.81% | 2.32% | 2.193% | 2.304% | 2.728% | 0.5777% | 2.163% | 0.6755% | 0.9018% | 1.448% | 3.282% | 1.827% | 0.6007% | 3.407% | 0.7143%
146 | LFB |  | 6/9/2023 4:15:06 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.500 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.267% | 2.333% | 1.665% | 4.59% | 0.9123% | 2.308% | 2.478% | 3.987% | 1.208% | 3.174% | 1.434% | 0.9372% | 3.586% | 1.869% | 0.5544% | 2.492% | 0.8433% | 0.7069% | 0.3707% | 1.31% | 6.33% | 1.837% | 1.574% | 1.318% | 6.322% | 1.548% | 1.687% | 3.353% | 9.506% | 0.8252% | 1.804% | 1.636% | 2.248% | 1.032% | 0.7409% | 1.497% | 1.141% | 1.784% | 2.098% | 0.4071% | 1.829% | 1.052%
147 | LFB |  | 6/9/2023 4:21:22 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.501 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.553% | 2.501% | 2.116% | 10.92% | 1.532% | 2.906% | 2.137% | 3.905% | 1.577% | 2.587% | 1.014% | 2.079% | 1.764% | 4.656% | 2.86% | 2.969% | 2.78% | 0.9114% | 0.9252% | 2.411% | 1.866% | 2.988% | 1.726% | 1.241% | 4.285% | 1.91% | 5.002% | 4.217% | 1.87% | 1.178% | 1.369% | 2.188% | 1.864% | 1.661% | 0.4768% | 0.788% | 0.5475% | 3.237% | 1.071% | 1.2% | 1.17% | 1.935%
148 | CCV 10 ppb MAJORS |  | 6/9/2023 4:27:10 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.502 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.11% | 3.868% | 0.9511% | 8.901% | 1.298% | 2.816% | 0.7094% | 1.277% | 0.4146% | 1.635% | 3.369% | 4.432% | 1.638% | 2.733% | 0.759% | 1.798% | 0.7789% | 4.127% | 3.006% | 2.747% | 2.633% | 2.504% | 2.171% | 3.991% | 2.795% | 1.634% | 2.507% | 5.793% | 3.108% | 3.612% | 1.237% | 1.966% | 1.544% | 1.217% | 2.329% | 1.18% | 0.3998% | 3.078% | 3.368% | 3.529% | 2.302% | 2.04%
149 | CCV 10 ppb MAJORS |  | 6/9/2023 4:33:27 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.503 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.392% | 3.228% | 0.7064% | 9.903% | 2.306% | 6.935% | 2.877% | 1.182% | 2.727% | 2.176% | 11.78% | 10.55% | 0.9578% | 2.87% | 3.331% | 3.46% | 1.377% | 11.85% | 11.69% | 0.7601% | 9.008% | 2.215% | 1.059% | 1.849% | 2.615% | 0.6121% | 1.758% | 1.211% | 3.391% | 9.984% | 2.509% | 1.114% | 0.5826% | 2.441% | 7.878% | 0.8698% | 0.4604% | 2.582% | 11.17% | 11.58% | 1.179% | 2.21%
150 | MT 33 D |  | 6/9/2023 4:39:15 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\MT 33 D.504 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 52.13% |  | 4.847% | 47.86% | 2.337% | 28.62% | 285.5% | 995.8% | 240.8% | 29.38% | 21.07% | 22.49% | 38.84% | 44.99% | 2403% | 185% | 28.25% | 3.438% | 14% | 46.68% | 0% | 29.56% | 22.97% | 10.97% | 0% | 327.1% | 466.8% | 79.39% | 86.6% | 2.504% | 9.286% | 1.579% | 0.7729% | 2.11% | 0.5099% | 0.5528% | 1.372% | 2.005% | 40.18% | 38.64% | 4.591% | 6.053%
151 | MT 33 D LSP |  | 6/9/2023 4:45:03 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\MT 33 D LSP.505 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.012% | 2.017% | 4.688% | 4.711% | 2.294% | 1.813% | 0.8547% | 3.775% | 3.519% | 2.647% | 1.142% | 1.157% | 1.047% | 2.799% | 4.233% | 4.57% | 3.197% | 1.67% | 0.9631% | 5.126% | 3.14% | 4.088% | 0.2414% | 6.048% | 2.287% | 1.934% | 0.753% | 4.196% | 2.109% | 0.8136% | 1.142% | 3.993% | 0.7362% | 2.683% | 1.168% | 0.9815% | 1.674% | 3.671% | 0.284% | 1.059% | 4.042% | 0.6455%
152 | cal blank |  | 6/9/2023 4:50:51 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\cal blank.506 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 108.7% |  | 8.668% | 0% | 10.91% | 101.6% | 206% | 0% | 295.6% | 5777% | 186.3% | 32.59% | 111.7% | 63.47% | 83.83% | 107.4% | 58.13% | 27.47% | 44.45% | 178.4% | 862.4% | 3.325% | 6.726% | 38.02% | 105.1% | 83.25% | 161.6% | 90.13% | 86.61% | 6.63% | 7.84% | 2.059% | 0.871% | 1.863% | 1.028% | 1.469% | 0.6951% | 2.1% | 100.2% | 143.8% | 1736% | 207.3%
153 | STD 0.01 |  | 6/9/2023 4:56:40 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 0.01.507 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 16.43% | 54.52% | 1.143% | 27.91% | 2.353% | 33.23% | 90.51% | 42.02% | 79.81% | 24.14% | 10.31% | 8.21% | 6.88% | 23.45% | 19.44% | 12.5% | 30.42% | 3.082% | 1.158% | 52.11% | 0% | 7.2% | 7.263% | 22.89% | 53.51% | 38.95% | 66.92% | 213.3% | 99.46% | 4.872% | 4.365% | 1.81% | 1.032% | 0.8952% | 1.351% | 1.093% | 1.372% | 3.022% | 7.28% | 12.56% | 9.581% | 1.183%
154 | std 0.05 |  | 6/9/2023 5:02:29 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\std 0.05.508 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.792% | 13.11% | 2.535% | 28.28% | 2.034% | 78.11% | 37.72% | 23.67% | 33.94% | 13.59% | 2.85% | 5.021% | 4.173% | 36.85% | 11.65% | 16.61% | 18.24% | 6.284% | 2.158% | 16.84% | 44.14% | 0.6188% | 5.874% | 40.53% | 80.57% | 17.52% | 9.465% | 90.23% | 55.39% | 5.604% | 5.395% | 1.582% | 2.929% | 1.151% | 1.193% | 15.76% | 1.92% | 4.475% | 4.929% | 4.466% | 2.75% | 15.97%
155 | std 0.1 |  | 6/9/2023 5:08:17 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\std 0.1.509 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 5.938% | 24.96% | 1.838% | 54.53% | 1.72% | 41.87% | 9.535% | 6.232% | 8.927% | 11.33% | 3.256% | 4.969% | 2.764% | 5.885% | 2.284% | 36.13% | 4.677% | 2.675% | 5.136% | 36.72% | 18.12% | 7.676% | 5.374% | 10.92% | 37.61% | 6.643% | 12.28% | 109.8% | 21.69% | 4.551% | 1.715% | 0.3687% | 0.8705% | 1.407% | 0.9109% | 2.12% | 1.988% | 2.297% | 3.611% | 1.199% | 3.463% | 1.967%
156 | std 0.5 |  | 6/9/2023 5:14:05 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\std 0.5.510 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 5.56% | 7.63% | 2.236% | 23.31% | 1.174% | 18.94% | 6.591% | 16.62% | 4.998% | 5.129% | 2.814% | 0.3843% | 1.163% | 3.535% | 4.157% | 5.561% | 4.972% | 1.652% | 4.203% | 5.209% | 19.27% | 2.856% | 2.217% | 16.71% | 10.03% | 13.03% | 12.65% | 9.602% | 9.26% | 1.091% | 2.738% | 0.203% | 1.675% | 3.227% | 1.293% | 0.5741% | 2.32% | 1.435% | 3.039% | 2.129% | 2.721% | 0.528%
157 | std 1 |  | 6/9/2023 5:19:54 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\std 1.511 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.885% | 8.573% | 1.295% | 4.121% | 2.704% | 19.2% | 6.59% | 5.933% | 5.492% | 2.748% | 3.635% | 3.351% | 0.3344% | 1.96% | 4.667% | 4.931% | 3.807% | 1.058% | 1.528% | 1.665% | 20.17% | 3.665% | 3.799% | 4.161% | 17.75% | 7.91% | 7.124% | 5.186% | 8.941% | 2.564% | 1.892% | 0.515% | 1.524% | 2.639% | 2.548% | 1.507% | 1.548% | 1.744% | 2.016% | 1.645% | 0.8354% | 2.544%
158 | std 5 |  | 6/9/2023 5:25:42 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\std 5.512 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.318% | 5.443% | 3.619% | 13.24% | 1.676% | 4.854% | 2.207% | 5.146% | 1.662% | 2.096% | 6.05% | 6.978% | 1.73% | 1.383% | 3.054% | 4.281% | 2.02% | 6.499% | 6.543% | 3.358% | 6.292% | 0.9736% | 1.206% | 3.804% | 6.603% | 0.4275% | 2.9% | 4.559% | 5.024% | 6.217% | 0.7318% | 1.261% | 0.3108% | 1.144% | 5.056% | 0.4924% | 2.468% | 1.677% | 5.796% | 6.584% | 1.134% | 1.183%
159 | std 10 |  | 6/9/2023 5:31:30 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\std 10.513 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.052% | 4.614% | 1.984% | 14.32% | 1.273% | 3.627% | 2.983% | 4.585% | 0.6465% | 3.424% | 1.185% | 2.247% | 1.108% | 1.048% | 0.8256% | 1.141% | 1.299% | 1.578% | 0.8354% | 1.84% | 1.907% | 1.844% | 0.6585% | 4.085% | 0.1667% | 2.459% | 0.3875% | 1.603% | 3.888% | 0.9712% | 0.2732% | 1.543% | 0.7209% | 1.525% | 1.776% | 1.106% | 0.3193% | 2.394% | 1.382% | 1.151% | 1.538% | 1.682%
160 | LBLANK |  | 6/9/2023 5:37:19 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.514 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 146.7% | 173.2% | 4.65% | 2.437% | 25.78% | 33.24% | 149.8% | 0% | 887.7% | 59% | 45.65% | 41.58% | 55.51% | 39.19% | 508.4% | 209.4% | 181.9% | 50.87% | 187.2% | 74.12% | 289.9% | 61629% | 439.6% | 28.63% | 0% | 346% | 94717% | 186.5% |  | 5.632% | 27.34% | 1.101% | 1.003% | 1.789% | 1.206% | 4.265% | 1.353% | 1.771% | 833% | 541.1% | 179.6% | 57.08%
161 | CCV 1ppb TRACE |  | 6/9/2023 5:43:07 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.515 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10.12% | 9.084% | 0.6757% | 36.85% | 8.174% | 32.5% | 9.302% | 14.18% | 7.857% | 5.326% | 1.704% | 0.5185% | 0.9292% | 15.1% | 6.149% | 6.556% | 11.58% | 2.673% | 0.03827% | 2.337% | 13.99% | 1.283% | 1.432% | 5.834% | 18.26% | 2.682% | 5.523% | 12.92% | 8.36% | 1.59% | 3.093% | 1.367% | 0.4955% | 6.739% | 0.1628% | 1.393% | 1.109% | 0.5009% | 2.413% | 0.9285% | 1.604% | 1.666%
162 | CCV 1ppb TRACE |  | 6/9/2023 5:49:25 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.516 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4.754% | 1.456% | 1.701% | 4.489% | 4.138% | 3.853% | 4.034% | 5.848% | 9.357% | 3.856% | 2.96% | 3.373% | 6.404% | 2.961% | 4.632% | 2.075% | 2.632% | 1.94% | 3.19% | 3.729% | 12.36% | 4.732% | 6.069% | 12.12% | 15.27% | 3.652% | 7.175% | 10.24% | 18.3% | 0.7115% | 5.686% | 1.563% | 5.064% | 0.8824% | 1.512% | 1.7% | 1.828% | 3.902% | 2.439% | 0.4241% | 2.508% | 1.262%
163 | CCV 10 ppb MAJORS |  | 6/9/2023 5:55:13 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.517 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.203% | 3.019% | 1.351% | 13.47% | 0.8056% | 1.388% | 2.403% | 1.37% | 1.625% | 0.8472% | 1.039% | 1.393% | 0.7801% | 2.502% | 1.411% | 1.708% | 1.421% | 0.5703% | 0.5729% | 1.434% | 1.329% | 4.029% | 0.4614% | 1.958% | 3.18% | 1.652% | 0.7759% | 2.06% | 6.387% | 0.9612% | 0.5107% | 0.7089% | 0.2961% | 1.216% | 1.158% | 23.71% | 1.911% | 4.489% | 1.654% | 1.262% | 1.312% | 23.35%
164 | CCV 10 ppb MAJORS |  | 6/9/2023 6:01:30 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.518 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4.955% | 6.134% | 0.524% | 13.67% | 1.171% | 7.529% | 6.15% | 4.889% | 6.249% | 4.351% | 0.1003% | 0.7045% | 0.7305% | 5.235% | 4.062% | 6.237% | 5.864% | 1.385% | 0.9966% | 1.672% | 2.445% | 2.224% | 0.4336% | 3.267% | 2.526% | 2.61% | 2.189% | 2.921% | 0.8906% | 0.5571% | 1.047% | 1.532% | 0.8206% | 4.924% | 0.5745% | 0.7944% | 1.476% | 1.656% | 0.5124% | 0.1178% | 1.098% | 1.932%
165 | std 25 |  | 6/9/2023 6:07:19 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\std 25.519 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.174% | 3.637% | 1.53% | 7.201% | 0.9272% | 8.219% | 2.578% | 6.161% | 3.298% | 3.003% | 0.9202% | 0.3979% | 1.07% | 5.395% | 3.13% | 2.405% | 4.92% | 1.605% | 0.9603% | 0.7192% | 5.706% | 0.7227% | 0.5919% | 2.898% | 2.05% | 2.215% | 0.2515% | 0.4846% | 1.589% | 1.03% | 0.9129% | 1.375% | 1.127% | 3.284% | 0.9746% | 0.9442% | 0.6637% | 0.3031% | 1.71% | 0.5836% | 0.7221% | 0.8773%
166 | std 50 |  | 6/9/2023 6:13:06 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\std 50.520 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.7462% | 0.3291% | 2.657% | 2% | 1.219% | 2.409% | 1.939% | 1.503% | 1.572% | 2.351% | 0.9776% | 1.216% | 3.199% | 0.9709% | 2.075% | 1.414% | 2.094% | 0.2198% | 1.52% | 1.457% | 3.145% | 3.173% | 3.342% | 0.7243% | 4.44% | 2.618% | 2.156% | 1.402% | 1.497% | 0.6337% | 2.721% | 0.5775% | 2.68% | 1.784% | 1.458% | 0.6424% | 1.353% | 2.914% | 0.4949% | 1.412% | 2.071% | 1.59%
167 | std 50 |  | 6/9/2023 6:19:23 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\std 50.521 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.949% | 3.752% | 3.153% | 1.944% | 0.6589% | 5.026% | 3.279% | 5.203% | 5.087% | 5.581% | 9.292% | 8.592% | 1.528% | 5.282% | 4.579% | 4.257% | 3.682% | 8.47% | 9.237% | 2.333% | 2.533% | 3.734% | 0.4274% | 1.147% | 1.902% | 2.324% | 1.006% | 2.921% | 1.835% | 8.666% | 1.038% | 1.117% | 1.113% | 3.678% | 7.884% | 1.085% | 1.03% | 2.887% | 8.366% | 8.807% | 1.437% | 0.7407%
168 | std 100 |  | 6/9/2023 6:25:11 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\std 100.522 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.768% | 2.825% | 0.8935% | 2.472% | 0.9555% | 3.458% | 1.161% | 2.353% | 3.27% | 3.302% | 0.8056% | 1.217% | 0.7577% | 2.587% | 2.587% | 2.338% | 1.253% | 1.729% | 1.249% | 0.9518% | 4.169% | 1.417% | 0.6251% | 1.051% | 2.572% | 1.7% | 1.15% | 1.106% | 1.769% | 0.2885% | 0.5067% | 1.628% | 0.7071% | 2.297% | 1.782% | 0.6804% | 2.241% | 1.334% | 1.259% | 0.8655% | 1.199% | 0.291%
169 | std 100 |  | 6/9/2023 6:31:28 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\std 100.523 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.599% | 2.293% | 0.7583% | 2.196% | 0.3955% | 3.063% | 1.315% | 3.064% | 3.353% | 3.193% | 1.015% | 0.304% | 0.1405% | 1.992% | 2.443% | 2.066% | 2.468% | 0.3275% | 0.4098% | 1.041% | 3.069% | 1.335% | 0.4331% | 1.422% | 0.899% | 1.1% | 0.728% | 0.8381% | 1.541% | 1.058% | 1.215% | 1.481% | 1.029% | 1.933% | 0.4558% | 0.8705% | 1.048% | 0.8905% | 0.4738% | 0.5588% | 1.111% | 9.818%
170 | LBLANK |  | 6/9/2023 6:37:16 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.524 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 65.86% | 86.97% | 10.26% | 42.22% | 11.7% | 43.85% | 523.1% | 0% | 58.46% | 26.58% | 39.57% | 63.33% | 24.74% | 53.24% | 53.67% | 31.57% | 82.08% | 23.23% | 489.8% | 549.8% | 65.53% | 13.83% | 10.28% | 27.66% | 138.2% | 51.5% | 32.32% | 712.4% |  | 1.226% | 5.015% | 2.775% | 1.653% | 5.343% | 0.9078% | 1.72% | 0.4196% | 2.227% | 85.88% | 24.87% | 102% | 105.5%
171 | PE ICV |  | 6/9/2023 6:43:04 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\PE ICV.525 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.502% | 4.956% | 2.055% | 4.662% | 0.7821% | 3.413% | 7.106% | 6.104% | 5.739% | 5.29% | 0.6197% | 0.3096% | 2.602% | 6.52% | 5.396% | 5.125% | 6.092% | 0.8306% | 1.924% | 1.822% | 5.665% | 1.205% | 0.6762% | 2.991% | 1.015% | 2.585% | 3.962% | 3.46% | 4.794% | 1.056% | 0.8876% | 1.004% | 1.806% | 5.588% | 1.091% | 0.3502% | 1.686% | 1.103% | 1.421% | 0.6274% | 1.566% | 2.404%
172 | 5ppb CCV |  | 6/9/2023 6:48:52 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5ppb CCV.526 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.3489% | 3.44% | 5.305% | 9.212% | 0.6911% | 4.222% | 1.3% | 7.163% | 1.577% | 0.241% | 0.5577% | 4.677% | 1.359% | 0.8222% | 0.6919% | 0.4689% | 2.303% | 1.04% | 1.666% | 5.803% | 6.144% | 1.556% | 0.52% | 3.893% | 3.441% | 2.455% | 2.571% | 8.023% | 0.5439% | 0.2865% | 0.1953% | 4.711% | 2.128% | 0.7945% | 0.3058% | 1.024% | 1.975% | 3.108% | 0.8444% | 1.062% | 5.077% | 1.216%
173 | 25ppb CCV |  | 6/9/2023 6:54:40 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\25ppb CCV.527 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4.588% | 4.052% | 1.953% | 8.011% | 0.9191% | 3.443% | 2.432% | 4.882% | 2.416% | 2.809% | 1.636% | 2.574% | 0.5221% | 4.838% | 2.79% | 2.735% | 2.062% | 1.218% | 1.054% | 0.7677% | 4.782% | 0.3508% | 1.006% | 1.787% | 1.971% | 1.972% | 0.5984% | 1.36% | 0.5854% | 0.7182% | 1.029% | 1.474% | 1.228% | 1.598% | 1.323% | 0.5134% | 1.473% | 0.6235% | 0.4157% | 0.9199% | 1.413% | 0.9299%
174 | 8_30_21_W_1 |  | 6/9/2023 7:00:27 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_W_1.528 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 6.975% | 5.044% | 2.448% | 7.08% | 16.28% | 2.085% | 15.87% | 2.395% | 21.36% | 3.028% | 11.71% | 2.662% | 0.2721% | 3.608% | 38.31% | 6.265% | 4.235% | 1.98% | 3.571% | 4.89% | 17.35% | 1.833% | 3.029% | 22.78% | 163.7% | 80.61% | 57.43% | 268.4% | 114.1% | 2.604% | 1.169% | 1.916% | 1.18% | 2.608% | 2.37% | 1.791% | 1.311% | 1.401% | 20.89% | 12.61% | 4.632% | 4.508%
175 | LBLANK |  | 6/9/2023 7:06:16 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.529 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 290.5% | 173.2% | 1.726% | 52.34% | 5.556% | 31.29% | 311.8% | 0% | 143.2% | 151.6% | 52.4% | 54.92% | 77.28% | 42.18% | 57.01% | 73.51% | 227.9% | 27.52% | 99.56% | 162.9% | 900.2% | 48.62% | 44.16% | 19.93% | 43.36% | 438.9% | 47.68% | 25.53% | 173.2% | 7.305% | 40.32% | 4.98% | 1.15% | 2.6% | 0.7163% | 1.393% | 1.442% | 2.139% | 139.2% | 26.55% | 3113% | 16.06%
176 | CCV 1ppb TRACE |  | 6/9/2023 7:12:04 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.530 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.069% | 8.437% | 1.224% | 22.34% | 7.493% | 12.31% | 9.531% | 3.922% | 5.278% | 5.184% | 2.634% | 5.268% | 12.42% | 9.467% | 2.276% | 1.034% | 2.126% | 3.16% | 0.5945% | 6.518% | 18.14% | 7.092% | 10.07% | 4.797% | 5.913% | 15.74% | 14.71% | 4.347% | 13% | 2.263% | 10.58% | 1.685% | 8.847% | 0.06391% | 0.4677% | 1.674% | 3.025% | 5.982% | 2.145% | 1.544% | 1.156% | 0.5521%
177 | CCV 1ppb TRACE |  | 6/9/2023 7:18:22 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.531 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 60.4% |  | 8.138% | 0% | 1.296% | 2.954% | 77.82% | 0% | 0% | 413.9% | 83.57% | 19.21% | 96.88% | 4.288% | 181% | 71.76% | 91.72% | 114.5% | 4.641% | 2.127% | 1280% | 5.274% | 6.731% | 2.35% | 41.73% | 345.8% | 89.12% | 122.5% |  | 3.386% | 60.21% | 2.553% | 1.899% | 5.267% | 2.289% | 4.948% | 7.384% | 4.465% | 76.88% | 139.1% | 39.48% | 4.24%
178 | LFB |  | 6/9/2023 7:24:08 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.532 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.991% | 4.654% | 1.578% | 1.575% | 1.031% | 5.943% | 1.741% | 3.062% | 3.287% | 1.969% | 0.202% | 0.968% | 0.9896% | 3.008% | 1.946% | 2.801% | 4.39% | 1.547% | 1.567% | 2.926% | 1.102% | 3.428% | 2.385% | 2.258% | 3.13% | 0.4101% | 2.009% | 3.484% | 2.506% | 2.31% | 0.9774% | 3.142% | 1.134% | 2.348% | 1.089% | 0.4019% | 1.092% | 4.363% | 1.284% | 1.563% | 2.896% | 1.129%
179 | LFB |  | 6/9/2023 7:30:24 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.533 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4.061% | 1.711% | 0.3319% | 5.527% | 0.5318% | 8.531% | 2.851% | 2.735% | 4.502% | 5.463% | 0.5772% | 0.1881% | 2.561% | 4.76% | 2.152% | 3.469% | 3.614% | 0.583% | 0.462% | 2.157% | 5.118% | 1.873% | 2.513% | 4.595% | 6.22% | 1.491% | 2.875% | 2.791% | 2.017% | 1.236% | 3.542% | 0.9668% | 1.803% | 2.092% | 0.8352% | 0.1633% | 0.9283% | 0.8912% | 2.159% | 0.2405% | 0.1737% | 0.8278%
180 | CCV 10 ppb MAJORS |  | 6/9/2023 7:36:12 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.534 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4.472% | 4.113% | 0.79% | 5.14% | 0.1794% | 6.856% | 4.669% | 6.928% | 6.074% | 3.249% | 1.391% | 1.98% | 1.189% | 5.418% | 4.702% | 5.062% | 4.554% | 1.037% | 1.744% | 2.274% | 8.045% | 2.085% | 0.4987% | 0.9213% | 4.924% | 1.252% | 1.098% | 4.077% | 6.471% | 1.092% | 0.2536% | 1.766% | 0.9397% | 3.918% | 1.11% | 1.656% | 0.3% | 2.833% | 0.9885% | 1.076% | 1.259% | 1.852%
181 | CCV 10 ppb MAJORS |  | 6/9/2023 7:42:30 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.535 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 5.285% | 4.539% | 0.2934% | 3.043% | 1.237% | 5.533% | 6.342% | 4.157% | 5.461% | 5.05% | 0.9091% | 0.6791% | 1.977% | 7.434% | 8.372% | 3.609% | 6.838% | 1.136% | 0.9393% | 1.503% | 9.194% | 1.227% | 0.9483% | 2.489% | 5.747% | 1.848% | 1.837% | 3.55% | 2.105% | 0.9951% | 2.406% | 0.6312% | 1.746% | 4.298% | 1.738% | 4.305% | 1.668% | 2.034% | 1.246% | 1.013% | 0.6263% | 4.329%
182 | 8_30_21_W_6 |  | 6/9/2023 7:48:18 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_W_6.536 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 15.28% | 14.89% | 0.2502% | 12.14% | 2.327% | 18.9% | 13.75% | 12.48% | 14.22% | 25.24% | 7.095% | 2.502% | 1.796% | 11.06% | 19.63% | 12.22% | 8.108% | 3.246% | 4.475% | 8.979% | 20.78% | 1.562% | 0.7093% | 10.31% | 69.97% | 10.49% | 32.24% | 281.8% | 13.96% | 2.041% | 0.631% | 1.034% | 0.6132% | 9.924% | 3.099% | 2.839% | 1.362% | 2.763% | 63.77% | 16.48% | 2.697% | 4.4%
183 | 8_30_21_W_6_LD |  | 6/9/2023 7:54:05 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_W_6_LD.537 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.36% | 5.87% | 2.259% | 9.088% | 1.486% | 10.87% | 15.93% | 2.709% | 17.29% | 11.18% | 7.503% | 3.517% | 6.703% | 0.6885% | 12.04% | 6.888% | 3.169% | 3.401% | 0.5665% | 6.425% | 16.7% | 6.461% | 4.44% | 26.85% | 1123% | 26.17% | 17.12% | 41.21% | 45.83% | 3.503% | 7.52% | 1.231% | 5.461% | 1.367% | 2.412% | 2.559% | 1.457% | 7.367% | 23.69% | 12.46% | 5.504% | 0.9557%
184 | 8_30_21_W_6_LD |  | 6/9/2023 8:00:20 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_W_6_LD.538 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.032% | 12.1% | 0.4975% | 3.032% | 0.354% | 3.749% | 20.03% | 7.572% | 19.39% | 43.29% | 12.6% | 1.371% | 1.163% | 5.275% | 12.39% | 16.37% | 6.682% | 1.259% | 0.6886% | 0.7204% | 6.057% | 3.143% | 2.666% | 30.09% | 224.5% | 47.42% | 35.78% | 44.78% | 66.78% | 1.674% | 1.416% | 1.037% | 0.2317% | 5.866% | 0.7221% | 11.92% | 1.205% | 0.8942% | 15.27% | 19.76% | 4.789% | 13.57%
185 | 8_30_21_W_6_DUP |  | 6/9/2023 8:06:08 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_W_6_DUP.539 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.432% | 3.748% | 2.137% | 6.13% | 1.329% | 6.165% | 3.933% | 6.347% | 19.54% | 35.88% | 9.309% | 1.742% | 1.895% | 3.696% | 26.57% | 11.26% | 0.6826% | 2.003% | 1.111% | 3.452% | 9.84% | 3.47% | 2.09% | 12.43% | 302.5% | 61.23% | 41.32% | 59.57% | 115.5% | 1.376% | 2.633% | 3.656% | 1.327% | 0.5923% | 1.53% | 1.609% | 1.009% | 4.07% | 23.46% | 10.55% | 0.97% | 4.998%
186 | 8_30_21_W_6 |  | 6/9/2023 8:11:56 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_W_6.540 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 6.407% | 5.123% | 2.33% | 5.632% | 12.32% | 3.64% | 18.35% | 5% | 5.329% | 32.1% | 11.83% | 1.26% | 1.035% | 4.421% | 18.9% | 46.31% | 6.088% | 2.099% | 1.379% | 1.698% | 6.258% | 2.203% | 2.661% | 11.23% | 48.28% | 103.8% | 51.57% | 119.9% | 86.62% | 0.4071% | 1.676% | 1.139% | 1.143% | 2.507% | 0.671% | 29.4% | 8.17% | 3.673% | 12.97% | 26.23% | 4.71% | 30.27%
187 | 8_30_21_W_6_DUP_LSP |  | 6/9/2023 8:17:43 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_W_6_DUP_LSP.541 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.999% | 3.334% | 26.5% | 8.255% | 0.5911% | 6.258% | 5.855% | 4.972% | 4.713% | 6.374% | 1.388% | 0.7665% | 0.7098% | 4.185% | 5.192% | 4.995% | 4.628% | 0.7681% | 1.001% | 26.43% | 8.011% | 0.9033% | 0.3974% | 30.87% | 1.619% | 1.383% | 2.027% | 24.73% | 2.694% | 0.4596% | 1.385% | 23.01% | 0.3699% | 3.282% | 0.9642% | 1.085% | 2.727% | 0.9985% | 0.8698% | 0.6222% | 26.52% | 1.845%
188 | 8_30_21_W_6_DUP_LSP |  | 6/9/2023 8:24:01 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_W_6_DUP_LSP.542 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.49% | 3.516% | 2.976% | 6.096% | 0.8709% | 6.991% | 2.352% | 3.659% | 4.048% | 3.548% | 1.192% | 1.495% | 0.5414% | 3.507% | 4.575% | 3.43% | 4.799% | 0.8833% | 0.6906% | 0.8587% | 4.563% | 2.501% | 1.286% | 0.5024% | 0.7656% | 2.404% | 1.345% | 1.829% | 2.647% | 1.319% | 0.4471% | 0.3017% | 0.6686% | 3.76% | 0.3664% | 0.1975% | 1.051% | 1.714% | 1.804% | 0.9136% | 1.416% | 1.536%
189 | 8_30_21_W_8 |  | 6/9/2023 8:29:49 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_W_8.543 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 13.76% | 6.516% | 0.82% | 11.42% | 6.186% | 14.02% | 12.65% | 12.25% | 2.575% | 9.068% | 11.27% | 2.769% | 0.7798% | 10.55% | 35.08% | 14.11% | 10.44% | 2.473% | 3.108% | 4.238% | 26.97% | 4.323% | 0.4838% | 7.725% | 132.2% | 21.51% | 24.5% | 25.42% | 50.03% | 2.53% | 1.596% | 0.3451% | 0.3964% | 8.808% | 3.149% | 0.07555% | 5.605% | 3.734% | 41.16% | 20.76% | 3.059% | 3.209%
190 | 8_30_21_W_8 |  | 6/9/2023 8:36:07 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_W_8.544 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.617% | 10.55% | 2.203% | 10.03% | 1.443% | 9.818% | 14.63% | 3.362% | 2.358% | 14.61% | 20% | 1.25% | 0.5972% | 1.378% | 45.53% | 27.17% | 4.974% | 0.7509% | 1.55% | 0.3888% | 20.31% | 4.481% | 1.449% | 20.98% | 296.3% | 13.7% | 31.54% | 46.17% | 86.61% | 3.178% | 1.538% | 0.8612% | 0.6009% | 1.655% | 1.862% | 1.852% | 1.763% | 2.611% | 47.41% | 13.53% | 3.066% | 3.254%
191 | 8_30_21_W_8_SP |  | 6/9/2023 8:41:55 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_W_8_SP.545 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 7.933% | 10.85% | 0.3813% | 5.637% | 0.6476% | 14.73% | 10.08% | 11.4% | 16.52% | 9.63% | 3.908% | 5.488% | 1.085% | 8.899% | 5.063% | 4.442% | 10.93% | 5.215% | 3.919% | 0.3933% | 8.804% | 2.417% | 1.965% | 6.034% | 15.82% | 7.024% | 5.997% | 13.76% | 11.93% | 4.421% | 0.3609% | 0.8077% | 0.8928% | 5.821% | 4.121% | 1.318% | 1.526% | 1.585% | 2.617% | 5.141% | 1.675% | 1.124%
192 | LBLANK |  | 6/9/2023 8:47:44 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.546 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 814.3% |  | 5.175% | 135.9% | 8.395% | 77.07% | 27.11% | 0% | 0% | 57.38% | 42.97% | 56.66% | 91.35% | 188.1% | 859.6% | 27.02% | 475.4% | 34.26% | 200% | 125.2% | 642% | 2462% | 601.7% | 9.692% | 50.58% | 641.5% | 79.1% | 271.4% |  | 12.99% | 18.44% | 0.3913% | 7.999% | 4.817% | 2.381% | 0.7951% | 2.896% | 7.852% | 197.9% | 298.6% | 115% | 19.86%
193 | CCV 1ppb TRACE |  | 6/9/2023 8:53:33 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.547 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 6.933% | 18.21% | 5.234% | 3.866% | 6.277% | 18.27% | 7.758% | 19.23% | 2.288% | 8.684% | 2.849% | 1.051% | 2.133% | 8.486% | 9.438% | 9.151% | 9.397% | 1.884% | 2.116% | 7.275% | 7.996% | 2.718% | 2.214% | 3.773% | 11.86% | 12.11% | 2.414% | 13.16% | 3.644% | 1.572% | 0.6477% | 4.213% | 0.6761% | 6.604% | 0.6657% | 2.555% | 2.122% | 1.8% | 2.576% | 2.797% | 5.073% | 1.592%
194 | CCV 1ppb TRACE |  | 6/9/2023 8:59:50 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.548 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 5.976% | 8.76% | 2.411% | 17.5% | 4.739% | 20.61% | 3.433% | 2.625% | 5.682% | 4.835% | 3.643% | 0.8706% | 0.6325% | 11.34% | 5.805% | 2.523% | 3.878% | 1.322% | 2.765% | 5.306% | 19.62% | 5.778% | 2.105% | 16.58% | 13.3% | 5.181% | 3.549% | 10.13% | 10.53% | 0.6171% | 1.994% | 0.51% | 0.9168% | 4.647% | 1.172% | 0.8966% | 0.5803% | 3.044% | 2.002% | 1.948% | 0.4502% | 1.06%
195 | CCV 10 ppb MAJORS |  | 6/9/2023 9:05:39 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.549 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 7.733% | 7.727% | 1.605% | 6.212% | 1.176% | 3.157% | 5.497% | 7.454% | 6.524% | 4.361% | 0.3647% | 1.711% | 2.53% | 7.098% | 5.524% | 9.39% | 8.196% | 0.8343% | 1.01% | 1.951% | 7.911% | 2.565% | 2.845% | 0.8697% | 2.829% | 4.349% | 3.244% | 3.952% | 2.009% | 0.3176% | 2.219% | 2.923% | 2.596% | 5.936% | 1.276% | 0.3865% | 0.8899% | 2.217% | 1.541% | 0.9357% | 2.822% | 1.808%
196 | CCV 10 ppb MAJORS |  | 6/9/2023 9:11:56 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.550 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.251% | 3.32% | 1.619% | 6.512% | 1.385% | 2.126% | 3.322% | 3.153% | 0.2663% | 0.5944% | 1.361% | 0.8384% | 1.329% | 0.6812% | 0.437% | 1.272% | 0.459% | 1.423% | 0.8229% | 1.028% | 6.151% | 5.145% | 0.5289% | 1.172% | 5.417% | 3.064% | 2.069% | 4.146% | 8.437% | 0.5502% | 0.3463% | 0.8004% | 1.302% | 0.9883% | 1.996% | 1.524% | 1.028% | 4.687% | 0.7146% | 1.249% | 1.135% | 0.4572%
197 | 8_30_21_W_9 |  | 6/9/2023 9:17:44 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_W_9.551 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4.452% | 3.879% | 4.819% | 8.394% | 0.7646% | 13.86% | 12.29% | 4.751% | 11.53% | 25.88% | 16.63% | 0.7945% | 2.04% | 1.317% | 35.16% | 40.7% | 6.465% | 0.9346% | 2.827% | 7.145% | 12.79% | 4.505% | 3.869% | 66.51% | 756.1% | 109.8% | 49.92% | 103.3% | 92.28% | 0.1862% | 2.214% | 4.426% | 1.337% | 1.676% | 0.335% | 1.3% | 0.8568% | 1.434% | 28.65% | 27.71% | 7.838% | 4.06%
198 | 8_30_21_W_10 |  | 6/9/2023 9:23:32 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_W_10.552 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.788% | 3.34% | 2.458% | 10.17% | 3.886% | 5.933% | 11.47% | 4.933% | 15.96% | 13.73% | 8.164% | 4.259% | 3.377% | 3.914% | 24.21% | 27.62% | 7.644% | 2.344% | 6.035% | 7.496% | 22.99% | 1.385% | 2.082% | 21.08% | 184.3% | 16.51% | 32.38% | 18.43% | 69.51% | 2.704% | 3.233% | 3.153% | 1.325% | 3.178% | 2.383% | 0.4817% | 3.266% | 1.412% | 49.56% | 2.317% | 1.822% | 1.382%
199 | 8_30_21_W_11 |  | 6/9/2023 9:29:20 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_W_11.553 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4.35% | 0.6584% | 2.835% | 4.573% | 1.647% | 7.847% | 3.015% | 4.69% | 7.312% | 43.41% | 10.98% | 9.493% | 2.514% | 3.376% | 24.82% | 35.94% | 0.884% | 7.355% | 10.71% | 6.156% | 15.7% | 1.468% | 1.544% | 26.8% | 146.2% | 34.63% | 29.64% | 57.74% | 34.8% | 8.21% | 2.625% | 1.405% | 0.9277% | 0.7839% | 6.11% | 3.35% | 1.79% | 0.9791% | 55.52% | 5.927% | 1.62% | 3.544%
200 | 8_30_21_F_1 |  | 6/9/2023 9:35:08 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_F_1.554 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.814% | 4.363% | 1.489% | 11.24% | 2.106% | 3.498% | 5.916% | 3.347% | 11.51% | 10.28% | 3.962% | 1.024% | 1.696% | 6.767% | 19.13% | 4.763% | 5.981% | 1.56% | 0.5023% | 0.3422% | 4.484% | 3.153% | 0.521% | 16.56% | 90.32% | 26.69% | 31.02% | 12.94% | 49.59% | 2.076% | 2.077% | 1.713% | 0.5226% | 2.909% | 0.521% | 0.8469% | 0.8928% | 3.183% | 16.82% | 9.357% | 4.764% | 1.425%
201 | 8_30_21_F_1 |  | 6/9/2023 9:41:25 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_F_1.555 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 7.516% | 8.558% | 2.738% | 7.296% | 2.864% | 9.855% | 11.29% | 10.81% | 15.47% | 10.73% | 5.105% | 2.825% | 0.3343% | 7.643% | 5.368% | 9.009% | 6.954% | 0.5289% | 0.9389% | 1.257% | 16.98% | 1.415% | 1.644% | 19.87% | 19.24% | 31.48% | 32.55% | 20.1% | 23.19% | 1.14% | 0.4185% | 0.9547% | 0.4684% | 7.26% | 0.4881% | 0.6859% | 2.913% | 1.754% | 3.256% | 5.265% | 3.315% | 1.153%
202 | 8_30_21_F_1_DUP |  | 6/9/2023 9:47:13 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_F_1_DUP.556 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 6.537% | 5.512% | 15.57% | 12.08% | 1.836% | 5.267% | 5.946% | 4.308% | 5.646% | 2.612% | 1.938% | 0.6102% | 0.8682% | 7.914% | 6.95% | 10.13% | 4.204% | 0.6846% | 0.8578% | 14.16% | 10.45% | 1.974% | 1.672% | 52.28% | 80.18% | 39.83% | 9.036% | 24.57% | 68.24% | 0.1151% | 1.044% | 13.09% | 0.6088% | 4.189% | 0.942% | 1.752% | 1.458% | 1.68% | 26.24% | 18.94% | 12.78% | 1.402%
203 | 8_30_21_F_1_DUP |  | 6/9/2023 9:53:30 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_F_1_DUP.557 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 5.548% | 4.181% | 6.588% | 2.581% | 1.474% | 2.897% | 6.923% | 5.906% | 4.256% | 13.27% | 9.286% | 0.7746% | 1.146% | 1.897% | 32.07% | 7.376% | 2.617% | 0.5388% | 0.3471% | 6.305% | 4.045% | 1.129% | 1.205% | 22.96% | 61.11% | 13.35% | 50.26% | 6.83% | 33.52% | 1.221% | 0.8517% | 4.542% | 1.189% | 2.739% | 0.4843% | 0.9174% | 1.666% | 1.62% | 28.92% | 9.612% | 8.102% | 4.939%
204 | 8_30_21_F_6 |  | 6/9/2023 9:59:18 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_F_6.558 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.934% | 6.988% | 1.055% | 3.123% | 1.536% | 3.697% | 4.966% | 4.132% | 5.628% | 12.24% | 5.38% | 1.066% | 1.89% | 3.185% | 18.77% | 8.834% | 4.513% | 1.605% | 1.661% | 0.6157% | 6.318% | 5.215% | 1.499% | 10.22% | 87.37% | 21.26% | 20.64% | 8.698% | 23.43% | 1.587% | 1.744% | 0.9022% | 1.272% | 3.06% | 1.354% | 5.655% | 1.483% | 5.335% | 10.92% | 10.36% | 1.251% | 6.919%
205 | 8_30_21_F_6 |  | 6/9/2023 10:05:38 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_F_6.559 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10.16% | 11.94% | 1.004% | 12.94% | 1.666% | 14.88% | 14.13% | 12.69% | 15.29% | 22.47% | 2.655% | 1.184% | 1.948% | 16.94% | 11.16% | 9.59% | 13.66% | 1.623% | 0.3663% | 2.066% | 6.789% | 5.497% | 1.288% | 10.66% | 97.41% | 19.38% | 8.88% | 19.43% | 0.8799% | 1.021% | 0.8879% | 0.2916% | 0.8758% | 12.23% | 0.794% | 1.8% | 0.2391% | 5.492% | 12.14% | 11.87% | 2.93% | 1.446%
206 | 8_30_21_F_8 |  | 6/9/2023 10:11:26 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_F_8.560 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.25% | 3.743% | 2.221% | 0.8069% | 1.252% | 2.011% | 5.953% | 2.206% | 10.54% | 27.32% | 4.18% | 1.104% | 8.007% | 5.862% | 26.28% | 19.7% | 4.324% | 0.9325% | 2.907% | 0.5441% | 10.08% | 7.616% | 8.553% | 33.59% | 77.05% | 34.42% | 15.49% | 5.109% | 59.25% | 1.081% | 8.531% | 1.055% | 6.394% | 3.857% | 0.8661% | 1.578% | 0.2328% | 5.7% | 13.26% | 3.111% | 3.841% | 1.652%
207 | 8_30_21_F_8 |  | 6/9/2023 10:17:43 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_F_8.561 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.689% | 4.817% | 1.612% | 17.16% | 2.101% | 3.376% | 2.21% | 1.806% | 2.676% | 5.587% | 1.923% | 1.407% | 5.149% | 2.436% | 6.082% | 13.84% | 0.2129% | 1.561% | 0.7723% | 3.097% | 6.825% | 4.004% | 5.114% | 7.64% | 17.56% | 20.71% | 25.53% | 13.54% | 4.644% | 0.5013% | 5.511% | 0.171% | 4.617% | 0.9205% | 1.178% | 1.598% | 1.177% | 3.342% | 21.84% | 7.946% | 2.1% | 2.356%
208 | LBLANK |  | 6/9/2023 10:23:31 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.562 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 134.2% |  | 14.17% | 40.3% | 2.651% | 49.49% | 2.666% | 0% | 593.2% | 58% | 70.53% | 163.8% | 14.07% | 114.3% | 0% | 21.41% | 247.3% | 21.62% | 154% | 121% | 0% | 24.8% | 19.33% | 15.71% | 0% | 3.489% | 45.61% | 425.3% | 86.61% | 9.54% | 49.59% | 2.063% | 1.995% | 0.6249% | 12.27% | 2.261% | 1.466% | 1.347% | 416.5% | 78.32% | 35.62% | 13.46%
209 | CCV 1ppb TRACE |  | 6/9/2023 10:29:20 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.563 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 8.007% | 7.716% | 1.945% | 24.43% | 3.053% | 5.76% | 5.128% | 4.164% | 2.095% | 4.862% | 0.8625% | 2.387% | 1.654% | 3.028% | 5.328% | 1.82% | 2.144% | 1.893% | 2.281% | 7.78% | 11.39% | 0.647% | 2.248% | 15.04% | 14.08% | 2.588% | 0.1877% | 9.939% | 14.61% | 1.087% | 1.679% | 0.9334% | 0.8896% | 1.613% | 1.208% | 0.4991% | 1.182% | 0.9958% | 1.905% | 0.238% | 2.126% | 0.3493%
210 | CCV 1ppb TRACE |  | 6/9/2023 10:35:37 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.564 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.709% | 4.366% | 2.15% | 12.75% | 3.67% | 15.32% | 7.461% | 14.91% | 3.008% | 3.399% | 1.47% | 3.717% | 2.741% | 11.23% | 6.748% | 6.213% | 4.885% | 0.6766% | 1.587% | 4.387% | 18.81% | 3.902% | 2.151% | 1.103% | 10.61% | 5.086% | 5.151% | 4.562% | 3.441% | 0.5693% | 2.721% | 0.7658% | 1.496% | 5.835% | 0.5145% | 0.7479% | 1.299% | 2.912% | 0.4479% | 1.526% | 1.028% | 1.183%
211 | LFB |  | 6/9/2023 10:41:25 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.565 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.3717% | 0.7064% | 0.3751% | 2.135% | 1.406% | 2.98% | 2.29% | 2.792% | 2.073% | 0.7586% | 1.42% | 2.91% | 1.005% | 0.2596% | 0.8899% | 2.966% | 1.507% | 1.62% | 1.858% | 0.6748% | 0.5152% | 1.234% | 1.094% | 1.588% | 0.8954% | 2.448% | 3.068% | 3.435% | 3.982% | 1.434% | 1.84% | 0.6444% | 0.7507% | 0.3039% | 0.3871% | 0.5654% | 1.014% | 1.665% | 1.486% | 1.009% | 0.4027% | 1.691%
212 | LFB |  | 6/9/2023 10:47:41 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.566 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.361% | 0.9532% | 0.4572% | 9.781% | 2.958% | 4.164% | 1.907% | 4.107% | 3.204% | 0.586% | 0.8714% | 0.7675% | 2.463% | 1.738% | 0.9473% | 0.957% | 1.793% | 1.939% | 0.674% | 1.997% | 6.534% | 4.829% | 1.458% | 1.275% | 1.138% | 2.432% | 2.577% | 0.8307% | 5.091% | 0.8667% | 1.181% | 1.263% | 1.28% | 1.553% | 1.24% | 0.9009% | 1.28% | 3.83% | 1.021% | 1.875% | 1.392% | 1.396%
213 | CCV 10 ppb MAJORS |  | 6/9/2023 10:53:35 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.567 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.112% | 1.413% | 1.736% | 10.47% | 3.297% | 1.893% | 2.379% | 2.874% | 3.5% | 1.75% | 0.9934% | 1.571% | 2.835% | 4.975% | 2.677% | 6.457% | 1.977% | 1.54% | 0.9571% | 1.99% | 3.831% | 1.792% | 2.167% | 3.422% | 1.664% | 2.875% | 1.486% | 6.631% | 3.832% | 0.6266% | 1.669% | 0.9541% | 1.625% | 2.554% | 1.471% | 1.049% | 0.7704% | 1.192% | 1.844% | 1.14% | 0.754% | 1.823%
214 | LBLANK |  | 6/9/2023 11:28:06 AM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.568 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 173.2% |  | 5.623% | 25.56% | 3.835% | 8.406% | 69.11% |  | 173.2% | 24.44% | 35.21% | 1.77% | 44.69% | 3.58% | 173.2% | 26.21% | 56.31% | 21.33% | 13.11% | 11.7% |  | 14.53% | 12.8% | 2.015% |  | 99.72% |  | 71.02% |  | 40.62% | 45.31% | 1.137% | 0.473% | 1.171% | 0.9224% | 0.9045% | 1.856% | 3.571% | 173.2% | 65.59% | 38.82% | 24.87%
215 | Cal blank |  | 6/9/2023 11:33:55 AM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\Cal blank.569 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth |  |  | 6.481% | 46.25% | 4.184% | 7.571% | 71.39% |  | 173.2% | 2.29% | 8.423% | 1.081% | 11.71% | 5.529% | 36.24% | 28.95% | 37.2% | 4.233% | 7.668% | 3.729% | 173.2% | 25.6% | 25.85% | 9.63% |  | 173.2% | 1.815% | 78.38% | 86.6% | 92.36% | 18.19% | 1.732% | 1.833% | 2.313% | 0.9524% | 0.5893% | 1.579% | 3.016% | 44.36% | 56.15% | 43.86% | 17.15%
216 | CAL BLANK |  | 6/9/2023 11:39:43 AM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CAL BLANK.570 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 173.2% | 173.2% | 8.782% | 26.98% | 5.427% | 12.22% | 39.47% |  |  | 16.14% | 22.53% | 3.065% | 7.161% | 1.735% | 99.52% | 36.4% | 16.87% | 19.78% | 33.22% | 5.147% |  | 34.89% | 34.35% | 7.336% |  | 86.61% | 43.76% | 41.04% |  | 27.15% | 86.61% | 0.6623% | 0.5934% | 0.8174% | 1.564% | 0.7852% | 1.523% | 0.8993% |  | 56.13% | 53.25% | 24.9%
217 | LOW 0.01 |  | 6/9/2023 11:45:32 AM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LOW 0.01.571 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 44.22% | 72.63% | 2.945% | 29.04% | 3.095% | 622.4% | 76.48% | 0.9429% | 87.22% | 6.344% | 26.21% | 4.042% | 7.425% | 62.89% | 21.85% | 111.2% | 25.62% | 4.115% | 14.1% | 1988% |  | 29.51% | 31.76% | 188.3% | 70.72% | 31.84% | 19.14% | 214.7% | 99.17% | 8.029% | 8.648% | 0.3673% | 1.84% | 0.9474% | 0.7389% | 1.658% | 0.8092% | 1.431% | 3.584% | 8.652% | 8.831% | 3.871%
218 | LOW 0.05 |  | 6/9/2023 11:51:20 AM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LOW 0.05.572 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10.99% | 21.44% | 6.743% | 27.07% | 3.276% | 190% | 20.18% | 32.89% | 8.59% | 17.66% | 8.979% | 5.156% | 7.23% | 13% | 4.82% | 39.88% | 3.014% | 2.612% | 95.53% | 56.91% | 50.24% | 14.79% | 16.57% | 53.31% | 79.06% | 38.52% | 27.31% | 1030% | 45.65% | 10.33% | 0.72% | 1.424% | 1.024% | 0.635% | 0.6804% | 0.3827% | 1.29% | 1.942% | 4.113% | 5.416% | 2.09% | 1.846%
219 | LOW 0.1 |  | 6/9/2023 11:57:09 AM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LOW 0.1.573 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 7.594% | 9.012% | 4.73% | 299.3% | 2.253% | 452.3% | 8.493% | 15.91% | 14.68% | 5.354% | 3.763% | 8.444% | 1.102% | 14.01% | 5.37% | 3.412% | 5.655% | 4.172% | 180% | 21.96% | 32.32% | 5.783% | 4.714% | 17.5% | 20.77% | 41.93% | 5.859% | 53.72% | 11.27% | 4.875% | 0.6822% | 1.825% | 1.283% | 1.076% | 0.2006% | 1.706% | 1.319% | 1.875% | 4.984% | 0.793% | 5.076% | 1.828%
220 | LOW 0.5 |  | 6/9/2023 12:02:57 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LOW 0.5.574 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.611% | 9.013% | 6.566% | 27.95% | 2.236% | 43.67% | 6.989% | 6.202% | 10.36% | 6.481% | 8.856% | 9.607% | 2.157% | 4.148% | 7.698% | 4.605% | 0.187% | 10.14% | 13.84% | 13.11% | 9.942% | 2.011% | 3.789% | 26.11% | 5.784% | 1.844% | 4.409% | 33.31% | 21.56% | 11.17% | 3.324% | 2.946% | 0.292% | 0.5819% | 7.365% | 14.01% | 0.7157% | 2.644% | 9.039% | 8.022% | 5.303% | 16.36%
221 | STD 1 |  | 6/9/2023 12:08:45 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 1.575 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.126% | 2.821% | 5.231% | 40.88% | 1.86% | 0.8036% | 6.086% | 13.43% | 5.889% | 5.576% | 2.275% | 1.265% | 1.034% | 7.778% | 2.82% | 8.318% | 4.835% | 1.409% | 2.009% | 2.033% | 13.5% | 3.385% | 2.096% | 8.621% | 19.87% | 2.663% | 8.004% | 6.669% | 8.422% | 0.9603% | 0.6267% | 0.6116% | 0.7099% | 2.769% | 0.7788% | 1.58% | 2.355% | 2.789% | 2.471% | 0.7997% | 1.013% | 1.955%
222 | STD 5 |  | 6/9/2023 12:14:33 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 5.576 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.496% | 1.471% | 1.806% | 7.315% | 0.8108% | 6.078% | 3.107% | 2.435% | 0.866% | 1.669% | 0.7553% | 3.333% | 1.986% | 3.772% | 1.084% | 1.334% | 0.8532% | 0.94% | 1.115% | 1.896% | 13.46% | 2.889% | 2.072% | 3.45% | 6.281% | 4.371% | 5.439% | 5.101% | 2.534% | 0.9237% | 1.91% | 0.6203% | 2.11% | 1.185% | 0.5942% | 1.504% | 1.625% | 2.701% | 1.595% | 0.8761% | 1.426% | 1.644%
223 | STD 10 |  | 6/9/2023 12:20:21 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 10.577 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.101% | 2.518% | 1.817% | 9.859% | 2.085% | 9.48% | 4.128% | 2.687% | 3.772% | 4.21% | 1.472% | 1.357% | 1.064% | 4.713% | 2.225% | 4.069% | 3.543% | 0.4021% | 2.048% | 1.364% | 2.012% | 1.037% | 2.17% | 3.649% | 0.8595% | 0.1069% | 2.793% | 1.299% | 4.377% | 0.9702% | 1.556% | 0.9618% | 0.9593% | 3.411% | 0.9061% | 0.4659% | 0.1279% | 1.906% | 1.43% | 0.8213% | 0.6473% | 0.8373%
224 | STD 25 |  | 6/9/2023 12:26:09 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 25.578 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4.02% | 3.075% | 7.911% | 10.77% | 0.5818% | 5.767% | 5.581% | 5.786% | 2.015% | 3.802% | 0.5974% | 0.2123% | 0.9862% | 4.517% | 4.959% | 3.855% | 4.568% | 0.5214% | 0.5819% | 6.214% | 5.134% | 1.966% | 1.327% | 8.496% | 6.168% | 0.7183% | 1.339% | 6.471% | 0.1881% | 0.6199% | 0.8597% | 4.113% | 1.222% | 3.954% | 0.7926% | 1.09% | 1.664% | 2.126% | 1.297% | 0.2604% | 5.145% | 2.293%
225 | STD 50 |  | 6/9/2023 12:31:57 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 50.579 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.925% | 2.767% | 0.9692% | 3.751% | 2.421% | 1.346% | 3.821% | 1.877% | 1.943% | 1.457% | 2.304% | 2.215% | 0.9085% | 2.538% | 1.949% | 2.43% | 2.054% | 2.115% | 2.204% | 0.4331% | 3.287% | 2.933% | 1.018% | 3.145% | 0.5953% | 1.237% | 2.025% | 2.315% | 1.558% | 1.891% | 0.672% | 0.4796% | 1.006% | 1.65% | 0.6768% | 6.87% | 1.442% | 2.755% | 2.178% | 2.294% | 0.4625% | 8.177%
226 | STD 100 |  | 6/9/2023 12:37:44 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 100.580 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9.918% | 9.294% | 15.53% | 7.39% | 0.6614% | 8.824% | 10.14% | 7.953% | 9.682% | 9.771% | 1.553% | 0.3677% | 0.6402% | 9.759% | 9.027% | 9.717% | 9.128% | 1.564% | 0.4483% | 15.53% | 8.464% | 1.679% | 0.7627% | 16.44% | 2.786% | 0.4332% | 1.536% | 14.59% | 0.843% | 1.819% | 0.5625% | 13.82% | 0.7724% | 7.11% | 1.121% | 0.3927% | 0.5599% | 1.78% | 1.332% | 1.163% | 15.49% | 0.2585%
228 | LBLANK |  | 6/9/2023 12:43:32 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.581 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 29.18% | 248501% | 40.71% | 57.63% | 2.145% | 72.39% | 249.1% |  | 107.6% | 118.6% | 44.21% | 247.4% | 31.51% | 52.16% | 11.72% | 101.2% | 22.56% | 40.77% | 2.341% | 126.4% | 86.61% | 35.43% | 37.01% | 42.72% |  | 74.5% | 59.2% | 899.7% | 99.58% | 8.034% | 5.597% | 15.71% | 0.7355% | 0.8566% | 0.6528% | 0.7962% | 0.8173% | 2.215% | 138% | 43.61% | 119.4% | 7697%
229 | LBLANK |  | 6/9/2023 1:13:17 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.583 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 45.37% | 1.47e-06% | 214.8% | 135.3% | 7.968% | 137.4% | 45.61% |  | 115.5% | 145.2% | 41.12% | 96.49% | 355.8% | 299844% | 93.35% | 147.2% | 75.14% | 39.29% | 139.2% | 1729% | 173.2% | 45.03% | 40.47% | 244.7% | 173.2% | 72.91% | 61.51% | 287.5% |  | 1.341% | 23.91% | 15.24% | 0.4477% | 3.995% | 0.9617% | 0.8764% | 0.6516% | 3.854% | 173.2% | 114.7% | 841.2% | 57.57%
230 | LBLANK |  | 6/9/2023 1:19:06 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.584 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 41.89% | 1.47e-06% | 1.71% | 20.72% | 15.54% | 36.2% | 91.09% |  | 173.2% | 3.222% | 148.4% | 53.86% | 37538% | 96.66% | 93.74% | 258.7% | 50.54% | 512.8% | 3.788% | 50.55% |  | 79.74% | 77.19% | 1258% |  | 73.55% | 147.1% | 102% |  | 11.36% | 17.68% | 1.166% | 0.3643% | 2.559% | 1.811% | 0.9652% | 2.079% | 0.8685% | 173.2% | 117.3% | 313.6% | 38.31%
231 | ICV PE |  | 6/9/2023 1:24:53 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\ICV PE.585 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.311% | 2.246% | 3.544% | 14.4% | 1.636% | 5.565% | 3.41% | 2.925% | 4.118% | 4.079% | 1.665% | 1.395% | 2.227% | 6.224% | 5.719% | 3.303% | 3.495% | 0.3585% | 0.7487% | 3.383% | 3.237% | 2.505% | 1.688% | 2.224% | 3.263% | 1.556% | 2.541% | 2.413% | 4.47% | 0.6036% | 0.4698% | 2.723% | 1.423% | 4.253% | 0.5821% | 0.7338% | 0.491% | 2.63% | 0.7523% | 0.9611% | 2.958% | 0.6398%
232 | ICV PE |  | 6/9/2023 1:31:10 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\ICV PE.586 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.151% | 4.142% | 1.482% | 4.694% | 0.6776% | 3.746% | 1.862% | 4.506% | 4.333% | 3.264% | 19.43% | 20.27% | 1.21% | 2.841% | 4.485% | 4.414% | 2.815% | 18.25% | 16.58% | 1.434% | 7.032% | 2.963% | 0.4041% | 2.649% | 2.546% | 0.4836% | 0.4685% | 0.5261% | 3.112% | 18.56% | 1.68% | 1.911% | 1.198% | 2.819% | 14.57% | 0.7403% | 0.8267% | 1.843% | 19.59% | 17.08% | 1.614% | 0.3481%
233 | CCV 1ppb TRACE |  | 6/9/2023 1:36:58 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.587 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10.03% | 18.28% | 1.074% | 9.007% | 4.579% | 14.24% | 10.72% | 8.431% | 11.88% | 9.939% | 1.72% | 2.01% | 1.992% | 25.15% | 11.44% | 13.99% | 14.75% | 1.667% | 3.815% | 4.872% | 13.01% | 1.11% | 1.248% | 18.37% | 10.46% | 2.909% | 2.425% | 5.598% | 13.17% | 1.847% | 1.165% | 0.7307% | 0.5671% | 9.904% | 1.524% | 0.368% | 0.9798% | 1.813% | 0.6765% | 0.3887% | 1.768% | 2.33%
234 | CCV 1ppb TRACE |  | 6/9/2023 1:43:15 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.588 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.307% | 5.557% | 6.918% | 16.54% | 2.718% | 7.568% | 5.368% | 10.27% | 1.088% | 2.616% | 2.226% | 1.72% | 2.191% | 1.658% | 1.72% | 6.822% | 5.652% | 2.409% | 3.155% | 4.181% | 4.568% | 1.433% | 1.161% | 19.09% | 5.19% | 5.244% | 5.73% | 17.72% | 4.88% | 1.122% | 1.154% | 5.78% | 0.6314% | 2.097% | 1.011% | 0.4695% | 1.109% | 1.02% | 0.9877% | 0.1704% | 8.222% | 0.3218%
235 | LFB |  | 6/9/2023 1:49:03 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.589 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.976% | 2.696% | 1.949% | 2.302% | 1.849% | 1.784% | 2.146% | 3.916% | 2.24% | 2.932% | 0.5658% | 1.923% | 1.221% | 3.009% | 0.8418% | 2.322% | 0.9352% | 0.3315% | 1.611% | 2.124% | 4.809% | 2.449% | 1.219% | 2.714% | 2.709% | 0.699% | 1.755% | 3.197% | 5.087% | 0.2898% | 1.027% | 1.275% | 1.445% | 1.39% | 0.4243% | 1.086% | 1.012% | 2.508% | 0.623% | 0.2284% | 0.9031% | 1.008%
236 | LLOQ 0.01 |  | 6/9/2023 1:54:51 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LLOQ 0.01.590 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.191% | 33.36% | 1.786% | 60.41% | 5.341% | 39.48% | 41.16% | 68.05% | 18.3% | 53.9% | 13.97% | 7.347% | 5.022% | 30.78% | 13.04% | 168.1% | 18.45% | 6.124% | 35.62% | 24.01% | 173.2% | 17.07% | 17.62% | 197.5% | 98.78% | 117.6% | 121.2% | 158.9% | 1.901% | 6.46% | 6.443% | 1.702% | 1.889% | 1.14% | 1.595% | 1.349% | 0.7058% | 4.215% | 8.618% | 13.19% | 1.428% | 4.012%
237 | CCV 10 ppb MAJORS |  | 6/9/2023 2:00:40 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.591 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.006% | 5.194% | 1.384% | 10.73% | 1.344% | 5.325% | 1.012% | 5.194% | 1.939% | 3.536% | 4.361% | 5.097% | 1.713% | 3.71% | 2.351% | 2.646% | 2.89% | 4.029% | 4.579% | 0.8549% | 5.819% | 1.497% | 1.392% | 2.496% | 8.092% | 3.491% | 1.345% | 3.228% | 6.122% | 3.165% | 0.9916% | 1.26% | 0.4545% | 3.227% | 3.636% | 1.393% | 1.728% | 2.451% | 5.346% | 4.441% | 2.353% | 1.364%
238 | CCV 10 ppb MAJORS |  | 6/9/2023 2:06:57 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.592 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.908% | 4.632% | 0.4286% | 8.496% | 2.183% | 4.035% | 3.698% | 1.557% | 1.112% | 0.4332% | 0.3089% | 0.5794% | 0.7885% | 3.081% | 2.254% | 0.4193% | 1.105% | 0.2631% | 1.21% | 0.3956% | 4.142% | 1.71% | 0.9411% | 1.417% | 7.258% | 0.3245% | 1.71% | 1.232% | 2.402% | 0.3062% | 0.9475% | 0.2296% | 0.8037% | 1.232% | 0.336% | 0.8345% | 0.7409% | 1.433% | 1.182% | 0.304% | 0.8782% | 0.7929%
239 | LLOQ 0.1 |  | 6/9/2023 2:12:45 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LLOQ 0.1.593 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 13.27% | 2.551% | 3.428% | 10.69% | 6.491% | 1236% | 12.46% | 41.04% | 6.054% | 6.927% | 2.674% | 1.125% | 3.639% | 19.04% | 6.605% | 21.6% | 4.252% | 4.629% | 57.51% | 15.61% | 62.43% | 2.185% | 2.913% | 38.42% | 47.64% | 10.65% | 23.39% | 223.1% | 29.77% | 1.286% | 3.269% | 2.084% | 0.8149% | 1.134% | 0.3456% | 1.352% | 0.6818% | 1.147% | 5.219% | 1.954% | 1.414% | 1.835%
240 | 8_30_21_F_9 |  | 6/9/2023 2:19:27 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_F_9.594 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.293% | 3.069% | 2.248% | 7.932% | 1.779% | 4.524% | 5.561% | 3.671% | 4.264% | 18.4% | 5.541% | 2.305% | 1.777% | 0.7703% | 32.25% | 5.572% | 2.073% | 1.848% | 2.56% | 2.435% | 4.715% | 3.586% | 1.681% | 31.03% | 33.45% | 22.77% | 16.84% | 7.432% | 37.72% | 2.225% | 1.166% | 1.214% | 0.5334% | 0.7421% | 0.7146% | 0.5728% | 1.536% | 2.363% | 10.01% | 3.963% | 4.316% | 3.503%
241 | 8_30_21_F_10 |  | 6/9/2023 2:25:14 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_F_10.595 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.808% | 4.983% | 1.319% | 7.112% | 3.534% | 5.125% | 7.397% | 5.807% | 7.817% | 27.12% | 11.34% | 5.475% | 0.02904% | 8.806% | 22.18% | 20.31% | 5.454% | 6.409% | 5.577% | 1.453% | 4.773% | 2.042% | 1.079% | 34.35% | 28.58% | 34.58% | 20.67% | 17.74% | 29.03% | 6.338% | 0.9455% | 0.6525% | 0.3767% | 3.578% | 4.496% | 1.659% | 3.656% | 1.553% | 31.17% | 16.43% | 3.547% | 3.599%
242 | 8_30_21_F_11 |  | 6/9/2023 2:31:03 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_F_11.596 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.084% | 1.658% | 1.952% | 10.35% | 3.392% | 3.832% | 6.572% | 3.473% | 4.775% | 10.67% | 1.101% | 2.598% | 1.022% | 2.66% | 16.2% | 10.01% | 1.535% | 2.087% | 1.235% | 1.535% | 4.885% | 1.54% | 0.7656% | 10.49% | 7.577% | 17.94% | 12.58% | 29.7% | 33.4% | 1.178% | 1.024% | 0.5886% | 0.6527% | 1.832% | 1.149% | 9.273% | 1.648% | 2.532% | 8.162% | 9.073% | 1.94% | 11.07%
243 | 8_30_21_UF_1 |  | 6/9/2023 2:36:51 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_UF_1.597 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 5.896% | 9.456% | 2.799% | 19.94% | 0.9417% | 9.69% | 6.292% | 7.746% | 5.303% | 23.52% | 2.93% | 2.183% | 0.8765% | 9.036% | 16.97% | 8.036% | 6.161% | 1.469% | 1.09% | 1.186% | 7.23% | 3.691% | 1.051% | 4.758% | 24.67% | 18.93% | 10.24% | 24.74% | 77.68% | 1.527% | 0.6846% | 1.437% | 0.5175% | 5.774% | 1.353% | 1.215% | 1.494% | 4.192% | 15.82% | 6.653% | 5.212% | 2.578%
244 | 8_30_21_UF_6 |  | 6/9/2023 2:42:39 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_UF_6.598 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.2087% | 2.063% | 4.506% | 73.76% | 1.812% | 0.7957% | 2.438% | 2.184% | 3.282% | 19.07% | 20.1% | 18.78% | 11.71% | 9.251% | 17.85% | 13.34% | 0.1619% | 13.81% | 18.07% | 1.52% | 9.837% | 13.94% | 11.39% | 14.05% | 60.49% | 19.64% | 7.237% | 14.94% | 28.22% | 14.54% | 10.97% | 0.5587% | 10.53% | 0.8394% | 11.8% | 0.9709% | 1.227% | 13.27% | 19.53% | 24.75% | 2.379% | 2.121%
245 | 8_30_21_UF_8 |  | 6/9/2023 2:48:27 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_UF_8.599 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.492% | 3.288% | 2.062% | 20.6% | 0.6201% | 3.62% | 4.891% | 1.36% | 6.683% | 11.68% | 8.722% | 6.891% | 0.8893% | 8.315% | 38.93% | 4.922% | 1.791% | 0.9534% | 3.092% | 0.8367% | 3.054% | 1.166% | 0.2311% | 17.28% | 24.8% | 21.83% | 16.38% | 42.92% | 38.92% | 1.116% | 0.2928% | 0.9796% | 0.3013% | 1.754% | 0.4663% | 0.7182% | 1.557% | 1.112% | 27.24% | 15.04% | 5.901% | 1.906%
246 | 8_30_21_UF_9 |  | 6/9/2023 2:54:15 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_UF_9.600 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4.45% | 4.705% | 7.158% | 22.77% | 0.7446% | 8.438% | 3.88% | 7.355% | 6.289% | 15.53% | 2.686% | 0.9806% | 3.215% | 11.5% | 4.93% | 9.837% | 8.915% | 1.406% | 18.88% | 3.927% | 20.32% | 1.06% | 2.037% | 47.02% | 27.05% | 15.3% | 9.225% | 9.697% | 13.25% | 0.9531% | 3.688% | 3.237% | 2.818% | 5.445% | 0.6707% | 0.5695% | 1.097% | 2.003% | 3.829% | 15.19% | 3.874% | 1.629%
247 | 8_30_21_UF_9_LD |  | 6/9/2023 3:00:03 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_UF_9_LD.601 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.952% | 1.47% | 3.457% | 9.167% | 1.04% | 5.905% | 4.489% | 5.826% | 10.69% | 8.479% | 13.57% | 4.545% | 1.035% | 5.82% | 28.63% | 7.227% | 3.406% | 7.582% | 22.39% | 3.58% | 6.92% | 0.8003% | 0.2034% | 23.93% | 89.36% | 4.169% | 39.76% | 17.52% | 35.58% | 7.746% | 0.5262% | 2.52% | 0.3991% | 2.459% | 6.431% | 0.4269% | 1.181% | 1.131% | 30.96% | 14.77% | 5.316% | 0.8127%
248 | 8_30_21_UF_9_LD |  | 6/9/2023 3:06:19 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_UF_9_LD.602 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.17% | 6.425% | 2.6% | 19.6% | 0.957% | 3.608% | 7.702% | 3.029% | 5.704% | 4.845% | 5.318% | 2.862% | 1.072% | 9.103% | 24.65% | 8.581% | 2.415% | 1.41% | 7.024% | 1.799% | 12.43% | 1.511% | 1.44% | 38.61% | 61.85% | 28.03% | 13.84% | 6.757% | 18.13% | 1.092% | 0.7622% | 1.372% | 0.204% | 3.743% | 1.448% | 0.7068% | 0.9383% | 0.7132% | 4.416% | 13.2% | 5.009% | 4.004%
249 | 8_30_21_UF_10 |  | 6/9/2023 3:12:07 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_UF_10.603 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.495% | 1.93% | 0.6525% | 19.97% | 0.725% | 2.441% | 4.652% | 1.332% | 4.453% | 12.15% | 9.127% | 5.52% | 1.5% | 1.015% | 16.09% | 3.014% | 4.499% | 0.9226% | 3.005% | 0.4146% | 8.098% | 3.658% | 2.566% | 19.13% | 65.12% | 25.16% | 22.08% | 10.38% | 32.55% | 1.615% | 1.319% | 0.2173% | 1.153% | 1.461% | 1.049% | 1.291% | 0.1052% | 2.285% | 15.36% | 2.842% | 3.173% | 3.67%
250 | LBLANK |  | 6/9/2023 3:17:56 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.604 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3415% | 19698% | 10.75% | 16.79% | 1.584% | 223.2% | 113.5% | 86.89% |  | 329.3% | 49.83% | 85.3% | 604.5% | 597% | 173.2% | 372.2% | 135.8% | 287.6% | 1.214% | 97.32% | 47.56% | 7.595% | 6.413% | 102.7% | 173.2% | 202.4% | 481.5% | 427.8% | 173.2% | 24.54% | 33.67% | 8.876% | 0.5208% | 4.792% | 0.2839% | 0.8184% | 1.055% | 2.787% | 173.2% | 247.5% | 356.9% | 98.27%
251 | CCV 1ppb TRACE |  | 6/9/2023 3:23:44 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.605 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 6.675% | 7.985% | 1.534% | 18.49% | 2.227% | 50.38% | 3.591% | 19.74% | 10.47% | 10.42% | 1.183% | 5.032% | 2.1% | 12.93% | 6.944% | 4.609% | 4.218% | 1.937% | 2.898% | 2.106% | 13.74% | 0.9451% | 1.815% | 4.088% | 9.731% | 2.214% | 4.42% | 9.313% | 5.293% | 1.419% | 0.1166% | 1.358% | 0.5748% | 3.713% | 1.627% | 5.649% | 0.6799% | 1.078% | 2.146% | 2.977% | 0.84% | 4.603%
252 | CCV 1ppb TRACE |  | 6/9/2023 3:30:02 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.606 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 6.216% | 7.331% | 0.8398% | 17.03% | 2.75% | 31.23% | 4.654% | 7.103% | 6.105% | 4.742% | 4.051% | 2.707% | 1.782% | 4.635% | 1.372% | 4.606% | 4.724% | 2.698% | 4.659% | 2.517% | 4.029% | 1.354% | 1.594% | 9.378% | 0.2108% | 8.168% | 3.13% | 2.966% | 5.839% | 3.209% | 1.081% | 1.574% | 0.211% | 0.6129% | 1.745% | 0.8026% | 1.759% | 2.836% | 1.885% | 2.382% | 1.249% | 1.575%
253 | CCV 10 ppb MAJORS |  | 6/9/2023 3:35:51 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.607 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.138% | 0.9681% | 1.629% | 9.378% | 0.604% | 4.723% | 3.653% | 3.564% | 1.743% | 1.433% | 0.7049% | 4.275% | 11.34% | 0.5299% | 1.372% | 0.6123% | 1.443% | 2.077% | 0.745% | 0.9149% | 5.105% | 11.73% | 10.66% | 5.473% | 7.266% | 10.92% | 8.995% | 2.106% | 8.769% | 1.454% | 10.59% | 0.8051% | 9.267% | 1.287% | 1.042% | 0.7017% | 2.35% | 10.17% | 0.8175% | 1.27% | 0.715% | 0.7479%
254 | CCV 10 ppb MAJORS |  | 6/9/2023 3:42:08 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.608 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.38% | 0.9318% | 13.31% | 7.83% | 21.23% | 3.669% | 2.194% | 3.944% | 2.522% | 0.4128% | 0.4045% | 0.3972% | 0.4864% | 2.439% | 0.5053% | 0.6539% | 0.7622% | 1.023% | 0.4301% | 10.62% | 3.479% | 1.385% | 0.4422% | 16.54% | 3.23% | 2.912% | 1.412% | 11.86% | 5.785% | 0.9524% | 0.4917% | 11.06% | 0.7684% | 1.187% | 0.3132% | 1.553% | 18.64% | 1.259% | 1.853% | 0.2559% | 12.55% | 2.723%
255 | 8_30_21_UF_10_LSP |  | 6/9/2023 3:47:56 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_UF_10_LSP.609 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.646% | 1.947% | 6.793% | 7.824% | 1.465% | 1.836% | 1.336% | 1.363% | 3.221% | 2.648% | 1.201% | 0.6748% | 3.437% | 1.887% | 1.858% | 0.7756% | 2.126% | 0.7694% | 1.231% | 5.98% | 3.233% | 4.828% | 3.515% | 7.285% | 5.317% | 3.944% | 3.414% | 7.716% | 6.92% | 0.419% | 3.357% | 6.686% | 2.493% | 1.972% | 0.7269% | 4.064% | 0.2291% | 3.865% | 0.9557% | 0.7731% | 7.81% | 4.877%
256 | 8_30_21_UF_10_DUP_SP |  | 6/9/2023 3:53:45 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_UF_10_DUP_SP.610 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.741% | 2.046% | 0.8263% | 4.998% | 1.224% | 3.75% | 2.918% | 1.575% | 3.03% | 2.672% | 2.763% | 3.394% | 0.05165% | 2.372% | 1.028% | 2.586% | 1.478% | 2.813% | 3.833% | 0.6864% | 5.876% | 1.21% | 0.1754% | 0.8687% | 2.581% | 2.054% | 0.8062% | 5.357% | 6.056% | 4.209% | 0.1833% | 0.5516% | 0.4731% | 1.335% | 2.533% | 1.251% | 0.691% | 1.585% | 2.764% | 2.607% | 0.5095% | 0.9613%
257 | 8_30_21_UF_11 |  | 6/9/2023 3:59:33 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_UF_11.611 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.3659% | 2.616% | 6.384% | 12.66% | 1.143% | 1.754% | 2.676% | 0.6746% | 10.73% | 6.967% | 10.12% | 0.6418% | 0.5169% | 1.698% | 41.85% | 6.929% | 1.997% | 2.518% | 2.424% | 4.776% | 5.008% | 3.742% | 1.607% | 17.32% | 10% | 25.69% | 3.814% | 58.03% | 28.72% | 0.6067% | 1.468% | 2.354% | 0.2224% | 0.6777% | 0.3621% | 0.8093% | 0.6421% | 2.235% | 9.939% | 6.465% | 3.632% | 0.5628%
258 | 8_30_21_BB |  | 6/9/2023 4:05:21 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_BB.612 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 27.61% | 161.1% | 1.146% | 3.065% | 2.12% | 84.87% | 35.66% | 100.4% | 49.37% | 31.88% | 7.801% | 12.92% | 13.77% | 1.227% | 368.5% | 153.8% | 4.459% | 12.49% | 25.47% | 40.76% | 0.7135% | 30.68% | 30.68% | 96.87% |  | 56.37% | 456.4% | 182.6% |  | 25.71% | 34.93% | 0.814% | 0.7483% | 0.7162% | 9.986% | 0.3897% | 0.8353% | 0.4636% | 86.61% | 63.12% | 1.089% | 0.8344%
259 | 8_30_21_FB |  | 6/9/2023 4:11:09 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_FB.613 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 28.67% | 76.16% | 0.8716% | 1.335% | 3.248% | 83.14% | 110.9% | 39.05% | 114.6% | 9.187% | 7.552% | 2.371% | 3.357% | 7.507% | 47.9% | 0.7307% | 2.857% | 1.759% | 2.367% | 36.38% |  | 12.98% | 13.59% | 51.83% | 86.61% | 142.2% | 79.71% | 35.84% | 173.2% | 15.44% | 9.365% | 1.721% | 0.9475% | 0.1432% | 1.18% | 0.5146% | 3.437% | 2.342% | 14.37% | 8.378% | 4.91% | 2.59%
260 | 8_30_21_LAB_A |  | 6/9/2023 4:16:56 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_LAB_A.614 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 142.8% | 1.47e-06% | 2.913% | 43.13% | 6.285% | 226.3% | 230.8% |  | 33.64% | 24.4% | 5.022% | 3.458% | 16.72% | 9.817% | 1322% | 16.21% | 6.842% | 2.285% | 11.99% | 23.39% |  | 19.82% | 17.91% | 15.42% |  | 1380% | 69.08% | 92.06% |  | 5.625% | 5.507% | 1.746% | 1.115% | 1.406% | 0.9107% | 2.348% | 2.376% | 1.211% | 86.61% | 285% | 1.431% | 6.922%
261 | 8_30_21_LAB_B |  | 6/9/2023 4:22:44 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_LAB_B.615 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 113.1% | 1.47e-06% | 3.213% | 49.86% | 2.47% | 338.8% | 77.36% | 100.9% |  | 35.41% | 7.758% | 9.684% | 21.55% | 20.13% | 219.2% | 9.109% | 7.013% | 7.53% | 8.362% | 63.5% | 173.2% | 29.2% | 28.43% | 14.73% |  | 243.2% | 109.5% | 94.48% | 173.2% | 62.88% | 68.22% | 1.641% | 1.699% | 1.699% | 3.759% | 0.2761% | 0.878% | 3.94% |  | 849.1% | 1.794% | 0.5518%
262 | 9_13_21_W_1 |  | 6/9/2023 4:28:32 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_W_1.616 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.455% | 3.549% | 0.7172% | 9.604% | 0.963% | 9.448% | 9.641% | 2.86% | 14.32% | 36.56% | 10.09% | 0.9553% | 0.4879% | 6.718% | 35.29% | 20.79% | 7.356% | 3.43% | 59.74% | 3.58% | 13.5% | 3.546% | 0.2565% | 50.63% | 34.53% | 65.81% | 87.17% | 533.6% | 173.2% | 1.333% | 0.5703% | 1.711% | 0.1605% | 4.198% | 1.216% | 0.3187% | 0.7964% | 3.785% | 50.14% | 31.56% | 4.795% | 1.816%
263 | 9_13_21_W_6 |  | 6/9/2023 4:34:20 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_W_6.617 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.306% | 6.062% | 1.839% | 4.741% | 4.387% | 4.812% | 19.5% | 3.58% | 8.651% | 14.47% | 8.434% | 1.093% | 0.6201% | 0.6052% | 20.13% | 4.374% | 4.024% | 2.27% | 3.029% | 5.365% | 24.63% | 3.145% | 1.123% | 27.89% | 86.6% | 1.178% | 12.02% | 41.88% | 49.12% | 0.1893% | 1.602% | 0.7042% | 1.049% | 0.39% | 0.4836% | 0.6744% | 2.204% | 2.677% | 39.59% | 13.34% | 0.8033% | 2.041%
264 | 9_13_21_W_8 |  | 6/9/2023 4:40:07 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_W_8.618 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.934% | 2.818% | 2.702% | 8.936% | 1.382% | 17.04% | 6.582% | 3.042% | 21.72% | 18.47% | 33.23% | 30.86% | 2.982% | 1.43% | 10.83% | 21.81% | 4.572% | 32.98% | 63.47% | 2.682% | 18.79% | 5.732% | 5.453% | 18.7% | 42.74% | 24.67% | 40.72% | 108.1% | 32.34% | 31.22% | 2.556% | 2.482% | 2.347% | 1.263% | 23.81% | 0.8854% | 1.545% | 3.874% | 22.34% | 25.78% | 3.093% | 3.228%
265 | LBLANK |  | 6/9/2023 4:45:55 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.619 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 148.2% | 1.47e-06% | 1.502% | 44.8% | 3.863% | 208% | 141.1% |  | 173.2% | 400.7% | 32.87% | 27.02% | 43.15% | 106.8% | 95.39% | 160.8% | 108.3% | 32.61% | 0.5108% | 22.05% |  | 64.97% | 65.03% | 64.02% |  | 230.9% | 657.2% | 124.9% |  | 36.95% | 42.16% | 0.8423% | 0.8979% | 2.629% | 1.447% | 0.8619% | 0.213% | 2.021% | 99.78% | 133.1% | 50.3% | 21.79%
266 | CCV 1ppb TRACE |  | 6/9/2023 4:51:44 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.620 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4.555% | 5.98% | 2.324% | 5.491% | 3.058% | 5.43% | 2.811% | 8.771% | 7.131% | 2.536% | 1.489% | 2.648% | 2.901% | 6.303% | 2.753% | 10.46% | 3.787% | 0.4675% | 2.367% | 11.36% | 18.16% | 2.219% | 1.472% | 7.91% | 9.481% | 8.233% | 9.409% | 27.35% | 7.842% | 1.471% | 1.558% | 1.05% | 0.6168% | 2.365% | 0.7681% | 1.512% | 0.7936% | 1.078% | 3.039% | 0.5117% | 1.338% | 0.243%
267 | CCV 1ppb TRACE |  | 6/9/2023 4:58:02 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.621 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.343% | 4.539% | 0.729% | 15.26% | 3.266% | 21.1% | 3.027% | 6.803% | 1.306% | 0.5016% | 2.108% | 3.728% | 1.41% | 4.212% | 1.051% | 5.309% | 0.9388% | 1.628% | 1.542% | 2.005% | 16% | 1.481% | 1.274% | 18.58% | 4.861% | 3.643% | 4.735% | 11.92% | 16.22% | 1.955% | 1.566% | 1.26% | 0.9046% | 1.009% | 0.4907% | 0.8966% | 1.375% | 1.213% | 0.9199% | 0.3998% | 0.6297% | 1.022%
268 | LFB |  | 6/9/2023 5:03:50 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.622 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.376% | 2.068% | 0.8337% | 3.475% | 1.342% | 2.098% | 4.053% | 6.14% | 0.2166% | 0.678% | 1.158% | 1.287% | 0.7749% | 0.8081% | 3.031% | 1.757% | 0.8985% | 1.664% | 2.034% | 2.055% | 8.013% | 1.038% | 0.9986% | 4.026% | 4.162% | 1.46% | 0.702% | 3.083% | 3.278% | 0.2506% | 1.327% | 2.218% | 0.7422% | 1.246% | 1.692% | 1.31% | 0.7621% | 1.379% | 0.4804% | 1.302% | 1.585% | 1.491%
269 | CCV 10 ppb MAJORS |  | 6/9/2023 5:09:38 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.623 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4.715% | 5.563% | 3.558% | 10.15% | 1.338% | 4.68% | 4.232% | 2.513% | 3.35% | 3.832% | 10.19% | 10.37% | 5.882% | 3.857% | 5.516% | 4.89% | 4.172% | 9.828% | 9.979% | 2.89% | 3.335% | 3.536% | 5.355% | 7.625% | 4.176% | 7.169% | 5.38% | 3.649% | 11.65% | 10.24% | 4.989% | 3.166% | 4.009% | 2.617% | 8.037% | 3.562% | 0.3541% | 2.135% | 8.806% | 9.727% | 4.396% | 4.963%
270 | CCV 10 ppb MAJORS |  | 6/9/2023 5:15:55 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.624 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.065% | 5.247% | 0.7762% | 3.303% | 1.259% | 3.444% | 3.381% | 2.236% | 1.938% | 0.8447% | 0.5241% | 0.7478% | 1.08% | 1.897% | 0.3734% | 0.8579% | 1.029% | 1.068% | 0.6711% | 1.331% | 1.876% | 1.089% | 1.303% | 1.516% | 3.021% | 2.027% | 2.176% | 3.037% | 1.003% | 0.5007% | 0.4501% | 1.002% | 0.8197% | 0.786% | 0.3963% | 0.5005% | 1.964% | 1.09% | 0.9651% | 1.003% | 1.073% | 1.295%
271 | 9_13_21_W_9 |  | 6/9/2023 5:21:43 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_W_9.625 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.953% | 2.314% | 16.3% | 15.57% | 1.826% | 9.853% | 2.742% | 4.788% | 19.64% | 20.98% | 6.92% | 1.145% | 1.782% | 1.353% | 51.21% | 50.41% | 1.056% | 2.149% | 22.53% | 9.819% | 18.83% | 1.946% | 2.935% | 60.68% | 86.61% | 54.68% | 20.14% | 539.5% | 25.2% | 0.565% | 2.103% | 13.73% | 0.5977% | 1.193% | 0.7479% | 0.8537% | 0.6691% | 2.345% | 9.364% | 31.86% | 10.8% | 1.728%
272 | 9_13_21_W_10 |  | 6/9/2023 5:27:31 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_W_10.626 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.152% | 3.136% | 1.156% | 9.113% | 1.349% | 4.455% | 9.043% | 3.252% | 7.171% | 33.34% | 18.18% | 1.673% | 0.4021% | 1.464% | 17.07% | 15.15% | 1.539% | 2.109% | 0.6291% | 6.859% | 8.237% | 3.523% | 1.918% | 71.53% | 173.2% | 11.73% | 24.95% | 110.6% | 173.2% | 0.8795% | 2.424% | 0.587% | 0.4249% | 0.1409% | 0.7192% | 1.239% | 0.4237% | 2.127% | 40.72% | 11.55% | 1.725% | 2.915%
273 | 9_13_21_W_10_DUP |  | 6/9/2023 5:33:20 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_W_10_DUP.627 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 6.275% | 2.071% | 2.934% | 11.68% | 1.553% | 7.672% | 3.328% | 4.911% | 6.043% | 50.03% | 15.1% | 1.926% | 1.176% | 1.359% | 47.16% | 14.71% | 2.171% | 0.8991% | 8.929% | 1.249% | 28.11% | 3.295% | 2.744% | 47.15% | 91.03% | 13.83% | 63.56% | 41.3% | 92.31% | 1.431% | 1.843% | 0.5151% | 1.36% | 1.166% | 0.4041% | 2.168% | 1.155% | 1.937% | 15.89% | 26.14% | 2.196% | 4.526%
274 | 9_13_21_W_11 |  | 6/9/2023 5:39:08 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_W_11.628 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.008% | 9.382% | 8.798% | 1.17% | 1.012% | 11.7% | 5.132% | 2.093% | 11.65% | 19.62% | 5.624% | 0.6265% | 0.4469% | 0.6594% | 13.45% | 25.19% | 2.279% | 1.879% | 1.174% | 5.369% | 14.62% | 2.416% | 4.588% | 80.27% | 115.4% | 13.59% | 49.18% | 800.3% | 43.89% | 1.518% | 3.932% | 8.16% | 1.464% | 0.6951% | 0.5997% | 1.677% | 0.1948% | 1.709% | 47.08% | 10.45% | 5.748% | 1.195%
275 | 9_13_21_F_1 |  | 6/9/2023 5:44:56 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_F_1.629 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.2507% | 3.902% | 2.527% | 13.05% | 1.358% | 2.633% | 2.462% | 2.542% | 12.41% | 15.02% | 2.661% | 0.5236% | 1.459% | 4.318% | 16.83% | 4.421% | 2.808% | 1.141% | 2.789% | 4.893% | 7.836% | 1.889% | 1.354% | 11.72% | 77.48% | 27.25% | 25.07% | 27.69% | 24.11% | 1.879% | 0.7186% | 4.279% | 1.298% | 0.9488% | 0.6295% | 0.8144% | 1.132% | 2.898% | 32.78% | 11.13% | 0.7464% | 1.933%
276 | 9_13_21_F_6 |  | 6/9/2023 5:50:44 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_F_6.630 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 7.012% | 6.699% | 1.023% | 5.892% | 0.6032% | 8.914% | 4.756% | 8.285% | 7.255% | 15.67% | 6.834% | 2.545% | 2.167% | 8.719% | 5.471% | 13.46% | 10.2% | 2.031% | 2.007% | 0.3202% | 14.27% | 4.006% | 1.485% | 17.47% | 73.15% | 17.25% | 15.34% | 10.73% | 59.77% | 1.844% | 0.9214% | 0.2528% | 1.13% | 6.022% | 2.067% | 1.781% | 1.131% | 3.901% | 14.37% | 3.813% | 5.312% | 1.611%
277 | 9_13_21_F_8 |  | 6/9/2023 5:56:32 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_F_8.631 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 6.671% | 7.977% | 2.351% | 12.41% | 2.409% | 6.364% | 6.473% | 3.939% | 5.578% | 7.968% | 7.232% | 1.663% | 0.8427% | 6.51% | 32.83% | 10.11% | 6.28% | 1.17% | 0.277% | 0.4955% | 11.09% | 1.129% | 0.641% | 20.14% | 31.29% | 26.8% | 6.851% | 14.32% | 53.01% | 0.864% | 1.056% | 1.081% | 0.1801% | 5.054% | 0.8072% | 0.9484% | 1.494% | 1.856% | 5.568% | 4.706% | 2.121% | 1.825%
278 | 9_13_21_F_9 |  | 6/9/2023 6:02:20 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_F_9.632 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 7.574% | 5.608% | 1.275% | 6.133% | 1.66% | 6.724% | 7.123% | 6.961% | 7.098% | 11.57% | 1.917% | 3.261% | 3.322% | 8.49% | 30.34% | 9.495% | 0.2445% | 2.949% | 3.129% | 0.6002% | 11.64% | 7.827% | 5.158% | 9.858% | 67.14% | 26.07% | 27.08% | 32.83% | 41.41% | 3.416% | 5.354% | 0.3624% | 3.666% | 5.978% | 1.724% | 0.8754% | 1.262% | 6.462% | 25.36% | 4.868% | 1.953% | 1.822%
279 | 9_13_21_F_10 |  | 6/9/2023 6:08:08 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_F_10.633 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.271% | 2.652% | 1.129% | 10.67% | 1.935% | 1.691% | 2.453% | 2.766% | 3.794% | 6.586% | 11.53% | 7.127% | 0.6899% | 2.874% | 41.99% | 6.654% | 2.004% | 9.608% | 11.61% | 2.148% | 5.536% | 0.5167% | 1.309% | 16.35% | 12.89% | 30.23% | 41.5% | 20.96% | 46.88% | 10.01% | 0.4176% | 0.383% | 0.8397% | 3.106% | 7.523% | 0.6527% | 0.4728% | 0.09927% | 12.36% | 32.18% | 5.284% | 2.073%
280 | 9_13_21_F_11 |  | 6/9/2023 6:13:56 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_F_11.634 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.693% | 3.904% | 1.036% | 8.307% | 1.935% | 5.229% | 7.977% | 3.637% | 1.422% | 18.7% | 10.74% | 4.511% | 0.8228% | 3.159% | 28.13% | 24.6% | 2.834% | 4.422% | 6.282% | 2.189% | 12.51% | 3.173% | 0.6123% | 21.28% | 50.08% | 6.757% | 25.55% | 14.03% | 48.15% | 4.491% | 0.388% | 1.915% | 0.2078% | 3.012% | 3.366% | 2.789% | 1.788% | 3.673% | 21.63% | 7.306% | 1.856% | 5.004%
281 | 9_13_21_F_11_LD |  | 6/9/2023 6:19:44 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_F_11_LD.635 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 7.18% | 6.497% | 4.744% | 13.19% | 1.406% | 4.56% | 6.654% | 6.587% | 2.494% | 12.34% | 3.018% | 2.453% | 2.51% | 8.846% | 15.45% | 18.24% | 6.255% | 1.572% | 1.707% | 3.463% | 10.93% | 3.146% | 3.795% | 11.79% | 86.65% | 20.9% | 25.66% | 9.808% | 31.26% | 2.295% | 3.169% | 2.377% | 2.108% | 5.02% | 0.06788% | 1.014% | 1.819% | 0.9668% | 4.028% | 6.153% | 5.541% | 3.31%
282 | 9_13_21_F_11_LD |  | 6/9/2023 6:26:00 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_F_11_LD.636 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.493% | 1.858% | 2.173% | 2.525% | 3.418% | 2.386% | 4.039% | 1.844% | 6.005% | 21.18% | 7.461% | 0.6966% | 1.297% | 1.469% | 25.88% | 10.32% | 1.416% | 0.9904% | 1.132% | 0.8949% | 4.433% | 5.165% | 0.9381% | 16.3% | 27.07% | 22.09% | 21.83% | 36.06% | 19.87% | 0.469% | 1.416% | 1.663% | 0.1337% | 0.9135% | 1.239% | 1.175% | 1.141% | 4.819% | 23.98% | 2.322% | 2.307% | 0.1788%
283 | LBLANK |  | 6/9/2023 6:31:48 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.637 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 307.5% | 1.47e-06% | 0.1297% | 95.96% | 23.01% | 99.2% | 38.46% |  | 173.2% | 139.2% | 145.1% | 26.69% | 60.21% | 73.44% | 104.7% | 83.25% | 113.6% | 17621% | 0.8029% | 7.341% |  | 75.15% | 82.37% | 72.36% | 173.2% | 235.7% | 120.8% | 947% | 173.2% | 55.21% | 16.41% | 5.394% | 1.186% | 1.318% | 0.04205% | 1.383% | 1.301% | 3.179% | 173.2% | 440.7% | 101.7% | 22.43%
284 | CCV 1ppb TRACE |  | 6/9/2023 6:37:37 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.638 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.692% | 13.46% | 0.9091% | 30.97% | 1.155% | 40.75% | 3.945% | 4.408% | 1.362% | 1.413% | 2.015% | 2.701% | 1.379% | 3.05% | 2.788% | 4.494% | 2.725% | 1.303% | 4.514% | 2.855% | 13.81% | 4.069% | 1.188% | 8.306% | 4.61% | 1.288% | 6.886% | 14.03% | 11.91% | 0.2512% | 1.581% | 0.604% | 0.5305% | 1.872% | 0.5734% | 3.518% | 0.521% | 3.35% | 1.59% | 1.352% | 1.453% | 2.552%
285 | CCV 1ppb TRACE |  | 6/9/2023 6:43:54 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.639 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 8.09% | 12.72% | 3.413% | 46.27% | 6.555% | 10.33% | 11.73% | 7.608% | 6.547% | 4.71% | 2.076% | 1.801% | 1.488% | 4.698% | 3.123% | 10.39% | 3.557% | 1.926% | 2.736% | 7.899% | 21.14% | 4.732% | 3.241% | 4.464% | 4.146% | 7.181% | 6.196% | 10.71% | 5.012% | 1.073% | 0.8437% | 1.363% | 1.239% | 3.178% | 0.7482% | 14.15% | 0.2544% | 2.792% | 1.936% | 2.151% | 2.93% | 16.69%
286 | CCV 10 ppb MAJORS |  | 6/9/2023 6:49:43 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.640 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.749% | 5.022% | 13.28% | 4.743% | 0.9859% | 1.705% | 1.259% | 1.922% | 0.9033% | 1.415% | 1.565% | 0.9826% | 7.055% | 2.107% | 2.346% | 1.239% | 1.061% | 0.8632% | 1.558% | 11% | 4.708% | 8.9% | 7.185% | 11.74% | 5.944% | 9.171% | 4.744% | 10.41% | 7.393% | 0.9807% | 6.832% | 9.981% | 6.897% | 0.4346% | 0.357% | 0.6787% | 0.253% | 8.823% | 1.072% | 0.66% | 12.15% | 1.204%
287 | CCV 10 ppb MAJORS |  | 6/9/2023 6:56:01 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.641 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.867% | 2.13% | 1.666% | 11.62% | 1.083% | 2.279% | 2.572% | 1.493% | 0.1942% | 0.5896% | 4.96% | 3.374% | 2.187% | 1.478% | 1.089% | 1.486% | 1.024% | 3.637% | 3.904% | 1.275% | 3.403% | 0.3541% | 1.051% | 0.5195% | 3.976% | 0.6026% | 2.795% | 4.33% | 1.808% | 3.564% | 0.8848% | 1.495% | 0.8839% | 1.35% | 1.917% | 10.99% | 0.7604% | 0.3437% | 4.11% | 3.562% | 0.4474% | 11.43%
288 | 9_13_21_F_11_DUP |  | 6/9/2023 7:01:49 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_F_11_DUP.642 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.11% | 4.266% | 16.3% | 2.613% | 1.321% | 2.996% | 6.977% | 3.393% | 5.586% | 16.42% | 3.868% | 1.38% | 0.6906% | 0.7655% | 11.77% | 3.487% | 4.575% | 0.4814% | 1.463% | 13.77% | 17.58% | 2.186% | 1.021% | 76.11% | 41.48% | 24.28% | 22.79% | 15.22% | 19.63% | 2.483% | 1.727% | 13.54% | 0.3303% | 0.9009% | 0.3685% | 0.5912% | 2.615% | 2.573% | 3.55% | 12.18% | 18.65% | 0.2852%
289 | 9_13_21_F_11_DUP_LSP |  | 6/9/2023 7:07:37 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_F_11_DUP_LSP.643 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.089% | 2.18% | 2.899% | 10.37% | 0.8233% | 2.204% | 3.642% | 0.8349% | 2.303% | 0.5409% | 0.4646% | 1.044% | 6.761% | 1.61% | 1.299% | 0.587% | 1.853% | 1.906% | 1.407% | 1.565% | 2.553% | 3.812% | 6.886% | 1.67% | 10.89% | 6.689% | 8.195% | 1.291% | 5.421% | 1.256% | 6.105% | 1.282% | 6.093% | 1.678% | 1.42% | 2.223% | 0.2122% | 2.894% | 1.232% | 2.247% | 0.9981% | 1.468%
290 | 9_13_21_UF_1 |  | 6/9/2023 7:13:24 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_UF_1.644 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.342% | 5.07% | 14.61% | 31.23% | 1.627% | 5.208% | 6.487% | 3.015% | 5.31% | 31.89% | 6.817% | 3.654% | 1.314% | 3.098% | 8.763% | 4.772% | 9.039% | 0.7656% | 3.516% | 11.55% | 8.957% | 2.464% | 1.933% | 47.33% | 24.64% | 15.59% | 21.92% | 37.49% | 88% | 0.8515% | 1.421% | 9.958% | 1.597% | 2.334% | 0.3272% | 2.583% | 0.5097% | 2.582% | 11.49% | 22.41% | 15.18% | 5.644%
291 | 9_13_21_UF_6 |  | 6/9/2023 7:19:12 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_UF_6.645 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.787% | 4.198% | 3.823% | 28.35% | 1.926% | 1.733% | 1.696% | 1.241% | 5.54% | 9.273% | 7.879% | 7.529% | 3.904% | 6.013% | 14.05% | 7.884% | 0.7564% | 3.836% | 3.246% | 1.606% | 7.846% | 6.418% | 4.537% | 26.08% | 12.57% | 11.34% | 27.49% | 11.56% | 74.42% | 3.206% | 2.97% | 0.8797% | 2.995% | 0.8532% | 2.878% | 1.589% | 1.367% | 4.814% | 8.295% | 2.995% | 7.465% | 1.163%
292 | 9_13_21_UF_6_DUP |  | 6/9/2023 7:25:00 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_UF_6_DUP.646 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.738% | 1.048% | 3.517% | 26.95% | 1.333% | 2.703% | 4.173% | 0.4658% | 11.06% | 7.634% | 4.762% | 1.794% | 7.987% | 13.39% | 8.217% | 10.52% | 3.942% | 2.524% | 2.055% | 0.7252% | 8.409% | 5.291% | 9.388% | 31.99% | 72.31% | 29.48% | 15.15% | 19.59% | 52.36% | 1.768% | 8.532% | 0.7562% | 7.21% | 1.161% | 0.6748% | 1.753% | 0.9393% | 3.764% | 5.133% | 14.92% | 2.643% | 1.646%
293 | 9_13_21_UF_8 |  | 6/9/2023 7:30:49 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_UF_8.647 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.976% | 2.278% | 8.755% | 24.84% | 0.8619% | 2.284% | 4.323% | 2.42% | 8.897% | 21.21% | 5.701% | 3.903% | 1.929% | 2.746% | 11.37% | 11.99% | 0.7855% | 0.9351% | 4.405% | 5.27% | 3.268% | 1.578% | 2.866% | 59.18% | 33.96% | 28.69% | 11.09% | 6.977% | 104.3% | 0.9257% | 1.777% | 3.854% | 1.74% | 1.457% | 0.741% | 0.6348% | 2.173% | 1.886% | 25.04% | 11.89% | 3.998% | 2.618%

## Concentrations
| Sample Id | R | Acquisition Time | QC Status | Dataset File | Method File | Na KED 23/23 Helium KED_Low (ppb) | Mg KED 24/24 Helium KED_Low (ppb) | Al 27/27 (ppb) | Al KED 27/27 Helium KED_Low (ppb) | P Oshift 31/47 Oxygen DRC (ppb) | S (34) -1 34/34 Helium KED_Low (ppb) | K KED 39/39 Helium KED_Low (ppb) | Ca KED 43/43 Helium KED_Low (ppb) | V KED 51/51 Helium KED_Low (ppb) | Cr (52) 52/52 Helium KED_Low (ppb) | Cr (53) NH3 53/53 Ammonia DRC (ppb) | Fe (54) NH3 54/54 Ammonia DRC (ppb) | Mn Oshift 55/71 Oxygen DRC (ppb) | Fe (56) KED 56/56 Helium KED_Low (ppb) | Co KED 59/59 Helium KED_Low (ppb) | Ni KED 60/60 Helium KED_Low (ppb) | Cu (63) KED 63/63 Helium KED_Low (ppb) | Cu (65) NH3 65/65 Ammonia DRC (ppb) | Zn NH3 66/66 Ammonia DRC (ppb) | As 75/75 (ppb) | As KED 75/75 Helium KED_Low (ppb) | As RH Oshift 75/91 Oxygen DRC (ppb) | As Oshift 75/91 Oxygen DRC (ppb) | Se (77) 77/77 (ppb) | Se (77) Oshift 77/93 Oxygen DRC (ppb) | Se (78) Oshift 78/94 Oxygen DRC (ppb) | Se (80) Oshift 80/96 Oxygen DRC (ppb) | Se (82) 82/82 (ppb) | Se (82) Oshift 82/98 Oxygen DRC (ppb) | Mo (95) NH3 95/95 Ammonia DRC (ppb) | Mo (95) OShift 95/127 Oxygen DRC (ppb) | Rh (IS1) 103/103 (IS) | Rh (IS2) 103/103 (IS) Oxygen DRC | Rh (IS3) 103/103 (IS) Helium KED_Low | Rh (IS4) 103/103 (IS) Ammonia DRC | Rh (IS5) 103/103 (IS) Ammonia DRC | Rh (IS6) 103/103 (IS) Oxygen DRC | Rh OShift (IS1) 103/119 (IS) Oxygen DRC | Cd (111) NH3 111/111 Ammonia DRC (ppb) | Cd (114) NH3 114/114 Ammonia DRC (ppb) | Pb 208/208 (ppb) | Pb NH3 208/208 Ammonia DRC (ppb)
1 | Sac Blank |  | 06/08/2023 13:23:24 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\Sac Blank.356 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth
2 | sac blank cal |  | 06/08/2023 13:29:13 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\sac blank cal.357 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth
3 | CAL BLANK |  | 06/08/2023 13:35:02 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CAL BLANK.358 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | [31 empty] | 100% | 100% | 100% | 100% | 100% | 100% | 100%
4 | LOW 0.01 |  | 06/08/2023 13:40:50 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LOW 0.01.359 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | [31 empty] | 99.46% | 99.5% | 98.99% | 98.96% | 99.02% | 97.28% | 98.84%
5 | LOW 0.05 |  | 06/08/2023 13:46:39 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LOW 0.05.360 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth |  |  |  |  | 0.05 |  |  |  | 0.05 | 0.05 | 0.05 |  | 0.05 |  | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | [6 empty] | 0.05 | 0.05 | 96.71% | 97.54% | 96.66% | 96.06% | 99.78% | 97.35% | 96.72% | 0.05 | 0.05 | 0.05 | 0.05
6 | LOW 0.1 |  | 06/08/2023 13:52:27 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LOW 0.1.361 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth |  |  |  |  | 0.08067 |  |  |  | 0.096063 | 0.090199 | 0.097504 |  | 0.12719 |  | 0.10617 | 0.094887 | 0.098589 | 0.097288 | 0.22882 | 0.077063 | 0.098368 | 0.10104 | 0.10364 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.11127 | 0.10952 | 98.13% | 92.78% | 98.04% | 98.58% | 99.13% | 96.3% | 92.28% | 0.09857 | 0.097719 | 0.10179 | 0.10211
7 | LOW 0.5 |  | 06/08/2023 13:58:15 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LOW 0.5.362 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth |  |  |  |  | 0.15937 |  |  |  | 0.53897 | 0.49784 | 0.48804 | 0.5 | 0.60519 | 0.5 | 0.52415 | 0.51285 | 0.51696 | 0.49928 | 0.6686 | 0.48375 | 0.42073 | 0.49958 | 0.50119 | 0.46184 | 0.45275 | 0.49141 | 0.50515 | 0.6874 | 0.47689 | 0.55331 | 0.55418 | 98.47% | 95.89% | 93.43% | 97.42% | 97.81% | 93.98% | 95.72% | 0.48389 | 0.49637 | 0.49775 | 0.508
8 | STD 1 |  | 06/08/2023 14:04:03 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 1.363 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1000 | 1000 | 1 | 1 | 0.19292 |  | 1000 | 1000 | 1.0447 | 0.98556 | 0.97875 | 1.0097 | 1.1676 | 0.96268 | 1.0039 | 1.0365 | 1.01 | 0.98232 | 1.1638 | 1.0379 | 0.89754 | 1.0048 | 0.99778 | 0.89607 | 0.98087 | 0.95947 | 0.97561 | 1.1989 | 0.90082 | 1.0975 | 1.0879 | 97.03% | 95.25% | 95.49% | 95.17% | 95.56% | 93.62% | 92.09% | 0.96845 | 0.99495 | 0.98312 | 1.0259
9 | STD 5 |  | 06/08/2023 14:09:52 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 5.364 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4979.8464 | 4887.6514 | 4.8266 | 4.8375 | 0.96382 | 5000 | 5009.9906 | 5128.1752 | 4.7585 | 4.7674 | 4.8955 | 3.7294 | 5.5985 | 3.3639 | 4.8693 | 4.9649 | 4.6991 | 4.7033 | 5.4062 | 5.3697 | 4.1764 | 5.0013 | 4.9421 | 4.9 | 5.0152 | 4.8055 | 4.9715 | 5.7652 | 4.9769 | 5.327 | 5.3031 | 87.4% | 91.43% | 94.1% | 90.52% | 93.86% | 91.41% | 86.68% | 4.833 | 4.9158 | 5.127 | 4.8222
10 | STD 10 |  | 06/08/2023 14:15:40 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 10.365 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9973.9677 | 9988.6715 | 9.4441 | 9.4057 | 2.1142 | 9955.8343 | 10200.131 | 10523.2366 | 10.0664 | 9.6701 | 10.0256 | 8.2264 | 10.9119 | 7.4698 | 10.1718 | 9.9669 | 9.4727 | 9.4746 | 10.5784 | 10.1302 | 9.3372 | 9.9509 | 9.8405 | 9.2726 | 9.8747 | 9.2901 | 9.7564 | 10.2538 | 9.4619 | 10.7162 | 10.3566 | 91.65% | 91.28% | 90.53% | 89.4% | 91.31% | 89.1% | 88.07% | 9.8037 | 9.9269 | 9.4953 | 9.7702
11 | STD 25 |  | 06/08/2023 14:21:28 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 25.366 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 25315.1502 | 24630.9065 | 23.7289 | 24.0879 | 6.1982 | 26128.5984 | 25828.9422 | 26778.3022 | 25.9364 | 24.9548 | 25.6624 | 21.9204 | 28.0206 | 20.3561 | 25.5579 | 25.1387 | 24.2331 | 23.9183 | 26.678 | 26.3194 | 22.9203 | 25.2615 | 25.9924 | 23.6255 | 25.999 | 25.5927 | 25.9781 | 26.8225 | 25.666 | 27.2951 | 27.1937 | 87.69% | 84.4% | 86.12% | 83.56% | 92.98% | 82.26% | 80.14% | 25.1116 | 25.1135 | 23.4982 | 22.1433
12 | STD 50 |  | 06/08/2023 14:27:16 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 50.367 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 49447.0918 | 48504.3884 | 51.7875 | 48.3518 | 13.535 | 50362.4194 | 51311.0486 | 52594.7474 | 52.3862 | 50.3191 | 51.2753 | 44.7757 | 55.9035 | 41.9217 | 50.6641 | 49.4133 | 47.4587 | 47.916 | 52.5765 | 52.2393 | 46.1585 | 50.1529 | 52.7285 | 47.5563 | 51.763 | 51.4843 | 52.3632 | 53.9298 | 53.109 | 55.2276 | 55.6385 | 85.02% | 79.93% | 83.63% | 80.44% | 80.75% | 79.72% | 77.01% | 49.4126 | 49.4946 | 45.7608 | 47.8458
13 | STD 100 |  | 06/08/2023 14:33:03 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 100.368 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 100079.1431 | 93962.2249 | 103.1824 | 94.9518 | 28.971 | 101591.0065 | 103348.6199 | 106494.6693 | 107.8549 | 104.035 | 104.0079 | 90.2413 | 112.5456 | 86.3357 | 104.4566 | 98.7554 | 94.638 | 92.8021 | 100.6699 | 105.0888 | 91.5374 | 100.4066 | 103.1636 | 94.7812 | 105.5344 | 104.1968 | 104.8864 | 108.9352 | 105.6718 | 113.8257 | 113.1398 | 81.4% | 74.98% | 78.53% | 76.12% | 76.6% | 73.82% | 70.48% | 97.3575 | 98.2282 | 88.3585 | 98.0191
14 | Calibration Curves |  | 06/08/2023 14:33:03
15 | LBLANK |  | 06/08/2023 14:38:52 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.369 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | -11.7416 | -41.0609 | -0.69918 | -0.76605 | -0.23825 | 228.7424 | 3.2182 | -30.7443 | 0.0068477 | 0.005656 | 0.0027179 | -0.26883 | -0.024393 | -0.17974 | 0.002683 | 0.022413 | -0.004841 | -0.0062042 | -0.11381 | -0.041553 | 0.0084722 | 0.16761 | 0.0085315 | 0.15939 | -0.018378 | 0.0039245 | 0.0059578 | 0.025924 | 0.00737 | 0.026908 | 8.4559e-05 | 102.9% | 100.2% | 103.7% | 100.8% | 92.28% | 98.09% | 99.95% | 0.0015638 | 0.0006736 | -0.0017921 | -0.0013804
16 | LBLANK |  | 06/08/2023 14:57:55 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.370 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | -9.9038 | -42.7414 | -0.67285 | -0.78437 | -2.0623 | 141.26 | 2.8997 | -30.7443 | 0.0056732 | 0.0069391 | 0.0030874 | -0.51549 | -0.025186 | -0.55089 | 0.0021117 | 0.0043573 | -0.0080637 | -0.0072868 | -0.15372 | -0.021739 | 0.023137 | 0.15941 | -0.00034732 | 0.085389 | -0.010775 | 0.006613 | 0.00092024 | -0.022014 | 0.0047935 | -0.0059356 | -0.0081888 | 95.56% | 94.59% | 89.02% | 95.99% | 98.16% | 92.08% | 96.41% | 0.00028119 | -0.00046423 | -0.0022749 | -0.0019186
17 | ICV PE |  | 06/08/2023 15:03:43 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\ICV PE.371 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9910.2106 | 9986.4089 | 9.9707 | 9.3487 | 19.0658 | 9589.0762 | 9988.7685 | 10156.8568 | 9.8933 | 10.0135 | 9.74 | 10.0444 | 10.4635 | 10.2 | 10.099 | 10.2574 | 9.7402 | 9.5326 | 9.5867 | 10.1386 | 8.2954 | 9.5977 | 10.0923 | 7.5385 | 10.1506 | 9.9938 | 10.0369 | 10.9266 | 10.1441 | 10.035 | 10.1813 | 88.2% | 87.61% | 90.95% | 88.43% | 95.41% | 86.78% | 85.58% | 9.7246 | 9.8909 | 9.9805 | 9.4044
18 | ICV PE |  | 06/08/2023 15:09:59 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\ICV PE.372 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9801.7516 | 10242.1682 | 9.4473 | 9.6363 | 19.5113 | 9856.9094 | 9816.7945 | 10287.4771 | 10.0018 | 10.086 | 9.964 | 10.3213 | 10.5281 | 10.472 | 10.1863 | 10.3198 | 9.8091 | 9.7791 | 9.8334 | 9.7628 | 8.4124 | 9.1723 | 10.0533 | 7.3008 | 9.9368 | 9.6553 | 9.9805 | 10.0546 | 9.6627 | 10.0993 | 10.2349 | 91.96% | 87.46% | 90.05% | 87.41% | 90.23% | 85.93% | 89.14% | 9.9102 | 10.114 | 9.601 | 9.8706
19 | CCV 1ppb TRACE |  | 06/08/2023 15:15:47 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.373 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 963.5918 | 970.3544 | 1.0188 | 0.95078 | -0.56547 | 1344.678 | 976.128 | 962.5777 | 1.0279 | 1.0071 | 0.99805 | 0.59771 | 1.0075 | 0.59076 | 1.0292 | 1.0562 | 1.0438 | 1.0291 | 0.85937 | 0.94406 | 0.80569 | 1.0664 | 0.99222 | 1.0858 | 0.96297 | 0.99192 | 1.03 | 1.0712 | 0.85187 | 1.0403 | 1.0379 | 96.74% | 94.33% | 94.85% | 95.68% | 96.16% | 90.36% | 94.25% | 0.9723 | 1.0016 | 0.98748 | 1.0285
20 | CCV 1ppb TRACE |  | 06/08/2023 15:22:03 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.374 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1016.4433 | 1053.5242 | 1.0224 | 0.89479 | -0.98102 | 1072.081 | 986.0152 | 905.2676 | 0.98098 | 0.99507 | 0.99288 | 0.66355 | 1.0327 | 0.6398 | 1.0325 | 1.066 | 1.0819 | 1.0381 | 0.86327 | 1.0362 | 0.91812 | 1.0845 | 1.0114 | 1.0792 | 0.98803 | 0.94675 | 0.98258 | 1.0218 | 1.0772 | 1.0419 | 1.032 | 94.29% | 90.06% | 94.61% | 91.67% | 94.59% | 88.07% | 90.02% | 1.0026 | 1.0096 | 0.98362 | 1.0027
21 | LFB |  | 06/08/2023 15:27:51 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.375 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9591.8017 | 9708.532 | 8.8195 | 9.0889 | 9.9079 | 10108.5395 | 9525.6146 | 9366.2479 | 9.761 | 9.5988 | 10.0447 | 10.0076 | 10.2607 | 9.7319 | 10.0643 | 10.1776 | 9.6645 | 9.8918 | 10.289 | 9.5474 | 9.0222 | 9.4115 | 10.1802 | 9.4066 | 9.9294 | 9.7458 | 9.9836 | 9.9387 | 9.7442 | 10.648 | 10.4375 | 89.3% | 83.99% | 88.41% | 83.1% | 87.21% | 84.01% | 84.41% | 9.9167 | 10.1139 | 9.1342 | 9.5711
22 | LLOQ 0.01 |  | 06/08/2023 15:33:39 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LLOQ 0.01.376 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | -1.6627 | -30.1172 | -0.23747 | -0.2305 | -1.0393 | 358.4147 | 18.1369 | -16 | 0.01364 | 0.012039 | 0.010911 | -0.37113 | 0.0079147 | -0.38952 | 0.01076 | 0.01227 | 0.00791 | 0.011354 | -0.038124 | 0.033303 | 0.0045282 | 0.1668 | 0.0078158 | 0.013396 | 0.004558 | 0.015157 | 0.0084529 | -0.0978 | 0.017665 | 0.014279 | 0.012314 | 99.38% | 92.53% | 94.36% | 94% | 97.31% | 90.98% | 93.77% | 0.010372 | 0.010054 | 0.0099926 | 0.0090555
23 | CCV 10 ppb MAJORS |  | 06/08/2023 15:39:28 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.377 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | -12.8383 | -41.1886 | -0.73032 | -0.80788 | -2.3731 | -1892.7639 | -3.8651 | -28.9428 | 0.0032821 | 0.0021771 | 0.0008629 | -0.53724 | -0.025739 | -1.1955 | 0.0040634 | -0.00086363 | -0.0091461 | -0.0092637 | -0.1815 | 0.53901 | -0.0057419 | 0.15706 | -0.0030757 | -1.1802 | -0.01521 | 0.0015557 | 0.00049202 | -0.049839 | 0.0041782 | -0.011053 | -0.013216 | 107.1% | 103.8% | 108.8% | 112.7% | 111.4% | 98.11% | 101.1% | 0.00036437 | -0.00088094 | -0.0026241 | -0.0023583
24 | CCV 10 ppb MAJORS |  | 06/08/2023 15:45:42 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.378 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10570.1578 | 10848.3085 | 8.4876 | 9.8854 | 7.878 | 10722.0026 | 10934.0449 | 10683.3747 | 10.7684 | 10.625 | 9.9121 | 9.8704 | 10.3925 | 10.5423 | 10.752 | 10.8388 | 10.5791 | 9.6384 | 9.7375 | 9.9639 | 9.7083 | 9.3648 | 9.9952 | 9.2899 | 9.8888 | 9.6532 | 9.6943 | 10.0276 | 9.5757 | 10.4552 | 10.571 | 88.57% | 83.78% | 83.11% | 85.07% | 86.08% | 82.86% | 83.09% | 9.7097 | 9.8231 | 9.3278 | 9.859
25 | LLOQ 0.1 |  | 06/08/2023 15:51:31 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LLOQ 0.1.379 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 99.835 | 75.8092 | 0.065718 | -0.051843 | -0.56644 | 279.1031 | 103.7612 | 103.8505 | 0.096542 | 0.11222 | 0.10337 | -0.098896 | 0.089628 | -0.1612 | 0.097228 | 0.11966 | 0.11388 | 0.10932 | 0.10192 | 0.15234 | 0.05636 | 0.24575 | 0.095374 | 0.14192 | 0.036329 | 0.11013 | 0.10514 | 0.1039 | 0.12918 | 0.09714 | 0.096442 | 94.78% | 90.18% | 96.15% | 92.76% | 94.03% | 88.56% | 91.39% | 0.094455 | 0.09834 | 0.094652 | 0.099408
26 | STD 0.5 |  | 06/08/2023 15:57:19 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 0.5.380 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 501.3734 | 488.7631 | 0.46956 | 0.75969 | 0.31372 | 724.5439 | 471.3807 | 482.3687 | 0.51427 | 0.54837 | 0.56053 | 0.69189 | 0.51624 | 0.53141 | 0.51407 | 0.527 | 0.55145 | 0.60878 | 0.61458 | 0.53443 | 0.43089 | 0.61411 | 0.50895 | 0.60207 | 0.44687 | 0.53376 | 0.52509 | 0.45274 | 0.40338 | 0.57524 | 0.51412 | 91.97% | 88.13% | 92.05% | 82.4% | 92.73% | 87.25% | 90.2% | 0.55709 | 0.55807 | 0.49293 | 0.51428
27 | 8_2_21_UF_1 |  | 06/08/2023 16:03:07 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_UF_1.381 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 5502.7123 | 5515.3958 | 1.5052 | 1.5284 | 5.452 | 12329.3522 | 1218.5089 | 28419.4625 | 0.396 | 0.062235 | 0.058805 | 0.32235 | 15.1885 | 0.62754 | 0.020633 | 0.24011 | 0.93314 | 0.90947 | 1.4522 | 4.3289 | 3.7252 | 4.3594 | 4.3801 | 0.32472 | 0.035145 | 0.049722 | 0.048206 | 0.30367 | 0.035435 | 5.2008 | 5.2956 | 83.67% | 79.94% | 84.3% | 81.16% | 83.78% | 80.21% | 76.14% | 0.006901 | 0.0061183 | 0.013707 | 0.015818
28 | 8_2_21_UF_6 |  | 06/08/2023 16:08:55 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_UF_6.382 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10673.5634 | 6993.106 | 0.19846 | 0.11781 | 3.1604 | 12165.0275 | 2095.5035 | 34992.2024 | 1.0089 | 0.1001 | 0.089526 | 0.4944 | 13.7741 | 0.7065 | 0.022205 | 0.20966 | 2.1545 | 2.1857 | 2.2887 | 5.8023 | 4.8226 | 5.8373 | 6.1364 | 0.35773 | 0.090448 | 0.11082 | 0.11838 | 0.40007 | 0.11578 | 3.4543 | 3.4772 | 81.81% | 77.88% | 82% | 77.37% | 81.77% | 79.76% | 76.9% | 0.018245 | 0.015568 | 0.032344 | 0.034105
29 | 8_2_21_UF_6_SP |  | 06/08/2023 16:14:43 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_UF_6_SP.383 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 16267.1399 | 12850.6642 | 5.5262 | 5.6011 | 10.7677 | 17182.7234 | 7761.035 | 40810.4893 | 6.4601 | 5.409 | 5.5981 | 6.5072 | 19.9571 | 6.7763 | 5.5375 | 5.6914 | 7.3306 | 7.3327 | 7.4577 | 10.9728 | 10.2813 | 10.6111 | 11.5188 | 5.2769 | 5.4806 | 5.3618 | 5.4247 | 5.9916 | 5.4688 | 9.0518 | 9.3038 | 81.37% | 76.81% | 80.53% | 77.39% | 79.56% | 69.11% | 77.34% | 5.2923 | 5.283 | 5.1818 | 5.5548
30 | STD 1 |  | 06/08/2023 16:20:31 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 1.384 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1061.9021 | 1022.3951 | 1.1752 | 1.0968 | -0.058185 | 1483.1061 | 1098.78 | 1068.6497 | 1.0982 | 1.0578 | 0.99866 | 1.5526 | 1.0694 | 1.292 | 1.0498 | 1.0304 | 1.0501 | 1.0395 | 0.9972 | 1.0075 | 0.94017 | 1.1671 | 1.0899 | 1.1396 | 1.0234 | 0.97816 | 1.0129 | 1.068 | 0.92071 | 1.0446 | 1.0856 | 88.61% | 83% | 87.74% | 85.13% | 89.22% | 81.6% | 82.16% | 1.0151 | 1.0189 | 1.0151 | 1.0442
31 | 8_2_21_UF_6_SP_LD |  | 06/08/2023 16:26:19 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_UF_6_SP_LD.385 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 16322.174 | 13028.9898 | 6.4476 | 6.0626 | 10.7892 | 17646.209 | 7587.6511 | 41803.4979 | 6.4397 | 5.4944 | 5.6204 | 5.9489 | 19.4195 | 6.1644 | 5.3805 | 5.673 | 7.4264 | 7.3524 | 7.5718 | 11.1056 | 10.0499 | 10.7755 | 11.521 | 5.2366 | 5.1493 | 5.2769 | 5.5188 | 6.3339 | 5.4734 | 9.1674 | 9.1538 | 81.42% | 75.68% | 80.92% | 76.64% | 79.86% | 75.82% | 75.03% | 5.3426 | 5.34 | 5.1497 | 5.4843
32 | 8_2_21_UF_6_SP_LSP |  | 06/08/2023 16:32:06 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_UF_6_SP_LSP.386 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 73946.3549 | 66735.5321 | 20.0491 | 21.1067 | 26.6788 | 72722.0693 | 64321.6622 | 102925.7942 | 21.5038 | 19.8781 | 19.0814 | 20.7021 | 33.8219 | 22.1281 | 19.6579 | 19.4269 | 20.0057 | 19.2621 | 20.0408 | 23.9243 | 22.4755 | 22.9167 | 24.685 | 17.6732 | 18.1695 | 18.2202 | 18.7741 | 20.1524 | 18.5208 | 23.4714 | 23.1529 | 75.9% | 69.82% | 69.7% | 69.47% | 73.23% | 68.88% | 69.17% | 17.3447 | 17.5556 | 16.0727 | 17.2485
33 | 8_2_21_UF_8 |  | 06/08/2023 16:37:54 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_UF_8.387 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 11514.0465 | 7057.0144 | 0.13691 | -0.039635 | 9.1591 | 10630.6795 | 2365.1317 | 32571.7492 | 1.365 | 0.17028 | 0.15347 | 1.1273 | 13.7759 | 1.3776 | 0.021529 | 0.1843 | 2.1273 | 2.1268 | 0.93688 | 6.4983 | 5.8167 | 6.3577 | 6.8399 | 0.41872 | 0.062266 | 0.1176 | 0.12602 | 0.53474 | 0.10292 | 2.8528 | 2.9017 | 83.22% | 79.02% | 83.93% | 79.34% | 81.91% | 79.23% | 79.69% | 0.013118 | 0.011266 | 0.039558 | 0.042149
34 | 8_2_21_UF_9 |  | 06/08/2023 16:43:42 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_UF_9.388 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 8448.7078 | 6376.0207 | 0.13432 | 0.1801 | 7.3222 | 8589.4569 | 1837.3845 | 28046.7606 | 0.96084 | 0.088481 | 0.073535 | 0.1373 | 6.7035 | 0.37745 | 0.020649 | 0.1692 | 1.4687 | 1.4832 | 0.76957 | 4.6263 | 3.8894 | 4.5789 | 4.8122 | 0.34548 | 0.068057 | 0.089647 | 0.084482 | 0.32505 | 0.1046 | 1.9726 | 1.9953 | 82.4% | 77.74% | 82.35% | 78.66% | 82.28% | 79.25% | 77.31% | 0.010172 | 0.0089667 | 0.044858 | 0.047201
35 | LBLANK |  | 06/08/2023 16:49:31 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.389 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | -10.0872 | -42.7414 | -0.6818 | -0.77198 | -2.1827 | 321.7818 | 2.5662 | -28.4866 | 0.0045481 | 0.0012304 | 0.0030117 | -0.50857 | -0.025009 | -0.61299 | 0.0020636 | 0.013225 | -0.0081338 | -0.0073653 | -0.16558 | 0.015361 | 0.021366 | 0.18901 | 0.032126 | 0.09901 | -0.00054305 | 0.014658 | 0.0074398 | 0.010049 | 0.012979 | -0.0097971 | -0.011526 | 86.27% | 81.95% | 89.6% | 83.17% | 85.6% | 80.24% | 82.36% | 0.00024118 | -0.00038652 | -0.0023727 | -0.0017448
36 | CCV 1ppb TRACE |  | 06/08/2023 16:55:19 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.390 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 958.2467 | 954.1637 | 1.0925 | 1.0315 | -1.1519 | 1193.1936 | 960.3951 | 1029.761 | 0.99255 | 0.95997 | 1.0347 | 0.69564 | 1.0961 | 0.45499 | 0.99437 | 1.0008 | 1.0031 | 1.0593 | 0.92424 | 0.98373 | 0.86731 | 1.132 | 1.0571 | 1.0665 | 0.98164 | 0.94561 | 1.0226 | 0.9751 | 1.0735 | 1.0733 | 1.1139 | 85.61% | 80.25% | 90.93% | 82.14% | 84.6% | 78.3% | 79.73% | 1.0537 | 1.0673 | 1.0081 | 1.077
37 | CCV 1ppb TRACE |  | 06/08/2023 17:01:37 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.391 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1012.0937 | 1074.9191 | 1.0866 | 0.89436 | -1.1615 | 1573.0823 | 1035.5289 | 1096.9143 | 1.0931 | 1.0022 | 1.0463 | 0.68502 | 1.0739 | 0.60055 | 1.054 | 1.1651 | 1.0659 | 1.0682 | 0.91587 | 0.96788 | 0.96679 | 1.1457 | 1.0544 | 1.0106 | 1.0155 | 1.0507 | 1.103 | 0.99177 | 1.1424 | 1.0737 | 1.0832 | 85.88% | 81.63% | 86.3% | 82.83% | 84.84% | 79.32% | 79.74% | 1.0343 | 1.0425 | 0.99995 | 1.0728
38 | CCV 10 ppb MAJORS |  | 06/08/2023 17:07:25 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.392 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10449.6951 | 10827.3927 | 8.8754 | 9.385 | 8.2383 | 10773.6911 | 10424.1421 | 10695.1868 | 10.6484 | 10.3125 | 10.4227 | 10.1615 | 10.7974 | 10.6119 | 10.7221 | 10.7614 | 10.186 | 10.1361 | 10.2327 | 10.0147 | 9.7206 | 9.9877 | 10.6143 | 9.7898 | 10.0967 | 9.9586 | 10.3682 | 10.7447 | 10.154 | 10.8832 | 10.8953 | 81.72% | 75.72% | 80.3% | 77.33% | 80.42% | 75.71% | 74.68% | 10.0934 | 10.2379 | 9.6027 | 10.3107
39 | CCV 10 ppb MAJORS |  | 06/08/2023 17:13:42 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.393 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10344.4684 | 10497.8461 | 8.8338 | 9.7847 | 8.2239 | 10880.8404 | 10208.6794 | 10646.9174 | 10.4707 | 10.391 | 10.2279 | 10.1228 | 10.7894 | 10.4285 | 10.7659 | 10.7974 | 10.1789 | 9.9797 | 10.0794 | 9.9816 | 9.4301 | 9.5176 | 10.4027 | 9.9573 | 9.9292 | 10.162 | 10.2351 | 10.7463 | 10.1786 | 10.762 | 10.8587 | 83.52% | 77.58% | 81.94% | 78.91% | 81.81% | 76.17% | 78.78% | 9.9789 | 10.0871 | 9.5894 | 10.074
40 | 8_2_21_UF_10 |  | 06/08/2023 17:19:31 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_UF_10.394 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9598.4697 | 9886.1598 | 0.19407 | 0.19832 | 10.4002 | 16615.9048 | 2344.256 | 40010.3505 | 1.0446 | 0.07078 | 0.055195 | 0.2058 | 5.8855 | 0.62071 | 0.025608 | 0.21285 | 1.4485 | 1.3266 | 0.96015 | 5.5178 | 5.256 | 5.2426 | 5.626 | 0.44897 | 0.06676 | 0.1211 | 0.11179 | 0.44487 | 0.12025 | 2.0369 | 2.0181 | 78.69% | 75.06% | 75.54% | 74.57% | 78.54% | 72.26% | 75.88% | 0.007992 | 0.0080177 | 0.062174 | 0.064752
41 | 8_2_21_UF_11 |  | 06/08/2023 17:25:19 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_UF_11.395 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9051.3898 | 10150.4829 | 0.10919 | 0.090119 | 9.0303 | 16970.7616 | 2386.5172 | 41748.6807 | 1.1614 | 0.90429 | 0.85462 | 7.1801 | 6.5406 | 8.2762 | 0.032921 | 0.2131 | 1.2657 | 1.2265 | 0.77614 | 5.4533 | 5.5926 | 5.3587 | 5.7281 | 0.44271 | 0.058287 | 0.095793 | 0.093982 | 0.40543 | 0.078401 | 1.9382 | 1.964 | 77.13% | 73.95% | 74.45% | 74.16% | 77.13% | 72.86% | 74.44% | 0.00627 | 0.0052155 | 0.10511 | 0.11039
42 | 8_2_21_UF_11_DUP |  | 06/08/2023 17:31:09 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_UF_11_DUP.396 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 8241.29 | 9336.6655 | 0.078931 | -0.065985 | 8.828 | 14500.3266 | 2120.1682 | 36333.7762 | 0.97916 | 0.057679 | 0.053634 | -0.045502 | 6.4859 | 0.19076 | 0.021321 | 0.15092 | 1.1536 | 1.2326 | 0.81072 | 5.3027 | 4.4926 | 5.3103 | 5.5546 | 0.39162 | 0.087352 | 0.10057 | 0.098759 | 0.41687 | 0.14238 | 1.9005 | 1.8997 | 78.39% | 73.7% | 79.34% | 74.04% | 78.14% | 72.97% | 72.62% | 0.0051917 | 0.0049684 | 0.10233 | 0.10706
43 | 8_2_21_F_1 |  | 06/08/2023 17:36:57 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_F_1.397 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4792.7967 | 6094.6747 | 2.7795 | 2.8282 | 3.7619 | 14307.4564 | 1236.5063 | 30800.8928 | 0.38222 | 0.05173 | 0.047766 | 6.5388 | 16.9092 | 6.9413 | 0.020599 | 0.29549 | 1.2505 | 1.2564 | 1.9154 | 4.5765 | 3.9524 | 4.8132 | 4.9396 | 0.28602 | 0.034128 | 0.0648 | 0.056557 | 0.32561 | 0.073003 | 5.9789 | 6.0645 | 79.4% | 75.04% | 80.44% | 75.97% | 78.95% | 74.3% | 72.83% | 0.0075556 | 0.0077194 | 0.04365 | 0.045834
44 | 8_2_21_F_6 |  | 06/08/2023 17:44:00 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_F_6.398 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9870.3663 | 7479.4748 | 2.0716 | 1.8968 | 1.9888 | 13679.8545 | 2229.7006 | 37592.376 | 0.98798 | 0.41592 | 0.43471 | 10.283 | 13.7263 | 10.4716 | 0.02997 | 0.27013 | 4.0298 | 4.0867 | 3.4979 | 6.2147 | 5.5282 | 5.5969 | 5.9565 | 0.33265 | 0.11378 | 0.1182 | 0.099046 | 0.51878 | 0.1419 | 3.6802 | 3.4071 | 77.16% | 80.91% | 79.25% | 73.9% | 77.36% | 72% | 80.85% | 0.01971 | 0.017708 | 0.095176 | 0.09544
45 | 8_2_21_F_8 |  | 06/08/2023 17:49:48 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_F_8.399 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 11499.5573 | 7495.2387 | 2.2742 | 2.8496 | 8.6639 | 11775.331 | 2639.3459 | 36017.8876 | 1.4132 | 0.12737 | 0.13059 | 6.1418 | 14.823 | 6.7343 | 0.025819 | 0.21305 | 4.3378 | 4.2282 | 2.2153 | 7.0102 | 6.3684 | 6.6055 | 6.9809 | 0.43361 | 0.065256 | 0.12849 | 0.096311 | 0.60259 | 0.097011 | 3.2096 | 3.1028 | 78.35% | 77.11% | 77.21% | 74.18% | 78.67% | 72.72% | 76.21% | 0.011777 | 0.013197 | 0.09834 | 0.10214
46 | 8_2_21_F_8_DUP |  | 06/08/2023 17:55:36 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_F_8_DUP.400 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 13578.8846 | 8806.7453 | 2.264 | 2.6888 | 8.3294 | 14228.6245 | 2957.4 | 43547.8511 | 1.6198 | 0.062983 | 0.054222 | 5.78 | 15.1571 | 7.5621 | 0.034725 | 0.24438 | 5.1421 | 4.3268 | 2.4091 | 6.9238 | 6.9503 | 6.8477 | 7.2295 | 0.45331 | 0.12318 | 0.10519 | 0.13791 | 0.59814 | 0.14086 | 3.2218 | 3.126 | 79.75% | 74.63% | 69.41% | 73.19% | 78.86% | 74.67% | 73.69% | 0.017314 | 0.01522 | 0.10086 | 0.10677
47 | 8_2_21_F_8_DUP_LD |  | 06/08/2023 18:01:24 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_F_8_DUP_LD.401 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 11009.9034 | 7560.4159 | 4.0612 | 4.4295 | 12.9022 | 11281.9472 | 2477.2675 | 35525.7403 | 1.4023 | 0.08326 | 0.076935 | 13.1077 | 15.4068 | 13.7768 | 0.028096 | 0.23385 | 4.3116 | 4.1097 | 2.5675 | 7.0105 | 5.7353 | 6.6985 | 7.103 | 0.44583 | 0.10884 | 0.11354 | 0.12107 | 0.60585 | 0.12467 | 3.05 | 3.2086 | 77.71% | 75.01% | 80.04% | 77.67% | 78.93% | 63.97% | 74.41% | 0.013827 | 0.012919 | 0.10932 | 0.11128
48 | 8_2_21_F_8_DUP_LD |  | 06/08/2023 18:07:40 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_F_8_DUP_LD.402 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 11456.6737 | 7453.9698 | 4.1634 | 4.643 | 10.5713 | 11715.5566 | 2542.546 | 35658.55 | 1.4098 | 0.089804 | 0.077437 | 13.4773 | 15.5996 | 14.521 | 0.026216 | 0.22741 | 4.4075 | 4.2584 | 2.6589 | 6.951 | 5.5812 | 7.0372 | 7.2644 | 0.3823 | 0.12407 | 0.11039 | 0.11858 | 0.50839 | 0.077863 | 3.0866 | 3.1976 | 79.23% | 74.27% | 78.96% | 76% | 78.71% | 72.6% | 71.65% | 0.014075 | 0.014319 | 0.10668 | 0.11434
49 | 8_2_21_F_8_DUP_LSP |  | 06/08/2023 18:13:28 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_F_8_DUP_LSP.403 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 70322.5868 | 62147.1141 | 18.0702 | 18.9009 | 27.2479 | 69112.0548 | 60451.4375 | 94052.3193 | 15.8954 | 14.0499 | 14.2974 | 21.5694 | 30.6798 | 21.9376 | 13.9673 | 13.88 | 17.0964 | 16.909 | 15.5205 | 20.4712 | 18.7385 | 19.3883 | 21.182 | 13.1075 | 13.9567 | 13.6072 | 13.7101 | 14.8969 | 13.5431 | 17.8698 | 17.8033 | 72.34% | 65.82% | 70.29% | 66.15% | 69.13% | 64.67% | 66.24% | 12.8337 | 12.8337 | 11.7547 | 12.5433
50 | LBLANK |  | 06/08/2023 18:19:16 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.404 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | -6.8622 | -38.834 | -0.69722 | -0.77278 | -1.8744 | 551.2187 | 6.2902 | -30.7443 | 0.0034021 | 0.0031958 | 0.0038855 | -0.51479 | -0.024432 | -0.56598 | 0.0020661 | 0.017722 | -0.0067889 | -0.0060706 | -0.16573 | -0.022463 | 0.0055661 | 0.1642 | 0.0051691 | 0.17826 | -0.021858 | 0.0011072 | -0.00032446 | 0.010958 | 0.009361 | -0.0063377 | -0.0099329 | 87.41% | 81.96% | 87.76% | 86.09% | 86.87% | 79.05% | 85.09% | 0.00022846 | -0.00016505 | -0.0018503 | -0.0016199
51 | CCV 1ppb TRACE |  | 06/08/2023 18:25:05 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.405 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1121.0672 | 1101.9409 | 1.1446 | 0.94174 | -1.0595 | 1850.787 | 1160.7013 | 1166.2788 | 1.1208 | 1.1122 | 1.0651 | 0.72696 | 1.147 | 0.83038 | 1.1214 | 1.1956 | 1.183 | 1.077 | 0.91968 | 1.0707 | 1.0168 | 1.1313 | 1.0815 | 1.0937 | 1.1362 | 1.0544 | 1.1351 | 1.1711 | 1.0063 | 1.1082 | 1.1224 | 86.4% | 80.37% | 84.44% | 83.57% | 85.82% | 76.89% | 81.88% | 1.0277 | 1.0435 | 0.99686 | 1.0715
52 | CCV 1ppb TRACE |  | 06/08/2023 18:31:23 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.406 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1074.0423 | 1035.7773 | 1.2255 | 1.0572 | -1.1292 | 1627.1314 | 1053.138 | 1080.062 | 1.1531 | 1.0691 | 1.0641 | 0.70203 | 1.1187 | 0.78154 | 1.0968 | 1.1118 | 1.1253 | 1.093 | 0.90988 | 1.0676 | 0.83794 | 1.1339 | 1.0942 | 1.0592 | 1.0457 | 1.0296 | 1.0514 | 1.0642 | 0.96887 | 1.0689 | 1.1021 | 86.52% | 80.71% | 87.14% | 83.11% | 85.91% | 79.77% | 82.79% | 1.0265 | 1.0467 | 1.0068 | 1.0562
53 | LFB |  | 06/08/2023 18:37:11 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.407 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10161.1608 | 10441.4137 | 9.3182 | 9.2826 | 11.0197 | 11171.1939 | 10037.7278 | 10717.8098 | 10.3496 | 10.2957 | 10.3642 | 10.4413 | 11.0351 | 10.3389 | 10.7057 | 10.6916 | 10.2054 | 10.2777 | 10.5841 | 10.1772 | 9.3958 | 9.9527 | 10.7524 | 10.1355 | 10.3593 | 10.1254 | 10.6399 | 11.1844 | 10.7475 | 10.7279 | 11.067 | 82.41% | 75.56% | 83.18% | 77.18% | 80.41% | 74.51% | 75.8% | 10.0358 | 10.1841 | 9.3228 | 9.9476
54 | CCV 10 ppb MAJORS |  | 06/08/2023 18:42:59 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.408 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10458.6717 | 10458.0202 | 8.9889 | 9.0553 | 8.7825 | 10989.2516 | 10201.0749 | 10682.2509 | 10.4222 | 10.3563 | 10.7701 | 10.755 | 11.2439 | 10.4457 | 10.6393 | 10.8685 | 10.4075 | 10.5617 | 10.5997 | 10.1685 | 9.2452 | 9.8126 | 10.6672 | 9.9977 | 10.7093 | 10.3419 | 10.5523 | 10.5936 | 10.8071 | 11.173 | 11.1479 | 83.63% | 76.25% | 84.05% | 77.71% | 80.9% | 73.25% | 77% | 10.2947 | 10.3651 | 9.4158 | 10.1256
55 | CCV 10 ppb MAJORS |  | 06/08/2023 18:49:21 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.409 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 11226.1341 | 11171.3778 | 8.0549 | 10.1827 | 8.6149 | 11890.9129 | 11159.7109 | 11518.3926 | 11.3565 | 11.4204 | 10.7401 | 10.5869 | 11.286 | 11.3798 | 11.5004 | 11.9248 | 11.1092 | 10.4923 | 10.4743 | 9.3896 | 10.3233 | 9.9616 | 10.7008 | 8.8304 | 10.6567 | 10.204 | 10.5193 | 9.8819 | 10.6587 | 11.0819 | 11.1202 | 91.58% | 76.22% | 78.94% | 78.44% | 80.23% | 73.93% | 76% | 10.2097 | 10.2815 | 8.5632 | 10.1849
56 | LBLANK |  | 06/08/2023 19:26:31 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.410 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | -12.5329 | -42.7414 | -0.6932 | -0.76964 | -2.2011 | 88.4759 | 5.7847 | -30.7443 | 0.0034369 | 0.0049507 | 0.003945 | -0.5048 | -0.024582 | -0.52786 | 0.0018916 | 0.0051999 | -0.0063505 | -0.0070341 | -0.14708 | 0.080995 | 0.00023996 | 0.16015 | 0.00082783 | 0.18601 | -0.017423 | 0.0054974 | 0.0025319 | 0.064297 | -0.0014165 | -0.009219 | -0.010154 | 85.03% | 77.73% | 82.56% | 81.98% | 85.25% | 77.23% | 84.3% | 0.00024033 | -0.00033309 | -0.0022167 | -0.0018331
57 | PE ICV |  | 06/08/2023 19:32:19 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\PE ICV.411 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10665.6053 | 10833.5347 | 9.3065 | 10.4375 | 20.2114 | 11365.3898 | 10528.3076 | 10918.0171 | 10.7558 | 10.6705 | 10.6362 | 10.865 | 11.4466 | 11.0672 | 10.9086 | 11.0381 | 10.6237 | 10.457 | 10.4061 | 9.9477 | 8.808 | 9.8602 | 10.7122 | 7.0581 | 10.1976 | 10.5756 | 10.7054 | 10.733 | 10.6431 | 10.7202 | 10.777 | 83.75% | 74.18% | 80.37% | 76.29% | 80.84% | 72.14% | 74.96% | 10.2901 | 10.398 | 9.3585 | 10.0806
58 | 8_2_21_F_9 |  | 06/08/2023 19:38:07 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_F_9.412 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 7735.7764 | 6747.8197 | 1.9095 | 2.3241 | 6.9195 | 9781.4137 | 2025.2098 | 31186.209 | 1.0385 | 0.13673 | 0.11855 | 5.3315 | 7.8434 | 5.846 | 0.021986 | 0.19577 | 3.3724 | 3.1887 | 1.5009 | 5.034 | 4.5046 | 4.8741 | 5.2868 | 0.42041 | 0.039221 | 0.078945 | 0.08581 | 0.35005 | 0.085852 | 2.2429 | 2.2707 | 81.13% | 74.33% | 79.28% | 75.2% | 78.6% | 72.59% | 76.21% | 0.012202 | 0.012582 | 0.21189 | 0.22323
59 | 8_2_21_F_10 |  | 06/08/2023 19:43:55 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_F_10.413 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 8124.2163 | 9838.9341 | 3.54 | 3.8248 | 9.6386 | 17244.5966 | 2387.7536 | 41618.5546 | 1.0745 | 0.065264 | 0.055426 | 4.3578 | 7.2034 | 5.2629 | 0.034649 | 0.23057 | 2.6651 | 2.4919 | 2.4205 | 5.6501 | 5.5331 | 5.5438 | 5.9772 | 0.54537 | 0.050007 | 0.084771 | 0.11108 | 0.49156 | 0.12348 | 2.1848 | 2.2009 | 79.7% | 73.11% | 74.76% | 73.77% | 79.08% | 71.89% | 74.13% | 0.0097465 | 0.010742 | 0.13597 | 0.14464
60 | 8_2_21_F_10_DUP |  | 06/08/2023 19:49:43 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_F_10_DUP.414 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 7995.6742 | 9324.7709 | 3.539 | 3.4212 | 9.2696 | 16499.2578 | 2261.4864 | 39124.606 | 0.98698 | 0.26796 | 0.25495 | 5.5447 | 7.3102 | 5.939 | 0.027246 | 0.2237 | 2.5216 | 2.5323 | 2.3194 | 5.7496 | 5.1627 | 5.3895 | 5.9424 | 0.45504 | 0.073821 | 0.10134 | 0.10775 | 0.49512 | 0.13646 | 2.2017 | 2.226 | 78.69% | 72.95% | 77.93% | 73.86% | 78.88% | 72.47% | 75.74% | 0.011615 | 0.0098388 | 0.13669 | 0.14047
61 | 8_2_21_F_11 |  | 06/08/2023 19:55:32 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_F_11.415 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 7542.4887 | 9632.7621 | 3.8746 | 3.9782 | 8.3791 | 16444.6123 | 2247.1013 | 39599.1672 | 1.0704 | 0.08026 | 0.06355 | 4.7243 | 8.0376 | 5.0257 | 0.027594 | 0.22973 | 2.7439 | 2.7966 | 1.5078 | 5.6662 | 4.7211 | 5.4801 | 5.9019 | 0.4609 | 0.12744 | 0.11729 | 0.10513 | 0.45718 | 0.09178 | 2.0434 | 2.0547 | 79.33% | 72.73% | 78.09% | 73.39% | 71.6% | 71.19% | 73.69% | 0.0069367 | 0.0062619 | 0.39355 | 0.44953
62 | 8_2_21_F_11_SP |  | 06/08/2023 20:01:20 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_F_11_SP.416 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 13668.1948 | 15539.8554 | 9.4804 | 10.2143 | 14.5156 | 22293.1503 | 8195.0915 | 45968.7414 | 6.9717 | 5.976 | 6.1323 | 12.2207 | 14.2297 | 12.0601 | 5.7339 | 5.8872 | 8.0042 | 7.9533 | 6.9627 | 11.3835 | 9.9409 | 10.7033 | 11.5608 | 5.5694 | 5.6856 | 5.6063 | 5.5626 | 6.361 | 5.455 | 8.0623 | 8.1128 | 75.99% | 71.74% | 76.7% | 72.21% | 75.83% | 71.41% | 71.83% | 5.5233 | 5.5237 | 5.7947 | 5.9916
63 | 8_2_21_W_1 |  | 06/08/2023 20:07:08 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_W_1.417 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 961.8842 | 1198.8293 | 6.7818 | 8.247 | 1.0777 | 3362.1385 | 278.7989 | 6094.6939 | 0.085682 | 0.036386 | 0.032159 | 7.9725 | 5.1767 | 8.2102 | 0.012373 | 0.098668 | 0.44846 | 0.46944 | 0.13885 | 1.0415 | 0.76934 | 1.1233 | 1.0536 | 0.20236 | 0.010864 | 0.015953 | 0.016704 | 0.036815 | 0.014058 | 1.1677 | 1.2008 | 83.52% | 75.72% | 82.01% | 79.01% | 83.77% | 75.26% | 75.6% | 0.0025601 | 0.0024969 | 0.0357 | 0.035077
64 | 8_2_21_W_6 |  | 06/08/2023 20:12:56 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_W_6.418 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1956.6511 | 1441.6878 | 12.6143 | 14.3995 | 1.4443 | 3147.5132 | 499.7574 | 7569.36 | 0.23612 | 0.054883 | 0.050306 | 15.6104 | 6.2543 | 16.3252 | 0.015591 | 0.04863 | 1.5903 | 1.5115 | 1.009 | 1.3483 | 1.2493 | 1.4302 | 1.3798 | 0.27795 | 0.01459 | 0.024116 | 0.024613 | 0.10433 | 0.036505 | 0.73193 | 0.72366 | 83.19% | 77.23% | 80.86% | 78.56% | 81.83% | 74.98% | 76.63% | 0.007499 | 0.0063291 | 0.12876 | 0.13115
65 | 8_2_21_W_6_DUP |  | 06/08/2023 20:18:44 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_W_6_DUP.419 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1946.0628 | 1473.4633 | 11.1201 | 12.242 | 1.3729 | 3088.0709 | 372.356 | 7293.0718 | 0.21385 | 0.044112 | 0.040668 | 13.4954 | 6.0153 | 14.4264 | 0.011063 | 0.080583 | 1.4585 | 1.4259 | 0.85122 | 1.3073 | 1.038 | 1.3848 | 1.3692 | 0.19814 | -0.012608 | 0.036488 | 0.023081 | 0.079564 | 0.0063549 | 0.71941 | 0.72878 | 82.95% | 75.64% | 80.32% | 78.17% | 81.33% | 74.33% | 77.22% | 0.0074353 | 0.0056539 | 0.1184 | 0.11826
66 | LBLANK |  | 06/08/2023 20:24:33 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.420 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | -13.4076 | -42.7414 | -0.68761 | -0.75313 | -2.0219 | 429.4725 | 9.1767 | -30.7443 | 0.0040659 | 0.0044984 | 0.0028794 | -0.51172 | -0.025153 | -0.53747 | 0.0018928 | -0.0020792 | -0.0090217 | -0.0066424 | -0.16337 | 0.043694 | 0.00034057 | 0.159 | -0.00088352 | 0.15064 | -0.017448 | 0.0012374 | 0.0018458 | 0.016555 | -0.0014165 | -0.011277 | -0.013054 | 83.07% | 78.1% | 83.15% | 79.97% | 83.25% | 76.59% | 78.61% | 0.00024654 | -0.00035722 | -0.0022487 | -0.0019217
67 | CCV 1ppb TRACE |  | 06/08/2023 20:30:21 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.421 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1103.0697 | 1020.7486 | 1.2577 | 1.2835 | -1.1208 | 1470.9034 | 1089.9895 | 1119.9024 | 1.1846 | 1.0814 | 1.0739 | 0.75535 | 1.1138 | 0.73685 | 1.1072 | 1.1212 | 1.1034 | 1.0981 | 0.93956 | 1.1154 | 1.0021 | 1.1325 | 1.0663 | 1.3082 | 0.90907 | 1.0418 | 1.0974 | 1.1586 | 1.0195 | 1.1305 | 1.0976 | 80.52% | 77.33% | 82.86% | 78.91% | 83.14% | 75.74% | 77.4% | 1.0732 | 1.0758 | 1.0851 | 1.0828
68 | CCV 1ppb TRACE |  | 06/08/2023 20:36:39 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.422 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1055.8381 | 1030.8016 | 1.1158 | 1.1625 | -1.1697 | 1574.9501 | 1005.5786 | 1065.5611 | 1.1263 | 1.0592 | 1.0742 | 0.72877 | 1.1246 | 0.81247 | 1.1487 | 1.1242 | 1.1338 | 1.1124 | 0.91885 | 1.0537 | 1.0406 | 1.1137 | 1.086 | 1.1575 | 0.93309 | 0.99768 | 1.0916 | 1.0536 | 1.1759 | 1.0975 | 1.1172 | 84.95% | 77.28% | 84.06% | 80.48% | 82.62% | 75.24% | 80.37% | 1.058 | 1.0696 | 1.0132 | 1.0869
69 | CCV 10 ppb MAJORS |  | 06/08/2023 20:42:27 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.423 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10394.4355 | 10123.1069 | 8.9606 | 10.4524 | 8.5693 | 10636.0102 | 9949.1138 | 10421.0072 | 10.3064 | 10.1256 | 10.7798 | 10.8039 | 11.3846 | 10.0072 | 10.3051 | 10.4742 | 10.0267 | 10.5788 | 10.747 | 10.0081 | 8.8278 | 9.7108 | 10.7809 | 9.6899 | 10.5624 | 10.2631 | 10.5892 | 10.53 | 10.8058 | 11.2381 | 11.3551 | 82.58% | 72.99% | 83.29% | 74.95% | 79.58% | 71.87% | 75.24% | 10.4028 | 10.5105 | 9.7307 | 10.2385
70 | CCV 10 ppb MAJORS |  | 06/08/2023 20:48:44 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.424 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10712.17 | 10386.3941 | 8.9557 | 9.6579 | 8.7643 | 11081.0986 | 10424.5086 | 10876.6374 | 10.7028 | 10.5454 | 10.7862 | 10.679 | 11.3673 | 10.6017 | 11.0031 | 10.9919 | 10.4889 | 10.4813 | 10.6064 | 10.2466 | 9.8342 | 9.9471 | 10.8315 | 10.1333 | 10.4448 | 10.5841 | 10.8201 | 10.7146 | 10.7312 | 11.1392 | 11.438 | 82.38% | 73.71% | 81.22% | 75.78% | 80.19% | 71.8% | 74.5% | 10.2411 | 10.4511 | 9.6987 | 10.0626
71 | 8_2_21_W_8 |  | 06/08/2023 20:54:33 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_W_8.425 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2176.9882 | 1423.0907 | 7.0443 | 7.4512 | 2.8816 | 2663.1656 | 452.5192 | 7023.7502 | 0.28368 | 0.075693 | 0.076502 | 11.0021 | 7.7723 | 10.6713 | 0.011914 | 0.057119 | 1.2964 | 1.3245 | 0.578 | 1.5218 | 1.2487 | 1.5534 | 1.5343 | 0.25968 | 0.00077636 | 0.037115 | 0.030831 | 0.097545 | 0.029071 | 0.63858 | 0.61072 | 83.85% | 77.21% | 83.89% | 75.63% | 81.35% | 77.24% | 77.69% | 0.0049097 | 0.005197 | 0.07127 | 0.075417
72 | 8_2_21_W_9 |  | 06/08/2023 21:00:21 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_W_9.426 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1497.4462 | 1281.5104 | 6.1538 | 6.76 | 2.494 | 2142.3378 | 357.7206 | 6170.1572 | 0.19047 | 0.026515 | 0.029448 | 7.9639 | 4.8736 | 8.1376 | 0.009638 | 0.053089 | 0.83868 | 0.87619 | 0.29957 | 1.0537 | 0.87712 | 1.0987 | 1.0594 | 0.14221 | 0.0054482 | 0.022974 | 0.015819 | 0.0042289 | 0.010149 | 0.43297 | 0.43482 | 85.28% | 76.76% | 83.16% | 77.62% | 82.42% | 74.59% | 79.09% | 0.004109 | 0.0031299 | 0.040506 | 0.044479
73 | 8_2_21_W_9_DUP |  | 06/08/2023 21:06:09 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_W_9_DUP.427 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1478.2644 | 1266.3299 | 5.7879 | 7.3462 | 1.9683 | 2086.8939 | 363.3827 | 5698.61 | 0.19705 | 0.021123 | 0.01551 | 7.0819 | 3.7761 | 7.186 | 0.0074554 | 0.05669 | 0.84133 | 0.82631 | 0.30295 | 1.0958 | 0.8746 | 1.1157 | 1.0425 | 0.36652 | -0.012923 | 0.025362 | 0.017463 | 0.032585 | 0.032311 | 0.43016 | 0.42212 | 84.11% | 78.05% | 83.55% | 79.17% | 82.5% | 75.05% | 77.96% | 0.0032005 | 0.0029198 | 0.039574 | 0.041808
74 | 8_2_21_W_10 |  | 06/08/2023 21:11:57 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_W_10.428 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1513.6007 | 1788.0188 | 21.8921 | 24.7885 | 3.4727 | 3404.979 | 463.9478 | 7556.6481 | 0.21654 | 0.038975 | 0.038709 | 20.3872 | 7.0353 | 21.2392 | 0.019055 | 0.060138 | 1.0792 | 1.0593 | 1.07 | 1.1847 | 0.99959 | 1.2698 | 1.236 | 0.2779 | 0.01064 | 0.017367 | 0.029442 | 0.10327 | 0.014176 | 0.40585 | 0.41931 | 85.45% | 75.63% | 81.43% | 77.61% | 78.2% | 74.45% | 76.88% | 0.0068326 | 0.0056922 | 0.1182 | 0.13465
75 | 8_2_21_W_11 |  | 06/08/2023 21:17:45 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_W_11.429 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1401.6033 | 1801.2425 | 37.3112 | 37.0185 | 4.1537 | 3478.566 | 395.2253 | 7375.8407 | 0.2625 | 0.051438 | 0.054972 | 32.0054 | 11.7222 | 31.78 | 0.022709 | 0.061571 | 1.3499 | 1.401 | 2.0755 | 1.2916 | 0.94269 | 1.2932 | 1.2748 | 0.33352 | -0.0077042 | 0.019186 | 0.027715 | 0.082626 | 0.041586 | 0.39542 | 0.39518 | 83.93% | 75.13% | 81.73% | 76.88% | 81.44% | 74.15% | 77.17% | 0.011556 | 0.0086608 | 0.18973 | 0.19869
76 | 5_10_23_DB |  | 06/08/2023 21:23:33 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_10_23_DB.430 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | -11.218 | -42.7414 | -0.27088 | -0.32272 | 2.5741 | 426.4266 | 1.2341 | -30.7443 | 0.0028168 | 0.0222 | 0.013699 | -0.34176 | -0.0233 | -0.41047 | 0.0018855 | 0.0066639 | 0.035305 | 0.031233 | -0.16475 | 0.037983 | 0.00031299 | 0.15968 | 3.2582e-05 | 0.23609 | -0.017271 | 0.0013219 | 0.00062222 | 0.0048561 | 0.0102 | -0.012152 | -0.013247 | 82.22% | 76.04% | 82.79% | 79.24% | 82.54% | 73.76% | 78.84% | 6.6371e-05 | -0.00077686 | 0.00060337 | 0.0011691
77 | 5_10_23_MFB |  | 06/08/2023 21:29:21 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_10_23_MFB.431 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 5365.9413 | 5422.2039 | 4.5214 | 5.1516 | 10.7006 | 5647.27 | 5228.4258 | 5402.1154 | 5.5095 | 5.5473 | 5.3196 | 5.3852 | 5.6264 | 5.6972 | 5.5512 | 5.4855 | 5.4827 | 5.218 | 5.1295 | 4.877 | 4.6087 | 4.8073 | 5.2235 | 4.9652 | 5.0612 | 5.2307 | 5.2507 | 5.146 | 5.2214 | 5.3179 | 5.4742 | 84.38% | 73.56% | 77.52% | 76.48% | 80.43% | 71.93% | 75.51% | 5.1028 | 5.115 | 4.8536 | 5.247
78 | 5_10_23_T-239 |  | 06/08/2023 21:35:10 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_10_23_T-239.432 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 51632.5505 | 13752.5643 | 117.5695 | 122.9852 | 14.2376 | 1319.9589 | 4101.9059 | 54941.2253 | 2.0979 | 1.51 | 1.4545 | 45.1447 | 4.0369 | 47.3263 | 1.5975 | 1.8349 | 0.81993 | 0.84752 | 6.1572 | 2.4492 | 2.1755 | 2.4346 | 2.5447 | 4.809 | 2.0123 | 2.0016 | 2.071 | 2.3242 | 2.1223 | 1.1548 | 1.1845 | 79.23% | 68.64% | 74.57% | 72.51% | 75.09% | 66.64% | 70.15% | 3.4745 | 3.3905 | 1.2032 | 1.2695
79 | 5_10_23_DB |  | 06/08/2023 21:40:58 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_10_23_DB.433 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | -9.9328 | -42.7414 | -0.32624 | -0.24404 | 3.028 | 465.0854 | 2.7504 | -23.6358 | 0.0046158 | 0.00478 | 0.0019313 | -0.44476 | -0.024751 | -0.49712 | 0.0018717 | 0.005688 | 0.017386 | 0.016541 | -0.17072 | 0.10607 | 0.0062134 | 0.15884 | -0.00081494 | 0.056404 | -0.021858 | -0.00020938 | 0.0030083 | -0.019409 | 0.0021361 | -0.011058 | -0.013027 | 96.34% | 80.96% | 84.07% | 82.88% | 86.37% | 77.22% | 86.13% | 0.00032773 | -0.00079552 | -0.00090674 | 1.8169e-05
80 | 8_16_21_UF_1 |  | 06/08/2023 21:46:46 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_UF_1.434 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10571.0714 | 6203.0055 | 1.2856 | 1.6222 | 24.7266 | 26924.9868 | 2288.3182 | 40780.7317 | 0.49953 | 0.082009 | 0.080994 | 0.17123 | 14.5453 | 0.61986 | 0.033682 | 0.36497 | 0.74356 | 0.74384 | 0.80466 | 6.3777 | 5.7381 | 6.0383 | 6.713 | 0.58587 | 0.031463 | 0.061469 | 0.06928 | 0.41636 | 0.068041 | 19.8514 | 20.2361 | 81.28% | 72.09% | 77.54% | 74.13% | 77.55% | 71.28% | 75.18% | 0.0084327 | 0.0079057 | 0.013877 | 0.013417
81 | LBLANK |  | 06/08/2023 21:52:35 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.435 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | -12.1461 | -42.7414 | -0.68933 | -0.75516 | -1.5152 | 289.9104 | 3.389 | -26.044 | 0.0040317 | 0.0038581 | 0.0037226 | -0.50852 | -0.025251 | -0.57173 | 0.0016624 | 0.0070594 | -0.008451 | -0.0066872 | -0.16251 | 0.057302 | 3.4913e-05 | 0.16914 | 0.010652 | 0.16521 | -0.012874 | 0.0054712 | -0.0001752 | -0.021405 | 0.0098861 | -0.0096086 | -0.010268 | 85.67% | 78.08% | 85.19% | 80.56% | 84.13% | 76.74% | 80.29% | -2.3879e-05 | -0.00067103 | -0.0024194 | -0.0019407
82 | CCV 1ppb TRACE |  | 06/08/2023 21:58:24 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.436 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1083.8139 | 1074.0632 | 1.1396 | 1.6152 | -0.91991 | 1415.5344 | 1073.1178 | 1100.4318 | 1.1266 | 1.1065 | 1.0918 | 0.77772 | 1.1691 | 0.76269 | 1.1289 | 1.1687 | 1.1202 | 1.1199 | 0.93425 | 1.0681 | 0.96265 | 1.1286 | 1.0969 | 1.1809 | 1.0509 | 1.0511 | 1.053 | 1.1198 | 1.1705 | 1.118 | 1.1342 | 85.43% | 76.63% | 83.13% | 79.22% | 83.49% | 75.77% | 79.2% | 1.0708 | 1.065 | 1.0249 | 1.0864
83 | CCV 1ppb TRACE |  | 06/08/2023 22:04:41 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.437 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1069.9421 | 1040.364 | 1.1282 | 1.1078 | -1.1396 | 1451.9629 | 1027.3233 | 1061.703 | 1.1249 | 1.0939 | 1.0874 | 0.73885 | 1.1369 | 0.78214 | 1.1261 | 1.0755 | 1.18 | 1.1194 | 0.95502 | 1.0204 | 0.95972 | 1.1353 | 1.0911 | 1.1983 | 1.1021 | 1.0175 | 1.0921 | 1.0745 | 1.0357 | 1.1005 | 1.1456 | 87.62% | 76.9% | 84.26% | 79.71% | 79.71% | 81.03% | 78.6% | 1.0518 | 1.0477 | 1.0283 | 1.1594
84 | LFB |  | 06/08/2023 22:10:29 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.438 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10540.991 | 10387.9602 | 9.1304 | 10.6155 | 11.2237 | 11092.9045 | 9792.7414 | 10652.0255 | 10.5019 | 10.2952 | 10.547 | 10.6313 | 11.1506 | 10.2824 | 10.7343 | 10.6871 | 10.2373 | 10.3789 | 10.6509 | 10.1541 | 9.6477 | 9.7924 | 10.7244 | 9.98 | 10.4568 | 10.6225 | 10.7832 | 11.0253 | 10.9244 | 11.0404 | 11.11 | 82.63% | 73.1% | 79.36% | 76.15% | 78.63% | 72.02% | 74.32% | 10.1415 | 10.2407 | 9.5039 | 10.0777
85 | CCV 10 ppb MAJORS |  | 06/08/2023 22:16:17 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.439 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10649.4991 | 10628.4209 | 8.8167 | 9.9724 | 9.1292 | 10967.8857 | 10420.3227 | 11135.6009 | 10.6916 | 10.8152 | 11.012 | 10.9275 | 11.4545 | 10.5824 | 11.0306 | 11.03 | 10.424 | 10.6278 | 10.662 | 10.0995 | 9.4107 | 9.7515 | 10.7147 | 10.0855 | 10.2857 | 10.1452 | 10.6641 | 10.4133 | 10.75 | 11.2823 | 11.3646 | 83.51% | 74.06% | 81.09% | 75.51% | 79.25% | 72.62% | 75.56% | 10.2776 | 10.3963 | 9.593 | 10.2827
86 | CCV 10 ppb MAJORS |  | 06/08/2023 22:22:34 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.440 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10400.6221 | 10452.9318 | 8.8292 | 9.3935 | 8.7749 | 10752.0465 | 10239.9338 | 10814.7052 | 10.4035 | 10.4614 | 11.0016 | 10.8312 | 11.4635 | 10.3305 | 10.8005 | 10.5844 | 10.2011 | 10.6095 | 10.6924 | 10.2414 | 9.5224 | 9.7114 | 10.8677 | 10.2623 | 10.7228 | 10.7197 | 10.6732 | 10.8037 | 10.4807 | 11.253 | 11.5059 | 83.26% | 73.61% | 82.67% | 76.59% | 79.12% | 72.31% | 76.52% | 10.2741 | 10.3278 | 9.6288 | 10.2904
87 | 8_16_21_UF_6 |  | 06/08/2023 22:28:23 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_UF_6.441 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 11079.104 | 6645.0905 | 0.42058 | 0.74148 | 13.1253 | 18395.6562 | 2034.406 | 37191.7853 | 0.86095 | 0.060879 | 0.054929 | -0.00851 | 12.7497 | 0.30446 | 0.024697 | 0.75079 | 1.6312 | 1.679 | 2.1443 | 5.7462 | 5.0336 | 5.5346 | 6.0317 | 0.65247 | 0.090114 | 0.081445 | 0.10752 | 0.54473 | 0.076008 | 9.8908 | 10.0246 | 80.98% | 72.1% | 79.33% | 73.46% | 76.4% | 72.25% | 73.91% | 0.015344 | 0.013751 | 0.10338 | 0.11209
88 | 8_16_21_UF_8 |  | 06/08/2023 22:34:11 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_UF_8.442 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 11254.9186 | 7224.9877 | 0.41146 | 0.33161 | 17.9922 | 15457.1844 | 2225.5802 | 35532.3597 | 1.0786 | 0.099927 | 0.084717 | 0.24423 | 13.017 | 0.50941 | 0.022213 | 0.43559 | 1.8104 | 1.8638 | 1.2773 | 5.7752 | 5.4595 | 5.5413 | 5.9833 | 0.5607 | 0.060924 | 0.070799 | 0.082938 | 0.49311 | 0.10492 | 7.8815 | 7.9463 | 80.31% | 71.87% | 76.72% | 72.9% | 75.73% | 72.02% | 72.98% | 0.012273 | 0.010406 | 0.045325 | 0.050077
89 | 8_16_21_UF_9 |  | 06/08/2023 22:39:59 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_UF_9.443 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9423.6986 | 6758.0814 | 0.81516 | 0.93307 | 17.9306 | 13195.6134 | 1737.2278 | 31442.2764 | 0.89715 | 0.062894 | 0.062216 | -0.08536 | 7.1378 | 0.10518 | 0.018465 | 0.79868 | 1.3202 | 1.3772 | 0.5968 | 4.4339 | 3.3166 | 4.3841 | 4.7403 | 0.53314 | 0.051313 | 0.08474 | 0.068754 | 0.34282 | 0.047698 | 5.1546 | 5.35 | 81.13% | 71.85% | 78.27% | 73.29% | 75.67% | 71.38% | 73.63% | 0.006824 | 0.0066198 | 0.026635 | 0.027798
90 | 8_16_21_UF_10 |  | 06/08/2023 22:45:47 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_UF_10.444 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9344.4695 | 9408.4553 | 1.3491 | 1.347 | 19.8839 | 15807.3026 | 2031.4214 | 39179.9577 | 0.93195 | 0.097035 | 0.098869 | 0.48985 | 4.0983 | 0.77223 | 0.017635 | 0.31017 | 1.2177 | 1.2489 | 0.98653 | 4.7718 | 3.9598 | 4.6339 | 4.9976 | 0.56954 | 0.10879 | 0.10099 | 0.090313 | 0.39261 | 0.11255 | 2.5288 | 2.541 | 80.51% | 72.32% | 76.2% | 72.48% | 74.62% | 71.52% | 73.79% | 0.008228 | 0.0076852 | 0.049217 | 0.05518
91 | 8_16_21_UF_11 |  | 06/08/2023 22:51:35 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_UF_11.445 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 8209.4572 | 9033.2918 | 0.24229 | 0.42865 | 17.8392 | 15456.9286 | 2007.2819 | 35032.4647 | 0.88307 | 0.048478 | 0.042249 | -0.22491 | 2.7931 | 0.030921 | 0.012948 | 0.14347 | 0.95225 | 1.0058 | 0.87123 | 4.6843 | 3.9403 | 4.4481 | 4.8207 | 0.57822 | 0.076297 | 0.093077 | 0.084714 | 0.32334 | 0.097566 | 2.0428 | 2.0711 | 78.84% | 71.54% | 78.4% | 72.24% | 74.68% | 70.71% | 73.49% | 0.0031023 | 0.0030534 | 0.056075 | 0.05859
92 | 8_16_21_F_1 |  | 06/08/2023 22:57:23 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_F_1.446 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9353.5218 | 6316.9708 | 15.8961 | 16.1631 | 14.6921 | 25774.455 | 2228.6304 | 39524.5311 | 0.44209 | 0.069238 | 0.065122 | 5.6426 | 16.9444 | 5.6578 | 0.037541 | 0.36186 | 1.0085 | 1.0988 | 1.0366 | 6.6904 | 5.892 | 6.7298 | 7.3683 | 0.62193 | 0.044597 | 0.069379 | 0.077477 | 0.44857 | 0.080519 | 21.3321 | 22.3451 | 79.74% | 68.56% | 77.83% | 71.85% | 75.65% | 66.94% | 70.36% | 0.0082081 | 0.006458 | 0.024809 | 0.02789
93 | 8_16_21_F_1_DUP |  | 06/08/2023 23:03:11 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_F_1_DUP.447 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9489.7858 | 6532.8029 | 1.6941 | 1.9541 | 13.0395 | 26862.9716 | 2258.6811 | 41132.2518 | 0.44466 | 0.073441 | 0.070815 | 5.5444 | 16.048 | 5.6965 | 0.035867 | 0.35031 | 1.155 | 1.1962 | 1.022 | 6.5484 | 5.8171 | 6.4902 | 6.9223 | 0.45405 | 0.041546 | 0.072526 | 0.076258 | 0.37299 | 0.06414 | 21.0106 | 20.9236 | 79.68% | 71.76% | 75.48% | 71.8% | 75.62% | 72.01% | 71.69% | 0.0073884 | 0.006771 | 0.031263 | 0.034907
94 | 8_16_21_F_1_DUP_LD |  | 06/08/2023 23:08:58 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_F_1_DUP_LD.448 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 8935.6098 | 5978.774 | 3.0139 | 2.5672 | 14.828 | 25592.7418 | 2143.3792 | 38856.5913 | 0.40873 | 0.07541 | 0.077194 | 5.6818 | 16.0144 | 5.3947 | 0.041254 | 0.36392 | 1.1173 | 1.2298 | 1.103 | 6.8043 | 5.539 | 6.4043 | 7.0189 | 0.46777 | 0.032424 | 0.053156 | 0.061496 | 0.4353 | 0.081662 | 20.6859 | 21.3088 | 78.88% | 70.91% | 78.83% | 72.16% | 75.25% | 71.24% | 72.79% | 0.0077573 | 0.0077656 | 0.031747 | 0.03639
95 | 8_16_21_F_1_DUP_LD |  | 06/08/2023 23:15:14 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_F_1_DUP_LD.449 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9368.8164 | 6397.8173 | 3.0205 | 2.9167 | 15.0735 | 26206.0947 | 2267.7375 | 40682.0768 | 0.46082 | 0.083791 | 0.077624 | 5.7046 | 16.0094 | 6.0862 | 0.03838 | 0.41682 | 1.1822 | 1.2214 | 1.0836 | 6.6369 | 5.9217 | 6.1786 | 6.8985 | 0.5778 | 0.031894 | 0.069253 | 0.061492 | 0.49295 | 0.10087 | 20.7664 | 21.104 | 79.27% | 71.76% | 76.26% | 71.95% | 76.41% | 70.31% | 75.11% | 0.0098196 | 0.0078055 | 0.035166 | 0.03638
96 | 8_16_21_F_1_DUP_LSP |  | 06/08/2023 23:21:03 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_F_1_DUP_LSP.450 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 68464.0809 | 61915.0589 | 17.5294 | 18.0157 | 33.1685 | 83965.9752 | 59788.2915 | 101798.4975 | 15.5557 | 14.3962 | 14.9793 | 22.302 | 31.7038 | 21.6202 | 14.5808 | 14.407 | 14.3922 | 14.3345 | 14.5715 | 20.3666 | 18.6223 | 19.6715 | 21.4438 | 13.506 | 14.1598 | 13.8613 | 14.1908 | 15.1711 | 14.2148 | 37.0989 | 36.6893 | 72.73% | 63.69% | 67.84% | 64.36% | 74.27% | 62.38% | 63.93% | 12.9117 | 13.0751 | 11.6411 | 11.4061
97 | LBLANK |  | 06/08/2023 23:26:52 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.451 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | -9.9378 | -41.697 | -0.69387 | -0.78747 | -1.8712 | 699.2796 | 7.3142 | -28.3815 | 0.00343 | 0.0018182 | 0.0030687 | -0.50681 | -0.024898 | -0.48754 | 0.0023115 | 0.012101 | -0.0051764 | -0.0076033 | -0.16429 | -0.0042435 | 0.00017371 | 0.1647 | 0.0060638 | 0.49126 | -0.013193 | 0.0038676 | 0.0043455 | 0.1072 | 0.0096492 | -0.00069744 | -0.0066249 | 84.29% | 80.77% | 83.44% | 84.6% | 86.43% | 79.26% | 86.39% | 0.00014763 | -7.1062e-05 | -0.0019687 | -0.0019592
98 | CCV 1ppb TRACE |  | 06/08/2023 23:32:40 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.452 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1075.7356 | 1065.6145 | 1.1826 | 1.3138 | -1.0627 | 1548.7991 | 1080.3294 | 1122.7302 | 1.1198 | 1.0793 | 1.0998 | 0.81134 | 1.1546 | 0.75138 | 1.101 | 1.1429 | 1.1454 | 1.1363 | 0.96015 | 1.0722 | 0.95401 | 1.1202 | 1.0949 | 1.3668 | 1.0533 | 1.0713 | 1.0358 | 1.0991 | 1.1439 | 1.1506 | 1.1329 | 88.12% | 79.46% | 84.84% | 81.42% | 85.03% | 77.07% | 82.75% | 1.0267 | 1.0695 | 1.0094 | 1.0642
99 | CCV 1ppb TRACE |  | 06/08/2023 23:38:58 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.453 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1075.3169 | 1055.3616 | 1.1189 | 1.2052 | -1.1385 | 1719.5689 | 1055.5793 | 1112.1303 | 1.1231 | 1.0601 | 1.0837 | 0.7447 | 1.1293 | 0.81715 | 1.1109 | 1.1305 | 1.1414 | 1.1071 | 0.93478 | 1.0399 | 1.019 | 1.104 | 1.0764 | 1.322 | 0.91512 | 1.0297 | 1.133 | 1.094 | 1.1194 | 1.117 | 1.1111 | 88.22% | 79.87% | 85.61% | 81.74% | 85.38% | 77.36% | 83.14% | 1.0459 | 1.049 | 1.0043 | 1.0623
100 | CCV 10 ppb MAJORS |  | 06/08/2023 23:44:46 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.454 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10522.2205 | 10253.3821 | 8.8086 | 9.7325 | 8.9314 | 10993.611 | 9852.9594 | 10755.8497 | 10.6614 | 10.8242 | 10.9445 | 10.8515 | 11.4439 | 10.5267 | 11.0466 | 11.0528 | 10.5473 | 10.6169 | 10.7119 | 10.0794 | 9.3788 | 9.6111 | 10.6378 | 10.2046 | 10.3166 | 10.3514 | 10.6686 | 10.8022 | 10.4152 | 11.2476 | 11.2625 | 84.9% | 75.42% | 81.01% | 76.37% | 79.11% | 72.69% | 77.56% | 10.2242 | 10.3082 | 9.5753 | 10.3607
101 | CCV 10 ppb MAJORS |  | 06/08/2023 23:51:04 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.455 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10195.537 | 10325.5385 | 8.7585 | 9.3852 | 8.878 | 10646.2834 | 10001.4775 | 10830.1385 | 10.6174 | 10.3862 | 11.0038 | 10.9492 | 11.5407 | 10.3497 | 10.9763 | 10.8296 | 10.279 | 10.5901 | 10.6735 | 10.2518 | 9.8277 | 9.6832 | 10.6885 | 10.3024 | 10.4745 | 10.3285 | 10.5242 | 10.828 | 10.9306 | 11.205 | 11.3154 | 84.93% | 74.76% | 82.61% | 76.88% | 80.24% | 74.11% | 76.62% | 10.1536 | 10.3769 | 9.6745 | 10.1861
102 | 8_16_21_F_6 |  | 06/08/2023 23:56:52 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_F_6.456 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10046.722 | 6960.341 | 6.2659 | 5.388 | 11.3099 | 20296.4111 | 2204.5815 | 38446.2251 | 0.82544 | 0.084315 | 0.074965 | 7.9437 | 14.1443 | 7.9625 | 0.033056 | 0.97451 | 3.4583 | 3.4931 | 5.4122 | 6.563 | 5.2572 | 5.8898 | 6.5463 | 0.7746 | 0.080178 | 0.1091 | 0.098123 | 0.64363 | 0.072227 | 11.0302 | 11.1073 | 78.09% | 72.02% | 78.26% | 72.93% | 76.54% | 72.31% | 75.18% | 0.016381 | 0.018191 | 7.3869 | 7.4684
103 | 8_16_21_F_6_SP |  | 06/09/2023 00:02:41 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_F_6_SP.457 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 16271.5237 | 12881.6732 | 10.3791 | 11.0115 | 19.4842 | 25868.5377 | 7584.8406 | 45633.0892 | 6.8078 | 5.783 | 5.8938 | 13.4726 | 19.6135 | 13.6655 | 5.9408 | 6.7165 | 8.8335 | 8.8356 | 11.0441 | 11.3394 | 10.4647 | 10.4667 | 11.7769 | 5.5791 | 5.5933 | 5.4287 | 5.4188 | 5.9908 | 5.766 | 16.653 | 16.5385 | 81.58% | 71.29% | 74.12% | 71.12% | 75.25% | 69.57% | 74.4% | 5.3931 | 5.4459 | 5.457 | 5.8963
104 | 8_16_21_F_8 |  | 06/09/2023 00:08:29 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_F_8.458 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10845.9529 | 7778.4499 | 2.5567 | 3.1618 | 17.6744 | 18220.9661 | 2295.8651 | 39163.8822 | 1.2316 | 0.074592 | 0.071498 | 7.0254 | 14.6386 | 7.4707 | 0.028631 | 0.56907 | 4.0763 | 3.8964 | 2.2587 | 6.2272 | 5.6937 | 5.8262 | 6.5425 | 0.5391 | 0.036695 | 0.1213 | 0.10541 | 0.54698 | 0.076143 | 8.5618 | 8.7818 | 82.36% | 72.11% | 75.84% | 74.64% | 76.83% | 72.39% | 76.02% | 0.0164 | 0.011852 | 0.14636 | 0.15587
105 | 8_16_21_F_8_DUP |  | 06/09/2023 00:14:18 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_F_8_DUP.459 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10608.858 | 7193.0931 | 2.6162 | 2.9868 | 17.2306 | 17236.1285 | 2205.7298 | 35965.26 | 1.0582 | 0.10975 | 0.098627 | 7.2888 | 14.1392 | 7.4583 | 0.03069 | 0.56016 | 3.8273 | 3.8112 | 2.4528 | 6.1559 | 5.1738 | 5.8211 | 6.4794 | 0.57169 | 0.091361 | 0.099505 | 0.091819 | 0.43962 | 0.11886 | 8.5976 | 8.688 | 81.15% | 71.17% | 75.96% | 72.61% | 77.88% | 70.33% | 74.4% | 0.014038 | 0.01252 | 0.15159 | 0.1558
106 | 8_16_21_F_9 |  | 06/09/2023 00:20:06 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_F_9.460 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 8916.1791 | 6734.3524 | 4.8287 | 5.9473 | 18.5836 | 15603.3094 | 1828.1554 | 33564.631 | 0.95039 | 0.25638 | 0.22948 | 6.394 | 7.9853 | 6.4815 | 0.026678 | 1.2971 | 3.0096 | 2.9994 | 1.3806 | 5.0254 | 4.3433 | 4.6357 | 5.1686 | 0.58311 | 0.09904 | 0.085734 | 0.096962 | 0.36557 | 0.1002 | 5.8187 | 5.8929 | 81.25% | 72.34% | 76.62% | 72.92% | 76.78% | 68.86% | 76.33% | 0.0084765 | 0.0091219 | 0.10136 | 0.10859
107 | 8_16_21_F_10 |  | 06/09/2023 00:25:54 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_F_10.461 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 8086.155 | 9531.6307 | 3.03 | 3.1353 | 17.296 | 17641.1863 | 2198.587 | 40772.3842 | 0.97611 | 0.051629 | 0.06129 | 3.9816 | 4.8606 | 4.355 | 0.022935 | 0.17813 | 2.3325 | 2.3191 | 1.5278 | 5.1293 | 4.2941 | 5.0224 | 5.5139 | 0.59498 | 0.047852 | 0.089891 | 0.12628 | 0.45934 | 0.094513 | 2.7255 | 2.8369 | 81.33% | 70.37% | 75.95% | 72.86% | 75.47% | 70.21% | 72.91% | 0.01189 | 0.0093661 | 0.10068 | 0.10757
108 | 8_16_21_F_10_DUP |  | 06/09/2023 00:31:42 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_F_10_DUP.462 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 7704.4669 | 9102.0476 | 3.0858 | 3.2736 | 16.4146 | 16706.5181 | 2074.1772 | 38514.9004 | 0.90655 | 0.071571 | 0.07208 | 3.6301 | 4.4657 | 4.0366 | 0.020522 | 0.50282 | 2.3572 | 2.2581 | 1.3391 | 5.0827 | 4.1434 | 4.6025 | 5.0523 | 0.61589 | 0.10037 | 0.11553 | 0.082062 | 0.44922 | 0.078478 | 2.5753 | 2.5964 | 81% | 73.9% | 77.04% | 74.87% | 77.69% | 72.62% | 76.75% | 0.0097216 | 0.0093537 | 0.10088 | 0.10362
109 | 8_16_21_F_11 |  | 06/09/2023 00:37:30 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_F_11.463 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 7740.7378 | 9437.2523 | 3.9765 | 4.8839 | 15.7305 | 17814.3672 | 2092.7175 | 38476.5404 | 0.97541 | 0.052122 | 0.062038 | 4.1038 | 3.5619 | 4.3378 | 0.024372 | 0.20454 | 2.2459 | 2.2737 | 1.5921 | 5.1555 | 4.1906 | 4.7951 | 5.3725 | 0.63921 | 0.10686 | 0.11108 | 0.11634 | 0.47189 | 0.11485 | 2.3272 | 2.3276 | 80.32% | 70.76% | 76.67% | 71.7% | 76.04% | 71.62% | 74.89% | 0.0046543 | 0.0046445 | 0.12161 | 0.12237
110 | 8_16_21_W_1 |  | 06/09/2023 00:43:18 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_W_1.464 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1765.1783 | 1154.9932 | 3.3159 | 3.5896 | 6.0955 | 5981.2027 | 398.683 | 7850.108 | 0.084031 | 0.016073 | 0.018885 | 5.7488 | 4.8776 | 4.9137 | 0.013053 | 0.07904 | 0.29607 | 0.35193 | 0.089181 | 1.3938 | 1.1784 | 1.4489 | 1.4348 | 0.33019 | -0.012584 | 0.02188 | 0.008381 | 0.1005 | 0.021963 | 4.6313 | 4.2108 | 84.24% | 75.52% | 83.36% | 70.63% | 80% | 71.32% | 76.77% | 0.0031552 | 0.0016466 | 0.019611 | 0.021345
111 | 8_16_21_W_6 |  | 06/09/2023 00:49:05 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_W_6.465 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2031.2873 | 1397.7537 | 8.4198 | 9.8428 | 1.7467 | 4389.6395 | 429.2917 | 8100.1365 | 0.18767 | 0.029422 | 0.024797 | 11.1882 | 4.9971 | 11.3048 | 0.0097791 | 0.065233 | 1.2216 | 1.1981 | 0.58714 | 1.2939 | 1.209 | 1.3557 | 1.3679 | 0.27836 | 0.015037 | 0.032131 | 0.018883 | 0.17089 | 0.025854 | 2.1625 | 2.1486 | 88.3% | 75.67% | 80.6% | 77.78% | 80.89% | 74.44% | 79.05% | 0.0049499 | 0.0040669 | 0.09995 | 0.1038
112 | LBLANK |  | 06/09/2023 00:54:53 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.466 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | -12.6241 | -42.7414 | -0.70214 | -0.75557 | -1.8183 | 384.997 | 0.44467 | -30.7443 | 0.0028168 | 0.001162 | 0.0027566 | -0.50728 | -0.025033 | -0.58416 | 0.0016626 | -6.3058e-06 | -0.0079475 | -0.0067071 | -0.16464 | 0.037153 | -0.0057419 | 0.15754 | -0.002414 | 0.19771 | -0.021858 | 0.0012851 | -0.00013577 | -0.05994 | 0.010051 | -0.01102 | -0.011667 | 85.85% | 77.29% | 86.75% | 79.27% | 80.4% | 75.99% | 79.6% | 6.7726e-05 | -0.00046338 | -0.0020734 | -0.0018346
113 | CCV 1ppb TRACE |  | 06/09/2023 01:00:42 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.467 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1059.0547 | 1023.5366 | 1.0866 | 1.2312 | -0.90312 | 1616.4827 | 1054.7718 | 1051.8822 | 1.1041 | 1.0617 | 1.0762 | 0.75827 | 1.1412 | 0.78446 | 1.1141 | 1.1303 | 1.1258 | 1.104 | 0.92732 | 0.99623 | 0.95554 | 1.1265 | 1.0652 | 1.2126 | 1.1471 | 1.0674 | 1.1117 | 0.99372 | 1.0214 | 1.1014 | 1.1344 | 88.8% | 74.55% | 81.81% | 79.24% | 82.36% | 72.64% | 75.09% | 1.0313 | 1.0543 | 1.0218 | 1.0961
114 | CCV 1ppb TRACE |  | 06/09/2023 01:06:59 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.468 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1061.3316 | 1049.1345 | 1.0498 | 1.2418 | -1.0603 | 1495.5844 | 993.5035 | 1083.4123 | 1.0719 | 1.0829 | 1.0739 | 0.77724 | 1.1211 | 0.66232 | 1.1061 | 1.1197 | 1.1104 | 1.0858 | 0.90841 | 1.0095 | 0.94692 | 1.0991 | 1.057 | 1.2019 | 1.0128 | 1.0089 | 1.0729 | 1.0827 | 1.1624 | 1.0886 | 1.1345 | 86.13% | 75.69% | 82.02% | 79.8% | 78.62% | 73.63% | 77.78% | 1.0284 | 1.0534 | 1.0357 | 1.1541
115 | LFB |  | 06/09/2023 01:12:47 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.469 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10411.443 | 10176.3145 | 8.551 | 10.5198 | 11.3376 | 11254.3572 | 10032.8819 | 10944.4045 | 10.4819 | 10.5276 | 10.5895 | 10.4133 | 11.249 | 10.4677 | 10.8796 | 10.6369 | 10.3729 | 10.3611 | 10.4729 | 9.8152 | 9.7479 | 9.3534 | 10.5721 | 9.653 | 10.4995 | 10.634 | 10.5591 | 10.5016 | 10.8548 | 10.8869 | 11.0634 | 84.86% | 70.85% | 76.43% | 73.71% | 75.98% | 70.25% | 74.42% | 10.1906 | 10.2388 | 9.4088 | 10.2993
116 | CCV 10 ppb MAJORS |  | 06/09/2023 01:18:35 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.470 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10506.043 | 10123.676 | 8.4551 | 9.7246 | 8.9667 | 10850.9075 | 10202.4914 | 10915.2735 | 10.9189 | 10.5881 | 10.8122 | 10.7201 | 11.4696 | 10.6245 | 11.0236 | 10.9675 | 10.5965 | 10.5167 | 10.5185 | 10.0309 | 10.0877 | 9.4417 | 10.7086 | 10.1537 | 10.663 | 10.3829 | 10.4867 | 10.428 | 10.6185 | 11.2173 | 11.3052 | 84.69% | 72.31% | 77.51% | 74.78% | 77.45% | 72.3% | 76.2% | 10.255 | 10.2996 | 9.7402 | 10.5455
117 | CCV 10 ppb MAJORS |  | 06/09/2023 01:24:53 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.471 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10686.6783 | 10343.3676 | 8.558 | 9.103 | 9.3316 | 11101.8856 | 10156.0171 | 10668.8401 | 10.7468 | 10.7033 | 10.9293 | 11.0091 | 11.5584 | 10.3537 | 11.069 | 11.1126 | 10.583 | 10.6431 | 10.6241 | 10.157 | 9.1867 | 9.286 | 10.754 | 10.1504 | 10.8812 | 10.5997 | 10.8801 | 10.7601 | 10.9323 | 11.2961 | 11.5315 | 84.3% | 71.72% | 78.96% | 75.01% | 77.43% | 70.24% | 77.18% | 10.3244 | 10.4178 | 9.784 | 10.4368
118 | 8_16_21_W_6_DUP |  | 06/09/2023 01:30:41 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_W_6_DUP.472 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1869.1166 | 1291.9295 | 8.5166 | 10.4253 | 1.1172 | 3971.9093 | 358.78 | 7432.3982 | 0.17342 | 0.070123 | 0.069617 | 11.6022 | 4.945 | 11.7512 | 0.012874 | 0.055684 | 1.1641 | 1.1669 | 0.64832 | 1.232 | 1.1248 | 1.3465 | 1.3334 | 0.31151 | -0.017154 | 0.026298 | 0.023408 | 0.064823 | 0.022271 | 2.0557 | 2.1198 | 87.28% | 74.94% | 83.28% | 77.88% | 80.88% | 72.97% | 76.92% | 0.0047591 | 0.004948 | 0.088961 | 0.092624
119 | 8_16_21_W_8 |  | 06/09/2023 01:36:29 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_W_8.473 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1878.916 | 1266.6595 | 8.7202 | 9.4988 | 3.2324 | 3119.5542 | 382.25 | 6749.4131 | 0.21912 | 0.019755 | 0.027086 | 11.2131 | 6.5776 | 10.6969 | 0.012795 | 0.047955 | 1.1116 | 1.1096 | 0.57888 | 1.2847 | 0.85359 | 1.3289 | 1.344 | 0.31885 | 0.025783 | 0.034648 | 0.023985 | 0.16236 | 0.018613 | 1.6847 | 1.7073 | 85.32% | 73.26% | 84.08% | 76.14% | 79.17% | 72.49% | 76.97% | 0.0051431 | 0.004435 | 0.079979 | 0.082931
120 | 8_16_21_W_9 |  | 06/09/2023 01:42:18 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_W_9.474 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1667.4427 | 1295.857 | 9.4626 | 10.3281 | 2.3028 | 3325.6333 | 347.2434 | 6734.4457 | 0.20531 | 0.038542 | 0.041481 | 13.6975 | 5.3736 | 11.8755 | 0.010681 | 0.052368 | 0.90005 | 1.0718 | 0.47365 | 1.0637 | 0.9429 | 1.0893 | 1.0749 | 0.27703 | 0.020186 | 0.019257 | 0.021395 | 0.12087 | 0.018463 | 1.348 | 1.1603 | 84.54% | 74.3% | 79.62% | 67.2% | 79.12% | 72.3% | 78.43% | 0.0054825 | 0.0032199 | 0.067566 | 0.071613
121 | 8_16_21_W_9_DUP |  | 06/09/2023 01:48:06 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_W_9_DUP.475 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1555.2133 | 1211.0954 | 10.4555 | 12.025 | 2.1509 | 2988.3617 | 332.3602 | 6131.1853 | 0.19029 | 0.031869 | 0.026556 | 12.8137 | 5.8979 | 12.7605 | 0.014257 | 0.053359 | 0.92282 | 0.92166 | 0.42129 | 1.0469 | 0.82014 | 1.1092 | 1.0602 | 0.2549 | -0.021858 | 0.031856 | 0.019689 | -0.0013693 | 0.026857 | 1.1181 | 1.1258 | 83.97% | 72.98% | 81.03% | 76.43% | 78.02% | 71.69% | 74.45% | 0.0032262 | 0.0030247 | 0.074924 | 0.078887
122 | 8_16_21_W_10 |  | 06/09/2023 01:53:55 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_W_10.476 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1502.8483 | 1793.4033 | 11.6409 | 13.5315 | 1.9675 | 3485.8097 | 401.2559 | 7672.3893 | 0.19916 | 0.10517 | 0.10773 | 11.8556 | 3.8165 | 12.3045 | 0.0085364 | 0.048524 | 0.68656 | 0.7205 | 0.39574 | 1.0844 | 0.77951 | 1.0742 | 1.0892 | 0.41556 | 0.016727 | 0.012109 | 0.032909 | 0.12215 | 0.022971 | 0.52376 | 0.53315 | 83.97% | 72.83% | 78.44% | 75.76% | 78.4% | 71.66% | 79.22% | 0.00392 | 0.0021605 | 0.072008 | 0.07607
123 | 8_16_21_W_11 |  | 06/09/2023 01:59:43 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_W_11.477 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1432.1919 | 1814.1799 | 5.5915 | 6.5667 | 1.0573 | 3683.2075 | 374.5935 | 7673.3255 | 0.17931 | 0.013845 | 0.016913 | 5.2821 | 1.8849 | 5.3094 | 0.0055511 | 0.039264 | 0.58629 | 0.59229 | 0.22478 | 1.0731 | 0.75443 | 1.0936 | 1.0763 | 0.28319 | 0.0024025 | 0.015336 | 0.028922 | 0.14949 | 0.027272 | 0.45219 | 0.45736 | 83.91% | 71.96% | 78.49% | 75.99% | 74.61% | 71.94% | 75.75% | 0.0019914 | 0.0012061 | 0.041342 | 0.04724
124 | 5_10_23_DB |  | 06/09/2023 02:05:32 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_10_23_DB.478 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | -13.8492 | -42.7414 | -0.35374 | -0.37978 | 2.6292 | 297.698 | -0.52529 | -30.7443 | 0.0028168 | 0.0063856 | 0.0046394 | -0.41847 | -0.024294 | -0.50829 | 0.0018825 | 0.0063744 | 0.013461 | 0.017011 | -0.17112 | 0.021821 | 0.00015562 | 0.15649 | -0.0035484 | 0.22942 | -0.017172 | -0.00010356 | -0.00070609 | -0.17498 | 0.0064924 | -0.012048 | -0.013071 | 85.06% | 74.36% | 81.66% | 78.32% | 78.76% | 71.89% | 79.7% | -2.3879e-05 | -0.00081166 | -0.00041208 | -0.00023984
125 | 5_10_23_MFB |  | 06/09/2023 02:11:19 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_10_23_MFB.479 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4939.388 | 4907.1418 | 4.2675 | 4.8362 | 9.5109 | 5401.6322 | 4689.4509 | 5154.9066 | 5.0984 | 5.1081 | 5.4579 | 5.5878 | 5.667 | 5.0288 | 5.2231 | 5.1693 | 5.0642 | 5.3721 | 5.2747 | 4.8911 | 4.3997 | 4.7282 | 5.214 | 5.1472 | 4.6424 | 5.1383 | 5.206 | 5.0991 | 5.2303 | 5.5275 | 5.4559 | 84.14% | 71.44% | 79.83% | 73.09% | 76.77% | 70.81% | 74.5% | 5.2499 | 5.3153 | 4.9975 | 5.3953
126 | 5_10_23_T-239 |  | 06/09/2023 02:17:07 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_10_23_T-239.480 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 48745.8116 | 12605.4585 | 123.6386 | 110.6863 | 12.9154 | 1174.3707 | 3746.0799 | 54176.4977 | 2.0217 | 1.4535 | 1.4889 | 44.152 | 4.0816 | 45.9021 | 1.5274 | 1.8146 | 0.75729 | 0.8243 | 6.0527 | 2.6704 | 2.0411 | 2.2712 | 2.5167 | 5.3918 | 2.1303 | 1.908 | 2.1157 | 2.5691 | 2.2025 | 1.1549 | 1.1861 | 74.93% | 65.18% | 72.38% | 69.51% | 71.63% | 64.34% | 70.95% | 3.4327 | 3.4631 | 1.3534 | 1.2989
127 | 5_10_23_DB |  | 06/09/2023 02:22:55 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_10_23_DB.481 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | -10.8239 | -41.6648 | -0.040188 | -0.11831 | 3.5216 | 477.3698 | 7.596 | -25.9111 | 0.0052826 | 0.017623 | 0.017607 | -0.26815 | -0.023836 | -0.27262 | 0.0016789 | -0.0028178 | 0.026063 | 0.030542 | -0.17011 | 0.02368 | -0.0057419 | 0.15837 | -0.0014099 | 0.24998 | -0.017364 | -0.0015974 | -0.0007842 | -0.059112 | 0.0062418 | -0.0099051 | -0.012984 | 91.66% | 76.59% | 83.05% | 82.23% | 81.78% | 73.63% | 81.57% | 0.00015425 | -0.0008318 | 0.00018792 | 0.0012048
128 | LBLANK |  | 06/09/2023 02:28:44 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.482 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | -12.9698 | -42.7414 | -0.70369 | -0.80379 | -1.1278 | 455.5397 | 3.1372 | -25.9281 | 0.0034437 | 0.0050085 | 0.002168 | -0.51708 | -0.025263 | -0.53706 | 0.0014541 | -0.00027793 | -0.0085932 | -0.0077548 | -0.1653 | 0.034399 | -0.0057419 | 0.15719 | -0.0027479 | 0.27724 | -0.012592 | 0.0013548 | 0.0027863 | -0.037514 | -0.0014165 | -0.010916 | -0.012314 | 88.57% | 75.15% | 82.06% | 84.53% | 82.5% | 73.8% | 79.73% | -2.3879e-05 | -0.00076703 | -0.0021525 | -0.0019502
129 | CCV 1ppb TRACE |  | 06/09/2023 02:34:32 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.483 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1052.5906 | 989.8255 | 1.0977 | 1.4093 | -0.93359 | 1502.2892 | 910.1179 | 939.8155 | 1.1138 | 1.1145 | 1.0856 | 0.73225 | 1.097 | 0.61284 | 1.0944 | 1.1226 | 1.0962 | 1.1169 | 0.9209 | 1.0637 | 0.9599 | 1.0416 | 1.0683 | 1.3095 | 0.90619 | 1.03 | 1.0523 | 1.1701 | 1.1188 | 1.109 | 1.0846 | 84.34% | 75.66% | 82.42% | 78.23% | 80.5% | 72.19% | 83.72% | 1.0414 | 1.0653 | 1.0831 | 1.1025
130 | CCV 1ppb TRACE |  | 06/09/2023 02:40:49 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.484 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1041.0015 | 957.3049 | 0.98491 | 1.1163 | -1.0897 | 1367.7174 | 970.1029 | 995.2081 | 1.105 | 1.0515 | 1.0771 | 0.74848 | 1.1382 | 0.66903 | 1.1051 | 1.1126 | 1.108 | 1.1062 | 0.93832 | 1.0186 | 0.81931 | 1.0692 | 1.086 | 1.3417 | 1.1282 | 1.0149 | 1.1185 | 1.0496 | 1.1053 | 1.1056 | 1.1531 | 87.33% | 73.82% | 83.14% | 79.11% | 80.11% | 73.23% | 80.6% | 1.0474 | 1.0586 | 1.0206 | 1.0894
131 | CCV 10 ppb MAJORS |  | 06/09/2023 02:46:38 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.485 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10356.5737 | 9983.1131 | 8.0889 | 9.4782 | 8.916 | 10990.5441 | 9940.5629 | 10699.272 | 10.8473 | 10.5968 | 10.6712 | 10.5622 | 11.5428 | 10.2954 | 10.9612 | 11.0111 | 10.5235 | 10.343 | 10.5023 | 9.9541 | 9.7851 | 9.3474 | 10.796 | 10.2314 | 10.7548 | 10.4601 | 10.8769 | 10.3907 | 10.8284 | 11.1769 | 11.5011 | 84.44% | 69.31% | 77.46% | 73.94% | 75.36% | 69.45% | 74.44% | 10.1741 | 10.3961 | 9.7876 | 10.6141
132 | CCV 10 ppb MAJORS |  | 06/09/2023 02:52:55 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.486 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10271.5016 | 9936.3669 | 8.4039 | 9.3585 | 9.4053 | 10912.3798 | 9802.8815 | 11043.2201 | 10.6373 | 10.476 | 10.8694 | 10.6794 | 11.2887 | 10.4824 | 10.9061 | 10.7761 | 10.5068 | 10.5846 | 10.5224 | 10.1115 | 10.0138 | 9.4949 | 10.6146 | 10.4333 | 10.5058 | 10.1291 | 10.3115 | 10.9188 | 10.3008 | 11.1412 | 11.3613 | 83.93% | 71.58% | 77.4% | 74.06% | 76.78% | 66.89% | 74.34% | 10.2073 | 10.2703 | 9.8385 | 10.4226
133 | MT 34 D |  | 06/09/2023 02:58:43 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\MT 34 D.487 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | -10.8277 | -42.7414 | 0.23901 | 0.43595 | 4.8407 | 543.6492 | 1.6478 | -14.5383 | 0.0058732 | 0.01294 | 0.014673 | -0.29073 | -0.022874 | -0.24256 | 0.0031352 | 0.033872 | 0.0080021 | 0.010263 | -0.10061 | 0.013508 | -0.0057419 | 0.15859 | -0.00097193 | 0.29121 | -0.021858 | 0.001239 | -0.0015406 | 0.011605 | -0.0014165 | 0.014307 | 0.01164 | 89.67% | 78.2% | 84.79% | 83.48% | 85.23% | 75.7% | 86.38% | 6.4543e-05 | 0.00021047 | 0.025358 | 0.026957
134 | MT 35 D |  | 06/09/2023 03:04:31 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\MT 35 D.488 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | -13.012 | -42.7414 | 0.43198 | 0.41328 | 5.7319 | 258.3589 | -1.995 | -21.5825 | 0.0034198 | 0.008369 | 0.0095954 | -0.40778 | -0.02306 | -0.45272 | 0.0016593 | 0.015904 | 0.0074941 | 0.010364 | -0.12058 | 0.00090049 | 7.4596e-05 | 0.1579 | -0.0018912 | 0.1822 | -0.017295 | -0.0015974 | -0.00078806 | -0.056265 | 0.014027 | -0.0075421 | -0.012234 | 89.64% | 76.49% | 86.53% | 81.07% | 83.04% | 73.75% | 82.34% | 6.7595e-05 | 5.5457e-05 | 0.026064 | 0.027221
135 | MT 36 D |  | 06/09/2023 03:10:18 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\MT 36 D.489 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | -11.711 | -42.7414 | 0.19463 | -0.058307 | 5.7752 | 520.1581 | 4.5843 | -21.1992 | 0.0028168 | 0.019254 | 0.015047 | -0.41859 | -0.022608 | -0.46609 | 0.0014541 | 0.0049225 | 0.0097397 | 0.012298 | -0.097188 | 0.035949 | -0.0057419 | 0.15833 | -0.0015026 | 0.13825 | -0.017301 | -0.00015644 | 0.00058315 | 0.12591 | 0.0024119 | -0.0071885 | -0.010263 | 89.66% | 76.79% | 83.93% | 80.58% | 83.29% | 73.4% | 80.04% | 0.00033736 | -0.00028153 | 0.024544 | 0.026796
136 | MT 37 D |  | 06/09/2023 03:16:06 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\MT 37 D.490 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | -9.1713 | -41.6987 | 0.37209 | 0.25843 | 6.115 | 313.3414 | -0.073759 | -21.1912 | 0.0028168 | 0.0095054 | 0.0089949 | -0.4562 | -0.023182 | -0.4998 | 0.00231 | 0.0056653 | 0.0052343 | 0.008013 | -0.10296 | 0.038575 | -0.0057419 | 0.15712 | -0.0027621 | 0.22358 | -0.021858 | 0.0013298 | 0.0020793 | -0.050644 | -0.0014165 | -0.0078164 | -0.011081 | 88.62% | 75.73% | 84.74% | 80.68% | 84.47% | 73.19% | 83.17% | 0.00033604 | -0.0003208 | 0.02506 | 0.027014
137 | MT 38 D |  | 06/09/2023 03:21:55 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\MT 38 D.491 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.9695 | -41.7114 | 0.73041 | 1.201 | 6.8098 | 269.772 | 11.3921 | -16.0414 | 0.0040534 | 0.012221 | 0.013329 | 3.6051 | -0.00028103 | 3.844 | 0.0029808 | 0.01975 | 0.094094 | 0.10123 | 0.28238 | 0.025252 | 0.012494 | 0.15797 | -0.0018506 | 0.24786 | -0.021858 | 0.007244 | 0.0006511 | 0.06173 | -0.0014165 | -0.0078569 | -0.0099949 | 88.81% | 75.6% | 81.73% | 79.93% | 82.04% | 69.9% | 79.86% | 0.00034106 | -0.0002362 | 0.025678 | 0.027844
138 | MT 39 D |  | 06/09/2023 03:27:43 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\MT 39 D.492 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | -11.7463 | -42.7414 | 0.27651 | 0.26248 | 6.2095 | 336.2363 | 2.4481 | -23.7593 | 0.0052261 | 0.014992 | 0.01479 | 0.64182 | -0.017731 | 0.36372 | 0.0016667 | 0.003218 | 0.015908 | 0.018678 | -0.10496 | 0.0025863 | -0.0057419 | 0.15708 | -0.0028449 | 0.29666 | -0.021858 | 0.0013109 | -0.0021686 | -0.035541 | -0.0014165 | 0.0034843 | 0.0010705 | 87.7% | 74.73% | 85.5% | 76.68% | 82.31% | 72.3% | 80.9% | 0.00077705 | 0.00028921 | 0.025475 | 0.027688
139 | MT 40 D |  | 06/09/2023 03:33:31 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\MT 40 D.493 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | -12.5533 | -42.7414 | 0.32587 | 0.37595 | 6.6491 | 451.7767 | 0.070073 | -23.5857 | 0.0028168 | 0.008623 | 0.0093602 | -0.40099 | -0.021118 | -0.41493 | 0.0033431 | 0.0086097 | 0.012825 | 0.013266 | -0.10572 | -0.0054084 | -0.0057419 | 0.15751 | -0.002378 | 0.20651 | -0.021858 | -0.00012784 | -6.3475e-05 | -0.032903 | -0.0014165 | -0.0096581 | -0.010228 | 89.42% | 76.09% | 84.63% | 80.47% | 79.06% | 73.59% | 80.9% | 0.00052045 | 2.7819e-05 | 0.027373 | 0.029017
140 | MT 32 D |  | 06/09/2023 03:39:20 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\MT 32 D.494 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | -11.6123 | -42.7414 | 0.66214 | 0.84125 | 5.9762 | 406.3225 | 0.6661 | -18.4341 | 0.004075 | 0.011567 | 0.010195 | -0.3752 | -0.016008 | -0.3927 | 0.0014541 | 0.0049401 | 0.044793 | 0.043742 | -0.097942 | 0.04521 | 0.00032675 | 0.15862 | -0.0010317 | 0.15393 | -0.012499 | 0.0013672 | 1.1729e-05 | -0.036469 | -0.0014165 | -0.0096691 | -0.010303 | 89.35% | 75.31% | 79.95% | 80.86% | 83.34% | 74.13% | 81.4% | 0.00015558 | -0.00013145 | 0.036823 | 0.040688
141 | MT 32 D LD |  | 06/09/2023 03:45:07 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\MT 32 D LD.495 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | -11.6557 | -42.7414 | 0.29154 | 0.27062 | 5.0144 | 402.165 | 5.8515 | -25.8977 | 0.0028168 | 0.016638 | 0.013089 | -0.42477 | -0.022634 | -0.47981 | 0.0016722 | 0.0052687 | 0.033957 | 0.033642 | -0.075663 | -0.0030021 | -0.0057419 | 0.1577 | -0.0021064 | 0.21335 | -0.017034 | -0.00011537 | -0.00070088 | -0.039275 | -0.0014165 | -0.0090254 | -0.0099686 | 87.18% | 74.06% | 82.2% | 81.97% | 77.46% | 83.34% | 80.6% | -2.3879e-05 | 0.00027349 | 0.035452 | 0.040847
142 | MT 32 D LD |  | 06/09/2023 03:51:24 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\MT 32 D LD.496 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | -12.0832 | -41.6832 | 0.26833 | 0.36646 | 6.3501 | 571.4285 | 5.9722 | -25.8863 | 0.0034447 | 0.014069 | 0.012714 | -0.43123 | -0.022927 | -0.49223 | 0.0014541 | 0.0061521 | 0.030312 | 0.033012 | -0.075527 | 0.03637 | -0.0057419 | 0.15795 | -0.0018326 | 0.15998 | -0.01241 | -0.0001179 | 0.0016397 | 0.0083569 | 0.014351 | -0.0092743 | -0.010083 | 86.79% | 72.71% | 81.65% | 79.3% | 79.96% | 70.65% | 78.06% | 0.00053082 | -0.00011403 | 0.035991 | 0.039001
143 | LBLANK |  | 06/09/2023 03:57:12 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.497 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | -12.5767 | -42.7414 | -0.70634 | -0.77097 | -1.131 | 313.9096 | 7.0578 | -28.3955 | 0.0034244 | 0.002147 | 0.0027951 | -0.51161 | -0.025165 | -0.56836 | 0.0014541 | 0.0024533 | -0.0027638 | -0.0070497 | -0.16616 | 0.051529 | -0.0057419 | 0.15721 | -0.0027403 | 0.18241 | -0.012245 | 0.0028582 | -0.00033899 | -0.03785 | -0.0014165 | -0.011642 | -0.013073 | 87.21% | 68.2% | 84.68% | 82.43% | 81.76% | 73.2% | 71.73% | 0.0002271 | -0.00080048 | -0.0021612 | -0.0019083
144 | CCV 1ppb TRACE |  | 06/09/2023 04:03:01 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.498 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1036.5984 | 966.2173 | 0.95107 | 0.86853 | -0.84618 | 1444.5943 | 937.8407 | 992.7871 | 1.112 | 1.0465 | 1.0663 | 0.73551 | 1.1112 | 0.63826 | 1.0987 | 1.1617 | 1.0716 | 1.1164 | 0.92656 | 1.0232 | 0.89782 | 1.0393 | 1.0705 | 1.2681 | 1.2022 | 1.049 | 1.0599 | 1.0012 | 1.1514 | 1.1211 | 1.1208 | 85.82% | 72.58% | 80.86% | 76.85% | 78.52% | 71.26% | 80.68% | 1.0339 | 1.0692 | 1.0381 | 1.1159
145 | CCV 1ppb TRACE |  | 06/09/2023 04:09:18 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.499 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 952.4867 | 920.3942 | 0.89949 | 1.3961 | -1.0523 | 1169.3073 | 894.8116 | 935.8916 | 1.0645 | 1.0183 | 1.0814 | 0.76452 | 1.116 | 0.58453 | 1.0508 | 1.0572 | 1.0662 | 1.1021 | 0.91888 | 0.99898 | 0.84647 | 1.0347 | 1.0648 | 1.1647 | 1.1172 | 1.0883 | 1.105 | 1.0828 | 0.94063 | 1.0975 | 1.14 | 88.6% | 72.51% | 83.52% | 76.55% | 79.08% | 70.34% | 80.6% | 1.0635 | 1.0543 | 0.99776 | 1.1161
146 | LFB |  | 06/09/2023 04:15:06 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.500 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9655.3023 | 9422.9929 | 7.968 | 9.4554 | 11.2017 | 10083.7246 | 9111.2353 | 10266.8773 | 9.9671 | 10.052 | 10.5737 | 10.432 | 11.2055 | 9.8508 | 10.3153 | 10.4705 | 10.2236 | 10.2428 | 10.4287 | 9.8336 | 9.127 | 9.4018 | 10.7556 | 9.6608 | 10.8031 | 10.6016 | 10.9115 | 10.3717 | 10.5244 | 10.869 | 11.2639 | 82.89% | 67.17% | 76.28% | 71.85% | 73.86% | 66.55% | 71.39% | 10.1629 | 10.2121 | 9.5371 | 10.3114
147 | LFB |  | 06/09/2023 04:21:22 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.501 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9656.6532 | 9273.0567 | 8.353 | 9.8519 | 11.0636 | 10003.9804 | 8980.1949 | 10216.0904 | 10.1707 | 9.9523 | 10.5156 | 10.364 | 11.1921 | 9.9731 | 10.4615 | 10.4314 | 10.0178 | 10.2409 | 10.5259 | 10.0852 | 9.556 | 9.1411 | 10.6072 | 10.1785 | 10.7257 | 10.5576 | 10.7993 | 10.7729 | 10.8282 | 10.8881 | 11.1021 | 82.78% | 69.07% | 77% | 72.89% | 74.25% | 69.35% | 74.53% | 10.101 | 10.267 | 9.6012 | 10.2737
148 | CCV 10 ppb MAJORS |  | 06/09/2023 04:27:10 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.502 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9994.5089 | 9616.8679 | 7.8722 | 9.2965 | 8.7872 | 10492.0982 | 9483.7382 | 10467.4739 | 10.5612 | 10.6635 | 11.1277 | 11.1114 | 11.464 | 10.1943 | 10.712 | 10.8583 | 10.2157 | 10.7439 | 10.69 | 9.8211 | 9.5118 | 9.2813 | 10.6774 | 10.0933 | 10.4667 | 10.6094 | 10.6429 | 10.0384 | 10.721 | 11.5107 | 11.4835 | 85.76% | 69.23% | 77.78% | 72.55% | 74.68% | 68.82% | 74.03% | 10.3355 | 10.4697 | 9.6241 | 10.5172
149 | CCV 10 ppb MAJORS |  | 06/09/2023 04:33:27 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.503 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10179.8295 | 9996.6224 | 8.1827 | 9.6363 | 8.8462 | 10714.2111 | 9654.1533 | 10923.3615 | 10.9967 | 10.6108 | 10.1528 | 9.9857 | 11.4208 | 10.6172 | 10.9366 | 10.8637 | 10.4136 | 9.9154 | 9.8559 | 10.0347 | 9.398 | 9.3666 | 10.6098 | 10.4625 | 10.2342 | 10.3869 | 10.5024 | 10.3776 | 10.6661 | 10.6211 | 11.3745 | 84.56% | 70.17% | 77.06% | 78.69% | 77.03% | 68.88% | 73.87% | 9.5247 | 9.7712 | 9.7723 | 10.3593
150 | MT 33 D |  | 06/09/2023 04:39:15 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\MT 33 D.504 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | -9.4496 | -42.7414 | 0.31682 | 0.43273 | 4.979 | 382.9278 | 6.6417 | -28.3349 | 0.0046516 | 0.010729 | 0.0099189 | -0.42763 | -0.021513 | -0.45584 | 0.0023092 | 0.0052866 | 0.039574 | 0.036291 | -0.12176 | 0.033194 | -0.0057419 | 0.15945 | -5.0384e-05 | 0.44776 | -0.021858 | -0.00014835 | -0.00074856 | 0.055243 | 0.0063868 | 0.0074551 | -0.0085949 | 87.27% | 75.8% | 82.51% | 81.02% | 81.4% | 73.34% | 81.8% | 0.00078518 | 0.00036522 | 0.029453 | 0.028925
151 | MT 33 D LSP |  | 06/09/2023 04:45:03 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\MT 33 D LSP.505 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 26044.1296 | 24339.7693 | 24.2298 | 27.2187 | 31.8764 | 25839.3101 | 25094.8771 | 27354.045 | 27.422 | 26.9878 | 28.0209 | 28.7198 | 29.5234 | 26.7482 | 27.2455 | 26.3297 | 25.5064 | 26.0889 | 25.668 | 25.809 | 21.7992 | 22.7301 | 26.1103 | 24.8809 | 25.3539 | 25.4079 | 25.9678 | 26.7067 | 25.5047 | 28.7165 | 29.4031 | 77.97% | 66.55% | 74.66% | 70.16% | 73.05% | 69.62% | 70.38% | 24.8957 | 25.0185 | 25.0866 | 25.2813
152 | cal blank |  | 06/09/2023 04:50:51 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\cal blank.506 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | -11.63 | -42.7414 | -0.72278 | -0.82141 | -1.4135 | 390.0964 | -0.44732 | -30.7443 | 0.0040603 | -0.00030561 | 0.00015366 | -0.52038 | -0.025176 | -0.51838 | 0.0018945 | 0.006169 | -0.0115 | -0.0078486 | -0.17168 | 0.00068128 | 0.00047342 | 0.16085 | 0.0014075 | 0.31405 | -0.012608 | 0.0085479 | -0.0021793 | 0.020526 | 0.0063098 | 0.025833 | -0.0089063 | 86.93% | 76.41% | 81.63% | 79.06% | 79.54% | 72.72% | 79.91% | 0.00043548 | -0.00062049 | -0.0018182 | -0.001494
153 | STD 0.01 |  | 06/09/2023 04:56:40 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 0.01.507 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | -0.88572 | -24.6378 | -0.21189 | -0.13153 | -0.85617 | 281.5926 | 9.01 | -18.3835 | 0.011137 | 0.018818 | 0.011412 | -0.36138 | 0.016297 | -0.36335 | 0.015899 | 0.010412 | 0.012683 | 0.01124 | -0.018936 | 0.028596 | -0.0057419 | 0.16948 | 0.01196 | 0.37887 | -0.012506 | 0.011739 | 0.0071279 | 0.028066 | 0.01034 | 0.032119 | 0.014964 | 87.51% | 74.66% | 80.67% | 78.8% | 80.79% | 70.79% | 81.32% | 0.0083913 | 0.0090535 | 0.0094101 | 0.010552
154 | std 0.05 |  | 06/09/2023 05:02:29 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\std 0.05.508 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 42.3443 | 25.1399 | -0.15424 | 0.041214 | -0.14317 | 318.4276 | 55.8965 | 45.6529 | 0.058316 | 0.064247 | 0.058441 | -0.33816 | 0.033607 | -0.40509 | 0.066578 | 0.06618 | 0.062487 | 0.05813 | -0.056326 | 0.092106 | 0.061691 | 0.20528 | 0.052789 | 0.29763 | 0.029765 | 0.062492 | 0.065505 | 0.084585 | 0.050272 | 0.055842 | 0.044431 | 86.64% | 74.58% | 80.85% | 77.86% | 78.57% | 72.79% | 78.63% | 0.055323 | 0.051647 | 0.049436 | 0.056223
155 | std 0.1 |  | 06/09/2023 05:08:17 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\std 0.1.509 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 90.5481 | 54.4037 | -0.44397 | -0.56456 | -1.1119 | 534.6699 | 91.0056 | 86.9093 | 0.11832 | 0.10389 | 0.10585 | -0.16717 | 0.093735 | -0.15083 | 0.11559 | 0.11577 | 0.12544 | 0.11139 | 0.031115 | 0.13689 | 0.10802 | 0.24482 | 0.10286 | 0.48144 | 0.10407 | 0.097214 | 0.11075 | 0.099131 | 0.10429 | 0.10617 | 0.10104 | 87.7% | 75.13% | 82.86% | 79.61% | 80.97% | 72.96% | 82.72% | 0.10189 | 0.10235 | 0.099535 | 0.1057
156 | std 0.5 |  | 06/09/2023 05:14:05 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\std 0.5.510 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 505.7336 | 488.3675 | 0.42288 | 0.41398 | 0.27733 | 980.7375 | 503.4289 | 646.0592 | 0.52765 | 0.56339 | 0.54525 | 0.63737 | 0.55691 | 0.62369 | 0.57313 | 0.61677 | 0.60285 | 0.57109 | 0.5516 | 0.52234 | 0.42378 | 0.61217 | 0.53526 | 0.73029 | 0.5429 | 0.54155 | 0.52642 | 0.5162 | 0.52131 | 0.55013 | 0.55823 | 87.63% | 74.97% | 81.12% | 79.84% | 79.99% | 72.76% | 80.96% | 0.51651 | 0.51024 | 0.51484 | 0.5379
157 | std 1 |  | 06/09/2023 05:19:54 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\std 1.511 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 993.727 | 948.1153 | 1.0216 | 1.1406 | 0.02795 | 1216.2986 | 921.8704 | 1088.4735 | 1.1009 | 1.0597 | 1.1338 | 1.1861 | 1.1349 | 1.0963 | 1.1082 | 1.1771 | 1.0809 | 1.1469 | 1.0834 | 0.9672 | 0.95137 | 1.0499 | 1.0523 | 1.1403 | 1.0078 | 1.0158 | 1.1041 | 0.92928 | 1.038 | 1.1582 | 1.1215 | 86.74% | 73.35% | 81.35% | 74.85% | 82.21% | 71.78% | 79.16% | 1.059 | 1.0748 | 1.0071 | 1.0193
158 | std 5 |  | 06/09/2023 05:25:42 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\std 5.512 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4918.4856 | 4819.3986 | 4.1051 | 4.7761 | 5.8422 | 5449.0828 | 4829.0861 | 5105.3724 | 5.2605 | 5.2326 | 5.7527 | 5.6957 | 5.6315 | 5.0438 | 5.3103 | 5.4015 | 5.1114 | 5.6779 | 5.7051 | 4.7948 | 4.5134 | 4.7657 | 5.2924 | 5.0166 | 5.1047 | 5.2234 | 5.2969 | 5.0031 | 5.2336 | 5.9163 | 5.5273 | 87.19% | 71.03% | 78.44% | 69.75% | 76.22% | 64.33% | 74.54% | 5.4796 | 5.5619 | 4.7133 | 5.2171
159 | std 10 |  | 06/09/2023 05:31:30 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\std 10.513 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10242.8307 | 9877.3074 | 8.4331 | 9.2918 | 10.2091 | 10556.8158 | 9404.0644 | 10529.487 | 10.6369 | 10.5505 | 10.83 | 10.776 | 11.3204 | 10.5035 | 10.9455 | 10.9796 | 10.4178 | 10.6069 | 10.5295 | 10.0338 | 9.5096 | 9.0234 | 10.4954 | 10.4257 | 10.3561 | 10.0568 | 10.3997 | 10.5136 | 10.327 | 11.2277 | 11.1581 | 82.68% | 69.68% | 75.99% | 72.84% | 74.97% | 67.03% | 75.37% | 10.1995 | 10.3067 | 9.7309 | 10.2976
160 | LBLANK |  | 06/09/2023 05:37:19 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.514 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | -12.9681 | -41.683 | -0.70315 | -0.76901 | -1.8662 | 442.0872 | -1.2013 | -30.7443 | 0.0034565 | 0.0032836 | 0.0039677 | -0.50335 | -0.025041 | -0.51864 | 0.0025447 | 0.0029418 | -0.0078968 | -0.0070702 | -0.1673 | 0.042346 | 0.0066003 | 0.1583 | -0.0013996 | 0.35019 | -0.021858 | -0.00012043 | -0.0014366 | 0.00057204 | -0.0014165 | 0.0058947 | -0.010263 | 85.94% | 74.52% | 81.59% | 79.6% | 77.08% | 71.74% | 80.86% | 0.00016021 | -0.00085873 | -0.0020506 | -0.0018641
161 | CCV 1ppb TRACE |  | 06/09/2023 05:43:07 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.515 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1095.1881 | 1052.2148 | 0.98874 | 1.4873 | -1.0254 | 1597.5271 | 1049.0373 | 1097.1047 | 1.2046 | 1.1554 | 1.0953 | 0.76683 | 1.1213 | 0.97591 | 1.2147 | 1.2524 | 1.1915 | 1.1161 | 0.91116 | 1.0394 | 1.2 | 1.0631 | 1.084 | 1.2623 | 0.98008 | 1.101 | 1.0509 | 1.0332 | 1.0577 | 1.1169 | 1.1325 | 88.39% | 73.37% | 77.01% | 78.45% | 79.2% | 72.65% | 80.37% | 1.0408 | 1.0519 | 0.99979 | 1.0841
162 | CCV 1ppb TRACE |  | 06/09/2023 05:49:25 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.516 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1024.0303 | 931.1209 | 0.98854 | 1.3009 | -1.1378 | 1394.8765 | 938.2511 | 944.6971 | 1.135 | 1.0838 | 1.1303 | 0.79247 | 1.201 | 0.76478 | 1.0956 | 1.0554 | 1.0791 | 1.147 | 0.94372 | 1.0306 | 0.90568 | 1.1293 | 1.1229 | 1.2505 | 1.0986 | 1.116 | 1.1163 | 1.1167 | 1.0692 | 1.1405 | 1.2174 | 86.4% | 70.84% | 81.22% | 75.38% | 78.45% | 71.86% | 74.84% | 1.0386 | 1.0706 | 1.0181 | 1.095
163 | CCV 10 ppb MAJORS |  | 06/09/2023 05:55:13 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.517 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10077.9948 | 9722.7396 | 8.0811 | 9.0427 | 8.8603 | 10644.8836 | 9633.7547 | 10401.7439 | 10.7421 | 10.5474 | 10.7787 | 10.5641 | 11.435 | 10.3795 | 10.753 | 10.4794 | 10.2912 | 10.486 | 10.3246 | 10.0781 | 9.8555 | 9.3768 | 10.6654 | 10.4149 | 9.7581 | 10.3503 | 10.4824 | 10.4773 | 10.3623 | 11.0465 | 11.3729 | 81.81% | 68.72% | 76.71% | 72.71% | 60.91% | 68.12% | 72.7% | 10.0787 | 10.2221 | 9.812 | 13.6664
164 | CCV 10 ppb MAJORS |  | 06/09/2023 06:01:30 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.518 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10127.2038 | 9660.7643 | 8.3293 | 9.0625 | 8.9396 | 10355.8633 | 9514.7639 | 10478.8626 | 10.4444 | 10.3642 | 10.8827 | 10.6084 | 11.6028 | 10.0025 | 10.6399 | 10.6506 | 10.14 | 10.5524 | 10.5135 | 10.1763 | 9.8881 | 9.3709 | 10.7637 | 10.3398 | 10.5144 | 10.6341 | 10.6843 | 10.4701 | 10.8723 | 11.1997 | 11.5358 | 81.89% | 68.26% | 76.6% | 72.22% | 74.51% | 67.44% | 72.86% | 10.1199 | 10.2084 | 9.744 | 10.3248
165 | std 25 |  | 06/09/2023 06:07:19 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\std 25.519 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 25398.1227 | 23629.8633 | 22.0588 | 23.81 | 28.0588 | 26197.1273 | 24586.3166 | 26664.5935 | 26.9583 | 26.2764 | 28.1751 | 28.7602 | 29.2513 | 25.9237 | 27.1549 | 26.2559 | 25.1535 | 26.6469 | 27.3665 | 26.3317 | 23.8956 | 24.1973 | 27.9375 | 25.7403 | 28.4751 | 28.1203 | 27.8609 | 27.8534 | 27.7892 | 29.1201 | 29.5809 | 78.83% | 64.82% | 73.23% | 68.12% | 68.83% | 64.83% | 68.8% | 25.5387 | 25.9901 | 23.8259 | 25.3839
166 | std 50 |  | 06/09/2023 06:13:06 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\std 50.520 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 51651.8561 | 48467.7539 | 55.9609 | 50.8192 | 60.0335 | 53601.7771 | 50084.3154 | 56040.8476 | 56.0659 | 54.2931 | 58.1352 | 54.7715 | 63.5806 | 53.763 | 54.8729 | 52.9413 | 50.4784 | 52.6223 | 54.5789 | 53.1201 | 48.1276 | 52.5513 | 60.13 | 51.8461 | 59.4216 | 59.8176 | 60.1396 | 55.8621 | 60.35 | 59.4817 | 63.8111 | 75.09% | 58.15% | 69.78% | 63.71% | 65.7% | 60.31% | 60.94% | 51.0029 | 51.2617 | 46.2789 | 49.1607
167 | std 50 |  | 06/09/2023 06:19:23 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\std 50.521 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 57514.8786 | 54098.408 | 57.5944 | 56.7123 | 62.1221 | 60725.4447 | 55251.255 | 61417.746 | 61.6791 | 59.963 | 58.3365 | 54.7018 | 60.2968 | 59.2017 | 59.5359 | 57.8086 | 55.5635 | 52.7725 | 54.6745 | 53.5853 | 52.5456 | 48.8933 | 56.3462 | 51.9037 | 55.7449 | 56.5226 | 57.2536 | 57.0671 | 57.2643 | 59.1213 | 60.6688 | 74.73% | 61.18% | 64.44% | 63.91% | 65.02% | 60.17% | 64.65% | 50.2574 | 51.0328 | 46.1231 | 48.9931
168 | std 100 |  | 06/09/2023 06:25:11 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\std 100.522 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 112871.2702 | 104438.8781 | 119.8729 | 108.1934 | 122.6976 | 120006.7696 | 112264.9622 | 122218.2656 | 123.8248 | 118.1572 | 122.1673 | 111.7135 | 125.6833 | 113.8752 | 117.0495 | 111.5168 | 106.2146 | 103.9908 | 108.2579 | 109.2092 | 102.5304 | 99.6412 | 113.5297 | 104.3797 | 117.7598 | 114.707 | 116.6555 | 115.5131 | 116.6196 | 123.3314 | 125.0769 | 67.03% | 55.16% | 61.17% | 56.92% | 59.11% | 54.45% | 57.49% | 100.9788 | 101.576 | 88.5866 | 95.0166
169 | std 100 |  | 06/09/2023 06:31:28 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\std 100.523 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 114581.4942 | 107924.3964 | 119.3748 | 109.6745 | 123.6732 | 124004.7376 | 113937.6549 | 126412.622 | 128.7455 | 122.1822 | 122.1227 | 113.51 | 124.9597 | 118.1089 | 119.4415 | 112.6024 | 109.9544 | 104.9428 | 107.6217 | 107.7039 | 105.1992 | 97.7652 | 113.9056 | 102.8396 | 116.2066 | 114.668 | 116.4887 | 114.8819 | 116.5623 | 123.9658 | 124.9704 | 68.64% | 55.86% | 60.32% | 57.32% | 59.31% | 54.65% | 59.54% | 100.0453 | 100.768 | 88.0225 | 107.195
170 | LBLANK |  | 06/09/2023 06:37:16 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.524 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | -8.5472 | -40.5879 | -0.69949 | -0.7321 | -1.7305 | 632.9079 | 8.8247 | -30.7443 | 0.0067175 | 0.0043399 | 0.0038576 | -0.50047 | -0.023891 | -0.43979 | 0.0063974 | 0.014635 | -0.0041995 | -0.005282 | -0.16664 | 0.0017407 | 0.051217 | 0.1841 | 0.029384 | 0.33492 | -0.0076645 | 0.01045 | 0.0087335 | -0.016551 | -0.0014165 | 0.10219 | 0.0082823 | 87.23% | 73.36% | 79.98% | 77.65% | 79.28% | 72.23% | 80.04% | 0.0014725 | 0.00047471 | -0.0020385 | -0.0012349
171 | PE ICV |  | 06/09/2023 06:43:04 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\PE ICV.525 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10784.0509 | 10806.6822 | 8.6196 | 10.2246 | 21.1985 | 11343.6038 | 10326.1352 | 11332.3659 | 11.0865 | 10.8954 | 11.1494 | 11.1034 | 11.7455 | 10.8905 | 11.1451 | 11.3409 | 10.8236 | 10.5072 | 10.28 | 10.0617 | 8.7484 | 9.1461 | 10.6438 | 7.6279 | 10.4446 | 10.6176 | 10.7366 | 10.6585 | 10.6997 | 10.8202 | 10.8647 | 83.11% | 68.83% | 73.01% | 71.8% | 74.61% | 67.94% | 74.45% | 10.0225 | 10.1664 | 9.4283 | 10.0802
172 | 5ppb CCV |  | 06/09/2023 06:48:52 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5ppb CCV.526 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 5099.1538 | 5060.1377 | 4.5765 | 5.4187 | 5.2908 | 5757.1128 | 4796.6496 | 5266.0258 | 5.3398 | 5.292 | 5.5083 | 5.529 | 5.7169 | 5.0515 | 5.4146 | 5.3646 | 5.2757 | 5.393 | 5.3424 | 5.3082 | 4.6621 | 4.5705 | 5.3399 | 5.6864 | 5.4002 | 5.4184 | 5.4412 | 5.5335 | 5.1843 | 5.5856 | 5.6512 | 81.57% | 70.52% | 78.34% | 73.51% | 74.88% | 68.99% | 77.98% | 4.9902 | 5.081 | 4.95 | 5.089
173 | 25ppb CCV |  | 06/09/2023 06:54:40 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\25ppb CCV.527 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 27052.5786 | 25739.4061 | 22.3597 | 26.7038 | 29.7251 | 28273.4832 | 25766.5741 | 28653.4629 | 28.9027 | 28.0587 | 28.8542 | 29.1463 | 29.5342 | 27.3924 | 28.4125 | 27.9263 | 27.1114 | 26.8724 | 27.6385 | 26.9164 | 25.8499 | 24.3267 | 28.1447 | 26.4734 | 27.7118 | 27.7448 | 28.189 | 28.8846 | 28.655 | 29.4376 | 29.4243 | 78.52% | 64.29% | 70.55% | 66.76% | 68.52% | 63.11% | 68.37% | 25.8446 | 26.068 | 23.658 | 25.2795
174 | 8_30_21_W_1 |  | 06/09/2023 07:00:27 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_W_1.528 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1445.2147 | 1107.9637 | 20.3931 | 22.5768 | 4.6327 | 4780.9248 | 358.5827 | 7056.5057 | 0.080743 | 0.035841 | 0.034346 | 6.7471 | 3.5623 | 6.6475 | 0.011779 | 0.083671 | 0.34181 | 0.33565 | 0.28282 | 0.99561 | 0.90543 | 1.0697 | 1.0591 | 0.40392 | -0.0066747 | 0.014316 | 0.0086444 | 0.022783 | 0.015337 | 3.2238 | 3.3373 | 85.99% | 69.67% | 77.6% | 75.02% | 76.37% | 67.46% | 74.03% | 0.0022111 | 0.001401 | 0.038595 | 0.041109
175 | LBLANK |  | 06/09/2023 07:06:16 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.529 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | -12.9562 | -41.6802 | -0.69969 | -0.75005 | -1.5046 | 495.6731 | 10.0863 | -30.7443 | 0.0041213 | 0.0038698 | 0.0020448 | -0.50349 | -0.024856 | -0.49308 | 0.0016759 | 0.0099529 | -0.0081047 | -0.0072256 | -0.16291 | -0.026293 | 0.00040846 | 0.16055 | 0.0012783 | 0.41103 | -0.016955 | -8.2658e-06 | 0.003197 | -0.074019 | 0.0028055 | 0.032649 | -0.008785 | 80.87% | 70.58% | 79.6% | 75.58% | 77.05% | 69.25% | 76.5% | 0.00036211 | -0.0005511 | -0.0018278 | -0.0018018
176 | CCV 1ppb TRACE |  | 06/09/2023 07:12:04 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.530 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1051.1689 | 1020.0044 | 0.97899 | 1.2627 | -0.93036 | 1309.7646 | 998.1394 | 1088.2421 | 1.155 | 1.0781 | 1.1414 | 0.83223 | 1.0851 | 0.744 | 1.0979 | 1.1357 | 1.1135 | 1.1581 | 0.9574 | 0.96994 | 0.82821 | 1.0186 | 1.0262 | 1.2978 | 1.0821 | 0.99643 | 1.0561 | 1.0974 | 0.92919 | 1.1757 | 1.0745 | 86.32% | 73.62% | 78.62% | 72.31% | 75.2% | 66.22% | 80.06% | 1.0438 | 1.0858 | 0.9751 | 1.0914
177 | CCV 1ppb TRACE |  | 06/09/2023 07:18:22 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.531 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | -12.226 | -42.7414 | -0.73088 | -0.82141 | -2.3965 | -1839.6045 | -6.5656 | -30.7443 | 0.0028168 | 0.00066113 | 0.0012662 | -0.53303 | -0.025781 | -1.2086 | 0.0032262 | -0.0040635 | -0.01051 | -0.0092415 | -0.18035 | 0.5533 | 2.4129e-06 | 0.15579 | -0.0043674 | -1.1983 | -0.017104 | 0.0013599 | -0.0028399 | -0.10368 | -0.0014165 | 0.024278 | -0.013457 | 91.81% | 74.5% | 89.85% | 84.11% | 85.48% | 67.7% | 79.78% | 0.00032415 | -0.00095384 | -0.0025274 | -0.0022858
178 | LFB |  | 06/09/2023 07:24:08 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.532 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10254.4035 | 9776.8176 | 7.8853 | 9.9854 | 11.6517 | 10403.5171 | 9463.4849 | 10132.481 | 10.6011 | 10.5571 | 10.8327 | 10.9165 | 11.4834 | 10.0498 | 10.6216 | 10.4949 | 10.2395 | 10.4629 | 10.6167 | 9.6694 | 9.2012 | 9.2974 | 10.7981 | 9.6129 | 10.6855 | 10.7273 | 10.9476 | 10.3157 | 10.2766 | 10.9803 | 11.2306 | 84.03% | 65.53% | 73.58% | 69.33% | 70.98% | 64.61% | 70.77% | 10.0459 | 10.144 | 8.9354 | 10.0646
179 | LFB |  | 06/09/2023 07:30:24 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.533 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10017.5401 | 9875.7031 | 8.3478 | 9.8424 | 11.7042 | 10429.6148 | 9441.887 | 10498.9096 | 10.5842 | 10.3864 | 11.0048 | 10.8546 | 11.5103 | 10.3028 | 10.6797 | 10.5309 | 10.309 | 10.4596 | 10.6408 | 10.151 | 9.3122 | 9.3152 | 10.8641 | 10.416 | 11.3736 | 10.5987 | 10.7255 | 10.4707 | 11.0097 | 11.1163 | 11.3147 | 80.51% | 65.68% | 73.13% | 68.88% | 71.45% | 62.84% | 71.16% | 9.9849 | 10.0809 | 9.3547 | 9.9893
180 | CCV 10 ppb MAJORS |  | 06/09/2023 07:36:12 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.534 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10782.0523 | 10337.5341 | 7.9763 | 10.0551 | 9.3865 | 11819.8229 | 10201.5805 | 11595.8589 | 11.5809 | 11.3385 | 11.0865 | 10.9656 | 11.8503 | 11.1059 | 11.603 | 11.524 | 11.058 | 10.6293 | 10.6038 | 9.8505 | 10.1951 | 9.3939 | 10.8976 | 10.0099 | 10.7302 | 10.6745 | 10.6631 | 10.3899 | 11.0971 | 11.2673 | 11.6711 | 82.83% | 66.02% | 71.09% | 70.1% | 71.7% | 64.48% | 71.18% | 10.2221 | 10.3652 | 9.41 | 10.2653
181 | CCV 10 ppb MAJORS |  | 06/09/2023 07:42:30 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.535 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10829.0372 | 10386.1016 | 8.1288 | 10.1892 | 9.1513 | 11204.1152 | 10091.0181 | 10646.2779 | 11.0517 | 10.9682 | 11.1525 | 11.1076 | 11.7863 | 10.9493 | 11.3019 | 11.151 | 10.829 | 10.7566 | 10.6965 | 10.0111 | 9.5612 | 9.4705 | 10.8046 | 10.4323 | 10.2157 | 10.3569 | 10.676 | 10.3559 | 10.4126 | 11.4757 | 11.5754 | 81.6% | 66.23% | 72.4% | 69.62% | 67.66% | 65.16% | 70.19% | 10.309 | 10.3199 | 9.5899 | 11.0177
182 | 8_30_21_W_6 |  | 06/09/2023 07:48:18 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_W_6.536 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1713.3194 | 1384.8936 | 8.5795 | 10.302 | 1.9582 | 4070.5121 | 373.4556 | 7446.5251 | 0.16625 | 0.033713 | 0.034457 | 9.0289 | 3.6429 | 9.0657 | 0.0099182 | 0.10284 | 0.96963 | 0.94765 | 0.55169 | 0.96941 | 0.72783 | 1.0821 | 1.0943 | 0.38403 | 0.024257 | 0.029292 | 0.017423 | 0.023127 | 0.054745 | 1.8237 | 1.8439 | 82.96% | 68.19% | 74.83% | 71.99% | 70.48% | 66.45% | 73.88% | 0.0042803 | 0.0038234 | 0.076649 | 0.088074
183 | 8_30_21_W_6_LD |  | 06/09/2023 07:54:05 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_W_6_LD.537 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1629.9337 | 1289.1539 | 11.4479 | 13.1858 | 2.9846 | 3923.8299 | 360.2321 | 7115.009 | 0.15889 | 0.024461 | 0.026272 | 8.9076 | 4.0091 | 8.7079 | 0.0090953 | 0.11234 | 0.91832 | 0.93419 | 0.54928 | 0.98709 | 0.6762 | 1.1936 | 1.1911 | 0.46986 | 0.00087091 | 0.032351 | 0.022948 | 0.19794 | 0.026533 | 1.7945 | 2.0344 | 83.63% | 62.19% | 75.01% | 71.73% | 75.86% | 66.25% | 65.53% | 0.0074877 | 0.005747 | 0.081385 | 0.083486
184 | 8_30_21_W_6_LD |  | 06/09/2023 08:00:20 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_W_6_LD.538 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1730.1363 | 1408.2995 | 11.0373 | 13.895 | 3.3329 | 4336.153 | 392.4985 | 7976.5083 | 0.17034 | 0.027728 | 0.025805 | 8.9557 | 3.6647 | 9.6923 | 0.0098082 | 0.1094 | 1.008 | 0.9406 | 0.56106 | 0.97469 | 0.84022 | 1.0791 | 1.0993 | 0.36121 | 0.0096757 | 0.025031 | 0.022728 | 0.10577 | 0.038369 | 1.8232 | 1.8528 | 80.61% | 66.61% | 70.56% | 69.79% | 63.66% | 64.8% | 72.73% | 0.0056183 | 0.0039239 | 0.083993 | 0.10159
185 | 8_30_21_W_6_DUP |  | 06/09/2023 08:06:08 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_W_6_DUP.539 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1644.1827 | 1291.7634 | 10.7325 | 12.8414 | 3.3864 | 3990.0734 | 362.3843 | 7128.1389 | 0.13476 | 0.02306 | 0.025681 | 9.7028 | 3.694 | 9.6211 | 0.0093403 | 0.096783 | 1.0186 | 1.0411 | 0.56784 | 0.95672 | 0.74399 | 1.0558 | 1.0624 | 0.33356 | -0.006285 | 0.010063 | 0.023626 | 0.13278 | 0.016409 | 1.821 | 1.8633 | 81.82% | 66.57% | 74.88% | 70.18% | 74.24% | 65.38% | 72.11% | 0.0044532 | 0.0041624 | 0.095798 | 0.093852
186 | 8_30_21_W_6 |  | 06/09/2023 08:11:56 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_W_6.540 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1651.8639 | 1367.7319 | 8.1547 | 10.0046 | 2.8575 | 4201.5693 | 369.6612 | 7649.4613 | 0.14564 | 0.030538 | 0.03397 | 8.9265 | 3.6284 | 9.1088 | 0.0085592 | 0.11404 | 0.9474 | 0.92896 | 0.54314 | 0.93433 | 1.0003 | 1.0903 | 1.0954 | 0.26328 | -0.01652 | 0.01178 | 0.015682 | 0.034683 | 0.025238 | 1.8272 | 1.8592 | 83.72% | 66.38% | 70.6% | 68.93% | 71.36% | 68.32% | 71.36% | 0.0054727 | 0.0049851 | 0.07855 | 0.08885
187 | 8_30_21_W_6_DUP_LSP |  | 06/09/2023 08:17:43 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_W_6_DUP_LSP.541 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 24410.2748 | 22811.342 | 29.2755 | 36.2223 | 26.8324 | 28535.2011 | 21352.6907 | 32565.2048 | 23.717 | 22.9373 | 22.4266 | 33.6525 | 27.1012 | 34.6269 | 23.114 | 23.4389 | 22.9904 | 21.925 | 21.8919 | 21.955 | 21.0684 | 19.1227 | 22.7403 | 20.8789 | 22.4305 | 21.4571 | 21.5996 | 21.79 | 22.1116 | 24.9539 | 25.2518 | 76.7% | 59.98% | 62.48% | 63.08% | 64.6% | 58.64% | 65.69% | 20.5686 | 20.6769 | 19.5534 | 20.5707
188 | 8_30_21_W_6_DUP_LSP |  | 06/09/2023 08:24:01 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_W_6_DUP_LSP.542 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 23297.3304 | 22080.5706 | 29.1137 | 34.8303 | 27.1284 | 27411.4871 | 20719.8489 | 31220.0385 | 23.0276 | 22.4004 | 22.4249 | 33.4705 | 27.5932 | 33.3037 | 22.5586 | 22.6137 | 22.2724 | 22.1001 | 22.1271 | 21.7824 | 21.1742 | 19.9588 | 23.1001 | 20.637 | 21.8095 | 21.8018 | 22.021 | 21.6341 | 22.2679 | 24.9207 | 25.629 | 76.04% | 60.61% | 66.26% | 64.13% | 65.73% | 59.9% | 64.6% | 20.3124 | 20.5536 | 18.8463 | 20.3084
189 | 8_30_21_W_8 |  | 06/09/2023 08:29:49 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_W_8.543 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1906.6865 | 1437.7037 | 8.7656 | 11.4516 | 5.2063 | 4353.9732 | 391.7948 | 7581.3015 | 0.17632 | 0.047681 | 0.039718 | 7.9169 | 4.0568 | 8.4035 | 0.011605 | 0.066095 | 1.0286 | 0.98613 | 0.35984 | 1.0637 | 0.95114 | 1.1036 | 1.1419 | 0.33572 | -0.011255 | 0.023672 | 0.023889 | 0.094173 | 0.025466 | 1.6545 | 1.6208 | 81.34% | 65.75% | 70.13% | 67.02% | 71.55% | 52.74% | 72.69% | 0.0046769 | 0.0037236 | 0.061714 | 0.065929
190 | 8_30_21_W_8 |  | 06/09/2023 08:36:07 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_W_8.544 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1732.795 | 1236.7628 | 8.898 | 10.3313 | 3.4164 | 3866.3181 | 359.0262 | 7250.1434 | 0.2007 | 0.039833 | 0.038235 | 7.6022 | 4.0946 | 7.4649 | 0.0077961 | 0.046629 | 0.96687 | 0.95082 | 0.36748 | 1.0399 | 0.89146 | 1.1477 | 1.1739 | 0.36254 | 0.016704 | 0.022729 | 0.017255 | 0.088108 | 0.0078195 | 1.5942 | 1.6539 | 79.07% | 63.78% | 73.13% | 67.84% | 70.28% | 62.36% | 69.25% | 0.0039508 | 0.0043502 | 0.061281 | 0.067404
191 | 8_30_21_W_8_SP |  | 06/09/2023 08:41:55 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_W_8_SP.545 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2886.0127 | 2390.8578 | 9.8296 | 12.4587 | 4.8817 | 4949.3681 | 1431.896 | 8732.8464 | 1.3548 | 1.1881 | 1.1646 | 8.2013 | 5.201 | 8.4444 | 1.1782 | 1.2798 | 2.0526 | 1.9756 | 1.4358 | 2.1199 | 2.015 | 2.1628 | 2.3512 | 1.3605 | 1.1318 | 1.179 | 1.1864 | 1.0825 | 1.2521 | 2.7271 | 2.8648 | 77.51% | 62.51% | 69.44% | 68.22% | 68.52% | 61.24% | 67.01% | 1.0803 | 1.0862 | 1.1258 | 1.2355
192 | LBLANK |  | 06/09/2023 08:47:44 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.546 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | -13.3912 | -42.7414 | -0.70319 | -0.78213 | -1.3829 | 379.4154 | -0.11352 | -30.7443 | 0.0028168 | 0.0030021 | 0.0032259 | -0.50412 | -0.024715 | -0.6075 | 0.0021995 | 0.019936 | -0.0095497 | -0.0076463 | -0.1644 | -0.032562 | 0.0010604 | 0.15838 | -0.0012533 | 0.27624 | -0.016323 | 0.0014387 | 0.0019614 | 0.023643 | -0.0014165 | -0.0080142 | -0.011205 | 81.11% | 66.86% | 74.61% | 74.78% | 72.67% | 61.11% | 73.79% | 0.00026628 | -0.00070797 | -0.0021271 | -0.001969
193 | CCV 1ppb TRACE |  | 06/09/2023 08:53:33 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.547 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1031.8537 | 953.3088 | 1.0689 | 1.3093 | -0.91718 | 1623.6707 | 1033.0956 | 1057.154 | 1.1153 | 1.1272 | 1.1152 | 0.76352 | 1.1597 | 0.63389 | 1.104 | 1.0999 | 1.0772 | 1.1377 | 0.92757 | 1.0033 | 0.96884 | 1.0672 | 1.0893 | 1.3203 | 1.0705 | 1.1046 | 1.14 | 1.0993 | 1.0261 | 1.1307 | 1.1511 | 77.24% | 65.05% | 73.54% | 68.94% | 69.99% | 62.75% | 71.3% | 1.0569 | 1.0482 | 1.0727 | 1.1084
194 | CCV 1ppb TRACE |  | 06/09/2023 08:59:50 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.548 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1079.761 | 984.0317 | 0.97784 | 1.0862 | -1.0429 | 1784.1959 | 987.648 | 1085.102 | 1.1324 | 1.0799 | 1.1164 | 0.80172 | 1.1584 | 0.7542 | 1.0848 | 1.1162 | 1.1407 | 1.134 | 0.94972 | 1.0102 | 0.90142 | 1.0622 | 1.0651 | 1.2202 | 1.0165 | 1.1738 | 1.1152 | 1.1001 | 1.1936 | 1.1413 | 1.1734 | 79.95% | 64.03% | 72.83% | 67.95% | 69.95% | 62.57% | 69.08% | 1.0182 | 1.0627 | 1.0167 | 1.103
195 | CCV 10 ppb MAJORS |  | 06/09/2023 09:05:39 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.549 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10992.7877 | 10545.2027 | 7.9606 | 11.007 | 9.3178 | 11988.486 | 10257.7245 | 11514.3664 | 11.4202 | 11.2655 | 11.0759 | 11.0225 | 11.7624 | 11.1772 | 11.4556 | 11.429 | 11.0224 | 10.6617 | 10.6764 | 10.1097 | 10.0917 | 9.1025 | 10.8117 | 10.5986 | 11.059 | 10.5726 | 10.6964 | 10.5923 | 10.839 | 11.4291 | 11.6291 | 76.5% | 60.91% | 66.04% | 65.35% | 66.32% | 59.03% | 67.24% | 10.2806 | 10.4932 | 9.6686 | 10.4674
196 | CCV 10 ppb MAJORS |  | 06/09/2023 09:11:56 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.550 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10389.8526 | 9967.1841 | 7.8848 | 9.9334 | 9.1764 | 11508.9672 | 9789.0791 | 10913.4873 | 10.8624 | 10.6877 | 11.1521 | 11.1438 | 11.8195 | 10.6551 | 10.862 | 10.8946 | 10.5392 | 10.7941 | 10.7343 | 10.1943 | 9.6135 | 9.5865 | 10.9978 | 10.3312 | 10.5074 | 10.5744 | 10.8123 | 10.8018 | 10.7438 | 11.5463 | 11.7391 | 77.86% | 61.65% | 70.2% | 65.34% | 67.53% | 60.04% | 65.78% | 10.354 | 10.5002 | 9.5143 | 10.3809
197 | 8_30_21_W_9 |  | 06/09/2023 09:17:44 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_W_9.551 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1519.6953 | 1264.3598 | 13.9836 | 17.0116 | 3.0075 | 3130.2946 | 343.3971 | 6979.7717 | 0.18591 | 0.034953 | 0.034332 | 15.494 | 5.5517 | 15.7576 | 0.014522 | 0.044203 | 1.0424 | 1.017 | 0.61058 | 0.86314 | 0.81785 | 0.98912 | 0.97815 | 0.3089 | 0.00011383 | 0.015647 | 0.027173 | 0.071097 | 0.021757 | 1.149 | 1.1669 | 78.85% | 63.81% | 72.88% | 68.37% | 71.61% | 63.73% | 68.72% | 0.0033906 | 0.0038938 | 0.091138 | 0.09389
198 | 8_30_21_W_10 |  | 06/09/2023 09:23:32 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_W_10.552 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1508.8286 | 1728.0829 | 12.4715 | 14.8945 | 4.0738 | 4020.6105 | 388.0208 | 8256.4271 | 0.19598 | 0.026693 | 0.031073 | 11.7787 | 3.588 | 12.3134 | 0.012576 | 0.045773 | 0.80954 | 0.76206 | 0.56606 | 0.91022 | 0.85357 | 1.0244 | 1.0021 | 0.19143 | 0.011741 | 0.030336 | 0.023721 | 0.10261 | 0.022207 | 0.72576 | 0.7619 | 79.54% | 62.38% | 69.42% | 67.23% | 68.5% | 62.1% | 66% | 0.0045009 | 0.0040684 | 0.079107 | 0.083173
199 | 8_30_21_W_11 |  | 06/09/2023 09:29:20 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_W_11.553 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1480.3533 | 1579.9002 | 14.717 | 16.97 | 4.2731 | 3884.0715 | 359.3443 | 7233.5574 | 0.21692 | 0.026761 | 0.027798 | 13.7296 | 4.0971 | 13.2398 | 0.010209 | 0.055881 | 0.779 | 0.82247 | 0.64369 | 0.9376 | 0.71423 | 1.0074 | 1.014 | 0.21525 | -0.010649 | 0.018059 | 0.024725 | 0.090709 | 0.022367 | 0.6933 | 0.69912 | 78.56% | 61.93% | 71.69% | 66% | 66.2% | 60.27% | 67.63% | 0.0046013 | 0.0058572 | 0.24568 | 0.27797
200 | 8_30_21_F_1 |  | 06/09/2023 09:35:08 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_F_1.554 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 7546.9993 | 5893.4038 | 4.1234 | 5.4336 | 8.2326 | 23277.9562 | 1918.9465 | 39001.853 | 0.36695 | 0.13047 | 0.13971 | 3.8402 | 12.858 | 4.2528 | 0.035297 | 0.4696 | 0.91798 | 0.92522 | 5.6603 | 4.7334 | 4.1202 | 4.211 | 5.0723 | 0.58023 | 0.043061 | 0.04526 | 0.051136 | 0.40318 | 0.028399 | 16.5593 | 17.217 | 75.23% | 59.18% | 65.74% | 63.28% | 65.05% | 59.99% | 67.7% | 0.0076911 | 0.006764 | 0.094584 | 0.1036
201 | 8_30_21_F_1 |  | 06/09/2023 09:41:25 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_F_1.555 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 7238.8248 | 5347.5665 | 4.3052 | 4.9601 | 7.5795 | 22524.8113 | 1795.6325 | 36153.42 | 0.36185 | 0.13422 | 0.14649 | 3.9721 | 12.7933 | 3.9015 | 0.034114 | 0.43107 | 0.86434 | 0.94563 | 5.7101 | 4.8098 | 4.1496 | 4.4161 | 5.0594 | 0.4978 | 0.053205 | 0.031363 | 0.044577 | 0.33395 | 0.095837 | 16.8223 | 17.1193 | 74.65% | 60.57% | 69.78% | 62.69% | 64.91% | 60.02% | 65.75% | 0.0057911 | 0.0074352 | 0.097379 | 0.10333
202 | 8_30_21_F_1_DUP |  | 06/09/2023 09:47:13 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_F_1_DUP.556 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 7752.4812 | 5762.7149 | 4.9108 | 5.3855 | 7.619 | 23318.8905 | 1795.3641 | 37666.5611 | 0.38171 | 0.13172 | 0.12302 | 3.8841 | 12.5475 | 4.2022 | 0.03405 | 0.47364 | 0.91571 | 0.94654 | 5.6642 | 5.1486 | 4.0809 | 4.4115 | 5.0008 | 0.72635 | 0.035867 | 0.056912 | 0.04277 | 0.4554 | 0.051839 | 16.6007 | 17.0399 | 69.54% | 60.69% | 67.16% | 63.29% | 66.22% | 59.43% | 65.19% | 0.0075863 | 0.006655 | 0.10496 | 0.10493
203 | 8_30_21_F_1_DUP |  | 06/09/2023 09:53:30 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_F_1_DUP.557 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 7650.2474 | 5679.3873 | 4.371 | 5.7819 | 7.4682 | 22870.8844 | 1842.7719 | 36766.0404 | 0.37471 | 0.12309 | 0.12712 | 3.8458 | 12.7895 | 3.9592 | 0.03163 | 0.45206 | 0.87886 | 0.94584 | 5.6941 | 4.6671 | 4.0617 | 4.4278 | 5.057 | 0.49434 | 0.018815 | 0.042613 | 0.04766 | 0.39072 | 0.067201 | 16.5282 | 17.0674 | 74.41% | 60.17% | 67.53% | 62.93% | 65.38% | 59.07% | 65.11% | 0.0080843 | 0.0059629 | 0.095829 | 0.10651
204 | 8_30_21_F_6 |  | 06/09/2023 09:59:18 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_F_6.558 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 8437.8915 | 6658.332 | 2.5772 | 2.7705 | 0.89808 | 19039.1386 | 1830.7598 | 36241.4473 | 0.72013 | 0.049118 | 0.063106 | 6.3566 | 12.4064 | 6.0683 | 0.028888 | 0.25293 | 2.7501 | 2.9508 | 3.3973 | 4.889 | 4.3966 | 4.629 | 5.2435 | 0.51356 | 0.053797 | 0.10099 | 0.083519 | 0.36134 | 0.13043 | 9.4547 | 9.5542 | 73.67% | 60.41% | 68.62% | 62.3% | 55.02% | 59.39% | 64.83% | 0.02032 | 0.016284 | 0.12611 | 0.16512
205 | 8_30_21_F_6 |  | 06/09/2023 10:05:38 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_F_6.559 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 8363.874 | 6860.6339 | 2.5257 | 2.7755 | 0.78046 | 19766.9928 | 1784.9782 | 38114.3138 | 0.71425 | 0.06332 | 0.064272 | 6.3954 | 12.7851 | 6.2713 | 0.027418 | 0.24759 | 2.9144 | 2.9855 | 3.4426 | 4.8534 | 4.1816 | 4.6165 | 5.2947 | 0.59542 | 0.077346 | 0.10724 | 0.092745 | 0.37676 | 0.10154 | 9.5752 | 9.6967 | 74.5% | 60.06% | 68.36% | 61.39% | 65.33% | 60.46% | 65.28% | 0.020889 | 0.0171 | 0.12263 | 0.13493
206 | 8_30_21_F_8 |  | 06/09/2023 10:11:26 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_F_8.560 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9010.4641 | 6602.8872 | 1.5135 | 1.8563 | 2.452 | 18134.7389 | 1923.611 | 36632.3517 | 0.90541 | 0.068773 | 0.062577 | 6.8386 | 15.1166 | 6.8628 | 0.026937 | 0.2484 | 3.1665 | 3.3355 | 1.6868 | 5.28 | 4.4256 | 5.3196 | 6.1959 | 0.5456 | 0.041655 | 0.099006 | 0.090011 | 0.42065 | 0.090699 | 8.5001 | 9.4844 | 73.74% | 55.67% | 68.63% | 61.82% | 64.27% | 59.12% | 61% | 0.017546 | 0.013855 | 0.10288 | 0.11192
207 | 8_30_21_F_8 |  | 06/09/2023 10:17:43 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_F_8.561 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9078.2803 | 6883.7996 | 1.5155 | 1.2477 | 2.3067 | 18594.5226 | 1935.3788 | 38345.1762 | 0.90178 | 0.067133 | 0.063359 | 6.8157 | 13.4181 | 7.1124 | 0.028799 | 0.22869 | 3.2746 | 3.2554 | 1.6816 | 5.3109 | 4.9143 | 4.8249 | 5.4584 | 0.59737 | 0.087629 | 0.09677 | 0.092591 | 0.45284 | 0.085654 | 8.4679 | 8.398 | 73.75% | 60.95% | 66.15% | 62.27% | 65.57% | 58.39% | 65.08% | 0.015406 | 0.014741 | 0.10661 | 0.11109
208 | LBLANK |  | 06/09/2023 10:23:31 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.562 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | -12.8327 | -42.7414 | -0.70053 | -0.72136 | -2.2825 | 603.3681 | 1.615 | -30.7443 | 0.003544 | 0.001236 | 0.0036236 | -0.49989 | -0.024707 | -0.50489 | 0.0014541 | 0.0085312 | -0.01003 | -0.0063536 | -0.16236 | -0.033331 | -0.0057419 | 0.16006 | 0.00081604 | 0.24527 | -0.021858 | 0.0035984 | 0.0038821 | -0.010501 | 0.0076898 | -0.008223 | -0.009355 | 78.28% | 64% | 71.19% | 64.15% | 68.39% | 61.56% | 70.86% | 7.9106e-05 | -0.00056883 | -0.0024448 | -0.0018272
209 | CCV 1ppb TRACE |  | 06/09/2023 10:29:20 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.563 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1084.9647 | 993.13 | 0.98655 | 1.2115 | -1.1015 | 1898.0804 | 976.3604 | 1229.5396 | 1.1575 | 1.1118 | 1.1371 | 0.81718 | 1.1555 | 0.81617 | 1.1495 | 1.1762 | 1.1615 | 1.1289 | 0.97656 | 0.98685 | 0.99232 | 1.0786 | 1.0923 | 1.1737 | 1.1536 | 1.0609 | 1.1133 | 1.0824 | 1.0107 | 1.1457 | 1.1698 | 78.7% | 63.42% | 70.25% | 67.2% | 69.34% | 60.44% | 68.81% | 1.0658 | 1.0737 | 0.9912 | 1.0887
210 | CCV 1ppb TRACE |  | 06/09/2023 10:35:37 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.564 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1091.3034 | 1077.1134 | 0.98195 | 1.379 | -1.1796 | 1868.8095 | 966.2532 | 1166.725 | 1.1809 | 1.1451 | 1.1464 | 0.81924 | 1.1395 | 0.86577 | 1.1539 | 1.1773 | 1.165 | 1.133 | 0.95846 | 0.95168 | 0.98756 | 1.0578 | 1.086 | 1.247 | 1.0386 | 1.1102 | 1.0989 | 1.1179 | 0.99256 | 1.1434 | 1.1772 | 79.66% | 63.97% | 70.52% | 67.27% | 69.7% | 61.8% | 70.66% | 1.0656 | 1.0671 | 0.9878 | 1.0953
211 | LFB |  | 06/09/2023 10:41:25 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.565 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9832.9683 | 9512.4877 | 7.8929 | 10.3227 | 12.0467 | 10952.736 | 9277.6341 | 10700.7818 | 10.3999 | 10.2808 | 10.8458 | 10.6165 | 11.4097 | 9.9622 | 10.3816 | 10.4985 | 10.1137 | 10.2807 | 10.6458 | 9.8879 | 9.0591 | 8.9786 | 10.8905 | 9.7347 | 10.7941 | 10.8602 | 10.8534 | 10.7314 | 11.0785 | 11.0507 | 11.4922 | 75.43% | 58.71% | 68.28% | 62.56% | 64.34% | 58.13% | 66.21% | 10.05 | 10.2693 | 9.0812 | 10.0593
212 | LFB |  | 06/09/2023 10:47:41 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.566 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9938.389 | 9253.4913 | 8.1182 | 9.945 | 12.5979 | 10929.352 | 9350.0958 | 10403.04 | 10.348 | 10.2445 | 10.8123 | 10.6193 | 11.4597 | 9.8958 | 10.5201 | 10.4879 | 10.1432 | 10.3584 | 10.6537 | 10.2057 | 9.4876 | 9.351 | 10.8188 | 10.1654 | 10.7933 | 10.7603 | 11.044 | 10.8664 | 10.9967 | 11.1718 | 11.446 | 75.3% | 59.84% | 69.47% | 63.35% | 65.19% | 56.98% | 64.4% | 10.1133 | 10.2828 | 9.1207 | 10.1033
213 | CCV 10 ppb MAJORS |  | 06/09/2023 10:53:35 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.567 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10573.7087 | 10097.63 | 7.9675 | 9.7495 | 9.4746 | 11710.2334 | 9917.0711 | 11235.9814 | 11.097 | 10.8697 | 11.2383 | 11.0506 | 12.0634 | 10.9045 | 11.0704 | 10.927 | 10.5331 | 10.6238 | 10.6343 | 10.1742 | 9.6743 | 9.2865 | 11.1364 | 10.4692 | 11.1464 | 10.938 | 11.154 | 10.8519 | 10.7872 | 11.5057 | 11.9297 | 78.2% | 59.96% | 69.85% | 65.14% | 66.94% | 59.47% | 66.8% | 10.2002 | 10.353 | 9.3466 | 10.2881
214 | LBLANK |  | 06/09/2023 11:28:06 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.568 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth
215 | Cal blank |  | 06/09/2023 11:33:55 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\Cal blank.569 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth
216 | CAL BLANK |  | 06/09/2023 11:39:43 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CAL BLANK.570 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | [31 empty] | 100% | 100% | 100% | 100% | 100% | 100% | 100%
217 | LOW 0.01 |  | 06/09/2023 11:45:32 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LOW 0.01.571 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | [31 empty] | 99.73% | 99.78% | 101.4% | 101% | 100.1% | 101% | 99.07%
218 | LOW 0.05 |  | 06/09/2023 11:51:20 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LOW 0.05.572 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | [8 empty] | 0.05 | 0.05 | 0.05 |  | 0.05 |  | 0.05 | 0.05 | 0.05 | 0.05 |  | 0.05 | 0.05 | 0.05 | 0.05 | [6 empty] | 0.05 | 0.05 | 100.4% | 98.87% | 98.01% | 100.3% | 100.1% | 101% | 100.4% | 0.05 | 0.05 | 0.05 | 0.05
219 | LOW 0.1 |  | 06/09/2023 11:57:09 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LOW 0.1.573 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | [8 empty] | 0.087104 | 0.10707 | 0.099403 |  | 0.12409 |  | 0.09896 | 0.091533 | 0.098328 | 0.098929 |  | 0.084465 | 0.083518 | 0.098642 | 0.096277 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.10618 | 0.11419 | 96.34% | 97.39% | 98.29% | 98.56% | 98.37% | 96.57% | 100.6% | 0.10143 | 0.099011 | 0.10268 | 0.097276
220 | LOW 0.5 |  | 06/09/2023 12:02:57 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LOW 0.5.574 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth |  |  |  |  | 0.5 |  |  |  | 0.49413 | 0.47486 | 0.48139 | 0.5 | 0.65836 | 0.5 | 0.48102 | 0.47385 | 0.50812 | 0.51569 | 0.5 | 0.37164 | 0.56696 | 0.50036 | 0.49609 | 0.36485 | 0.41213 | 0.5072 | 0.50701 | 0.46012 | 0.56548 | 0.55952 | 0.57619 | 97.53% | 98.18% | 99.78% | 98.22% | 84.07% | 97.44% | 100.8% | 0.49547 | 0.49443 | 0.50215 | 0.55518
221 | STD 1 |  | 06/09/2023 12:08:45 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 1.575 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1000 | 1000 | 1 | 1 | 0.55522 |  | 1000 | 1000 | 1.0279 | 1.0008 | 0.99475 | 0.91334 | 1.2405 | 0.88241 | 0.99038 | 0.92553 | 0.99065 | 0.98827 | 0.96184 | 0.81413 | 1.0616 | 1.0095 | 1.0216 | 0.77729 | 0.85579 | 0.9918 | 0.99198 | 0.97996 | 1.0996 | 1.0939 | 1.1454 | 98.97% | 95.61% | 98.46% | 96.89% | 97.01% | 95.76% | 96.52% | 1.0051 | 0.98941 | 0.98213 | 0.95279
222 | STD 5 |  | 06/09/2023 12:14:33 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 5.576 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4856.1427 | 4918.7394 | 4.7459 | 4.8287 | 3.53 | 5000 | 5077.9857 | 5012.7903 | 4.9915 | 4.8182 | 4.9916 | 3.8132 | 5.9735 | 3.7391 | 4.7624 | 4.7094 | 4.8572 | 4.7822 | 4.8327 | 4.228 | 4.9347 | 5.0035 | 5.1086 | 4.7663 | 4.4998 | 4.8919 | 4.9788 | 5.1468 | 5.2046 | 5.3354 | 5.4776 | 95.92% | 92.23% | 95.64% | 92.68% | 92.41% | 94.3% | 92.07% | 4.9252 | 4.9921 | 4.8438 | 4.675
223 | STD 10 |  | 06/09/2023 12:20:21 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 10.577 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10101.1051 | 10091.8913 | 9.7017 | 9.4484 | 6.9961 | 10349.7893 | 10447.8693 | 10249.2277 | 10.3488 | 10.0408 | 10.1648 | 8.3378 | 11.6951 | 8.0378 | 10.0127 | 10.1035 | 9.8306 | 9.7489 | 9.8438 | 8.8053 | 10.1776 | 10.0077 | 10.2766 | 9.5276 | 9.6872 | 9.763 | 9.8211 | 10.0907 | 10.3835 | 10.8496 | 10.9932 | 92.95% | 90.12% | 91.28% | 90.49% | 89.93% | 91.84% | 90.63% | 9.8264 | 10.0065 | 9.6466 | 9.4106
224 | STD 25 |  | 06/09/2023 12:26:09 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 25.578 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 25070.0782 | 24991.2149 | 27.3334 | 23.5337 | 19.5029 | 26155.6288 | 25647.6427 | 25044.7638 | 26.1792 | 25.2582 | 25.8833 | 22.2877 | 28.7751 | 21.3861 | 25.1612 | 24.6289 | 24.4845 | 24.5423 | 25.3559 | 25.6451 | 27.0333 | 25.1241 | 26.6622 | 27.4735 | 24.4891 | 26.0348 | 25.6789 | 28.557 | 27.3585 | 27.3519 | 27.7447 | 82.57% | 86.35% | 89.71% | 86.35% | 86.35% | 88.01% | 88.08% | 25.1122 | 25.2793 | 26.0154 | 22.9519
225 | STD 50 |  | 06/09/2023 12:31:57 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 50.579 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 51529.8733 | 50332.3498 | 59.2987 | 49.0069 | 41.4208 | 52095.888 | 52273.9954 | 51171.8421 | 53.8646 | 51.6358 | 52.4368 | 42.433 | 57.0428 | 45.2206 | 51.0892 | 50.3715 | 49.6855 | 48.2667 | 49.7924 | 45.9809 | 53.5519 | 50.614 | 52.7338 | 48.0343 | 50.456 | 51.9228 | 51.5557 | 50.7775 | 53.0673 | 54.2528 | 55.2244 | 87.36% | 83.91% | 84.54% | 82.75% | 76.52% | 83.74% | 80.6% | 48.8322 | 48.9903 | 44.7382 | 47.6419
226 | STD 100 |  | 06/09/2023 12:37:44 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 100.580 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 108498.323 | 105277.2281 | 121.9483 | 102.2174 | 84.9644 | 109393.8023 | 111350.031 | 107409.5804 | 117.1948 | 112.5209 | 113.0715 | 90.6799 | 120.057 | 95.0887 | 108.7277 | 104.4821 | 102.5525 | 98.3322 | 100.4032 | 95.2398 | 111.6938 | 100.136 | 107.2627 | 100.3927 | 104.0236 | 107.5178 | 106.6472 | 106.6085 | 109.4809 | 115.3563 | 114.7526 | 80.09% | 76.37% | 77.75% | 73.38% | 75.62% | 78.48% | 75.21% | 98.2079 | 98.9217 | 86.6002 | 84.4994
227 | Calibration Curves |  | 06/09/2023 12:37:44
228 | LBLANK |  | 06/09/2023 12:43:32 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.581 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 28.8843 | 13.3084 | -0.54623 | -0.71016 | -0.48328 | 469.8215 | 33.7033 | 65.8758 | 0.0062431 | 3.8361e-05 | -0.00024657 | -0.19936 | -0.02705 | -0.046952 | 0.0037233 | 0.0082987 | -0.0047314 | -0.0066264 | -0.13116 | -0.05553 | 0.023883 | 0.26669 | 0.0063026 | 0.29045 | 0.001725 | 0.008059 | 0.0085619 | 0.042051 | -0.029558 | 0.088502 | -0.0059357 | 93.48% | 100.1% | 100.8% | 102.9% | 102.2% | 105.6% | 104.3% | 0.0015747 | -0.0003043 | -0.0029503 | -0.0012646
229 | LBLANK |  | 06/09/2023 13:13:17 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.583 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 33.848 | 12.1963 | -0.63327 | -0.58786 | -0.67693 | -305.5873 | 35.2432 | 65.8758 | 0.0039552 | -0.00098948 | -0.00052821 | -0.51706 | 0.0012718 | -0.53421 | 0.0019316 | 0.0042903 | -0.0068075 | -0.0065981 | 0.032302 | -0.024094 | 0.024445 | 0.26285 | 0.0018791 | 0.18619 | -0.042869 | 0.0084746 | 0.0029837 | 0.030816 | -0.041654 | 0.049407 | 0.0070623 | 94.94% | 95.75% | 97.84% | 94.91% | 97.26% | 94.72% | 99.05% | 0.0011205 | -0.0010036 | -0.0033749 | -0.0018772
230 | LBLANK |  | 06/09/2023 13:19:06 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.584 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 28.0894 | 12.1963 | -0.72088 | -0.7574 | -0.84198 | 423.1132 | 35.3658 | 65.8758 | 0.0019891 | -0.0060389 | -0.0016394 | -0.53294 | 0.0010633 | -0.50355 | 0.0014738 | 0.0041766 | -0.009049 | -0.0076148 | -0.14069 | -0.0026811 | 0.010141 | 0.26237 | 0.0013302 | 0.044187 | -0.053078 | 0.0085013 | 0.0059206 | -0.0062396 | -0.041654 | 0.044702 | 0.003428 | 99.61% | 95.68% | 97.3% | 97.23% | 98.89% | 84.65% | 97.67% | 0.0010166 | -0.0011082 | -0.0033389 | -0.0020406
231 | ICV PE |  | 06/09/2023 13:24:53 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\ICV PE.585 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10402.4606 | 10790.0138 | 8.6359 | 9.6861 | 18.1854 | 10365.3137 | 10205.5983 | 10292.2368 | 10.733 | 10.5487 | 9.9666 | 10.139 | 10.057 | 10.548 | 10.6406 | 10.687 | 10.49 | 9.6931 | 9.4901 | 9.6037 | 9.5931 | 9.4309 | 10.1099 | 6.8086 | 10.1598 | 9.9229 | 9.7291 | 9.6263 | 10.1502 | 9.5231 | 9.4949 | 96.09% | 91.25% | 87.83% | 90.87% | 91.34% | 91.23% | 92.26% | 9.6305 | 9.7447 | 9.216 | 9.4524
232 | ICV PE |  | 06/09/2023 13:31:10 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\ICV PE.586 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10122.2283 | 10855.4755 | 8.9054 | 9.4452 | 18.1516 | 9683.2897 | 9953.6138 | 10338.8355 | 10.4088 | 10.3483 | 11.5413 | 11.7797 | 10.0963 | 10.3264 | 10.2743 | 10.5963 | 10.4382 | 11.2558 | 11.1597 | 9.7487 | 9.4015 | 9.4857 | 10.1739 | 7.0021 | 10.1279 | 9.7158 | 9.8288 | 9.6421 | 10.1201 | 10.9642 | 9.5224 | 94.95% | 90.59% | 90.59% | 82.69% | 94.66% | 91.91% | 91.64% | 11.1516 | 11.2557 | 9.3969 | 9.2748
233 | CCV 1ppb TRACE |  | 06/09/2023 13:36:58 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.587 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1012.0115 | 1031.7646 | 1.1072 | 1.1547 | 0.21933 | 1080.0053 | 1042.3484 | 1008.9618 | 1.0311 | 0.98852 | 1.0257 | 0.64356 | 1.0076 | 0.62252 | 0.97966 | 0.98319 | 1.0464 | 1.0593 | 0.89225 | 0.92828 | 1.0605 | 1.1684 | 1.015 | 1.1284 | 0.98134 | 0.92326 | 0.99144 | 0.90329 | 1.0864 | 1.0278 | 0.99404 | 99.4% | 96.96% | 102.2% | 96.39% | 96.78% | 96.31% | 99.47% | 1.01 | 1.0126 | 0.98687 | 1.0111
234 | CCV 1ppb TRACE |  | 06/09/2023 13:43:15 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.588 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1043.7191 | 1063.8988 | 1.0537 | 1.2045 | 0.072191 | 1221.5046 | 1066.8359 | 1057.3282 | 1.079 | 1.0358 | 1.0503 | 0.67741 | 1.0355 | 0.78205 | 1.0177 | 1.0107 | 1.0884 | 1.0503 | 0.88945 | 1.1188 | 1.1159 | 1.1659 | 1.0081 | 1.0799 | 0.97475 | 1.0336 | 1.0393 | 1.0669 | 0.95948 | 1.0445 | 1.0203 | 97.93% | 94.74% | 98.44% | 94.27% | 95.5% | 95.27% | 96.78% | 1.0321 | 1.027 | 0.96723 | 0.99754
235 | LFB |  | 06/09/2023 13:49:03 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.589 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9526.0582 | 9724.8269 | 8.4123 | 8.6757 | 10.5802 | 9700.8945 | 9346.9596 | 9322.9533 | 9.6882 | 9.7567 | 9.794 | 9.6263 | 9.7864 | 9.3537 | 9.5731 | 9.715 | 9.8361 | 9.7629 | 9.8083 | 9.8132 | 10.5522 | 9.1845 | 10.0564 | 9.7069 | 10.5715 | 9.8014 | 9.8874 | 9.8269 | 9.9559 | 9.7918 | 9.6215 | 94.3% | 89.18% | 92.98% | 89.97% | 90.47% | 92.42% | 92.17% | 9.6619 | 9.7692 | 8.9503 | 9.0958
236 | LLOQ 0.01 |  | 06/09/2023 13:54:51 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LLOQ 0.01.590 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 34.7293 | 24.036 | -0.18318 | -0.13095 | 0.33499 | -45.9485 | 49.555 | 76.3561 | 0.021211 | 0.0062971 | 0.0080325 | -0.36542 | 0.037102 | -0.41536 | 0.0093336 | 0.011693 | 0.0076755 | 0.010278 | 0.025539 | 0.061947 | 0.016909 | 0.27198 | 0.011849 | 0.11901 | -0.037815 | 0.014684 | 0.0096409 | 0.0045345 | -0.028951 | 0.043077 | 0.023431 | 97.94% | 95.36% | 103.2% | 97.37% | 96.99% | 96.06% | 95.85% | 0.012204 | 0.0080342 | 0.0083319 | 0.0096463
237 | CCV 10 ppb MAJORS |  | 06/09/2023 14:00:40 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.591 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10350.2821 | 10395.9406 | 8.2917 | 8.2791 | 8.8623 | 10539.3843 | 10188.7254 | 9940.4101 | 10.6015 | 10.5051 | 9.6156 | 9.4123 | 10.2109 | 10.2352 | 10.4859 | 10.5562 | 10.4632 | 9.3715 | 9.2352 | 10.1257 | 11.4212 | 9.7152 | 10.2558 | 9.7888 | 10.1078 | 9.6431 | 9.8179 | 9.8217 | 9.8025 | 9.5089 | 10.0476 | 93.59% | 88.01% | 88.92% | 93.93% | 88.94% | 90.2% | 87.55% | 9.3278 | 9.4768 | 9.4024 | 9.5241
238 | CCV 10 ppb MAJORS |  | 06/09/2023 14:06:57 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.592 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9896.7305 | 9948.2934 | 8.4312 | 8.5052 | 8.6271 | 9698.5241 | 9916.7378 | 9574.7797 | 10.2203 | 10.0061 | 10.3564 | 10.2811 | 10.3055 | 9.8619 | 10.1985 | 10.1822 | 10.0987 | 10.0408 | 9.7984 | 10.1363 | 10.6039 | 9.5365 | 10.2556 | 10.169 | 10.0833 | 10.0715 | 9.7977 | 9.7178 | 10.01 | 10.2348 | 10.1442 | 93.76% | 88.24% | 92.31% | 89.01% | 90.21% | 90.74% | 89.46% | 9.9873 | 10.0255 | 9.4366 | 9.4457
239 | LLOQ 0.1 |  | 06/09/2023 14:12:45 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LLOQ 0.1.593 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 133.0755 | 130.3552 | -0.42096 | -0.35051 | -0.37528 | 202.177 | 139.043 | 145.5491 | 0.10403 | 0.10336 | 0.10204 | -0.17154 | 0.10846 | -0.051706 | 0.11566 | 0.10179 | 0.11814 | 0.10638 | 0.45672 | 0.2031 | 0.15196 | 0.35565 | 0.10433 | 0.2092 | 0.094324 | 0.091263 | 0.1066 | 0.10935 | 0.072271 | 0.11944 | 0.10119 | 97.83% | 95.77% | 97.96% | 95.98% | 97.76% | 97.33% | 96.95% | 0.09904 | 0.097686 | 0.095693 | 0.095319
240 | 8_30_21_F_9 |  | 06/09/2023 14:19:27 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_F_9.594 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 7372.8488 | 6724.842 | 1.1053 | 1.5575 | 3.444 | 13230.5621 | 1643.2144 | 30115.7601 | 0.72611 | 0.057552 | 0.058717 | 4.0314 | 6.5002 | 4.3067 | 0.016903 | 0.16078 | 2.6419 | 2.6617 | 0.85072 | 4.3507 | 3.9003 | 4.2577 | 4.3082 | 0.38355 | -0.0023603 | 0.073499 | 0.083641 | 0.43748 | 0.056395 | 5.1663 | 5.1573 | 89.62% | 86.42% | 91.44% | 86.23% | 87.98% | 89.2% | 85.45% | 0.012089 | 0.010099 | 0.089304 | 0.091094
241 | 8_30_21_F_10 |  | 06/09/2023 14:25:14 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_F_10.595 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 7605.4317 | 8562.5889 | 2.4702 | 2.7483 | 5.653 | 15614.9794 | 2027.2265 | 35400.3667 | 0.823 | 0.071903 | 0.071788 | 3.8707 | 4.7868 | 4.3119 | 0.016704 | 0.16923 | 2.2124 | 2.1133 | 2.2426 | 4.6293 | 4.2277 | 4.4619 | 4.6144 | 0.32003 | 0.048495 | 0.085906 | 0.093583 | 0.50439 | 0.033279 | 3.3088 | 3.4063 | 88.04% | 86.25% | 88.37% | 86.77% | 84.62% | 89.9% | 86.88% | 0.010856 | 0.0087869 | 0.099545 | 0.10543
242 | 8_30_21_F_11 |  | 06/09/2023 14:31:03 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_F_11.596 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 7528.8781 | 9191.9043 | 2.542 | 2.5494 | 3.7502 | 15500.1986 | 1970.5247 | 34033.8756 | 0.90535 | 0.073176 | 0.069375 | 3.3053 | 3.7728 | 3.8449 | 0.018276 | 0.18032 | 2.1842 | 2.116 | 1.4155 | 4.7482 | 4.4507 | 4.4323 | 4.6714 | 0.54324 | 0.07571 | 0.085125 | 0.11545 | 0.51325 | 0.069863 | 3.1972 | 3.2026 | 88.9% | 86.97% | 88.18% | 85.33% | 92.91% | 87.37% | 89.32% | 0.0093834 | 0.0071903 | 0.10305 | 0.096277
243 | 8_30_21_UF_1 |  | 06/09/2023 14:36:51 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_UF_1.597 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 8320.4135 | 5831.4832 | 1.9205 | 2.0267 | 19.7294 | 19352.791 | 1924.7639 | 33123.7424 | 0.38764 | 0.058208 | 0.061593 | -0.081697 | 10.4533 | 0.32625 | 0.025743 | 0.37545 | 0.63925 | 0.63523 | 2.4048 | 4.672 | 4.4787 | 4.4706 | 4.5227 | 0.58761 | -0.013792 | 0.050874 | 0.03954 | 0.3523 | -0.0045232 | 13.6187 | 13.7337 | 88.05% | 86.76% | 87.28% | 86.34% | 89.97% | 90.92% | 85.55% | 0.0081768 | 0.0054316 | 0.076785 | 0.072067
244 | 8_30_21_UF_6 |  | 06/09/2023 14:42:39 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_UF_6.598 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9523.0812 | 6935.5368 | 0.85983 | 0.59981 | 15.3217 | 16554.712 | 1911.6149 | 34065.3098 | 0.65319 | 0.058027 | 0.056616 | 0.22156 | 10.8045 | 0.65174 | 0.023888 | 0.25852 | 1.7779 | 1.6438 | 1.3796 | 4.8385 | 4.3707 | 4.8383 | 4.9559 | 0.56287 | 0.041442 | 0.05797 | 0.10526 | 0.43462 | 0.025084 | 7.8126 | 8.2503 | 89.67% | 85.56% | 87.57% | 89.03% | 86.24% | 85.8% | 85.31% | 0.014868 | 0.014218 | 0.080954 | 0.087163
245 | 8_30_21_UF_8 |  | 06/09/2023 14:48:27 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_UF_8.599 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9534.3572 | 7103.6571 | 1.4054 | 1.5788 | 16.8898 | 15586.1169 | 2014.0234 | 32621.0707 | 0.83508 | 0.072469 | 0.052878 | -0.12639 | 11.2526 | 0.23476 | 0.020179 | 0.19625 | 1.8802 | 1.8273 | 0.85176 | 4.9927 | 4.8737 | 4.875 | 5.0245 | 0.57168 | 0.081949 | 0.099428 | 0.078658 | 0.47057 | 0.032948 | 6.9325 | 6.9145 | 89.77% | 86.53% | 88.41% | 85.44% | 87.93% | 89.35% | 86.4% | 0.015126 | 0.011403 | 0.06056 | 0.059831
246 | 8_30_21_UF_9 |  | 06/09/2023 14:54:15 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_UF_9.600 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 8404.8755 | 6524.6541 | 0.24392 | 0.23937 | 15.5168 | 12499.1699 | 1647.4951 | 29454.6348 | 0.71135 | 0.05764 | 0.052054 | -0.17014 | 6.1432 | 0.24216 | 0.019327 | 0.18918 | 1.7099 | 1.6538 | 3.1745 | 4.0082 | 4.4781 | 3.956 | 4.107 | 0.30994 | 0.015163 | 0.070336 | 0.072725 | 0.41079 | 0.01944 | 4.795 | 4.8001 | 90.49% | 85.92% | 86.44% | 84.72% | 86.58% | 92.66% | 87.54% | 0.010693 | 0.0074593 | 0.059537 | 0.06337
247 | 8_30_21_UF_9_LD |  | 06/09/2023 15:00:03 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_UF_9_LD.601 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 8031.3725 | 6515.491 | 1.0267 | 1.335 | 18.9922 | 12539.6398 | 1624.7286 | 29269.6609 | 0.70477 | 0.048263 | 0.053716 | -0.064952 | 5.9675 | 0.19273 | 0.017213 | 0.16256 | 1.5967 | 1.7289 | 1.5396 | 3.9674 | 3.57 | 3.9774 | 4.0085 | 0.30317 | 0.025474 | 0.068014 | 0.093703 | 0.32015 | 0.01881 | 5.0348 | 4.6685 | 89.92% | 86.85% | 87.73% | 80.71% | 85.95% | 84.16% | 85.89% | 0.011324 | 0.0079519 | 0.065401 | 0.067274
248 | 8_30_21_UF_9_LD |  | 06/09/2023 15:06:19 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_UF_9_LD.602 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 7965.0959 | 6546.9333 | 1.0362 | 1.2345 | 17.4724 | 12303.4553 | 1641.874 | 28653.0401 | 0.72786 | 0.058323 | 0.056416 | -0.10597 | 6.0293 | 0.18329 | 0.011504 | 0.18564 | 1.6287 | 1.6307 | 0.9205 | 3.9099 | 3.9799 | 3.8943 | 4.0846 | 0.4118 | 0.043376 | 0.079304 | 0.077703 | 0.29052 | 0.07594 | 4.7215 | 4.738 | 89.66% | 85.8% | 88.38% | 84.62% | 85.51% | 87.34% | 88.44% | 0.0096816 | 0.0089778 | 0.063337 | 0.065462
249 | 8_30_21_UF_10 |  | 06/09/2023 15:12:07 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_UF_10.603 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 8472.3085 | 8484.3442 | 1.9179 | 2.1762 | 17.9355 | 14571.7212 | 1948.3658 | 34763.217 | 0.81316 | 0.068043 | 0.06421 | 0.83048 | 4.2891 | 1.1461 | 0.020972 | 0.15756 | 1.3915 | 1.402 | 1.3438 | 4.394 | 3.9521 | 4.2927 | 4.4186 | 0.40296 | 0.027 | 0.066 | 0.10395 | 0.4738 | 0.067368 | 3.2265 | 3.2417 | 90.53% | 85.15% | 87.69% | 84.04% | 86.18% | 87.68% | 85.59% | 0.011232 | 0.0082623 | 0.07988 | 0.080989
250 | LBLANK |  | 06/09/2023 15:17:56 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.604 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 26.2709 | 13.319 | -0.69936 | -0.76018 | -0.37202 | 130.629 | 27.3721 | 70.2176 | 0.0013424 | -0.0047257 | -0.0006886 | -0.5246 | 0.0011045 | -0.54598 | 0.00060981 | 0.0032985 | -0.010553 | -0.0080581 | -0.15021 | 0.089651 | 0.038237 | 0.27621 | 0.016887 | 0.23529 | -0.042548 | 0.00072501 | 0.0060475 | 0.028391 | -0.037289 | 0.0026916 | 0.0037586 | 90.53% | 92.37% | 99.93% | 92.75% | 93.67% | 91.76% | 95.18% | 0.0010229 | -0.0013405 | -0.0033056 | -0.0018326
251 | CCV 1ppb TRACE |  | 06/09/2023 15:23:44 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.605 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1030.3002 | 1107.7589 | 1.0969 | 1.1722 | 0.19496 | 1257.8902 | 1038.9043 | 1095.2024 | 1.0234 | 1.069 | 1.0357 | 0.69805 | 1.0109 | 0.6622 | 1.0115 | 1.0752 | 1.0758 | 1.0485 | 0.86927 | 1.0899 | 1.1851 | 1.166 | 1.0262 | 1.1277 | 0.96804 | 0.92377 | 0.97121 | 1.1089 | 0.96313 | 1.0169 | 1.0125 | 95.42% | 91.58% | 92.72% | 90.03% | 85.05% | 91.66% | 95.23% | 1.0137 | 1.0271 | 1.0182 | 1.1293
252 | CCV 1ppb TRACE |  | 06/09/2023 15:30:02 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.606 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1048.8007 | 1147.2892 | 1.0941 | 0.72079 | 0.088427 | 1163.7187 | 1048.0881 | 1103.3704 | 1.0533 | 1.0081 | 1.0417 | 0.74341 | 1.0271 | 0.64882 | 1.0202 | 1.0339 | 1.0599 | 1.0761 | 0.89906 | 1.0757 | 1.0736 | 1.2161 | 1.0325 | 1.039 | 1.1258 | 1.0488 | 1.0353 | 0.93568 | 0.82707 | 1.0314 | 1.0313 | 93.91% | 89.22% | 92.56% | 88.15% | 90.05% | 90.55% | 88.47% | 1.0504 | 1.0656 | 1.0235 | 1.0292
253 | CCV 10 ppb MAJORS |  | 06/09/2023 15:35:51 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.607 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10079.2403 | 10296.5081 | 8.5051 | 9.6642 | 8.4516 | 10013.7953 | 9964.8395 | 9910.8312 | 10.4844 | 10.0899 | 10.2684 | 10.1233 | 11.6735 | 9.6451 | 9.9282 | 10.0368 | 10.0718 | 9.9945 | 9.8268 | 10.1057 | 10.6171 | 10.7958 | 11.7603 | 9.9018 | 11.5121 | 11.2319 | 11.3624 | 9.8211 | 11.813 | 10.2358 | 11.5954 | 91.38% | 76.63% | 87.42% | 84.92% | 86.81% | 86.51% | 78.54% | 10.1339 | 10.274 | 9.7834 | 9.9058
254 | CCV 10 ppb MAJORS |  | 06/09/2023 15:42:08 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.608 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10060.8796 | 10229.1112 | 8.2553 | 9.5989 | 5.9915 | 10299.688 | 9854.0791 | 10178.4009 | 10.3675 | 10.232 | 10.1164 | 10.1097 | 10.3263 | 9.9755 | 10.0747 | 10.3532 | 10.1758 | 9.8394 | 9.6371 | 9.7508 | 10.6018 | 9.6584 | 10.23 | 9.8167 | 10.3005 | 9.8925 | 9.9642 | 9.3935 | 9.948 | 10.0776 | 10.1505 | 94.68% | 86.93% | 88.77% | 88.2% | 87.97% | 113.8% | 86.77% | 9.9543 | 10.0537 | 9.5638 | 9.8555
255 | 8_30_21_UF_10_LSP |  | 06/09/2023 15:47:56 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_UF_10_LSP.609 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 17365.4278 | 17581.9066 | 11.1992 | 11.476 | 23.5128 | 23197.1074 | 11782.5598 | 40956.2313 | 10.9144 | 9.9621 | 10.2421 | 11.2112 | 13.1278 | 11.3792 | 9.8013 | 9.9271 | 10.8114 | 10.8244 | 10.8601 | 13.827 | 14.4124 | 12.3795 | 13.252 | 10.0911 | 9.5057 | 9.438 | 9.3648 | 10.1008 | 9.7307 | 12.9794 | 12.1851 | 90.66% | 88.21% | 86.65% | 82.57% | 86.44% | 85.9% | 88.5% | 9.8159 | 9.8213 | 9.3499 | 9.3586
256 | 8_30_21_UF_10_DUP_SP |  | 06/09/2023 15:53:45 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_UF_10_DUP_SP.610 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 14108.5777 | 14084.5612 | 5.2743 | 5.4299 | 24.5372 | 20751.6749 | 7570.4687 | 42227.5235 | 6.4444 | 5.5419 | 5.8325 | 5.7491 | 9.909 | 5.7674 | 5.4167 | 5.6421 | 6.7151 | 6.8213 | 6.7085 | 9.8112 | 10.3554 | 9.2842 | 9.9296 | 5.8174 | 5.5191 | 5.3775 | 5.356 | 5.5505 | 5.6788 | 9.0124 | 8.636 | 88.1% | 83.28% | 85.68% | 79.91% | 82.81% | 85.74% | 84.04% | 5.4407 | 5.5574 | 5.1485 | 5.3223
257 | 8_30_21_UF_11 |  | 06/09/2023 15:59:33 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_UF_11.611 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 8065.7075 | 8504.7753 | 0.61339 | 0.65688 | 17.1705 | 14908.8046 | 1987.4744 | 32827.2518 | 0.83344 | 0.067081 | 0.057486 | -0.10642 | 3.381 | 0.26452 | 0.015242 | 0.19224 | 1.439 | 1.4216 | 0.83114 | 4.545 | 4.5633 | 4.4215 | 4.5489 | 0.43427 | 0.045713 | 0.08479 | 0.10218 | 0.32428 | 0.059543 | 3.0612 | 3.0904 | 88.03% | 83.79% | 87.56% | 82.88% | 84.79% | 85.67% | 84.04% | 0.0099806 | 0.0049391 | 0.05923 | 0.062604
258 | 8_30_21_BB |  | 06/09/2023 16:05:21 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_BB.612 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 29.1599 | 14.6033 | 3.0502 | 2.9447 | 0.0016889 | 329.1188 | 47.4722 | 72.9402 | 0.0053734 | 0.024807 | 0.021117 | 0.73705 | 0.012399 | 0.95052 | 0.001103 | 0.0068628 | 0.10484 | 0.094086 | 0.0061283 | 0.074308 | 0.032691 | 0.27196 | 0.0117 | 0.086407 | -0.053078 | 0.015527 | 0.0061205 | 0.014641 | -0.041654 | 0.0041907 | 0.0047277 | 93.75% | 91.17% | 92.24% | 94.94% | 91.52% | 91.96% | 90.71% | 0.0011339 | -0.00085688 | 0.033817 | 0.035714
259 | 8_30_21_FB |  | 06/09/2023 16:11:09 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_FB.613 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 55.394 | 24.3595 | 1.427 | 1.9395 | -0.28418 | 303.1345 | 38.5668 | 103.9039 | 0.0040616 | 0.15976 | 0.1519 | 0.17986 | 0.13601 | 0.17232 | 0.007061 | 9.4736 | 0.12651 | 0.1175 | 5.5104 | 0.071241 | 0.010141 | 0.27043 | 0.010284 | 0.1428 | -0.042193 | 0.0074574 | 0.0085552 | -0.0093558 | -0.0371 | 0.0044411 | 0.004955 | 92.57% | 89.42% | 91.45% | 90.34% | 89.75% | 88.11% | 91.03% | 0.0023058 | -0.0001493 | 0.063451 | 0.065462
260 | 8_30_21_LAB_A |  | 06/09/2023 16:16:56 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_LAB_A.614 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 27.2854 | 12.1963 | -0.25959 | -0.11942 | -0.66346 | 253.0791 | 33.439 | 65.8758 | 0.0047203 | 0.070653 | 0.080143 | 0.092313 | 0.0088334 | 0.10053 | 0.00087106 | 0.28522 | 0.099523 | 0.096049 | 0.013109 | 0.072263 | 0.010141 | 0.267 | 0.006365 | 0.20081 | -0.053078 | 0.0025344 | 0.0030486 | 0.0012292 | -0.041654 | 0.023982 | 0.025113 | 90.71% | 88.13% | 91.64% | 90.05% | 90.09% | 88.47% | 88.58% | 0.0011326 | -0.0011926 | 0.032982 | 0.032868
261 | 8_30_21_LAB_B |  | 06/09/2023 16:22:44 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_LAB_B.615 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 28.2395 | 12.1963 | -0.22676 | -0.31875 | -0.6428 | 148.4438 | 37.4518 | 73.0302 | 0.0013424 | 0.021579 | 0.028696 | -0.091964 | 0.0051528 | -0.16986 | 0.00064181 | 0.26659 | 0.075653 | 0.092292 | -0.020739 | 0.057848 | 0.017818 | 0.26559 | 0.0047472 | 0.14761 | -0.053078 | 0.0041921 | 0.0038174 | -0.0094023 | -0.037078 | 0.0017026 | 0.0028624 | 92.41% | 89.2% | 92.1% | 84.07% | 89.48% | 88.32% | 88.17% | 0.0009187 | -0.0013025 | 0.032653 | 0.03178
262 | 9_13_21_W_1 |  | 06/09/2023 16:28:32 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_W_1.616 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1431.059 | 1177.3561 | 6.4934 | 8.0822 | 2.6273 | 4178.8156 | 453.1148 | 6784.2511 | 0.058839 | 0.028678 | 0.035264 | 7.2768 | 4.381 | 7.1902 | 0.0093741 | 0.068688 | 0.33552 | 0.32986 | 0.07969 | 0.88049 | 0.75282 | 1.0125 | 0.83515 | 0.141 | -0.02497 | 0.01121 | 0.01363 | 0.05912 | -0.027642 | 3.1849 | 3.1184 | 92.9% | 86.59% | 90.66% | 85.72% | 90.19% | 85.6% | 88.33% | 0.0032873 | 0.00089727 | 0.034703 | 0.034128
263 | 9_13_21_W_6 |  | 06/09/2023 16:34:20 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_W_6.617 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1482.1859 | 1322.6815 | 5.8551 | 6.3051 | 1.231 | 3068.5802 | 335.1227 | 6205.8169 | 0.12314 | 0.032214 | 0.029852 | 7.776 | 3.1643 | 7.7823 | 0.0059579 | 0.05734 | 0.8517 | 0.82494 | 0.49123 | 0.90553 | 0.90412 | 1.07 | 0.87993 | 0.22008 | -0.018736 | 0.025334 | 0.017873 | 0.15201 | -0.0087112 | 1.1615 | 1.157 | 90.47% | 85.59% | 90.77% | 85.35% | 86.25% | 86.27% | 85.41% | 0.0059896 | 0.0028704 | 0.060291 | 0.062624
264 | 9_13_21_W_8 |  | 06/09/2023 16:40:07 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_W_8.618 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1627.3536 | 1373.6259 | 5.9716 | 7.1458 | 1.6691 | 2720.0914 | 400.5724 | 6467.0618 | 0.17124 | 0.060711 | 0.076185 | 9.5979 | 2.8293 | 8.1304 | 0.0074319 | 0.1123 | 0.69324 | 0.83719 | 0.33101 | 1.0229 | 0.79213 | 1.0746 | 0.90938 | 0.12963 | -0.030413 | 0.025321 | 0.024499 | 0.10167 | 0.00061869 | 1.1729 | 0.98398 | 90.37% | 85.84% | 86.79% | 76.61% | 88.09% | 84.38% | 88.02% | 0.00515 | 0.0027356 | 0.048866 | 0.050529
265 | LBLANK |  | 06/09/2023 16:45:55 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.619 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 27.7253 | 12.1963 | -0.73915 | -0.75231 | -0.75297 | 113.8578 | 32.3785 | 65.8758 | 0.0020398 | -0.0028941 | -0.00087695 | -0.55435 | 0.00083728 | -0.55519 | 0.00041452 | -0.0014944 | -0.010466 | -0.008599 | -0.15804 | 0.10023 | 0.010141 | 0.26283 | 0.0018582 | 0.12282 | -0.053078 | 0.00082217 | 0.0046871 | -0.036165 | -0.041654 | 0.0019808 | 0.0031888 | 90.68% | 86.81% | 91.48% | 89.17% | 85.63% | 87.13% | 89.26% | 0.0012397 | -0.0011874 | -0.0037063 | -0.0022137
266 | CCV 1ppb TRACE |  | 06/09/2023 16:51:44 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.620 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 993.1619 | 1041.3404 | 0.48807 | 0.54156 | -0.019812 | 957.7847 | 999.2865 | 1065.9513 | 0.97358 | 0.96627 | 0.99705 | 0.77491 | 0.95789 | 0.78497 | 0.97097 | 0.96372 | 0.97835 | 0.98649 | 0.8434 | 0.99404 | 1.079 | 1.1465 | 0.99948 | 0.97419 | 1.0197 | 0.92347 | 1.0124 | 0.83096 | 0.9417 | 0.98114 | 0.98494 | 90.64% | 85.77% | 90.1% | 86.62% | 86.48% | 86.07% | 88.78% | 0.97995 | 0.99255 | 0.99272 | 1.0217
267 | CCV 1ppb TRACE |  | 06/09/2023 16:58:02 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.621 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 966.7869 | 971.3645 | 0.46476 | 0.64628 | -0.065615 | 1056.7536 | 949.6856 | 1108.4414 | 1.0172 | 0.98682 | 0.9931 | 0.79067 | 0.94275 | 0.74872 | 0.96006 | 0.9526 | 1.0304 | 0.98146 | 0.84948 | 1.0453 | 0.88978 | 1.1391 | 0.97712 | 1.0156 | 0.896 | 0.90304 | 1.0256 | 0.93563 | 0.97357 | 0.97889 | 0.98462 | 91.82% | 86.73% | 91.51% | 86.44% | 87.87% | 85.33% | 88.5% | 1.0051 | 1.0004 | 0.98972 | 1.0196
268 | LFB |  | 06/09/2023 17:03:50 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.622 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9262.4846 | 9471.6401 | 8.1244 | 9.0627 | 10.738 | 9258.761 | 9226.0767 | 9558.0387 | 9.5624 | 9.4227 | 9.5924 | 9.4899 | 9.6537 | 9.1613 | 9.3395 | 9.3215 | 9.4256 | 9.4542 | 9.5007 | 9.5848 | 10.4672 | 9.226 | 9.914 | 9.6342 | 10.2524 | 9.8394 | 9.8034 | 9.6658 | 10.2339 | 9.6754 | 9.7665 | 90.98% | 81.99% | 86.51% | 83.61% | 85.07% | 81.96% | 83.14% | 9.7837 | 9.8395 | 9.4397 | 9.6013
269 | CCV 10 ppb MAJORS |  | 06/09/2023 17:09:38 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.623 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9677.6664 | 9970.5728 | 8.16 | 8.6489 | 8.2377 | 10400.4483 | 9814.0069 | 9921.448 | 10.2958 | 9.8427 | 10.7959 | 10.7359 | 9.9488 | 9.6644 | 9.8859 | 10.0682 | 9.9351 | 10.4581 | 10.7515 | 9.9187 | 10.7892 | 9.1103 | 9.7958 | 10.2309 | 10.1506 | 9.8266 | 9.5771 | 9.5753 | 9.9265 | 10.6589 | 9.8607 | 91.13% | 84.78% | 87.87% | 80.8% | 85.92% | 84.74% | 85.97% | 10.6465 | 10.7446 | 9.7305 | 9.8364
270 | CCV 10 ppb MAJORS |  | 06/09/2023 17:15:55 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.624 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9642.6703 | 9692.0215 | 7.9438 | 8.8318 | 8.3094 | 9318.3186 | 9668.4219 | 10123.8807 | 9.798 | 9.5417 | 10.0094 | 9.9519 | 10.1888 | 9.5323 | 9.6938 | 9.8957 | 9.7993 | 9.7489 | 10.0579 | 9.7141 | 10.4829 | 9.2205 | 10.0749 | 9.5894 | 10.3581 | 9.6943 | 9.7155 | 9.4027 | 10.0792 | 9.9916 | 10.0249 | 91.84% | 84.24% | 90.44% | 85.32% | 86.31% | 86.15% | 86.85% | 9.8276 | 9.9611 | 9.4716 | 9.6015
271 | 9_13_21_W_9 |  | 06/09/2023 17:21:43 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_W_9.625 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1452.5202 | 1355.911 | 4.5469 | 4.6595 | 1.2511 | 2646.2283 | 331.7143 | 5866.5146 | 0.14078 | 0.018604 | 0.019793 | 4.5486 | 1.7309 | 4.569 | 0.0056605 | 0.035378 | 0.58871 | 0.57913 | 0.079389 | 0.90221 | 0.79963 | 0.97404 | 0.79185 | 0.37823 | -0.030474 | 0.016464 | 0.026782 | 0.064454 | -0.0088088 | 0.76106 | 0.75983 | 84.77% | 86.09% | 92.1% | 86.87% | 88.04% | 87.15% | 87.68% | 0.0031368 | 0.0011386 | 0.033869 | 0.033499
272 | 9_13_21_W_10 |  | 06/09/2023 17:27:31 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_W_10.626 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1353.6894 | 1691.1696 | 11.2151 | 12.1638 | 1.9732 | 2706.3437 | 362.0248 | 6470.5883 | 0.16784 | 0.013093 | 0.017975 | 10.2469 | 2.8566 | 9.9885 | 0.0079445 | 0.044731 | 0.60759 | 0.62425 | 0.38713 | 0.92803 | 0.75668 | 1.0908 | 0.92338 | 0.20922 | -0.047208 | 0.029727 | 0.025942 | 0.10217 | -0.031922 | 0.4947 | 0.48757 | 90.6% | 83.1% | 89.15% | 85.04% | 86.4% | 86.21% | 84.83% | 0.0038632 | 0.0020711 | 0.067946 | 0.065375
273 | 9_13_21_W_10_DUP |  | 06/09/2023 17:33:20 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_W_10_DUP.627 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1305.2801 | 1722.5311 | 9.5939 | 9.7557 | 2.0356 | 3004.2621 | 390.6448 | 6378.6575 | 0.15885 | 0.020603 | 0.027796 | 8.5626 | 2.5026 | 8.7277 | 0.0084008 | 0.035213 | 0.50826 | 0.55308 | 0.33577 | 0.92111 | 0.75867 | 1.0603 | 0.89266 | 0.21189 | -0.024034 | 0.02236 | 0.023347 | 0.14582 | -0.017446 | 0.48412 | 0.48843 | 90.39% | 83.58% | 89.09% | 84.78% | 85.69% | 84.4% | 85.61% | 0.0042131 | 0.0017923 | 0.079493 | 0.081176
274 | 9_13_21_W_11 |  | 06/09/2023 17:39:08 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_W_11.628 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1262.2545 | 1648.2715 | 22.8041 | 26.0731 | 2.0672 | 2849.4449 | 350.9965 | 6326.2494 | 0.1627 | 0.034828 | 0.040507 | 18.1507 | 3.584 | 18.1442 | 0.0093079 | 0.048495 | 0.84507 | 0.82269 | 0.80306 | 0.86672 | 0.78054 | 1.0454 | 0.86199 | 0.23744 | -0.029495 | 0.024452 | 0.025147 | 0.05197 | -0.022082 | 0.44564 | 0.44327 | 90.47% | 82.63% | 87.35% | 83.39% | 85.32% | 83.21% | 83.23% | 0.0056532 | 0.0033254 | 0.092055 | 0.090943
275 | 9_13_21_F_1 |  | 06/09/2023 17:44:56 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_F_1.629 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 7265.8993 | 5740.635 | 3.3542 | 3.4301 | 7.8391 | 20453.4331 | 1892.4248 | 34467.0894 | 0.31926 | 0.17214 | 0.18037 | 4.5784 | 13.2675 | 4.8002 | 0.033831 | 0.48595 | 0.82159 | 0.89081 | 1.5823 | 4.2432 | 3.6625 | 3.9073 | 4.0164 | 0.62471 | -0.0053425 | 0.054099 | 0.047152 | 0.32966 | 0.017664 | 15.9339 | 16.0864 | 82.52% | 81.53% | 86.01% | 82.69% | 84.22% | 84.03% | 82.36% | 0.0097698 | 0.0065376 | 0.12043 | 0.11513
276 | 9_13_21_F_6 |  | 06/09/2023 17:50:44 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_F_6.630 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 8090.0289 | 7439.234 | 2.5007 | 2.4975 | 1.4491 | 15734.8164 | 1885.3472 | 36122.8731 | 0.6942 | 0.10499 | 0.089059 | 4.3423 | 12.1021 | 5.0803 | 0.028459 | 0.40665 | 2.6074 | 2.4808 | 2.6743 | 4.1111 | 3.96 | 3.8671 | 4.1292 | 0.49143 | 0.036575 | 0.081188 | 0.098457 | 0.47141 | 0.046711 | 6.0134 | 6.0815 | 89.26% | 81.93% | 82.25% | 83.37% | 83.75% | 84.75% | 86.09% | 0.01602 | 0.015804 | 0.13682 | 0.14381
277 | 9_13_21_F_8 |  | 06/09/2023 17:56:32 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_F_8.631 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 8700.1091 | 7057.6205 | 2.8348 | 3.0194 | 2.7141 | 14805.0086 | 1876.6058 | 34430.3418 | 0.88873 | 0.062953 | 0.068192 | 5.221 | 10.8373 | 5.9127 | 0.023603 | 0.26911 | 2.7927 | 2.6821 | 2.0266 | 4.5302 | 4.663 | 4.2861 | 4.497 | 0.56704 | 0.011983 | 0.073642 | 0.080893 | 0.50217 | 0.031963 | 5.1192 | 5.0869 | 87.67% | 82.3% | 82.29% | 81.97% | 83.03% | 84.27% | 84.32% | 0.013257 | 0.011262 | 0.19062 | 0.19259
278 | 9_13_21_F_9 |  | 06/09/2023 18:02:20 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_F_9.632 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 7694.3349 | 7066.9083 | 6.851 | 7.2079 | 2.3701 | 12644.5093 | 1735.8204 | 30955.1119 | 0.71985 | 0.067375 | 0.067746 | 5.4635 | 3.9336 | 5.8622 | 0.018383 | 0.18979 | 2.354 | 2.2648 | 5.1581 | 3.8069 | 3.3942 | 3.5428 | 3.6982 | 0.50296 | 0.03041 | 0.07367 | 0.090924 | 0.38532 | 0.037234 | 3.7604 | 3.7511 | 87.62% | 82.68% | 82.95% | 81.38% | 82.13% | 85.86% | 85.61% | 0.0087501 | 0.005871 | 0.11027 | 0.11734
279 | 9_13_21_F_10 |  | 06/09/2023 18:08:08 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_F_10.633 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 6720.4679 | 8466.093 | 3.0675 | 3.5375 | 3.3289 | 13953.7113 | 1839.8809 | 33615.5471 | 0.76553 | 0.062824 | 0.065684 | 3.7798 | 4.6232 | 3.8619 | 0.01702 | 0.1833 | 1.783 | 1.8893 | 1.8002 | 4.2404 | 4.1209 | 4.1177 | 4.2793 | 0.44616 | 0.026058 | 0.092804 | 0.083112 | 0.3748 | 0.039088 | 2.5619 | 2.4679 | 85.79% | 79.97% | 84.07% | 77.78% | 81.22% | 81.88% | 81.35% | 0.0097031 | 0.0075681 | 0.11264 | 0.1188
280 | 9_13_21_F_11 |  | 06/09/2023 18:13:56 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_F_11.634 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 6546.1441 | 8488.2297 | 9.7718 | 10.8019 | 2.353 | 14529.1266 | 1821.4119 | 33490.9011 | 0.77635 | 0.064325 | 0.065044 | 4.3432 | 4.9411 | 4.2834 | 0.015147 | 0.17821 | 2.1329 | 2.2587 | 3.8006 | 4.1032 | 3.8567 | 4.0333 | 4.1773 | 0.45241 | 0.020084 | 0.083356 | 0.083279 | 0.44844 | 0.059348 | 2.4099 | 2.2889 | 86.75% | 79.91% | 83% | 76.35% | 80.45% | 79.93% | 81.19% | 0.011052 | 0.0079887 | 0.10648 | 0.11068
281 | 9_13_21_F_11_LD |  | 06/09/2023 18:19:44 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_F_11_LD.635 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 6606.4935 | 8650.7375 | 8.6731 | 8.9113 | 3.3946 | 14347.0358 | 1856.2064 | 33978.0754 | 0.72166 | 0.060198 | 0.059493 | 4.032 | 5.0291 | 4.2061 | 0.018694 | 0.17477 | 2.168 | 2.1982 | 3.5672 | 4.2156 | 3.7024 | 4.0482 | 4.2137 | 0.48908 | -0.0038044 | 0.088635 | 0.089946 | 0.33682 | 0.030011 | 2.3118 | 2.3086 | 84.47% | 78.61% | 82.92% | 78.68% | 81.52% | 80.36% | 80.17% | 0.010098 | 0.008941 | 0.10821 | 0.10923
282 | 9_13_21_F_11_LD |  | 06/09/2023 18:26:00 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_F_11_LD.636 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 6419.967 | 8408.2134 | 8.4553 | 9.3243 | 3.4651 | 13677.5987 | 1830.7418 | 32686.1297 | 0.71774 | 0.064445 | 0.055345 | 3.9319 | 4.9914 | 4.0516 | 0.017413 | 0.19412 | 2.1034 | 2.1427 | 3.4963 | 4.1084 | 3.8159 | 4.0614 | 4.134 | 0.4565 | 0.050707 | 0.090979 | 0.089569 | 0.42839 | 0.034225 | 2.2537 | 2.2884 | 86.37% | 79.78% | 85.25% | 79.44% | 81.22% | 81.93% | 79.7% | 0.011467 | 0.0067253 | 0.10273 | 0.10759
283 | LBLANK |  | 06/09/2023 18:31:48 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.637 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 26.8542 | 12.1963 | -0.73585 | -0.73226 | -0.78764 | 45.8085 | 40.4007 | 65.8758 | 0.0020443 | -0.005768 | -0.0012446 | -0.54655 | 0.00078471 | -0.5523 | 0.00042811 | 0.0049746 | -0.0099498 | -0.0077734 | -0.1623 | 0.04554 | 0.010141 | 0.262 | 0.00090522 | 0.17269 | -0.030217 | 0.0043397 | 0.0038854 | 0.057765 | -0.036988 | 0.0015626 | 0.0031974 | 89.92% | 86.31% | 88.49% | 87.12% | 86.62% | 84.91% | 85.72% | 0.0010293 | -0.0012204 | -0.0035345 | -0.0022193
284 | CCV 1ppb TRACE |  | 06/09/2023 18:37:37 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.638 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 990.5272 | 1004.2552 | 0.51466 | 0.59794 | -0.024345 | 1087.9378 | 1024.5979 | 998.689 | 1.0132 | 0.99961 | 0.99533 | 0.77654 | 0.97632 | 0.7694 | 0.98053 | 0.99964 | 0.96853 | 0.98302 | 0.81642 | 0.98861 | 0.99486 | 1.1792 | 0.98661 | 1.1434 | 0.97474 | 0.996 | 0.95995 | 0.95554 | 0.99607 | 0.97894 | 0.9928 | 90.89% | 84.8% | 87.41% | 85.84% | 76.51% | 85.46% | 83.62% | 0.97334 | 0.97579 | 1.0095 | 1.1754
285 | CCV 1ppb TRACE |  | 06/09/2023 18:43:54 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.639 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 978.5032 | 984.4313 | 0.50882 | 0.41605 | -0.075403 | 1341.7701 | 1067.8231 | 1044.0416 | 1.0319 | 1.0151 | 0.96791 | 0.79594 | 0.97529 | 0.87173 | 0.98077 | 0.99058 | 1.081 | 0.98512 | 0.84038 | 1.0323 | 1.0578 | 1.1271 | 0.98683 | 1.1375 | 0.93332 | 1.0285 | 0.9997 | 0.93953 | 1.0488 | 0.96639 | 0.97562 | 90.78% | 84.41% | 85.74% | 86.39% | 77.73% | 85.08% | 88.25% | 0.98266 | 0.98763 | 1.0123 | 1.1862
286 | CCV 10 ppb MAJORS |  | 06/09/2023 18:49:43 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.640 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9357.2943 | 9934.5988 | 8.0297 | 7.555 | 8.1385 | 9656.7771 | 9593.9796 | 9754.7959 | 9.9263 | 9.6521 | 9.8255 | 9.8189 | 10.8173 | 9.2958 | 9.6017 | 9.4213 | 9.6505 | 9.5815 | 9.7832 | 9.743 | 10.1818 | 10.2667 | 10.9586 | 9.6974 | 10.5995 | 10.6169 | 10.7044 | 9.5458 | 10.7835 | 9.7857 | 10.8307 | 88.8% | 74.18% | 84.58% | 81.52% | 82.34% | 83.03% | 74.64% | 9.7434 | 9.9244 | 9.7802 | 9.9968
287 | CCV 10 ppb MAJORS |  | 06/09/2023 18:56:01 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.641 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9359.2825 | 9758.1549 | 8.2043 | 8.6707 | 8.0438 | 9660.2451 | 9328.1657 | 9747.3858 | 9.7615 | 9.5959 | 10.1773 | 10.1279 | 10.1694 | 9.3479 | 9.6939 | 9.4181 | 9.5992 | 9.934 | 10.1681 | 10.0508 | 9.7274 | 9.4649 | 10.1579 | 9.8794 | 9.8433 | 9.6168 | 9.9831 | 9.6275 | 9.9185 | 10.1765 | 10.0221 | 88.17% | 80.45% | 87.35% | 80.9% | 89.66% | 85.46% | 81.4% | 9.986 | 10.1948 | 9.7992 | 9.0283
288 | 9_13_21_F_11_DUP |  | 06/09/2023 19:01:49 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_F_11_DUP.642 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 6319.7421 | 8200.8959 | 7.0696 | 7.8935 | 2.3763 | 14023.5136 | 1757.5997 | 33340.1601 | 0.70129 | 0.094072 | 0.11062 | 4.3052 | 4.9288 | 4.6348 | 0.019315 | 0.17221 | 2.1572 | 2.1897 | 3.3516 | 4.0512 | 3.7302 | 4.0078 | 4.1217 | 0.41564 | 0.015127 | 0.077222 | 0.091013 | 0.47286 | 0.076565 | 2.2531 | 2.2715 | 87.35% | 78.5% | 83.2% | 78.62% | 79.99% | 80.23% | 79.2% | 0.011208 | 0.0084817 | 0.11267 | 0.12116
289 | 9_13_21_F_11_DUP_LSP |  | 06/09/2023 19:07:37 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_F_11_DUP_LSP.643 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 15282.9974 | 17161.2892 | 15.887 | 17.6304 | 12.2822 | 21879.0557 | 11122.8162 | 39866.3127 | 10.731 | 9.8767 | 10.0676 | 14.3913 | 15.5426 | 14.3204 | 9.7356 | 9.7202 | 11.4251 | 11.2439 | 12.4523 | 13.1433 | 13.9219 | 13.2762 | 14.7963 | 9.7463 | 10.5485 | 10.3849 | 10.3646 | 9.5374 | 10.1288 | 11.9488 | 12.9806 | 85.73% | 72.76% | 80.51% | 77.71% | 77.58% | 77.79% | 75.68% | 9.5699 | 9.7408 | 9.4744 | 9.8775
290 | 9_13_21_UF_1 |  | 06/09/2023 19:13:24 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_UF_1.644 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 8406.4411 | 5776.4857 | 1.295 | 1.0836 | 6.5518 | 20677.0751 | 1914.2748 | 33226.2221 | 0.31017 | 0.052862 | 0.060253 | 0.44958 | 12.9073 | 0.74231 | 0.030787 | 0.48929 | 0.56333 | 0.62138 | 1.1934 | 4.3231 | 3.62 | 3.8744 | 3.9622 | 0.691 | -0.0097598 | 0.040885 | 0.047865 | 0.38443 | 0.0041033 | 15.4972 | 15.4994 | 77.6% | 78.69% | 82.25% | 78.62% | 80.14% | 86.09% | 79.13% | 0.0093707 | 0.0068681 | 0.051488 | 0.049953
291 | 9_13_21_UF_6 |  | 06/09/2023 19:19:12 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_UF_6.645 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 8590.7769 | 6678.1146 | 0.062114 | -0.063557 | 2.1955 | 14559.0137 | 1710.2953 | 33109.0591 | 0.65611 | 0.066997 | 0.059011 | 0.11675 | 11.413 | 0.44434 | 0.019112 | 0.22479 | 1.6633 | 1.7502 | 1.6765 | 4.0614 | 3.8797 | 3.8915 | 3.9651 | 0.51524 | 0.030526 | 0.085183 | 0.074196 | 0.35931 | 0.033624 | 5.855 | 5.6403 | 84.93% | 81.58% | 82.97% | 79.1% | 80.34% | 81.69% | 81.78% | 0.016644 | 0.014431 | 0.063749 | 0.068039
292 | 9_13_21_UF_6_DUP |  | 06/09/2023 19:25:00 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_UF_6_DUP.646 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 8727.8469 | 6832.6696 | 0.034985 | -0.045113 | 2.2106 | 14646.7774 | 1777.8245 | 32536.5065 | 0.6123 | 0.088386 | 0.081406 | 0.21757 | 13.4677 | 0.57943 | 0.021634 | 0.23609 | 1.7106 | 1.714 | 1.7324 | 4.0114 | 3.5668 | 4.5439 | 4.7133 | 0.57374 | 0.056426 | 0.10855 | 0.098715 | 0.38333 | 0.040242 | 5.7429 | 6.7444 | 84.83% | 70.21% | 82.99% | 79.69% | 79.75% | 81.71% | 70.61% | 0.016624 | 0.012808 | 0.064614 | 0.069286
293 | 9_13_21_UF_8 |  | 06/09/2023 19:30:49 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_UF_8.647 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9186.9545 | 7155.0816 | 0.0098375 | 0.42558 | 2.6685 | 13504.8571 | 1878.8146 | 32664.3532 | 0.82355 | 0.053444 | 0.049778 | 0.017128 | 10.3329 | 0.47986 | 0.020961 | 0.21961 | 1.5915 | 1.5671 | 0.95528 | 4.217 | 4.2853 | 4.3014 | 4.4073 | 0.44348 | 0.032998 | 0.072723 | 0.089949 | 0.35168 | 0.0091943 | 4.9401 | 4.9173 | 87.77% | 79.43% | 79.61% | 78.4% | 80.48% | 81.03% | 79.42% | 0.011493 | 0.0079384 | 0.050025 | 0.056634

## Concentrations RSDs
| Sample Id | R | Acquisition Time | QC Status | Dataset File | Method File | Na KED 23/23 Helium KED_Low | Mg KED 24/24 Helium KED_Low | Al 27/27 | Al KED 27/27 Helium KED_Low | P Oshift 31/47 Oxygen DRC | S (34) -1 34/34 Helium KED_Low | K KED 39/39 Helium KED_Low | Ca KED 43/43 Helium KED_Low | V KED 51/51 Helium KED_Low | Cr (52) 52/52 Helium KED_Low | Cr (53) NH3 53/53 Ammonia DRC | Fe (54) NH3 54/54 Ammonia DRC | Mn Oshift 55/71 Oxygen DRC | Fe (56) KED 56/56 Helium KED_Low | Co KED 59/59 Helium KED_Low | Ni KED 60/60 Helium KED_Low | Cu (63) KED 63/63 Helium KED_Low | Cu (65) NH3 65/65 Ammonia DRC | Zn NH3 66/66 Ammonia DRC | As 75/75 | As KED 75/75 Helium KED_Low | As RH Oshift 75/91 Oxygen DRC | As Oshift 75/91 Oxygen DRC | Se (77) 77/77 | Se (77) Oshift 77/93 Oxygen DRC | Se (78) Oshift 78/94 Oxygen DRC | Se (80) Oshift 80/96 Oxygen DRC | Se (82) 82/82 | Se (82) Oshift 82/98 Oxygen DRC | Mo (95) NH3 95/95 Ammonia DRC | Mo (95) OShift 95/127 Oxygen DRC | Rh (IS1) 103/103 (IS) | Rh (IS2) 103/103 (IS) | Rh (IS3) 103/103 (IS) | Rh (IS4) 103/103 (IS) | Rh (IS5) 103/103 (IS) | Rh (IS6) 103/103 (IS) | Rh OShift (IS1) 103/119 (IS) | Cd (111) NH3 111/111 Ammonia DRC | Cd (114) NH3 114/114 Ammonia DRC | Pb 208/208 | Pb NH3 208/208 Ammonia DRC
1 | Sac Blank |  | 6/8/2023 1:23:24 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\Sac Blank.356 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth
2 | sac blank cal |  | 6/8/2023 1:29:13 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\sac blank cal.357 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth
3 | CAL BLANK |  | 6/8/2023 1:35:02 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CAL BLANK.358 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | [31 empty] | 0.7181% | 1.453% | 1.271% | 0.3834% | 1.229% | 0.3919% | 1.067%
4 | LOW 0.01 |  | 6/8/2023 1:40:50 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LOW 0.01.359 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | [31 empty] | 1.278% | 1.13% | 1.195% | 1.267% | 0.8522% | 0.2996% | 1.297%
5 | LOW 0.05 |  | 6/8/2023 1:46:39 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LOW 0.05.360 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth |  |  |  |  | 6.275% |  |  |  | 15.65% | 9.644% | 15.52% |  | 14.63% |  | 8.141% | 36.54% | 3.336% | 6.207% | 18.78% | 20.26% | 38.05% | 3.751% | 5.993% | [6 empty] | 1.882% | 6.399% | 0.9218% | 1.429% | 1.689% | 0.6189% | 0.4923% | 2.372% | 0.5563% | 10.23% | 3.698% | 1.258% | 2.079%
6 | LOW 0.1 |  | 6/8/2023 1:52:27 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LOW 0.1.361 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth |  |  |  |  | 10.06% |  |  |  | 15.66% | 9.19% | 2.186% |  | 18.8% |  | 11.14% | 6.094% | 3.516% | 7.804% | 3.797% | 24.66% | 28.29% | 13.07% | 11.49% | 45.53% | 59.69% | 29.68% | 19.84% | 110.8% | 6.609% | 4.997% | 9.975% | 0.8594% | 5.115% | 4.619% | 1% | 1.438% | 1.021% | 7.014% | 2.256% | 2.592% | 5.602% | 2.987%
7 | LOW 0.5 |  | 6/8/2023 1:58:15 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LOW 0.5.362 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth |  |  |  |  | 3.211% |  |  |  | 7.294% | 5.339% | 2.876% | 4.257% | 3.082% | 11.78% | 5.47% | 7.07% | 8.174% | 2.209% | 2.811% | 8.057% | 9.094% | 2.652% | 4.915% | 12.66% | 17.72% | 9.016% | 7.782% | 9.143% | 5.739% | 2.516% | 4.959% | 0.2665% | 3.091% | 5.196% | 1.212% | 0.3302% | 0.4057% | 0.8037% | 0.7357% | 0.2068% | 0.5948% | 0.9214%
8 | STD 1 |  | 6/8/2023 2:04:03 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 1.363 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 5.709% | 9.065% | 3.022% | 32.43% | 1.055% |  | 6.5% | 7.496% | 5.915% | 3.947% | 2.452% | 2.945% | 1.081% | 2.503% | 1.695% | 2.092% | 1.819% | 1.552% | 0.9178% | 6.243% | 18.68% | 5.179% | 2.631% | 16.47% | 8.912% | 6.41% | 7.872% | 8.53% | 8.964% | 1.806% | 1.287% | 0.2087% | 2.187% | 0.8447% | 1.061% | 1.73% | 1.465% | 4.04% | 1.403% | 1.627% | 0.921% | 2.377%
9 | STD 5 |  | 6/8/2023 2:09:52 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 5.364 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4.027% | 3.074% | 4.345% | 4.885% | 2.672% | 7.342% | 3.345% | 4.025% | 2.977% | 1.644% | 0.1979% | 0.6448% | 1.054% | 2.232% | 2.131% | 3.092% | 1.819% | 1.04% | 0.5272% | 4.201% | 4.293% | 0.3876% | 0.05708% | 5.413% | 2.226% | 0.836% | 2.125% | 7.267% | 6.134% | 1.103% | 1.056% | 4.676% | 0.3397% | 0.7155% | 0.5782% | 0.3337% | 2.109% | 0.3813% | 0.2322% | 1.502% | 3.572% | 1.611%
10 | STD 10 |  | 6/8/2023 2:15:40 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 10.365 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.1572% | 1.206% | 1.173% | 6.207% | 0.8737% | 1.878% | 0.8517% | 0.8117% | 0.5883% | 1.355% | 1.682% | 1.956% | 1.721% | 1.396% | 1.961% | 0.9785% | 1.5% | 1.083% | 1.904% | 0.4183% | 3.99% | 1.276% | 1.932% | 1.463% | 1.85% | 5.599% | 4.143% | 3.022% | 3.043% | 2.358% | 1.405% | 0.5486% | 2.039% | 0.9279% | 2.81% | 1.823% | 0.4268% | 1.998% | 2.321% | 2.548% | 0.4662% | 0.9705%
11 | STD 25 |  | 6/8/2023 2:21:28 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 25.366 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.767% | 0.6075% | 1.496% | 2.796% | 1.543% | 2.628% | 0.6525% | 1.458% | 2.148% | 0.5357% | 1.085% | 1.091% | 1.291% | 0.8614% | 2.498% | 2.54% | 1.355% | 0.5547% | 0.9956% | 0.3397% | 1.867% | 5.392% | 0.36% | 1.453% | 2.322% | 1.182% | 0.8205% | 1.561% | 0.7889% | 1.568% | 0.6347% | 0.6048% | 1.045% | 0.5095% | 0.8364% | 4.03% | 1.321% | 4.63% | 0.4667% | 1.32% | 0.8075% | 1.922%
12 | STD 50 |  | 6/8/2023 2:27:16 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 50.367 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.253% | 5.374% | 1.047% | 5.422% | 0.6688% | 4.442% | 3.985% | 1.76% | 1.901% | 3.96% | 1.588% | 1.225% | 3.809% | 3.964% | 3.9% | 4.638% | 3.012% | 2.069% | 2.135% | 1.538% | 2.855% | 4.02% | 3.375% | 2.369% | 5.03% | 4.343% | 4.151% | 2.909% | 3.541% | 1.557% | 2.658% | 1.27% | 3.464% | 2.381% | 1.638% | 1.147% | 0.6748% | 3.989% | 1.12% | 1.662% | 1.068% | 0.713%
13 | STD 100 |  | 6/8/2023 2:33:03 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 100.368 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.076% | 0.8132% | 0.5738% | 1.024% | 1.04% | 1.8% | 1.881% | 1.895% | 1.198% | 1.972% | 1.473% | 2.091% | 1.691% | 1.024% | 1.352% | 1.36% | 1.241% | 0.4601% | 0.8982% | 0.7968% | 1.512% | 1.652% | 0.4559% | 0.4892% | 1.176% | 0.7752% | 0.7085% | 1.16% | 1.199% | 1.454% | 1.804% | 0.6864% | 0.6608% | 0.6473% | 1.208% | 1.065% | 1.602% | 1.482% | 0.7816% | 1.224% | 1.38% | 1.052%
15 | LBLANK |  | 6/8/2023 2:38:52 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.369 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9.396% | 3.545% | 0.03034% | 3.325% | 1.943% | 60.37% | 245.2% | 1.343e-06% | 47.25% | 21.69% | 52.88% | 2.043% | 5.178% | 17.3% | 46.33% | 21.09% | 81.76% | 3.692% | 1.725% | 70.82% | 164.6% | 0.6165% | 18.78% | 58.66% | 32.79% | 48.2% | 52.67% | 20.93% | 119.2% | 7.571% | 1299% | 0.826% | 0.5315% | 2.954% | 0.7769% | 1.662% | 0.9978% | 3.796% | 39.5% | 22.63% | 0.539% | 1.52%
16 | LBLANK |  | 6/8/2023 2:57:55 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.370 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 24.27% | 2.367e-06% | 6.025% | 8.18% | 0.6313% | 342.7% | 111.8% | 1.343e-06% | 32.39% | 98.63% | 42.15% | 1.44% | 0.9836% | 18.58% | 33.25% | 72.66% | 15.52% | 10.82% | 1.397% | 74.3% | 109.9% | 0.6711% | 348.4% | 80.42% | 102.9% | 81.58% | 379.2% | 77.11% | 224.4% | 20.19% | 6.253% | 0.6473% | 0.2683% | 11.92% | 0.8077% | 1.582% | 0.5199% | 1.146% | 93.96% | 13.22% | 11.84% | 5.801%
17 | ICV PE |  | 6/8/2023 3:03:43 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\ICV PE.371 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.223% | 0.8283% | 2.429% | 7.278% | 0.9562% | 4.827% | 2.313% | 3.424% | 1.68% | 1.887% | 1.548% | 0.1172% | 0.6104% | 1.271% | 1.646% | 3.179% | 3.409% | 1.286% | 1.249% | 3.058% | 5.003% | 2.139% | 1.177% | 1.791% | 2.544% | 2.087% | 0.355% | 2.413% | 5.166% | 0.6731% | 0.961% | 1.695% | 0.7044% | 1.85% | 1.344% | 6.61% | 2.238% | 2.261% | 0.7057% | 1.047% | 1.497% | 6.894%
18 | ICV PE |  | 6/8/2023 3:09:59 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\ICV PE.372 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.154% | 2.763% | 1.016% | 13.98% | 1.589% | 2.888% | 3.431% | 4.709% | 5.656% | 3.959% | 1.046% | 3.558% | 0.5038% | 4.834% | 2.037% | 2.63% | 2.991% | 1.795% | 0.7392% | 1.846% | 4.95% | 2.678% | 1.04% | 1.701% | 3.344% | 0.5396% | 2.422% | 0.7387% | 4.983% | 1.248% | 0.943% | 0.7618% | 0.9079% | 1.909% | 0.9095% | 1.115% | 2.288% | 2.925% | 2.22% | 1.873% | 0.02433% | 0.7608%
19 | CCV 1ppb TRACE |  | 6/8/2023 3:15:47 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.373 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.162% | 5.001% | 2.309% | 19.08% | 10.67% | 6.306% | 2.165% | 13.27% | 1.727% | 4.504% | 1.897% | 5.401% | 0.6934% | 8.199% | 3.01% | 2.485% | 2.425% | 1.376% | 2.029% | 3.459% | 12.09% | 0.4992% | 3.414% | 5.736% | 14.98% | 8.319% | 3.773% | 8.458% | 9.872% | 3.394% | 1.22% | 1.514% | 2.122% | 0.9313% | 1.015% | 0.5324% | 1.54% | 1.954% | 3.668% | 2.098% | 1.251% | 0.6195%
20 | CCV 1ppb TRACE |  | 6/8/2023 3:22:03 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.374 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.7425% | 2.604% | 2.975% | 75.11% | 2.806% | 15.97% | 6% | 10.71% | 6.019% | 0.4856% | 3.595% | 2.118% | 2.585% | 13.66% | 1.704% | 7.099% | 3.202% | 2.404% | 3.488% | 4.056% | 15.75% | 3.187% | 1.595% | 7.981% | 8.32% | 6.159% | 3.616% | 2.032% | 13.45% | 3.233% | 1.371% | 0.9699% | 1.165% | 0.4048% | 2.209% | 0.5599% | 0.9466% | 3.065% | 1.588% | 3.928% | 1.893% | 1.754%
21 | LFB |  | 6/8/2023 3:27:51 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.375 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.4299% | 4.166% | 1.49% | 3.203% | 2.599% | 4.081% | 0.39% | 2.886% | 2.349% | 2.248% | 3.542% | 4.888% | 1.18% | 2.608% | 2.204% | 3.981% | 0.5029% | 4.352% | 5.151% | 1.552% | 3.605% | 2.171% | 2.16% | 4.458% | 1.342% | 1.226% | 2.018% | 2.886% | 2.605% | 5.182% | 1.785% | 1.704% | 0.5031% | 0.6252% | 3.826% | 1.025% | 2.003% | 0.9338% | 5.601% | 4.302% | 1.217% | 1.099%
22 | LLOQ 0.01 |  | 6/8/2023 3:33:39 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LLOQ 0.01.376 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 308.6% | 12.31% | 5.451% | 41.13% | 0.9484% | 38.48% | 22.98% | 58.61% | 43.26% | 0.352% | 29.41% | 3.33% | 28.94% | 8.545% | 18.02% | 18.77% | 63.54% | 4.325% | 15.09% | 58.68% | 196.4% | 1.104% | 24.52% | 886.9% | 374.5% | 26.74% | 33.53% | 119.8% | 53.41% | 11.34% | 10.62% | 2.784% | 1.412% | 4.349% | 0.4254% | 1.262% | 0.3261% | 1.502% | 13.96% | 1.238% | 6.908% | 1.477%
23 | CCV 10 ppb MAJORS |  | 6/8/2023 3:39:28 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.377 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 8.007% | 6.53% | 0.05544% | 2.902% | 0.2924% | 1.745% | 85.87% | 10.78% | 24.56% | 84.48% | 75.38% | 0.743% | 0.5204% | 1.197% | 58.66% | 345.7% | 19.23% | 3.912% | 0.5899% | 0.559% | 1.756e-06% | 0.4532% | 22.48% | 2.887% | 37.89% | 198% | 324.1% | 50.84% | 116.1% | 2.355% | 0.6336% | 2.953% | 3.68% | 2.815% | 2.368% | 3.804% | 4.976% | 6.926% | 2.529% | 15.37% | 1.52% | 4.686%
24 | CCV 10 ppb MAJORS |  | 6/8/2023 3:45:42 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.378 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.924% | 3.826% | 1.675% | 7.245% | 0.858% | 5.07% | 4.18% | 7.165% | 5.051% | 3.04% | 0.3671% | 2.073% | 1.539% | 5.295% | 2.08% | 3.351% | 3.396% | 0.5702% | 1.153% | 1.958% | 5.212% | 1.719% | 0.582% | 2.74% | 2.749% | 3.096% | 2.426% | 2.537% | 5.963% | 1.111% | 0.6585% | 0.8399% | 0.918% | 2.645% | 1.254% | 0.8326% | 2.677% | 1.684% | 2.287% | 2.446% | 1.225% | 0.7778%
25 | LLOQ 0.1 |  | 6/8/2023 3:51:31 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LLOQ 0.1.379 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10% | 18.83% | 7.743% | 257.7% | 11.72% | 53.16% | 25.89% | 4.92% | 6.66% | 2.601% | 4.678% | 19.24% | 2.346% | 65.44% | 6.931% | 19.68% | 9.698% | 1.077% | 1.126% | 25.91% | 28.79% | 2.059% | 5.295% | 27.78% | 31.83% | 23.22% | 24.96% | 23.65% | 26.52% | 4.191% | 3.158% | 1.363% | 0.7574% | 1.414% | 0.5907% | 1.786% | 2.48% | 3.45% | 5.915% | 4.322% | 2.776% | 1.306%
26 | STD 0.5 |  | 6/8/2023 3:57:19 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 0.5.380 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.801% | 4.943% | 5.989% | 24.64% | 27.75% | 37.57% | 3.652% | 16.39% | 10.59% | 2.397% | 7.09% | 10.29% | 3.21% | 6.437% | 5.601% | 11.09% | 6.825% | 9.013% | 7.899% | 5.014% | 12.01% | 3.493% | 2.523% | 16.65% | 9.64% | 7.609% | 7.246% | 23.11% | 13.91% | 6.491% | 2.662% | 0.5691% | 1.706% | 1.928% | 5.45% | 0.3688% | 1.667% | 2.241% | 7.13% | 7.916% | 0.2498% | 2.513%
27 | 8_2_21_UF_1 |  | 6/8/2023 4:03:07 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_UF_1.381 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.7586% | 3.764% | 1.983% | 13.05% | 1.236% | 1.409% | 1.197% | 1.432% | 8.032% | 9.776% | 11.99% | 4.457% | 1.388% | 19.51% | 26.12% | 10.23% | 4.184% | 0.7583% | 1.159% | 2.543% | 4.38% | 1.939% | 2.275% | 42.74% | 114.7% | 21.47% | 4.419% | 15.04% | 47.78% | 0.613% | 1.045% | 1.21% | 0.4188% | 1.712% | 1.335% | 0.7152% | 2.049% | 0.07906% | 27.26% | 5.274% | 4.963% | 4.463%
28 | 8_2_21_UF_6 |  | 6/8/2023 4:08:55 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_UF_6.382 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.872% | 1.636% | 17.12% | 124.6% | 3.92% | 7.225% | 7.849% | 3.156% | 3.349% | 18.01% | 6.545% | 10.47% | 1.643% | 7.665% | 2.769% | 1.229% | 4.863% | 3.09% | 1.902% | 0.9884% | 11.1% | 3.043% | 2.105% | 21.03% | 8.63% | 11.78% | 22.2% | 13.04% | 31.48% | 2.829% | 0.8997% | 0.6371% | 0.02132% | 0.9061% | 1.816% | 1.39% | 0.8845% | 2.885% | 12% | 8.766% | 10.66% | 3.972%
29 | 8_2_21_UF_6_SP |  | 6/8/2023 4:14:43 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_UF_6_SP.383 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.487% | 4.31% | 1.488% | 2.321% | 6.489% | 3.105% | 3.779% | 1.607% | 1.624% | 3.01% | 2.568% | 3.463% | 0.8685% | 6.304% | 1.23% | 0.4725% | 1.19% | 1.749% | 1.409% | 1.12% | 2.508% | 2.947% | 0.9964% | 4.035% | 4.908% | 2.908% | 2.916% | 5.224% | 4.047% | 1.553% | 1.867% | 0.5132% | 0.1651% | 0.8316% | 1.552% | 1.917% | 5.294% | 1.965% | 2.354% | 1.769% | 0.5368% | 1.229%
30 | STD 1 |  | 6/8/2023 4:20:31 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 1.384 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.641% | 4.202% | 2.209% | 18.94% | 81.09% | 7.407% | 3.877% | 5.469% | 3.29% | 7.782% | 1.362% | 20.66% | 1.543% | 7.443% | 3.472% | 2.893% | 3.557% | 1.571% | 3.07% | 4.288% | 4.507% | 3.689% | 1.049% | 7.311% | 7.43% | 10.9% | 6.745% | 14.44% | 7.036% | 2.332% | 1.175% | 1.884% | 0.5413% | 2.148% | 1.09% | 1.041% | 1.707% | 5.861% | 1.916% | 1.929% | 3.645% | 1.468%
31 | 8_2_21_UF_6_SP_LD |  | 6/8/2023 4:26:19 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_UF_6_SP_LD.385 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4.416% | 3.781% | 1.743% | 11.09% | 0.9125% | 0.7852% | 4.985% | 2.346% | 1.937% | 1.765% | 0.4579% | 1.291% | 2.46% | 1.908% | 3.897% | 0.4767% | 2.437% | 0.7925% | 1.666% | 1.364% | 4.316% | 2.364% | 0.5882% | 0.2519% | 6.387% | 3.5% | 1.132% | 1.988% | 4.401% | 1.153% | 1.941% | 0.5446% | 0.8162% | 1.402% | 1.325% | 1.871% | 2.392% | 2.015% | 1.524% | 0.9371% | 1.307% | 1.489%
32 | 8_2_21_UF_6_SP_LSP |  | 6/8/2023 4:32:06 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_UF_6_SP_LSP.386 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.9675% | 1.472% | 3.206% | 10.68% | 1.922% | 0.1852% | 1.02% | 1.74% | 1.412% | 0.7039% | 0.6872% | 1.865% | 1.076% | 1.712% | 1.124% | 0.6035% | 1.194% | 1.289% | 0.6761% | 3.493% | 2.683% | 0.5734% | 0.7237% | 4.449% | 3.422% | 1.165% | 2.079% | 2.938% | 4.056% | 0.5877% | 0.5213% | 2.059% | 0.722% | 0.9421% | 0.9866% | 0.8148% | 0.9554% | 0.5294% | 1.146% | 0.6828% | 2.761% | 0.7277%
33 | 8_2_21_UF_8 |  | 6/8/2023 4:37:54 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_UF_8.387 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.873% | 2.457% | 6.413% | 287.5% | 1.025% | 5.717% | 1.049% | 4.954% | 2.691% | 7.496% | 2.493% | 3.569% | 0.8306% | 3.129% | 12.67% | 1.663% | 1.536% | 1.087% | 0.5794% | 1.486% | 10.63% | 2.967% | 0.8158% | 11.06% | 12.47% | 13.5% | 13.65% | 5.538% | 12.71% | 1.965% | 1.181% | 0.7286% | 0.1808% | 1.148% | 1.363% | 1.293% | 2.475% | 3.389% | 7.87% | 6.496% | 2.032% | 2.124%
34 | 8_2_21_UF_9 |  | 6/8/2023 4:43:42 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_UF_9.388 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.166% | 0.9676% | 12.07% | 181.9% | 0.4299% | 2.114% | 3.894% | 3.782% | 4.217% | 12.38% | 7.075% | 6.656% | 1.255% | 31.07% | 25.32% | 11.55% | 2.627% | 0.969% | 0.5814% | 2.085% | 7.035% | 1.145% | 1.532% | 15.06% | 59.92% | 29.12% | 22.21% | 6.284% | 24.89% | 1.594% | 0.2728% | 0.7503% | 0.7306% | 2.566% | 0.6562% | 0.7723% | 1.037% | 1.547% | 6.926% | 8.655% | 2.236% | 4.709%
35 | LBLANK |  | 6/8/2023 4:49:31 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.389 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 26.69% | 2.367e-06% | 0.4326% | 6.284% | 1.48% | 36.04% | 468.7% | 13.73% | 2.97% | 278.5% | 40.48% | 1.069% | 1.059% | 13.36% | 51.16% | 63.16% | 21.47% | 16.77% | 4.424% | 204.8% | 119.5% | 1.71% | 9.078% | 88.62% | 1342% | 48.76% | 39.8% | 406.4% | 48.76% | 9.044% | 7.349% | 0.6767% | 1.232% | 7.964% | 1.232% | 0.2253% | 0.9105% | 0.5329% | 110.6% | 34.45% | 3.21% | 6.163%
36 | CCV 1ppb TRACE |  | 6/8/2023 4:55:19 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.390 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.481% | 8.255% | 5.117% | 24.65% | 1.263% | 22.47% | 3.243% | 18.88% | 2.906% | 6.405% | 3.065% | 0.9744% | 1.944% | 20.8% | 4.163% | 6.897% | 4.492% | 0.8384% | 1.579% | 3.061% | 10.91% | 2.83% | 1.747% | 6.434% | 20.95% | 1.344% | 3.271% | 10.45% | 18.29% | 1.312% | 0.9194% | 1.163% | 0.909% | 3.291% | 0.5874% | 3.356% | 1.97% | 2.573% | 2.765% | 1.843% | 1.96% | 1.377%
37 | CCV 1ppb TRACE |  | 6/8/2023 5:01:37 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.391 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.657% | 5.919% | 2.587% | 24.3% | 4.86% | 12.09% | 9.893% | 4.535% | 4.429% | 5.222% | 0.9794% | 8.135% | 0.6665% | 8.366% | 1.646% | 3.511% | 1.756% | 2.095% | 1.308% | 2.774% | 8.875% | 1.071% | 2.712% | 2.807% | 5.17% | 6.982% | 3.569% | 11.4% | 12.91% | 2.18% | 1.317% | 1.534% | 0.7503% | 1.252% | 0.8184% | 0.7491% | 1.883% | 2.468% | 1.776% | 1.291% | 0.8901% | 0.335%
38 | CCV 10 ppb MAJORS |  | 6/8/2023 5:07:25 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.392 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.398% | 2.301% | 2.136% | 4.211% | 2.162% | 5.318% | 3.828% | 4.548% | 1.319% | 2.001% | 0.6263% | 2.353% | 2.792% | 1.879% | 1.781% | 2.393% | 0.7663% | 1.571% | 0.7093% | 2.017% | 6.087% | 1.938% | 1.772% | 5.454% | 6.628% | 2.415% | 2.156% | 2.226% | 2.962% | 1.783% | 1.95% | 2.064% | 1.555% | 1.419% | 0.6118% | 1.09% | 1.704% | 1.342% | 0.8862% | 1.062% | 2.642% | 0.7133%
39 | CCV 10 ppb MAJORS |  | 6/8/2023 5:13:42 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.393 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.163% | 3.52% | 2.496% | 1.908% | 3.549% | 3.049% | 3.207% | 3.507% | 2.41% | 1.684% | 1.536% | 2.545% | 1.108% | 1.507% | 1.972% | 0.9588% | 1.03% | 0.3916% | 0.6404% | 2.389% | 2.378% | 2.892% | 2.14% | 3.465% | 3.7% | 1.002% | 4.183% | 3.177% | 0.733% | 0.5248% | 1.059% | 1.514% | 0.9778% | 1.687% | 0.4846% | 0.4652% | 1.479% | 2.845% | 1.114% | 0.7689% | 1.817% | 2.85%
40 | 8_2_21_UF_10 |  | 6/8/2023 5:19:31 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_UF_10.394 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 5.511% | 3.424% | 12.16% | 127.7% | 0.6614% | 7.122% | 6.135% | 5.006% | 4.319% | 8.246% | 8.539% | 3.426% | 0.3492% | 20.37% | 28.15% | 13.89% | 3.328% | 0.181% | 0.7982% | 0.7264% | 4.582% | 1.55% | 0.855% | 13.03% | 13.46% | 37.69% | 20.8% | 24.01% | 11.9% | 2.054% | 0.4121% | 1.627% | 0.9795% | 5.576% | 0.8848% | 0.3215% | 0.8424% | 1.95% | 5.251% | 15.19% | 6.266% | 1.907%
41 | 8_2_21_UF_11 |  | 6/8/2023 5:25:19 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_UF_11.395 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 19.21% | 22.16% | 14.58% | 128.4% | 2.319% | 22.18% | 15.86% | 22.9% | 21.57% | 20.45% | 2.332% | 2.641% | 0.6581% | 23.63% | 5.6% | 20.27% | 20.99% | 1.07% | 2.663% | 0.5786% | 19.84% | 2.188% | 0.8812% | 17.26% | 59.75% | 9.563% | 7.623% | 22.05% | 36.02% | 1.99% | 1.678% | 0.4756% | 1.187% | 17.91% | 0.5905% | 1.555% | 1.831% | 2.621% | 7.705% | 6.741% | 1.823% | 3.088%
42 | 8_2_21_UF_11_DUP |  | 6/8/2023 5:31:09 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_UF_11_DUP.396 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.486% | 1.627% | 8.878% | 513% | 2.606% | 2.518% | 4.44% | 1.703% | 3.731% | 5.247% | 10.76% | 66.52% | 2.257% | 11.65% | 11.9% | 3.962% | 1.808% | 2.882% | 1.051% | 1.08% | 7.865% | 2.911% | 1.374% | 13.55% | 49.56% | 10.17% | 13.37% | 6.133% | 22.04% | 1.041% | 1.907% | 0.4259% | 1.515% | 1.876% | 1.526% | 0.4075% | 1.775% | 2.95% | 13.5% | 14.94% | 2.534% | 1.178%
43 | 8_2_21_F_1 |  | 6/8/2023 5:36:57 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_F_1.397 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.507% | 0.6963% | 2.578% | 3.035% | 3.486% | 2.213% | 4.066% | 0.8626% | 6.551% | 8.652% | 9.776% | 1.652% | 1.312% | 5.549% | 3.698% | 7.269% | 5.081% | 0.1523% | 0.4185% | 1.469% | 11.59% | 4.635% | 0.6284% | 45.83% | 108.9% | 17.28% | 37.83% | 10.36% | 48.94% | 0.0528% | 1.023% | 1.102% | 0.7922% | 2.018% | 0.5353% | 0.2488% | 2.767% | 3.616% | 18.52% | 23.95% | 6.716% | 1.942%
44 | 8_2_21_F_6 |  | 6/8/2023 5:44:00 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_F_6.398 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.8112% | 3.373% | 4.004% | 38.65% | 4.385% | 1.977% | 2.049% | 3.509% | 2.36% | 4.027% | 0.4673% | 0.7631% | 18.57% | 1.199% | 19.56% | 1.947% | 0.6562% | 0.2068% | 1.316% | 1.848% | 6.221% | 18.13% | 19.2% | 4.661% | 49.6% | 6.517% | 15.12% | 5.998% | 32.68% | 1.086% | 18.09% | 2.107% | 16.96% | 1.116% | 0.5041% | 0.4939% | 0.7912% | 16.41% | 4.248% | 10.31% | 2.777% | 0.4443%
45 | 8_2_21_F_8 |  | 6/8/2023 5:49:48 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_F_8.399 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.834% | 0.623% | 3.451% | 19.57% | 1.248% | 3.457% | 3.838% | 0.5094% | 5.906% | 0.4875% | 4.787% | 0.5265% | 6.456% | 4.857% | 16.35% | 9.475% | 1.748% | 0.1126% | 1.543% | 0.7869% | 7.539% | 3.757% | 5.088% | 13.82% | 56.93% | 31.02% | 20.24% | 10.42% | 37.65% | 2.084% | 6.609% | 2.065% | 4.363% | 0.4728% | 0.7762% | 0.5153% | 1.324% | 2.711% | 14.47% | 12.75% | 3.865% | 3.192%
46 | 8_2_21_F_8_DUP |  | 6/8/2023 5:55:36 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_F_8_DUP.400 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 20.9% | 22.42% | 3.654% | 29% | 1.414% | 28.91% | 18.69% | 23.33% | 21.21% | 18.52% | 5.062% | 4.085% | 1.888% | 26.28% | 7.764% | 9.096% | 20.36% | 1.487% | 1.146% | 0.7367% | 27.9% | 1.487% | 1.281% | 19.45% | 38.92% | 18.48% | 4.045% | 12.31% | 34.27% | 2.732% | 2.338% | 1.061% | 1.065% | 16.67% | 0.6848% | 0.3242% | 1.959% | 1.914% | 6.828% | 11.2% | 2.956% | 0.4421%
47 | 8_2_21_F_8_DUP_LD |  | 6/8/2023 6:01:24 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_F_8_DUP_LD.401 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.796% | 1.72% | 1.504% | 17.41% | 2.436% | 2.764% | 2.323% | 2.713% | 5.665% | 7.607% | 6.076% | 7.896% | 1.247% | 1.387% | 17.23% | 6.795% | 3.862% | 7.879% | 8.894% | 1.116% | 4.694% | 1.407% | 2.777% | 12.59% | 20.37% | 18.94% | 20.09% | 4.076% | 37.59% | 8.389% | 3.118% | 1.757% | 1.732% | 0.4066% | 7.31% | 1.44% | 1.493% | 2.337% | 19.67% | 22.03% | 0.06468% | 2.861%
48 | 8_2_21_F_8_DUP_LD |  | 6/8/2023 6:07:40 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_F_8_DUP_LD.402 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.509% | 3.788% | 2.15% | 6.066% | 2.3% | 5.551% | 4.769% | 5.022% | 1.251% | 22.58% | 6.987% | 2.97% | 1.315% | 6.762% | 10.23% | 6.648% | 1.274% | 2.786% | 4.365% | 0.9779% | 4.98% | 0.7934% | 0.5947% | 49.47% | 42.83% | 33.51% | 8.341% | 5.762% | 23.38% | 2.183% | 0.957% | 0.4498% | 0.5478% | 1.953% | 2.404% | 0.8697% | 0.7976% | 0.9844% | 17.49% | 5.179% | 2.423% | 2.547%
49 | 8_2_21_F_8_DUP_LSP |  | 6/8/2023 6:13:28 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_F_8_DUP_LSP.403 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.7742% | 0.3393% | 1.026% | 6.37% | 0.6693% | 1.59% | 0.7786% | 1.466% | 2.043% | 0.8003% | 0.7972% | 0.9252% | 1.027% | 2.955% | 1.528% | 1.722% | 1.174% | 0.7272% | 0.4149% | 1.182% | 4.357% | 1.934% | 0.4973% | 1.08% | 3.974% | 1.533% | 2.229% | 1.826% | 0.8757% | 0.4898% | 0.5241% | 1.162% | 0.03388% | 0.7448% | 0.4913% | 0.8345% | 1.718% | 2.177% | 2.057% | 1.221% | 0.829% | 1.051%
50 | LBLANK |  | 6/8/2023 6:19:16 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.404 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 37.58% | 4.333% | 0.08046% | 6.288% | 1.602% | 21.77% | 185.4% | 1.343e-06% | 29.8% | 23.12% | 46.7% | 0.9067% | 2.148% | 7.656% | 29.65% | 47.61% | 30.89% | 10.38% | 1.264% | 115.8% | 175.9% | 0.6536% | 21.33% | 60.7% | 0% | 211.6% | 690.9% | 254.8% | 1.054% | 17.55% | 3.078% | 0.4256% | 0.915% | 0.2225% | 1.344% | 2.11% | 0.9712% | 2.467% | 109.5% | 239% | 4.751% | 6.632%
51 | CCV 1ppb TRACE |  | 6/8/2023 6:25:05 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.405 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.331% | 8.691% | 1.699% | 43.67% | 2.929% | 6.333% | 6.957% | 10.58% | 9.599% | 2.52% | 1.772% | 2.354% | 3.355% | 4.001% | 5.016% | 7.016% | 2.914% | 1.615% | 3.697% | 4.285% | 24.01% | 6.136% | 2.489% | 8.225% | 5.311% | 3.005% | 1.337% | 14.44% | 5.332% | 1.43% | 0.5322% | 1.242% | 0.9058% | 3.012% | 2.135% | 1.906% | 1.345% | 4.724% | 1.55% | 3.184% | 0.9004% | 2.085%
52 | CCV 1ppb TRACE |  | 6/8/2023 6:31:23 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.406 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.46% | 5.725% | 15.48% | 21.47% | 5.102% | 18.01% | 5.245% | 13.64% | 4.191% | 3.896% | 2.844% | 3.08% | 1.599% | 7.866% | 3.734% | 3.788% | 1.877% | 1.785% | 3.312% | 11.78% | 5.957% | 1.98% | 0.9404% | 0.2673% | 9.309% | 0.2283% | 9.394% | 13.28% | 1.191% | 1.693% | 0.8009% | 6.156% | 1.851% | 1.638% | 0.8231% | 1.744% | 0.9822% | 0.5946% | 1.274% | 2.246% | 6.379% | 2.392%
53 | LFB |  | 6/8/2023 6:37:11 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.407 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.853% | 0.4268% | 1.647% | 5.207% | 0.9278% | 4.228% | 3.691% | 3.677% | 1.671% | 2.105% | 0.9434% | 1.748% | 2.922% | 1.424% | 0.8571% | 1.151% | 2.268% | 1.077% | 1.741% | 1.479% | 5.035% | 3.296% | 0.8654% | 1.933% | 5.543% | 4.132% | 4.198% | 0.6829% | 2.537% | 0.6451% | 1.835% | 0.9397% | 1.322% | 0.8294% | 0.7835% | 1.203% | 1.693% | 2.628% | 0.3681% | 0.9523% | 1.54% | 1.922%
54 | CCV 10 ppb MAJORS |  | 6/8/2023 6:42:59 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.408 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.112% | 4.906% | 1.872% | 10.72% | 1.448% | 9.433% | 3.579% | 4.774% | 4.357% | 3.016% | 1.021% | 1.484% | 0.7007% | 5.581% | 3.58% | 4.051% | 3.982% | 1.078% | 0.2076% | 1.703% | 7.327% | 3.373% | 0.8733% | 3.124% | 4.14% | 0.3015% | 1.247% | 2.745% | 1.721% | 1.023% | 1.688% | 1.526% | 0.3053% | 3.025% | 0.2478% | 0.4219% | 1.176% | 2.545% | 0.8978% | 0.5974% | 0.7201% | 1.521%
55 | CCV 10 ppb MAJORS |  | 6/8/2023 6:49:21 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.409 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 5.82% | 4.912% | 18.26% | 13.05% | 1.272% | 6.42% | 4.533% | 2.793% | 3.298% | 5.645% | 3.362% | 1.295% | 1.617% | 6.732% | 6.421% | 5.501% | 5.538% | 1.493% | 0.9544% | 15.33% | 7.289% | 1.999% | 1.809% | 16.68% | 3.088% | 1.778% | 1.037% | 16.24% | 1.469% | 0.7366% | 1.877% | 12.12% | 1.358% | 5.813% | 0.5542% | 0.6777% | 0.8079% | 1.634% | 0.6119% | 0.7003% | 14.85% | 2.218%
56 | LBLANK |  | 6/8/2023 7:26:31 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.410 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10.54% | 2.367e-06% | 0.03935% | 0.04888% | 0.3144% | 208.6% | 99.24% | 1.343e-06% | 31.25% | 111.3% | 13.28% | 1.264% | 2.646% | 6.75% | 40.06% | 69.73% | 31.26% | 12.98% | 7.173% | 35.74% | 4318% | 0.9715% | 240.9% | 52.75% | 44.09% | 117.5% | 239.1% | 126.7% | 0% | 7.142% | 8.137% | 2.344% | 1.316% | 0.7243% | 1.651% | 0.5655% | 0.742% | 4.12% | 109% | 123.1% | 2.604% | 13.99%
57 | PE ICV |  | 6/8/2023 7:32:19 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\PE ICV.411 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.888% | 2.989% | 2.786% | 5.92% | 1.687% | 1.728% | 0.3544% | 5.468% | 1.412% | 0.8395% | 0.8167% | 2.953% | 1.85% | 1.566% | 0.663% | 3.708% | 1.307% | 0.6357% | 1.208% | 2.829% | 2.919% | 5.926% | 0.5978% | 3.498% | 5.171% | 1.741% | 1.145% | 1.802% | 2.607% | 0.5339% | 0.855% | 0.6745% | 1.012% | 0.3535% | 0.5433% | 1.359% | 0.9893% | 5.248% | 1.501% | 1.209% | 1.556% | 1.506%
58 | 8_2_21_F_9 |  | 6/8/2023 7:38:07 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_F_9.412 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.341% | 2.71% | 1.635% | 7.542% | 1.253% | 1.595% | 2.773% | 1.472% | 5.201% | 10.93% | 1.875% | 2.628% | 1.706% | 3.747% | 10.75% | 2.141% | 3.532% | 1.924% | 0.9329% | 1.876% | 6.074% | 4.197% | 0.7183% | 22.96% | 82.09% | 10.35% | 29.71% | 33.9% | 22.24% | 0.6689% | 1.55% | 0.8986% | 1.027% | 2.668% | 0.5173% | 1.548% | 0.4929% | 4.268% | 24.72% | 11.13% | 1.349% | 0.3873%
59 | 8_2_21_F_10 |  | 6/8/2023 7:43:55 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_F_10.413 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10.94% | 7.09% | 1.271% | 3.34% | 2.251% | 12.73% | 6.346% | 10.01% | 10.49% | 5.895% | 11.23% | 2.747% | 0.7871% | 11.99% | 9.132% | 6.497% | 7.727% | 1.303% | 1.242% | 0.8435% | 4.371% | 1.962% | 0.5673% | 4.213% | 86.58% | 21.17% | 14.82% | 17.7% | 24.89% | 0.7901% | 1.719% | 1.2% | 0.2992% | 8.127% | 0.7232% | 0.7729% | 1.136% | 1.7% | 28.42% | 13.09% | 3.708% | 1.783%
60 | 8_2_21_F_10_DUP |  | 6/8/2023 7:49:43 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_F_10_DUP.414 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.178% | 2.22% | 2.768% | 3.029% | 1.704% | 2.074% | 3.348% | 2.285% | 4.352% | 4.088% | 1.664% | 2.538% | 9.477% | 1.621% | 22.74% | 4.452% | 1.379% | 1.23% | 1.119% | 0.7977% | 3.687% | 8.128% | 7.917% | 12.96% | 27% | 29.05% | 9.612% | 27.33% | 45.21% | 1.461% | 7.956% | 0.7766% | 7.909% | 1.151% | 0.4265% | 1.137% | 1.019% | 8.433% | 31.75% | 6.837% | 0.1897% | 2.262%
61 | 8_2_21_F_11 |  | 6/8/2023 7:55:32 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_F_11.415 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4.127% | 2.997% | 1.1% | 11.49% | 2.787% | 0.9975% | 8.34% | 0.3751% | 4.486% | 4.805% | 8.686% | 2.029% | 0.8366% | 0.4694% | 20.11% | 11.05% | 1.133% | 1.643% | 2.223% | 2.239% | 4.211% | 1.752% | 1.2% | 19.15% | 51.57% | 22% | 10.8% | 20.36% | 33.72% | 0.8168% | 0.9949% | 1.377% | 0.4379% | 0.7896% | 1.174% | 3.282% | 2.071% | 1.158% | 35.42% | 15.29% | 0.9764% | 4.01%
62 | 8_2_21_F_11_SP |  | 6/8/2023 8:01:20 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_F_11_SP.416 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.294% | 1.493% | 2.992% | 3.19% | 1.664% | 1.398% | 3.29% | 0.4992% | 0.6853% | 1.955% | 1.212% | 1.927% | 1.708% | 2.861% | 1.986% | 2.014% | 3.737% | 1.502% | 1.301% | 3.146% | 3.906% | 1.198% | 1.158% | 2.416% | 3.918% | 3.884% | 2.951% | 3.885% | 4.557% | 1.47% | 1.669% | 2.49% | 1.337% | 2.096% | 0.2242% | 0.5709% | 1.169% | 1.045% | 2.448% | 1.073% | 2.467% | 1.561%
63 | 8_2_21_W_1 |  | 6/8/2023 8:07:08 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_W_1.417 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.894% | 1.992% | 0.6638% | 6.58% | 7.162% | 6.079% | 5.995% | 4.228% | 5.323% | 7.213% | 4.259% | 0.7548% | 2.294% | 0.5671% | 24.34% | 22.46% | 4.593% | 1.522% | 6.275% | 5.831% | 24.6% | 1.579% | 2.156% | 61.27% | 269.3% | 27.27% | 42.5% | 72.82% | 45.37% | 1.532% | 0.9048% | 0.2326% | 1.946% | 0.5275% | 0.8372% | 0.6291% | 1.08% | 1.133% | 27.38% | 22.29% | 4.934% | 1.97%
64 | 8_2_21_W_6 |  | 6/8/2023 8:12:56 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_W_6.418 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9.384% | 6.502% | 1.766% | 3.783% | 3.84% | 5.813% | 4.15% | 4.962% | 6.391% | 27.12% | 3.999% | 0.5425% | 1.905% | 0.7556% | 24.07% | 30.16% | 3.803% | 2.252% | 0.165% | 4.032% | 15.6% | 3.396% | 1.526% | 17.7% | 144.6% | 45.06% | 13.29% | 103% | 45.71% | 0.9091% | 1.994% | 0.8327% | 1.807% | 2.276% | 1.053% | 1.456% | 1.445% | 1.505% | 28.53% | 17.65% | 3.324% | 0.6556%
65 | 8_2_21_W_6_DUP |  | 6/8/2023 8:18:44 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_W_6_DUP.419 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.6198% | 5.303% | 0.3709% | 8.626% | 4.925% | 7.704% | 10.4% | 2.197% | 6.123% | 4.336% | 5.518% | 0.8233% | 1.337% | 4.259% | 35.06% | 2.162% | 2.825% | 1.386% | 4.051% | 3.008% | 22.38% | 1.295% | 0.5506% | 98.84% | 127.1% | 18.27% | 31.85% | 84.2% | 105.9% | 1.63% | 2.131% | 2.209% | 0.3325% | 1.169% | 2.007% | 0.5503% | 0.5122% | 2.329% | 4.561% | 19.08% | 1.644% | 2.557%
66 | LBLANK |  | 6/8/2023 8:24:33 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.420 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 5.704% | 2.367e-06% | 0.8827% | 7.858% | 0.5153% | 33.21% | 66.34% | 1.343e-06% | 53.21% | 37.82% | 7.727% | 1.973% | 1.455% | 13.45% | 40.14% | 275.4% | 41.37% | 9.844% | 1.936% | 18.75% | 3093% | 0.7996% | 149.1% | 22.16% | 43.78% | 198.5% | 165% | 374.1% | 0% | 6.011% | 3.778% | 1.089% | 1.463% | 2.338% | 1.806% | 0.3488% | 1.454% | 1.954% | 108.8% | 77.68% | 13.71% | 2.259%
67 | CCV 1ppb TRACE |  | 6/8/2023 8:30:21 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.421 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4.784% | 2.119% | 8.358% | 47.74% | 1.377% | 6.556% | 4.448% | 9.7% | 6.646% | 3.352% | 0.7191% | 2.422% | 3.199% | 9.266% | 1.855% | 0.2962% | 6.259% | 1.079% | 3.591% | 4.944% | 4.509% | 0.525% | 1.33% | 16.38% | 11.62% | 7.884% | 8.88% | 7.251% | 4.437% | 1.609% | 1.764% | 4.546% | 1.629% | 2.687% | 0.3267% | 0.4631% | 0.595% | 0.7419% | 4.182% | 0.888% | 6.799% | 0.4381%
68 | CCV 1ppb TRACE |  | 6/8/2023 8:36:39 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.422 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.278% | 1.157% | 1.659% | 16.88% | 3.089% | 17.18% | 5.519% | 6.396% | 1.058% | 4.257% | 3.77% | 1.954% | 1.575% | 7.836% | 4.597% | 4.591% | 0.5582% | 1.714% | 0.4569% | 4.39% | 12.42% | 1.947% | 1.352% | 12.07% | 8.007% | 3.385% | 3.045% | 6.142% | 2.949% | 2.42% | 1.228% | 0.8958% | 1.013% | 1.382% | 0.6356% | 1.164% | 2.832% | 3.492% | 2.5% | 1.356% | 1.968% | 0.6185%
69 | CCV 10 ppb MAJORS |  | 6/8/2023 8:42:27 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.423 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9.652% | 10.23% | 1.133% | 20.94% | 1.913% | 10.51% | 8.22% | 8.157% | 8.658% | 11.09% | 0.6992% | 2.692% | 1.552% | 12.74% | 12.59% | 9.952% | 10.69% | 0.6192% | 1.034% | 3.569% | 9.276% | 1.619% | 2.043% | 1.661% | 3.644% | 3.145% | 1.369% | 2.275% | 3.362% | 0.9512% | 1.661% | 1.809% | 0.6531% | 8.73% | 0.5137% | 1.389% | 1.196% | 1.013% | 1.47% | 1.216% | 2.054% | 1.159%
70 | CCV 10 ppb MAJORS |  | 6/8/2023 8:48:44 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.424 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.531% | 1.324% | 2.353% | 5.286% | 0.4586% | 1.417% | 2.841% | 3.097% | 3.616% | 3.01% | 1.803% | 2.595% | 0.6656% | 5.096% | 1.286% | 1.73% | 2.156% | 1.508% | 1.473% | 2.335% | 6.449% | 0.4683% | 0.5822% | 3.065% | 4.766% | 1.313% | 2.078% | 5.488% | 1.611% | 1.623% | 0.745% | 2.857% | 0.5105% | 1.842% | 1.205% | 6.545% | 1.067% | 0.9986% | 1.333% | 0.7408% | 2.958% | 10.63%
71 | 8_2_21_W_8 |  | 6/8/2023 8:54:33 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_W_8.425 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.494% | 2.861% | 1.993% | 3.992% | 1.999% | 5.491% | 4.73% | 2.669% | 10.44% | 9.899% | 5.796% | 4.94% | 1.534% | 2.46% | 1.584% | 22.91% | 2.532% | 5.641% | 7.316% | 1.66% | 8.677% | 3.332% | 0.2755% | 28.33% | 1002% | 39.87% | 28.47% | 64.27% | 22.37% | 4.776% | 1.24% | 0.5854% | 0.6998% | 2.047% | 4.571% | 0.5382% | 1.08% | 3.225% | 40.71% | 26.8% | 3.789% | 2.279%
72 | 8_2_21_W_9 |  | 6/8/2023 9:00:21 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_W_9.426 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.652% | 2.37% | 1.603% | 11.86% | 4.238% | 9.79% | 13.06% | 3.11% | 3.21% | 11.03% | 1.617% | 3.956% | 2.604% | 2.67% | 16.51% | 3.657% | 5.788% | 2.731% | 5.885% | 4.825% | 8.003% | 0.9907% | 2.421% | 67.62% | 247.9% | 57.8% | 72.94% | 1581% | 114.2% | 2.363% | 1.281% | 0.9821% | 0.7714% | 2.336% | 2.308% | 2.456% | 1.057% | 2.274% | 20.57% | 28.26% | 2.535% | 2.001%
73 | 8_2_21_W_9_DUP |  | 6/8/2023 9:06:09 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_W_9_DUP.427 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.536% | 2.599% | 2.762% | 8.566% | 2.758% | 17.97% | 3.464% | 3.995% | 16.03% | 24.44% | 24.73% | 0.1683% | 0.6653% | 1.251% | 9.656% | 22.38% | 2.608% | 1.532% | 2.372% | 6.009% | 17.57% | 7.131% | 3.816% | 23.38% | 59.88% | 8.429% | 39.62% | 216% | 103.7% | 0.7935% | 2.834% | 2.159% | 1.14% | 0.4551% | 0.1802% | 0.19% | 0.8813% | 5.538% | 27.91% | 32.1% | 9.073% | 2.761%
74 | 8_2_21_W_10 |  | 6/8/2023 9:11:57 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_W_10.428 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 6.047% | 2.191% | 1.598% | 7.683% | 3.035% | 8.581% | 9.016% | 1.088% | 4.79% | 32.05% | 4.937% | 1.751% | 1.503% | 3.227% | 4.435% | 32.43% | 6.25% | 0.9124% | 1.594% | 2.361% | 12.45% | 1.236% | 3.536% | 26.34% | 201.2% | 56.64% | 16.1% | 65.92% | 48.17% | 0.8145% | 1.385% | 1.147% | 1.297% | 1.335% | 1.083% | 0.7649% | 2.394% | 2.133% | 16.45% | 9.87% | 3.326% | 4.329%
75 | 8_2_21_W_11 |  | 6/8/2023 9:17:45 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_W_11.429 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4.39% | 3.964% | 0.7892% | 4.414% | 3.742% | 17.18% | 7.571% | 5.275% | 4.302% | 11.71% | 10.8% | 1.115% | 2.905% | 2.418% | 18.36% | 19.94% | 1.707% | 2.572% | 1.224% | 0.2831% | 21.47% | 1.031% | 1.731% | 14.83% | 184.5% | 68.79% | 56.15% | 134.7% | 42.01% | 1.023% | 1.712% | 0.6368% | 1.979% | 2.123% | 0.8546% | 2.034% | 1.398% | 2.777% | 15.74% | 26.82% | 0.4687% | 1.028%
76 | 5_10_23_DB |  | 6/8/2023 9:23:33 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_10_23_DB.430 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 31.01% | 2.367e-06% | 3.986% | 23.05% | 1.757% | 77.76% | 263.3% | 1.343e-06% | 0% | 35.36% | 14.59% | 6.812% | 3.397% | 2.273% | 19.82% | 138.9% | 36.51% | 11.83% | 1.095% | 143.2% | 3351% | 0.3738% | 2400% | 26.89% | 46% | 382.5% | 582.2% | 2266% | 0.3006% | 1.816% | 0.009662% | 1.723% | 0.2643% | 1.344% | 1.887% | 0.5783% | 1.941% | 3.629% | 235.5% | 18.58% | 28.48% | 24.1%
77 | 5_10_23_MFB |  | 6/8/2023 9:29:21 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_10_23_MFB.431 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.535% | 2.118% | 1.723% | 1.167% | 1.311% | 5.277% | 1.379% | 1.057% | 3.344% | 2.62% | 2.092% | 2.596% | 0.6811% | 2.675% | 4.291% | 3.286% | 1.638% | 1.674% | 1.894% | 2.525% | 3.679% | 1.977% | 1.673% | 3.829% | 6.147% | 3.088% | 2.854% | 2.263% | 2.245% | 2.576% | 1.578% | 0.4281% | 0.5394% | 1.838% | 1.264% | 0.1646% | 1.932% | 1.84% | 1.252% | 1.059% | 0.71% | 2.884%
78 | 5_10_23_T-239 |  | 6/8/2023 9:35:10 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_10_23_T-239.432 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4.465% | 5.354% | 1.675% | 2.13% | 0.432% | 14.4% | 4.431% | 3.324% | 5.342% | 6.647% | 2.835% | 0.7299% | 1.035% | 6.117% | 4.624% | 3.941% | 5.833% | 2.726% | 1.074% | 2.021% | 19.8% | 3.32% | 1.995% | 2.99% | 6.228% | 5.091% | 2.051% | 6.357% | 6.925% | 1.008% | 2.716% | 0.7626% | 1.058% | 3.463% | 0.8974% | 1.372% | 0.4926% | 2.339% | 0.3804% | 2.281% | 1.387% | 0.9163%
79 | 5_10_23_DB |  | 6/8/2023 9:40:58 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_10_23_DB.433 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 27.44% | 2.367e-06% | 6.507% | 45.67% | 2.035% | 65.21% | 142.5% | 1.088% | 67.51% | 36.47% | 67.42% | 2.888% | 0.2896% | 11.14% | 19.32% | 14.32% | 20.54% | 16.99% | 1.193% | 36.24% | 166.9% | 0.3952% | 99.35% | 144.5% | 0% | 1148% | 167.4% | 152.9% | 288.1% | 0.5765% | 3.684% | 3.917% | 2.06% | 3.543% | 0.8157% | 2.021% | 1.893% | 2.819% | 46.14% | 43.85% | 23.45% | 1007%
80 | 8_16_21_UF_1 |  | 6/8/2023 9:46:46 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_UF_1.434 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 7.497% | 4.15% | 1.09% | 12.59% | 1.436% | 5.926% | 5.757% | 4.699% | 6.445% | 17.62% | 10.31% | 9.658% | 0.6703% | 7.697% | 15.57% | 13.38% | 4.244% | 0.923% | 2.875% | 2.572% | 6.379% | 0.8771% | 0.8378% | 14.33% | 106.6% | 34.99% | 24.74% | 4.169% | 21.11% | 2.087% | 1.148% | 2.166% | 0.2769% | 3.495% | 0.8223% | 1.747% | 0.8733% | 1.835% | 3.213% | 7.194% | 2.616% | 2.453%
81 | LBLANK |  | 6/8/2023 9:52:35 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.435 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 12.15% | 2.367e-06% | 0.435% | 10.06% | 0.9292% | 45.56% | 84.01% | 15.64% | 26.11% | 89.41% | 34.88% | 0.3115% | 1.193% | 3.925% | 21.71% | 23.36% | 6.993% | 10.28% | 2.406% | 8.043% | 28659% | 0.4781% | 11.38% | 35.14% | 60.43% | 118.3% | 1772% | 533% | 113.9% | 2.637% | 1.585% | 1.986% | 0.6458% | 1.782% | 1.358% | 0.7688% | 2.472% | 3.367% | 0% | 25.44% | 4.191% | 2.895%
82 | CCV 1ppb TRACE |  | 6/8/2023 9:58:24 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.436 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 6.005% | 7.083% | 1.648% | 20.41% | 8.354% | 12.66% | 8.62% | 7.982% | 4.33% | 4.201% | 0.9479% | 2.161% | 0.2774% | 4.848% | 0.5987% | 4.91% | 7.602% | 1.707% | 2.708% | 5.788% | 12.78% | 2.012% | 3.66% | 7.08% | 12.66% | 6.97% | 4.557% | 16.63% | 4.772% | 0.3417% | 3.779% | 1.714% | 0.365% | 2.303% | 0.6137% | 0.3899% | 1.344% | 1.048% | 2.849% | 0.9939% | 1.55% | 0.5666%
83 | CCV 1ppb TRACE |  | 6/8/2023 10:04:41 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.437 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4.588% | 1.511% | 4.976% | 22.13% | 3.95% | 23.19% | 15.54% | 8.873% | 0.4676% | 7.781% | 1.598% | 3.509% | 1.138% | 12.74% | 1.893% | 8.555% | 1.807% | 2.074% | 1.596% | 2.021% | 12.21% | 3.622% | 2.633% | 12.29% | 4.591% | 15.26% | 4.963% | 7.591% | 21.5% | 0.942% | 1.888% | 1.331% | 1.256% | 1.775% | 1.681% | 10.01% | 3.655% | 3.007% | 1.934% | 1.949% | 1.194% | 10.74%
84 | LFB |  | 6/8/2023 10:10:29 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.438 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.421% | 1.357% | 3.028% | 3.766% | 2.196% | 4.277% | 4.188% | 1.714% | 0.8205% | 0.985% | 0.2087% | 2.029% | 1.604% | 1.396% | 2.036% | 1.402% | 0.9228% | 1.134% | 0.197% | 1.735% | 4.899% | 1.59% | 1.929% | 5.72% | 2.765% | 3.396% | 1.75% | 2.753% | 3.042% | 1.7% | 2.259% | 2.722% | 1.402% | 0.7825% | 0.78% | 0.2444% | 0.838% | 0.7367% | 0.2889% | 1.944% | 1.933% | 1.276%
85 | CCV 10 ppb MAJORS |  | 6/8/2023 10:16:17 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.439 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 6.305% | 2.131% | 3.576% | 5.114% | 1.357% | 3.614% | 4.269% | 5.643% | 2.25% | 2.919% | 0.8148% | 1.892% | 2.444% | 2.07% | 3.013% | 4.064% | 4.566% | 1.909% | 1.324% | 2.204% | 9.699% | 1.49% | 1.926% | 0.9116% | 1.368% | 1.105% | 2.113% | 3.774% | 3.278% | 1.151% | 2.101% | 2.14% | 0.5004% | 2.31% | 0.6927% | 1.531% | 0.5834% | 2.649% | 0.8158% | 0.4726% | 0.7265% | 0.8947%
86 | CCV 10 ppb MAJORS |  | 6/8/2023 10:22:34 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.440 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.333% | 2.473% | 1.703% | 9.724% | 2.115% | 3.677% | 3.352% | 4.043% | 1.159% | 3.066% | 0.1845% | 2.35% | 0.6768% | 4.325% | 3.149% | 4.923% | 4.067% | 1.079% | 1.5% | 1.003% | 4.056% | 3.258% | 1.421% | 2.682% | 5.019% | 2.494% | 1.45% | 2.956% | 4.033% | 1.525% | 1.45% | 0.2995% | 1.173% | 1.881% | 1.318% | 1.333% | 1.045% | 2.697% | 1.205% | 0.8559% | 1.152% | 1.287%
87 | 8_16_21_UF_6 |  | 6/8/2023 10:28:23 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_UF_6.441 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.319% | 1.548% | 4.613% | 26.64% | 0.8048% | 3.985% | 2.546% | 1.896% | 0.08465% | 18.11% | 9.155% | 155.6% | 0.4696% | 28.39% | 31.57% | 4.503% | 4.687% | 0.2479% | 1.344% | 1.755% | 10.37% | 1.697% | 2.267% | 12.13% | 93.23% | 10.73% | 4.552% | 10.52% | 32.68% | 1.79% | 1.238% | 0.2729% | 1.009% | 0.316% | 1.029% | 0.999% | 1.135% | 3.033% | 29.04% | 4.542% | 3.109% | 1.091%
88 | 8_16_21_UF_8 |  | 6/8/2023 10:34:11 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_UF_8.442 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.69% | 3.462% | 3.156% | 71.77% | 1.866% | 3.971% | 4.892% | 4.202% | 3.108% | 26.92% | 2.769% | 9.141% | 1.908% | 13.43% | 19.25% | 4.265% | 4.529% | 2.691% | 1.451% | 0.6686% | 4.252% | 1.138% | 1.337% | 17.8% | 28.03% | 17.85% | 35.55% | 11.49% | 28.59% | 0.9432% | 1.821% | 0.3269% | 0.9418% | 1.159% | 0.3786% | 1.574% | 1.443% | 1.964% | 26.34% | 10.41% | 6.887% | 3.054%
89 | 8_16_21_UF_9 |  | 6/8/2023 10:39:59 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_UF_9.443 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 6.633% | 6.854% | 2.471% | 17.33% | 1.976% | 8.276% | 1.768% | 8.691% | 2.332% | 9.755% | 3.929% | 39.05% | 1.784% | 66.51% | 29.41% | 7.068% | 9.262% | 2.916% | 2.968% | 1.269% | 20.99% | 1.471% | 1.225% | 34.68% | 50.22% | 27.03% | 11.25% | 13.07% | 92.11% | 1.61% | 2.134% | 0.9286% | 1.333% | 6.494% | 2.079% | 1.027% | 0.6065% | 1.406% | 21.14% | 12.87% | 1.542% | 5.272%
90 | 8_16_21_UF_10 |  | 6/8/2023 10:45:47 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_UF_10.444 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 7.896% | 5.565% | 3.661% | 20.94% | 0.8188% | 5.31% | 7.59% | 7.986% | 3.192% | 13.23% | 1.949% | 5.299% | 1.005% | 13.53% | 8.165% | 6.244% | 6.185% | 2.616% | 3.258% | 0.8655% | 14.44% | 2.191% | 0.4732% | 19.63% | 13.71% | 26.98% | 11.06% | 39.94% | 12.21% | 1.291% | 0.8771% | 0.53% | 0.4721% | 4.474% | 0.6153% | 2.17% | 0.0812% | 2.17% | 11.16% | 10.61% | 3.931% | 3.625%
91 | 8_16_21_UF_11 |  | 6/8/2023 10:51:35 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_UF_11.445 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 7.477% | 10.28% | 9.021% | 69.17% | 1.495% | 13.89% | 10.84% | 12.79% | 11% | 18.75% | 2.911% | 7.105% | 1.014% | 331% | 11.15% | 7.387% | 8.515% | 2.014% | 2.701% | 1.488% | 15.53% | 3.201% | 1.242% | 3.126% | 60.19% | 23.72% | 17.86% | 6.77% | 23.16% | 1.189% | 1.381% | 0.6087% | 1.42% | 8.435% | 0.9888% | 1.059% | 2.268% | 3.491% | 13.92% | 29.59% | 1.5% | 2.158%
92 | 8_16_21_F_1 |  | 6/8/2023 10:57:23 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_F_1.446 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.936% | 2.345% | 1.717% | 3.4% | 4.183% | 1.291% | 4.751% | 3.642% | 9.651% | 10.43% | 5.762% | 2.878% | 2.663% | 5.058% | 12.34% | 4.135% | 4.799% | 2.21% | 5.159% | 1.466% | 2.537% | 5.675% | 2.218% | 8.451% | 22.5% | 11.75% | 15.26% | 16.53% | 38.91% | 1.507% | 3.317% | 1.826% | 2.627% | 1.547% | 2.628% | 1.181% | 4.479% | 6.253% | 12.7% | 4.93% | 3.685% | 0.7706%
93 | 8_16_21_F_1_DUP |  | 6/8/2023 11:03:11 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_F_1_DUP.447 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.064% | 2.131% | 1.263% | 11.91% | 3.097% | 3.122% | 4.726% | 2.066% | 10.1% | 16.14% | 7.01% | 2.023% | 1.506% | 2.096% | 2.04% | 13.6% | 1.273% | 1.015% | 1.668% | 0.7421% | 5.461% | 3.138% | 1.631% | 26.87% | 21.08% | 16.84% | 22.87% | 12.06% | 47.85% | 0.239% | 0.117% | 0.4696% | 0.4723% | 1.549% | 0.1372% | 0.9902% | 2.864% | 2.161% | 2.519% | 26.09% | 8.73% | 6.663%
94 | 8_16_21_F_1_DUP_LD |  | 6/8/2023 11:08:58 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_F_1_DUP_LD.448 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 7.231% | 7.141% | 2.146% | 9.5% | 2.418% | 5.914% | 6.27% | 5.928% | 4.451% | 19.24% | 6.159% | 4.914% | 0.7333% | 6.258% | 29.51% | 8.515% | 4.338% | 0.7923% | 1.52% | 1.022% | 8.663% | 1.53% | 1.039% | 19.35% | 26.7% | 19.33% | 8.731% | 21.68% | 35.4% | 0.9264% | 0.8884% | 0.3947% | 0.9927% | 6.076% | 1.261% | 0.9871% | 0.8203% | 1.144% | 19.93% | 16.74% | 2.789% | 5.606%
95 | 8_16_21_F_1_DUP_LD |  | 6/8/2023 11:15:14 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_F_1_DUP_LD.449 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4.365% | 4.766% | 1.57% | 6.969% | 0.5083% | 4.827% | 3.803% | 3.232% | 9.415% | 31.74% | 6.28% | 2.497% | 2.649% | 5.208% | 25.01% | 4.065% | 2.145% | 3.032% | 6.214% | 0.3673% | 5.757% | 0.4424% | 1.464% | 23.24% | 55.12% | 35.92% | 8.163% | 15.35% | 36.25% | 2.201% | 1.394% | 0.09173% | 1.145% | 3.099% | 2.351% | 0.7248% | 0.5667% | 0.773% | 13.16% | 13.13% | 2.691% | 3.063%
96 | 8_16_21_F_1_DUP_LSP |  | 6/8/2023 11:21:03 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_F_1_DUP_LSP.450 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.816% | 0.5159% | 1.211% | 1.941% | 1.375% | 2.028% | 2.335% | 2.567% | 0.1337% | 2.218% | 1.907% | 3.289% | 0.7815% | 0.8523% | 2.038% | 2.035% | 3.292% | 1.134% | 1.389% | 0.6031% | 1.181% | 1.171% | 0.7572% | 1.872% | 5.691% | 1.999% | 1.748% | 2.773% | 7.838% | 0.6827% | 1.597% | 0.7187% | 0.7526% | 1.516% | 1.123% | 1.947% | 1.145% | 0.8549% | 1.722% | 1.773% | 0.864% | 3.789%
97 | LBLANK |  | 6/8/2023 11:26:52 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.451 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 26.67% | 4.338% | 1.185% | 7.465% | 2.553% | 22.63% | 32.99% | 14.42% | 30.97% | 154% | 37.32% | 0.4176% | 1.601% | 6.832% | 15.84% | 48.63% | 72.29% | 3.381% | 1.444% | 994.4% | 5898% | 0.5689% | 19.1% | 37.94% | 56.9% | 58.77% | 48.55% | 87.95% | 115.4% | 254.5% | 10.65% | 9.862% | 1.758% | 1.371% | 0.7345% | 0.6461% | 0.6941% | 2.442% | 201.2% | 438.8% | 12.21% | 6.509%
98 | CCV 1ppb TRACE |  | 6/8/2023 11:32:40 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.452 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4.296% | 4.885% | 0.4031% | 47.52% | 1.08% | 7.823% | 4.816% | 10.16% | 1.514% | 6.874% | 2.262% | 3.731% | 1.667% | 25.96% | 3.567% | 8.211% | 1.17% | 3.462% | 3.128% | 4.291% | 11.71% | 2.282% | 0.5215% | 6.129% | 24.56% | 8.058% | 7.707% | 11.04% | 8.011% | 2.212% | 0.8737% | 1.712% | 1.207% | 1.071% | 1.81% | 1.001% | 1.833% | 2.31% | 2.793% | 2.241% | 0.6873% | 0.3072%
99 | CCV 1ppb TRACE |  | 6/8/2023 11:38:58 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.453 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.571% | 1.871% | 6.885% | 29.6% | 2.997% | 15.1% | 1.643% | 11.37% | 2.76% | 1.443% | 2.357% | 3.356% | 1.882% | 4.034% | 2.848% | 2.701% | 2.004% | 1.214% | 2.521% | 5.127% | 19.81% | 2.442% | 2.268% | 9.896% | 8.968% | 5.18% | 7.126% | 2.661% | 2.642% | 2.407% | 0.7582% | 1.523% | 0.9683% | 1.687% | 1.931% | 0.5873% | 1.156% | 1.603% | 1.193% | 1.476% | 1.204% | 1.45%
100 | CCV 10 ppb MAJORS |  | 6/8/2023 11:44:46 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.454 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.561% | 2.323% | 0.4597% | 0.5806% | 0.9748% | 0.7811% | 3.265% | 1.227% | 1.906% | 3.073% | 1.074% | 3.368% | 1.353% | 0.9957% | 1.185% | 2.515% | 1.951% | 1.408% | 1.422% | 1.694% | 8.074% | 2.646% | 1.566% | 3.528% | 3.817% | 0.9127% | 3.742% | 3.149% | 2.02% | 1.444% | 1.015% | 1.337% | 0.7502% | 1.837% | 1.87% | 0.8617% | 1.188% | 3.266% | 2.369% | 1.789% | 0.8116% | 0.666%
101 | CCV 10 ppb MAJORS |  | 6/8/2023 11:51:04 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.455 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.099% | 4.97% | 1.829% | 6.916% | 2.236% | 8.17% | 6.59% | 7.024% | 4.519% | 5.354% | 1.677% | 2.201% | 0.9493% | 5.057% | 5.132% | 6.707% | 5.054% | 1.648% | 0.8047% | 1.871% | 7.225% | 0.8138% | 2.029% | 1.196% | 4.677% | 2.953% | 1.326% | 1.363% | 4.715% | 0.7306% | 0.4558% | 1.83% | 0.8537% | 4.35% | 1.445% | 1.142% | 0.9676% | 0.6513% | 0.88% | 0.6275% | 1.501% | 1.009%
102 | 8_16_21_F_6 |  | 6/8/2023 11:56:52 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_F_6.456 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.252% | 0.9196% | 4.963% | 8.855% | 1.318% | 2.535% | 5.772% | 2.552% | 4.952% | 7.628% | 8.526% | 0.6888% | 1.68% | 4.32% | 16.12% | 3.814% | 1.791% | 2.574% | 0.9967% | 4.092% | 5.447% | 3.276% | 0.6582% | 21.34% | 31.63% | 44.56% | 7.249% | 13.57% | 17.72% | 0.7195% | 1.714% | 4.12% | 1.066% | 0.5566% | 1.839% | 0.4059% | 0.9384% | 3.185% | 5.569% | 7.272% | 5.315% | 1.064%
103 | 8_16_21_F_6_SP |  | 6/9/2023 12:02:41 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_F_6_SP.457 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.598% | 2.178% | 2.915% | 10.96% | 1.786% | 5.42% | 4.561% | 2.92% | 5.641% | 3.611% | 1.536% | 1.255% | 2.508% | 2.517% | 4.06% | 3.47% | 3.676% | 1.682% | 1.201% | 0.1546% | 3.341% | 2.065% | 1.551% | 0.685% | 9.477% | 6.1% | 1.332% | 3.643% | 5.647% | 1.154% | 2.255% | 0.9558% | 1.787% | 3.645% | 1.124% | 0.8169% | 1.191% | 1.719% | 1.383% | 2.12% | 0.6517% | 0.7939%
104 | 8_16_21_F_8 |  | 6/9/2023 12:08:29 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_F_8.458 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 5.921% | 7.932% | 2.468% | 2.317% | 2.745% | 4.511% | 10.05% | 5.189% | 3.65% | 23.64% | 3.492% | 1.851% | 1.128% | 8.057% | 26.89% | 9.874% | 4.243% | 2.084% | 1.491% | 1.89% | 3.42% | 0.6537% | 1.872% | 21.59% | 159.1% | 4.631% | 11.49% | 9.149% | 18.41% | 1.573% | 1.436% | 0.7318% | 1.297% | 4.947% | 1.349% | 1.596% | 0.583% | 1.253% | 13.93% | 7.664% | 1.509% | 3.277%
105 | 8_16_21_F_8_DUP |  | 6/9/2023 12:14:18 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_F_8_DUP.459 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.9245% | 4.95% | 1.875% | 4.881% | 2.274% | 3.059% | 4.464% | 1.97% | 4.683% | 11.1% | 4.68% | 1.86% | 1.069% | 3.621% | 2.256% | 7.581% | 2.556% | 1.424% | 1.438% | 1.029% | 7.129% | 2.605% | 2.127% | 8.399% | 41.51% | 18.14% | 8.466% | 13.75% | 40.67% | 0.8845% | 0.316% | 0.7788% | 0.937% | 1.43% | 1.107% | 2.573% | 1.037% | 2.029% | 5.351% | 6.245% | 3.003% | 1.761%
106 | 8_16_21_F_9 |  | 6/9/2023 12:20:06 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_F_9.460 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.5452% | 1.625% | 2.052% | 2.959% | 2.828% | 4.223% | 1.297% | 3.293% | 1.165% | 8.215% | 1.552% | 1.512% | 0.295% | 2.146% | 25.26% | 1.666% | 1.115% | 0.4251% | 2.163% | 2.289% | 9.431% | 2.741% | 1.178% | 1.941% | 8.445% | 10.94% | 16.49% | 7.773% | 45.83% | 1.206% | 1.052% | 1.521% | 0.5233% | 0.8233% | 0.204% | 0.2324% | 1.431% | 3.191% | 19.69% | 19.74% | 4.489% | 1.815%
107 | 8_16_21_F_10 |  | 6/9/2023 12:25:54 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_F_10.461 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.916% | 0.4392% | 1.166% | 5.579% | 1.418% | 2.982% | 3.125% | 1.807% | 4.716% | 15.04% | 5.355% | 1.511% | 1.188% | 3.755% | 22.58% | 9.504% | 1.989% | 1.175% | 1.764% | 0.5994% | 9.734% | 3.444% | 1.801% | 24.6% | 48.94% | 14.02% | 35.04% | 18.2% | 29.28% | 1.196% | 1.274% | 0.358% | 2.012% | 1.344% | 1.583% | 1.291% | 0.6617% | 3.59% | 10.83% | 19.39% | 1.277% | 2.78%
108 | 8_16_21_F_10_DUP |  | 6/9/2023 12:31:42 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_F_10_DUP.462 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4.485% | 2.142% | 0.4651% | 25.44% | 3.343% | 1.724% | 5.489% | 2.987% | 6.744% | 25.43% | 5.629% | 6.019% | 5.406% | 4.925% | 4.042% | 10.72% | 0.6844% | 3.68% | 5.439% | 0.7748% | 8.693% | 5.73% | 5.974% | 6.864% | 61.77% | 12.3% | 8.358% | 16.69% | 24.08% | 5.721% | 5.342% | 0.8373% | 4.53% | 1.341% | 3.78% | 1.365% | 2.745% | 4.607% | 4.71% | 8.453% | 1.361% | 1.083%
109 | 8_16_21_F_11 |  | 6/9/2023 12:37:30 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_F_11.463 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 6.934% | 7.176% | 0.7294% | 2.228% | 2.077% | 4.504% | 4.608% | 4.049% | 0.453% | 8.094% | 8.191% | 0.945% | 1.924% | 6.658% | 8.54% | 15.64% | 6.738% | 0.452% | 1.412% | 0.6052% | 4.059% | 1.02% | 0.9355% | 9.505% | 64.7% | 38.05% | 5.456% | 6.723% | 23.63% | 2.541% | 1.863% | 1.65% | 1.315% | 4.767% | 1.686% | 0.7693% | 1.186% | 0.6053% | 10.01% | 27.67% | 2.519% | 0.939%
110 | 8_16_21_W_1 |  | 6/9/2023 12:43:18 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_W_1.464 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10.26% | 9.403% | 1.677% | 14.09% | 3.084% | 10.02% | 9.998% | 4.441% | 14.5% | 40.37% | 20.6% | 12.65% | 0.7827% | 11.64% | 40.91% | 7.487% | 8.135% | 12.1% | 47.91% | 4.124% | 11.25% | 2.124% | 1.927% | 21.78% | 127.6% | 30.77% | 38.47% | 54.02% | 92.19% | 12.44% | 0.805% | 1.295% | 0.2771% | 5.899% | 10.25% | 0.9086% | 2.217% | 1.242% | 38.38% | 21.09% | 8.006% | 8.084%
111 | 8_16_21_W_6 |  | 6/9/2023 12:49:05 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_W_6.465 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 17.53% | 19.3% | 3.294% | 19.64% | 3.895% | 26.3% | 27.7% | 20.1% | 25.56% | 17.3% | 13.96% | 1.688% | 0.6171% | 23.93% | 29.76% | 16.74% | 18.49% | 0.3836% | 0.8229% | 2.719% | 32.78% | 2.26% | 0.3127% | 20.52% | 103.8% | 21.81% | 7.457% | 49.24% | 26.93% | 1.221% | 2.279% | 2.28% | 0.9384% | 16.91% | 1.167% | 0.8165% | 1.573% | 1.469% | 29.07% | 23.67% | 7.74% | 3.169%
112 | LBLANK |  | 6/9/2023 12:54:53 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.466 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 16.81% | 2.367e-06% | 0.2235% | 3.845% | 0.6843% | 47.1% | 1605% | 1.343e-06% | 0% | 110.8% | 18.08% | 1.911% | 0.9897% | 5.113% | 21.72% | 76495% | 33.17% | 4.954% | 3.209% | 63.76% | 1.756e-06% | 0.4787% | 36.09% | 90.11% | 0% | 194.3% | 3160% | 87.22% | 113.7% | 2.687% | 4.589% | 1.964% | 1.029% | 1.548% | 1.078% | 1.524% | 2.507% | 2.19% | 234.3% | 53.46% | 13.47% | 6.586%
113 | CCV 1ppb TRACE |  | 6/9/2023 1:00:42 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.467 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4.805% | 13.67% | 3.046% | 33.43% | 3.128% | 25.46% | 3.528% | 11.2% | 4.656% | 3.537% | 1.945% | 4.309% | 2.502% | 11.06% | 2.29% | 7.407% | 3.296% | 0.57% | 2.292% | 2.848% | 10.5% | 3.644% | 1.014% | 7.623% | 7.357% | 6.521% | 5.882% | 13.68% | 0.9607% | 1.228% | 0.5477% | 1.548% | 0.9687% | 3.05% | 1.976% | 0.9733% | 0.9663% | 3.391% | 3.054% | 0.8178% | 1.033% | 0.9342%
114 | CCV 1ppb TRACE |  | 6/9/2023 1:06:59 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.468 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 8.124% | 10.6% | 0.429% | 18.34% | 1.398% | 12.22% | 10.04% | 13.28% | 5.883% | 6.251% | 3.886% | 8.608% | 1.598% | 16.69% | 1.887% | 1.816% | 4.449% | 3.643% | 2.84% | 5.003% | 10.5% | 1.24% | 2.078% | 4.191% | 6.115% | 3.185% | 4.766% | 8.825% | 6.321% | 3.763% | 3.626% | 1.173% | 1.962% | 2.863% | 3.339% | 1.21% | 1.275% | 3.203% | 2.484% | 4.304% | 0.8858% | 0.1672%
115 | LFB |  | 6/9/2023 1:12:47 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.469 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.901% | 3.397% | 2.067% | 12.01% | 1.126% | 3.011% | 1.445% | 4.269% | 3.19% | 2.563% | 0.8825% | 1.437% | 0.7557% | 3.366% | 3.377% | 2.75% | 3.374% | 0.6168% | 1.154% | 0.6864% | 6.076% | 2.81% | 0.8805% | 2.949% | 3.06% | 5.303% | 1.387% | 1.993% | 2.697% | 1.372% | 1.329% | 2.58% | 1.871% | 1.793% | 0.279% | 1.137% | 2.477% | 2.268% | 0.6576% | 0.9302% | 0.8986% | 1.474%
116 | CCV 10 ppb MAJORS |  | 6/9/2023 1:18:35 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.470 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.9866% | 4.566% | 1.7% | 16.49% | 4.33% | 3.829% | 3.383% | 3.527% | 2.789% | 1.129% | 0.3295% | 0.7158% | 1.652% | 1.91% | 1.302% | 1.61% | 1.899% | 1.525% | 1.966% | 1.88% | 3.906% | 1.141% | 0.3744% | 2.105% | 2.899% | 2.222% | 2.557% | 2.685% | 4.619% | 1.626% | 0.8331% | 1.314% | 0.7835% | 1.261% | 1.627% | 0.504% | 0.3669% | 2.207% | 2.466% | 2.695% | 0.9918% | 0.4528%
117 | CCV 10 ppb MAJORS |  | 6/9/2023 1:24:53 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.471 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.734% | 3.184% | 1.989% | 0.218% | 3.262% | 3.701% | 4.617% | 2.429% | 1.451% | 0.57% | 0.7688% | 1.304% | 0.554% | 0.184% | 0.4318% | 2.314% | 2.807% | 1.181% | 0.31% | 2.289% | 2.125% | 0.889% | 0.3817% | 1.317% | 2.564% | 1.257% | 3.44% | 2.717% | 3.773% | 0.2432% | 0.9816% | 0.4798% | 0.7623% | 1.984% | 0.6755% | 1.288% | 1.023% | 0.9139% | 1.279% | 0.5765% | 0.5871% | 1.856%
118 | 8_16_21_W_6_DUP |  | 6/9/2023 1:30:41 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_W_6_DUP.472 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.5497% | 4.417% | 1.741% | 4% | 21.35% | 4.322% | 8.102% | 3.631% | 10.34% | 4.163% | 2.015% | 2.901% | 1.339% | 2.562% | 9.745% | 39.04% | 2.714% | 3.15% | 1.171% | 2.513% | 19.94% | 1.237% | 0.9085% | 26.44% | 47.5% | 67.77% | 24.73% | 44.6% | 54.08% | 1.971% | 1.243% | 0.8148% | 1.833% | 2.474% | 1.861% | 2.343% | 5.142% | 3.532% | 32.32% | 25.98% | 5.321% | 0.8762%
119 | 8_16_21_W_8 |  | 6/9/2023 1:36:29 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_W_8.473 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.906% | 10.66% | 3.131% | 5.734% | 0.9431% | 12.38% | 1.502% | 10.09% | 2.003% | 26.02% | 4.443% | 1.64% | 1.83% | 7.892% | 15.31% | 22.35% | 8.734% | 0.8924% | 1.666% | 2.25% | 17.36% | 3.629% | 1.097% | 10.81% | 83.32% | 64.83% | 27.04% | 28.5% | 98.82% | 0.8599% | 0.8691% | 0.8363% | 0.7829% | 4.818% | 1.071% | 1.594% | 1.305% | 3.35% | 14.04% | 12.38% | 1.156% | 0.6938%
120 | 8_16_21_W_9 |  | 6/9/2023 1:42:18 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_W_9.474 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.824% | 3.953% | 1.241% | 1.123% | 2.327% | 3.376% | 3.967% | 1.327% | 5.497% | 12.29% | 12.09% | 9.139% | 3.026% | 3.278% | 14.85% | 45.88% | 7.561% | 10.42% | 9.745% | 5.437% | 11.68% | 1.934% | 4.321% | 39.31% | 136.5% | 11.6% | 13.19% | 43.12% | 38.37% | 9.056% | 3.504% | 0.6867% | 1.875% | 1.902% | 8.411% | 0.327% | 0.898% | 3.381% | 26.5% | 44.27% | 1.124% | 1.233%
121 | 8_16_21_W_9_DUP |  | 6/9/2023 1:48:06 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_W_9_DUP.475 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4.009% | 3.722% | 1.725% | 7.413% | 2.387% | 9.293% | 9.262% | 1.973% | 0.5458% | 19.16% | 5.105% | 0.6885% | 0.7976% | 1.682% | 9.888% | 35.97% | 2.044% | 1.348% | 3.812% | 1.883% | 15.99% | 2.551% | 2.985% | 32.5% | 0% | 50.82% | 16.4% | 9871% | 52.32% | 2.51% | 1.519% | 2.945% | 0.568% | 2.032% | 1.667% | 1.656% | 0.7095% | 3.362% | 19.97% | 15.03% | 2.788% | 0.7418%
122 | 8_16_21_W_10 |  | 6/9/2023 1:53:55 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_W_10.476 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.539% | 6.027% | 1.902% | 8.211% | 2.277% | 7.032% | 6.897% | 5.24% | 10.38% | 8.232% | 3.119% | 0.612% | 0.7133% | 1.561% | 36.22% | 16.04% | 2.107% | 2.574% | 2.673% | 1.902% | 2.943% | 1.609% | 0.8385% | 34.33% | 133.5% | 76.08% | 22.33% | 65.78% | 106% | 2.882% | 0.8174% | 0.4395% | 1.074% | 0.4996% | 1.338% | 1.091% | 1.271% | 2.239% | 7.393% | 21.6% | 3.729% | 0.7027%
123 | 8_16_21_W_11 |  | 6/9/2023 1:59:43 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_W_11.477 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.077% | 1.355% | 0.7113% | 8.97% | 4.325% | 13.85% | 13% | 3.069% | 5.415% | 31.86% | 9.172% | 1.854% | 0.8563% | 1.957% | 42.03% | 39.41% | 5.664% | 0.7976% | 8.457% | 3.08% | 19.72% | 1.424% | 2.166% | 21.51% | 343.3% | 16.97% | 23.26% | 53.15% | 27.12% | 0.273% | 3.018% | 1.254% | 1.455% | 0.7518% | 0.8231% | 13.11% | 1.793% | 3.945% | 14.72% | 49.68% | 5.883% | 13.54%
124 | 5_10_23_DB |  | 6/9/2023 2:05:32 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_10_23_DB.478 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0% | 2.367e-06% | 2.713% | 50.34% | 1.302% | 114.8% | 945.4% | 1.343e-06% | 0% | 26.99% | 36.87% | 1.836% | 2.01% | 2.039% | 19.71% | 127.7% | 40.87% | 6.602% | 2.629% | 71.7% | 6564% | 0.2928% | 14.98% | 17.57% | 47.27% | 2498% | 175.7% | 28.84% | 105.5% | 3.544% | 1.249% | 0.6514% | 0.3468% | 4.246% | 1.196% | 0.8099% | 1.351% | 2.522% | 0% | 31.01% | 231.3% | 182.9%
125 | 5_10_23_MFB |  | 6/9/2023 2:11:19 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_10_23_MFB.479 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.638% | 3.035% | 0.7181% | 5.496% | 1.819% | 1.093% | 4.225% | 5.95% | 0.9729% | 2.478% | 2.408% | 1.599% | 1.521% | 3.956% | 3.017% | 2.918% | 2.342% | 3.415% | 1.663% | 0.8845% | 4.957% | 3.068% | 1.27% | 4.081% | 6.801% | 2.137% | 0.4932% | 2.119% | 4.289% | 2.754% | 0.7496% | 0.9238% | 1.177% | 0.4866% | 2.302% | 2.018% | 1.174% | 3.21% | 3.288% | 2.639% | 1.523% | 1.539%
126 | 5_10_23_T-239 |  | 6/9/2023 2:17:07 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_10_23_T-239.480 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.581% | 4.215% | 3.785% | 1.944% | 1.259% | 10.24% | 0.353% | 1.742% | 2.25% | 1.41% | 0.5081% | 0.5575% | 0.9284% | 2.238% | 3.608% | 5.184% | 3.267% | 0.31% | 0.315% | 5.238% | 15.41% | 1.861% | 1.456% | 4.998% | 5.128% | 6.747% | 2.816% | 9.344% | 5.27% | 3.411% | 0.7931% | 3.786% | 0.4834% | 1.372% | 0.2909% | 1.177% | 0.1404% | 1.484% | 0.3196% | 0.6152% | 3.398% | 0.7933%
127 | 5_10_23_DB |  | 6/9/2023 2:22:55 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_10_23_DB.481 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 17.69% | 4.476% | 107.6% | 39.33% | 2.734% | 36.2% | 134.8% | 16.18% | 19.02% | 31.12% | 23.33% | 6.35% | 0.1479% | 14.28% | 23.19% | 76.06% | 3.468% | 9.591% | 1.749% | 187.6% | 1.756e-06% | 1.048% | 125.4% | 31.5% | 44.82% | 0% | 156.1% | 187.5% | 106.3% | 7.284% | 1.826% | 5.084% | 1.435% | 3.572% | 0.4082% | 0.9682% | 1.095% | 3.749% | 200% | 8.08% | 55.55% | 19.07%
128 | LBLANK |  | 6/9/2023 2:28:44 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.482 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 5.874% | 2.367e-06% | 0.004685% | 3.797% | 2.125% | 43.53% | 215.9% | 16.09% | 31.53% | 53.22% | 49.81% | 1.053% | 2.341% | 3.732% | 0% | 895.1% | 39.14% | 9.129% | 4.35% | 66.47% | 1.756e-06% | 0.3225% | 23.35% | 54.61% | 127.5% | 377.4% | 75% | 217.9% | 0% | 3.237% | 0.8123% | 1.952% | 0.4323% | 1.415% | 1.112% | 1.548% | 0.4756% | 4.293% | 0% | 38.34% | 7.585% | 5.859%
129 | CCV 1ppb TRACE |  | 6/9/2023 2:34:32 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.483 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 6.406% | 3.733% | 17.19% | 7.225% | 2.413% | 10.93% | 3.829% | 7.783% | 6.758% | 2.08% | 1.39% | 6.3% | 2.164% | 7.075% | 1.084% | 0.6213% | 2.351% | 2.787% | 3.637% | 11.53% | 10.85% | 3.414% | 1.176% | 19.05% | 14.16% | 4.028% | 5.142% | 26.03% | 7.477% | 1.641% | 1.87% | 8.779% | 1.578% | 2.068% | 1.398% | 0.705% | 0.3807% | 2.378% | 1.806% | 3.353% | 10.89% | 1.05%
130 | CCV 1ppb TRACE |  | 6/9/2023 2:40:49 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.484 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.5636% | 4.067% | 1.493% | 13.46% | 1.843% | 19.4% | 4.469% | 14.11% | 9.084% | 1.688% | 3.437% | 3.464% | 3.53% | 8.405% | 0.09538% | 9.096% | 2.027% | 2.825% | 2.471% | 1.586% | 8.737% | 4.012% | 1.336% | 4.678% | 20.19% | 12.51% | 7.133% | 7.205% | 12.7% | 2.599% | 2.032% | 0.9302% | 2.163% | 1.547% | 1.608% | 0.2209% | 0.6941% | 5.292% | 1.852% | 0.9524% | 1.04% | 1.213%
131 | CCV 10 ppb MAJORS |  | 6/9/2023 2:46:38 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.485 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.24% | 3.91% | 0.4469% | 9.831% | 1.266% | 1.989% | 4.328% | 9.195% | 3.94% | 2.605% | 2.505% | 1.788% | 0.6731% | 1.792% | 4.482% | 3.456% | 2.423% | 1.088% | 1.004% | 1.593% | 7.148% | 3.508% | 0.7177% | 0.3774% | 1.179% | 1.739% | 1.423% | 3.432% | 2.419% | 2.421% | 0.6006% | 1.519% | 0.8867% | 3.707% | 0.735% | 0.6177% | 1.684% | 3.675% | 1.27% | 2.01% | 1.785% | 1.372%
132 | CCV 10 ppb MAJORS |  | 6/9/2023 2:52:55 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.486 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.792% | 3.826% | 8.315% | 3.488% | 2.104% | 4.433% | 1.268% | 3.389% | 1.482% | 1.028% | 1.719% | 0.2639% | 1.202% | 0.2959% | 0.985% | 1.616% | 1.573% | 0.2985% | 0.9905% | 8.438% | 3.851% | 1.265% | 0.5623% | 9.027% | 3.159% | 3.873% | 0.7131% | 8.988% | 4.598% | 1.193% | 0.7664% | 7.42% | 0.5892% | 0.8202% | 0.7148% | 0.689% | 0.8504% | 2.359% | 0.8013% | 0.7989% | 8.418% | 1.884%
133 | MT 34 D |  | 6/9/2023 2:58:43 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\MT 34 D.487 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 28.18% | 2.367e-06% | 2.487% | 56.52% | 4.054% | 48.66% | 409.9% | 99.05% | 48.32% | 16.96% | 27.01% | 2.039% | 3.053% | 50.78% | 22.02% | 48.72% | 48.65% | 28.05% | 5.937% | 40.87% | 1.756e-06% | 0.4254% | 82.21% | 45.84% | 0% | 198.3% | 130.7% | 86.63% | 0% | 11.29% | 19.89% | 0.8334% | 1.127% | 2.42% | 1.282% | 1.715% | 2.209% | 1.242% | 237.3% | 195.1% | 5.743% | 5.751%
134 | MT 35 D |  | 6/9/2023 3:04:31 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\MT 35 D.488 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 5.573% | 2.367e-06% | 2.735% | 18.51% | 1.718% | 51.97% | 203.5% | 73.53% | 30.54% | 33.7% | 2.458% | 2.684% | 4.051% | 6.155% | 21.42% | 45.22% | 55.08% | 13.82% | 7.493% | 2558% | 13505% | 0.8345% | 79.13% | 62.83% | 45.69% | 0% | 151.6% | 245.7% | 126.4% | 15.9% | 7.882% | 1.242% | 0.4161% | 1.153% | 1.601% | 0.6902% | 1.27% | 3.417% | 234.4% | 738.9% | 7.007% | 4.139%
135 | MT 36 D |  | 6/9/2023 3:10:18 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\MT 36 D.489 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 12.44% | 2.367e-06% | 0.5371% | 290.7% | 0.4763% | 35.85% | 64.93% | 39.01% | 0% | 8.322% | 2.757% | 1.276% | 0.6683% | 12.01% | 0% | 49.8% | 17.24% | 17.62% | 2.29% | 89.41% | 1.756e-06% | 0.6114% | 71.31% | 30.35% | 45.62% | 1595% | 354.6% | 58.71% | 274.9% | 17.41% | 5.056% | 0.8744% | 0.1385% | 1.874% | 1.032% | 1.176% | 1.504% | 3.88% | 45.55% | 85.78% | 5.194% | 5.833%
136 | MT 37 D |  | 6/9/2023 3:16:06 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\MT 37 D.490 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 15.8% | 4.331% | 6.05% | 92.26% | 0.5821% | 42.37% | 4962% | 52.52% | 0% | 19.37% | 11.24% | 2.471% | 4.68% | 5.354% | 32.1% | 92.81% | 163.8% | 12.97% | 4.131% | 28.29% | 1.756e-06% | 0.2068% | 13.5% | 25.32% | 0% | 190.7% | 214.5% | 182.6% | 0% | 14.74% | 2.164% | 0.5003% | 1.251% | 2.686% | 2.011% | 1.692% | 2.203% | 1.899% | 92.8% | 78.48% | 2.315% | 9.371%
137 | MT 38 D |  | 6/9/2023 3:21:55 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\MT 38 D.491 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 191.6% | 4.277% | 1.671% | 9.037% | 0.7177% | 18.94% | 33.53% | 47.18% | 52.84% | 27.16% | 11.1% | 3.725% | 453.1% | 4.308% | 24.92% | 50.47% | 8.9% | 3.509% | 1.791% | 144.4% | 2.658% | 0.625% | 62.81% | 57.77% | 0% | 161.6% | 857.8% | 158.1% | 0% | 18.65% | 5.636% | 0.7602% | 0.7168% | 1.825% | 0.895% | 1.248% | 1.445% | 0.5458% | 122.1% | 32.35% | 10.28% | 1.359%
138 | MT 39 D |  | 6/9/2023 3:27:43 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\MT 39 D.492 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 6.042% | 2.367e-06% | 3.676% | 50.99% | 1.672% | 35.38% | 266% | 0.4208% | 19.95% | 28.09% | 23.89% | 16.27% | 9.145% | 34.75% | 22.09% | 233.6% | 42.25% | 17.7% | 9.435% | 1611% | 1.756e-06% | 0.2457% | 15.15% | 19.25% | 0% | 384.3% | 111.2% | 199.1% | 0% | 74.72% | 273.2% | 0.8961% | 1.76% | 1.431% | 6.407% | 0.4438% | 0.8516% | 0.5133% | 121.8% | 159.6% | 1.976% | 2.536%
139 | MT 40 D |  | 6/9/2023 3:33:31 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\MT 40 D.493 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10.47% | 2.367e-06% | 2.123% | 48.09% | 0.9201% | 48.28% | 6265% | 30.76% | 0% | 31.87% | 12.87% | 0.9921% | 1.873% | 11.41% | 49.36% | 67.24% | 63.75% | 33.55% | 2.504% | 249.8% | 1.756e-06% | 0.2554% | 20.36% | 26.53% | 0% | 1991% | 1962% | 11.1% | 0% | 9.561% | 4.362% | 0.1994% | 1.359% | 2.365% | 0.7701% | 4.341% | 1.556% | 2.984% | 52.42% | 965.2% | 5.705% | 5.416%
140 | MT 32 D |  | 6/9/2023 3:39:20 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\MT 32 D.494 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 12.9% | 2.367e-06% | 1.754% | 45.93% | 1.525% | 77.54% | 980.8% | 20.39% | 53.48% | 65.85% | 16.2% | 0.6702% | 11.67% | 15.83% | 0% | 177.8% | 9.267% | 3.864% | 6.171% | 104% | 3217% | 0.2852% | 45.66% | 50.59% | 64.86% | 187.8% | 38284% | 166.9% | 0% | 6.636% | 2.026% | 1.434% | 1.689% | 5.538% | 0.9985% | 0.9056% | 2.31% | 2.124% | 99.9% | 140.4% | 12.37% | 1.884%
141 | MT 32 D LD |  | 6/9/2023 3:45:07 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\MT 32 D LD.495 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 6.524% | 2.367e-06% | 14.59% | 50.92% | 4.559% | 25.19% | 141.9% | 32.41% | 0% | 61.6% | 14.9% | 1.759% | 2.29% | 12.12% | 22.59% | 183.4% | 23.04% | 6.354% | 6.551% | 1093% | 1.756e-06% | 0.4899% | 41.08% | 19.46% | 49.05% | 2225% | 462.9% | 82.02% | 0% | 5.731% | 1.446% | 1.454% | 1.796% | 0.07034% | 0.9252% | 11.27% | 1.47% | 1.56% | 0% | 150.1% | 7.413% | 12.18%
142 | MT 32 D LD |  | 6/9/2023 3:51:24 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\MT 32 D LD.496 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 16.69% | 4.397% | 1.937% | 80.96% | 1.675% | 7.556% | 68.56% | 16.25% | 31.57% | 14.14% | 10.16% | 1.486% | 2.813% | 4.18% | 0% | 111.6% | 32.98% | 7.387% | 3.498% | 11.89% | 1.756e-06% | 0.244% | 24.5% | 17.3% | 131.9% | 2173% | 220.9% | 893.7% | 125.4% | 6.233% | 8.127% | 0.9039% | 4.251% | 0.6272% | 1.013% | 0.4387% | 2.442% | 7.927% | 91.49% | 420.4% | 3.566% | 1.95%
143 | LBLANK |  | 6/9/2023 3:57:12 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.497 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10.03% | 2.367e-06% | 0.1306% | 11.33% | 2.465% | 34.92% | 68.5% | 14.33% | 30.74% | 32.06% | 22.02% | 1.893% | 0.5321% | 13.12% | 0% | 268.9% | 100.7% | 6.673% | 0.1776% | 54.93% | 1.756e-06% | 0.243% | 16.03% | 54.31% | 68.05% | 270% | 839.9% | 143.6% | 0% | 2.56% | 0.2545% | 0.8985% | 12.04% | 1.277% | 5.761% | 1.625% | 2.086% | 11% | 191.4% | 26.57% | 7.004% | 9.276%
144 | CCV 1ppb TRACE |  | 6/9/2023 4:03:01 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.498 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.993% | 2.923% | 3.241% | 25.58% | 6.551% | 5.061% | 2.358% | 5.915% | 6.012% | 3.912% | 1.052% | 6.607% | 1.309% | 12.62% | 2.29% | 1.846% | 2.385% | 0.7465% | 2.126% | 4.449% | 13.11% | 3.834% | 2.868% | 10.61% | 5.143% | 3.461% | 1.915% | 19.97% | 17.34% | 1.25% | 1.462% | 2.215% | 0.591% | 1.411% | 0.5086% | 0.8483% | 1.025% | 1.69% | 1.115% | 1.361% | 1.296% | 0.5146%
145 | CCV 1ppb TRACE |  | 6/9/2023 4:09:18 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.499 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.3913% | 3.283% | 1.843% | 30.98% | 1.79% | 14.12% | 4.768% | 13.67% | 3.244% | 2.024% | 2.863% | 3.58% | 1.618% | 4.056% | 2.515% | 3.253% | 2.682% | 1.958% | 3.53% | 3.705% | 14.27% | 2.147% | 1.335% | 10.06% | 13.51% | 6.882% | 10.74% | 23.34% | 2.324% | 2.218% | 2.33% | 2.728% | 0.5777% | 2.163% | 0.6755% | 0.9018% | 1.448% | 3.282% | 1.826% | 0.6011% | 3.413% | 0.7153%
146 | LFB |  | 6/9/2023 4:15:06 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.500 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.269% | 2.343% | 1.815% | 4.982% | 1.072% | 2.292% | 2.476% | 3.998% | 1.207% | 3.174% | 1.434% | 0.9814% | 3.594% | 1.976% | 0.5543% | 2.492% | 0.8441% | 0.7075% | 0.3766% | 1.311% | 6.331% | 1.806% | 1.575% | 1.321% | 6.32% | 1.548% | 1.687% | 3.361% | 9.507% | 0.8262% | 1.807% | 1.636% | 2.248% | 1.032% | 0.7409% | 1.497% | 1.141% | 1.784% | 2.098% | 0.4071% | 1.829% | 1.052%
147 | LFB |  | 6/9/2023 4:21:22 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.501 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.556% | 2.512% | 2.297% | 11.82% | 1.804% | 2.885% | 2.136% | 3.916% | 1.577% | 2.587% | 1.014% | 2.177% | 1.768% | 4.919% | 2.86% | 2.969% | 2.783% | 0.9122% | 0.9398% | 2.412% | 1.867% | 2.937% | 1.726% | 1.244% | 4.284% | 1.91% | 5.003% | 4.226% | 1.87% | 1.18% | 1.37% | 2.188% | 1.864% | 1.661% | 0.4768% | 0.788% | 0.5475% | 3.237% | 1.071% | 1.201% | 1.171% | 1.935%
148 | CCV 10 ppb MAJORS |  | 6/9/2023 4:27:10 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.502 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.111% | 3.886% | 1.038% | 9.674% | 1.587% | 2.797% | 0.709% | 1.281% | 0.4145% | 1.635% | 3.369% | 4.628% | 1.641% | 2.885% | 0.7588% | 1.798% | 0.7796% | 4.131% | 3.053% | 2.749% | 2.633% | 2.461% | 2.171% | 4.001% | 2.794% | 1.634% | 2.507% | 5.808% | 3.108% | 3.616% | 1.239% | 1.966% | 1.544% | 1.217% | 2.329% | 1.18% | 0.3998% | 3.078% | 3.368% | 3.529% | 2.303% | 2.04%
149 | CCV 10 ppb MAJORS |  | 6/9/2023 4:33:27 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.503 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.395% | 3.242% | 0.7682% | 10.73% | 2.817% | 6.89% | 2.875% | 1.185% | 2.726% | 2.176% | 11.78% | 11.07% | 0.9599% | 3.023% | 3.331% | 3.46% | 1.378% | 11.86% | 11.88% | 0.7606% | 9.008% | 2.177% | 1.059% | 1.853% | 2.614% | 0.6121% | 1.758% | 1.214% | 3.392% | 9.996% | 2.512% | 1.114% | 0.5826% | 2.441% | 7.878% | 0.8698% | 0.4604% | 2.582% | 11.17% | 11.59% | 1.179% | 2.21%
150 | MT 33 D |  | 6/9/2023 4:39:15 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\MT 33 D.504 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 22.28% | 2.367e-06% | 15.81% | 137.1% | 3.258% | 23.39% | 58.51% | 14.73% | 68.32% | 30.05% | 20.02% | 3.42% | 7.241% | 10.68% | 42.16% | 163.7% | 34.65% | 4.323% | 5.103% | 54.6% | 1.756e-06% | 0.2128% | 730.6% | 11.6% | 0% | 1692% | 430.4% | 115% | 105.8% | 6.726% | 4.797% | 1.579% | 0.7729% | 2.11% | 0.5099% | 0.5528% | 1.372% | 2.005% | 33.94% | 127.2% | 4.878% | 6.387%
151 | MT 33 D LSP |  | 6/9/2023 4:45:03 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\MT 33 D LSP.505 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.013% | 2.02% | 4.827% | 4.851% | 2.435% | 1.808% | 0.8545% | 3.779% | 3.518% | 2.647% | 1.142% | 1.177% | 1.048% | 2.858% | 4.233% | 4.57% | 3.198% | 1.671% | 0.9693% | 5.127% | 3.14% | 4.06% | 0.2414% | 6.054% | 2.287% | 1.934% | 0.753% | 4.2% | 2.109% | 0.814% | 1.142% | 3.993% | 0.7362% | 2.683% | 1.168% | 0.9815% | 1.674% | 3.671% | 0.284% | 1.059% | 4.043% | 0.6456%
152 | cal blank |  | 6/9/2023 4:50:51 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\cal blank.506 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 17.37% | 2.367e-06% | 0.07457% | 0% | 4.234% | 83.4% | 2638% | 1.343e-06% | 53.05% | 1107% | 416.6% | 1.736% | 1.54% | 5.595% | 20.14% | 96.84% | 12.75% | 5.227% | 1.439% | 1653% | 2274% | 0.05265% | 14.62% | 41.13% | 127.1% | 77.22% | 54.9% | 199.1% | 106.1% | 9.856% | 3.635% | 2.059% | 0.871% | 1.863% | 1.028% | 1.469% | 0.6951% | 2.1% | 72.14% | 50.17% | 23.41% | 13.97%
153 | STD 0.01 |  | 6/9/2023 4:56:40 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 0.01.507 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 233.8% | 40.06% | 2.723% | 143.3% | 3.038% | 24.98% | 37.47% | 23.71% | 55.93% | 24.46% | 9.856% | 2.983% | 17.66% | 12.95% | 16.56% | 11.77% | 51.95% | 5.644% | 9.002% | 62.38% | 1.756e-06% | 0.475% | 8.267% | 24.45% | 64.77% | 36.89% | 80.43% | 401.9% | 113.1% | 6.779% | 8.167% | 1.81% | 1.032% | 0.8952% | 1.351% | 1.093% | 1.372% | 3.022% | 7.174% | 13.72% | 11.46% | 1.362%
154 | std 0.05 |  | 6/9/2023 5:02:29 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\std 0.05.508 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 5% | 35.39% | 9.242% | 582.1% | 25.84% | 60.96% | 34.16% | 38.57% | 32% | 13.64% | 2.826% | 2.294% | 7.342% | 14.46% | 11.24% | 16.46% | 20.86% | 7.293% | 4.206% | 17.87% | 44.7% | 0.1416% | 6.058% | 44.03% | 73.45% | 17.35% | 9.673% | 116.7% | 56.95% | 6.866% | 6.978% | 1.582% | 2.929% | 1.151% | 1.193% | 15.76% | 1.92% | 4.475% | 4.918% | 4.538% | 2.852% | 16.43%
155 | std 0.1 |  | 6/9/2023 5:08:17 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\std 0.1.509 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 6.823% | 44.57% | 1.129% | 23.43% | 1.315% | 36.39% | 8.981% | 8.294% | 8.675% | 11.36% | 3.241% | 9.676% | 3.516% | 16.12% | 2.237% | 35.94% | 5.012% | 2.9% | 32.56% | 38.23% | 18.25% | 2.713% | 5.461% | 11.51% | 36.66% | 6.6% | 12.44% | 137.3% | 21.99% | 5.09% | 1.936% | 0.3687% | 0.8705% | 1.407% | 0.9109% | 2.12% | 1.988% | 2.297% | 3.606% | 1.209% | 3.527% | 1.997%
156 | std 0.5 |  | 6/9/2023 5:14:05 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\std 0.5.510 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 5.708% | 8.298% | 6.026% | 68.75% | 9.481% | 17.59% | 6.522% | 17.36% | 4.966% | 5.131% | 2.811% | 0.6813% | 1.216% | 6.732% | 4.14% | 5.556% | 5.046% | 1.679% | 5.469% | 5.266% | 19.3% | 2.117% | 2.223% | 17.3% | 9.986% | 13.01% | 12.69% | 10.06% | 9.285% | 1.116% | 2.802% | 0.203% | 1.675% | 3.227% | 1.293% | 0.5741% | 2.32% | 1.435% | 3.038% | 2.133% | 2.73% | 0.5296%
157 | std 1 |  | 6/9/2023 5:19:54 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\std 1.511 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.924% | 8.959% | 2.203% | 7.036% | 192.5% | 18.1% | 6.552% | 6.089% | 5.476% | 2.748% | 3.634% | 4.743% | 0.3419% | 2.969% | 4.657% | 4.928% | 3.839% | 1.067% | 1.762% | 1.674% | 20.18% | 3.112% | 3.805% | 4.254% | 17.7% | 7.906% | 7.133% | 5.325% | 8.953% | 2.592% | 1.914% | 0.515% | 1.524% | 2.639% | 2.548% | 1.507% | 1.548% | 1.744% | 2.016% | 1.646% | 0.837% | 2.548%
158 | std 5 |  | 6/9/2023 5:25:42 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\std 5.512 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.322% | 5.491% | 4.251% | 15.48% | 2.239% | 4.792% | 2.204% | 5.175% | 1.661% | 2.096% | 6.05% | 7.581% | 1.738% | 1.538% | 3.053% | 4.28% | 2.023% | 6.51% | 6.733% | 3.362% | 6.293% | 0.9413% | 1.207% | 3.823% | 6.6% | 0.4274% | 2.901% | 4.581% | 5.026% | 6.23% | 0.7335% | 1.261% | 0.3108% | 1.144% | 5.056% | 0.4924% | 2.468% | 1.677% | 5.796% | 6.585% | 1.134% | 1.183%
159 | std 10 |  | 6/9/2023 5:31:30 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\std 10.513 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.055% | 4.634% | 2.152% | 15.57% | 1.518% | 3.603% | 2.981% | 4.597% | 0.6463% | 3.424% | 1.185% | 2.35% | 1.111% | 1.104% | 0.8255% | 1.141% | 1.3% | 1.579% | 0.8486% | 1.841% | 1.907% | 1.812% | 0.6586% | 4.095% | 0.1667% | 2.459% | 0.3876% | 1.606% | 3.889% | 0.9723% | 0.2736% | 1.543% | 0.7209% | 1.525% | 1.776% | 1.106% | 0.3193% | 2.394% | 1.382% | 1.151% | 1.538% | 1.682%
160 | LBLANK |  | 6/9/2023 5:37:19 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.514 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 5.887% | 4.398% | 0.08874% | 0.1209% | 1.321% | 27.98% | 808.3% | 1.343e-06% | 32.06% | 63.44% | 39.93% | 0.883% | 1.067% | 3.433% | 38.95% | 166.1% | 24.88% | 16.34% | 1.32% | 83.98% | 323.9% | 0.6162% | 79.5% | 30.72% | 0% | 2124% | 148% | 8275% | 0% | 17.64% | 7.384% | 1.101% | 1.003% | 1.789% | 1.206% | 4.265% | 1.353% | 1.771% | 199% | 13.69% | 18.21% | 8.247%
161 | CCV 1ppb TRACE |  | 6/9/2023 5:43:07 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.515 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10.24% | 9.453% | 1.165% | 56.84% | 7.465% | 31.08% | 9.255% | 14.56% | 7.835% | 5.327% | 1.703% | 0.8516% | 0.9503% | 23.83% | 6.137% | 6.553% | 11.67% | 2.695% | 0.04525% | 2.35% | 14% | 1.092% | 1.434% | 5.952% | 18.21% | 2.681% | 5.53% | 13.23% | 8.371% | 1.608% | 3.128% | 1.367% | 0.4955% | 6.739% | 0.1628% | 1.393% | 1.109% | 0.5009% | 2.413% | 0.9292% | 1.607% | 1.668%
162 | CCV 1ppb TRACE |  | 6/9/2023 5:49:25 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.516 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4.816% | 1.523% | 2.935% | 7.275% | 2.997% | 3.66% | 4.012% | 6.026% | 9.329% | 3.856% | 2.958% | 5.47% | 6.54% | 5.144% | 4.622% | 2.074% | 2.654% | 1.956% | 3.751% | 3.749% | 12.37% | 4.069% | 6.078% | 12.37% | 15.24% | 3.65% | 7.184% | 10.47% | 18.33% | 0.7193% | 5.747% | 1.563% | 5.064% | 0.8824% | 1.512% | 1.7% | 1.828% | 3.902% | 2.438% | 0.4244% | 2.513% | 1.264%
163 | CCV 10 ppb MAJORS |  | 6/9/2023 5:55:13 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.517 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.205% | 3.032% | 1.471% | 14.68% | 0.9839% | 1.378% | 2.402% | 1.374% | 1.625% | 0.8472% | 1.039% | 1.458% | 0.7819% | 2.638% | 1.411% | 1.708% | 1.422% | 0.5708% | 0.5821% | 1.435% | 1.329% | 3.961% | 0.4614% | 1.963% | 3.18% | 1.652% | 0.776% | 2.065% | 6.388% | 0.9623% | 0.5113% | 0.7089% | 0.2961% | 1.216% | 1.158% | 23.71% | 1.911% | 4.489% | 1.654% | 1.263% | 1.312% | 23.36%
164 | CCV 10 ppb MAJORS |  | 6/9/2023 6:01:30 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.518 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4.962% | 6.161% | 0.569% | 14.89% | 1.428% | 7.478% | 6.147% | 4.902% | 6.247% | 4.351% | 0.1003% | 0.7372% | 0.7321% | 5.53% | 4.061% | 6.237% | 5.87% | 1.386% | 1.012% | 1.673% | 2.445% | 2.186% | 0.4337% | 3.275% | 2.525% | 2.61% | 2.19% | 2.928% | 0.8907% | 0.5577% | 1.048% | 1.532% | 0.8206% | 4.924% | 0.5745% | 0.7944% | 1.476% | 1.656% | 0.5124% | 0.1178% | 1.098% | 1.932%
165 | std 25 |  | 6/9/2023 6:07:19 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\std 25.519 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.176% | 3.644% | 1.58% | 7.445% | 0.9921% | 8.197% | 2.578% | 6.168% | 3.298% | 3.003% | 0.9202% | 0.4047% | 1.071% | 5.513% | 3.129% | 2.405% | 4.922% | 1.606% | 0.9661% | 0.7193% | 5.706% | 0.718% | 0.592% | 2.901% | 2.05% | 2.215% | 0.2515% | 0.485% | 1.589% | 1.031% | 0.9133% | 1.375% | 1.127% | 3.284% | 0.9746% | 0.9442% | 0.6637% | 0.3031% | 1.71% | 0.5836% | 0.7221% | 0.8774%
166 | std 50 |  | 6/9/2023 6:13:06 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\std 50.520 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.7463% | 0.3294% | 2.691% | 2.032% | 1.259% | 2.406% | 1.939% | 1.504% | 1.572% | 2.351% | 0.9776% | 1.227% | 3.2% | 0.9811% | 2.075% | 1.414% | 2.094% | 0.2198% | 1.525% | 1.457% | 3.145% | 3.163% | 3.343% | 0.7246% | 4.44% | 2.618% | 2.156% | 1.403% | 1.497% | 0.6338% | 2.722% | 0.5775% | 2.68% | 1.784% | 1.458% | 0.6424% | 1.353% | 2.914% | 0.4949% | 1.412% | 2.071% | 1.59%
167 | std 50 |  | 6/9/2023 6:19:23 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\std 50.521 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.95% | 3.755% | 3.192% | 1.972% | 0.6797% | 5.02% | 3.278% | 5.205% | 5.087% | 5.581% | 9.292% | 8.669% | 1.528% | 5.332% | 4.579% | 4.257% | 3.683% | 8.471% | 9.265% | 2.334% | 2.533% | 3.722% | 0.4274% | 1.148% | 1.902% | 2.324% | 1.006% | 2.922% | 1.836% | 8.667% | 1.038% | 1.117% | 1.113% | 3.678% | 7.884% | 1.085% | 1.03% | 2.887% | 8.366% | 8.807% | 1.437% | 0.7407%
168 | std 100 |  | 6/9/2023 6:25:11 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\std 100.522 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.768% | 2.826% | 0.8988% | 2.491% | 0.9708% | 3.456% | 1.161% | 2.354% | 3.269% | 3.302% | 0.8056% | 1.222% | 0.7578% | 2.6% | 2.587% | 2.338% | 1.254% | 1.729% | 1.25% | 0.9519% | 4.169% | 1.415% | 0.6251% | 1.051% | 2.572% | 1.7% | 1.15% | 1.106% | 1.769% | 0.2886% | 0.5067% | 1.628% | 0.7071% | 2.297% | 1.782% | 0.6804% | 2.241% | 1.334% | 1.259% | 0.8655% | 1.199% | 0.291%
169 | std 100 |  | 6/9/2023 6:31:28 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\std 100.523 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.599% | 2.293% | 0.7628% | 2.212% | 0.4017% | 3.062% | 1.315% | 3.065% | 3.353% | 3.193% | 1.015% | 0.3053% | 0.1406% | 2.001% | 2.443% | 2.066% | 2.468% | 0.3275% | 0.4105% | 1.041% | 3.07% | 1.333% | 0.4331% | 1.422% | 0.899% | 1.1% | 0.728% | 0.8383% | 1.541% | 1.058% | 1.215% | 1.481% | 1.029% | 1.933% | 0.4558% | 0.8705% | 1.048% | 0.8905% | 0.4738% | 0.5588% | 1.111% | 9.818%
170 | LBLANK |  | 6/9/2023 6:37:16 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.524 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 38.08% | 4.615% | 0.2505% | 4.329% | 1.563% | 39% | 210.1% | 1.343e-06% | 29.47% | 28.1% | 34.47% | 0.9887% | 1.689% | 15.05% | 33.95% | 30.26% | 93.38% | 17.85% | 1.505% | 2329% | 66.52% | 1.938% | 10.85% | 29.78% | 185.7% | 48.45% | 37.64% | 355.3% | 0% | 1.377% | 12.91% | 2.775% | 1.653% | 5.343% | 0.9078% | 1.72% | 0.4196% | 2.227% | 78.77% | 68.73% | 9.792% | 30.73%
171 | PE ICV |  | 6/9/2023 6:43:04 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\PE ICV.525 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.507% | 4.976% | 2.226% | 5.03% | 0.8545% | 3.392% | 7.102% | 6.12% | 5.737% | 5.291% | 0.6197% | 0.3233% | 2.607% | 6.858% | 5.395% | 5.125% | 6.097% | 0.8313% | 1.955% | 1.823% | 5.666% | 1.184% | 0.6763% | 3.001% | 1.015% | 2.585% | 3.963% | 3.468% | 4.794% | 1.057% | 0.8887% | 1.004% | 1.806% | 5.588% | 1.091% | 0.3502% | 1.686% | 1.103% | 1.421% | 0.6274% | 1.567% | 2.404%
172 | 5ppb CCV |  | 6/9/2023 6:48:52 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5ppb CCV.526 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.3498% | 3.469% | 6.136% | 10.58% | 0.9473% | 4.171% | 1.299% | 7.202% | 1.576% | 0.241% | 0.5577% | 5.094% | 1.365% | 0.914% | 0.6916% | 0.4688% | 2.307% | 1.042% | 1.717% | 5.809% | 6.145% | 1.502% | 0.5202% | 3.91% | 3.439% | 2.455% | 2.572% | 8.059% | 0.5441% | 0.2872% | 0.1957% | 4.711% | 2.128% | 0.7945% | 0.3058% | 1.024% | 1.975% | 3.108% | 0.8443% | 1.062% | 5.079% | 1.217%
173 | 25ppb CCV |  | 6/9/2023 6:54:40 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\25ppb CCV.527 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4.59% | 4.059% | 2.015% | 8.253% | 0.9797% | 3.435% | 2.431% | 4.887% | 2.416% | 2.81% | 1.636% | 2.618% | 0.5226% | 4.937% | 2.789% | 2.735% | 2.062% | 1.218% | 1.06% | 0.7679% | 4.783% | 0.3485% | 1.006% | 1.789% | 1.971% | 1.972% | 0.5984% | 1.361% | 0.5854% | 0.7185% | 1.03% | 1.474% | 1.228% | 1.598% | 1.323% | 0.5134% | 1.473% | 0.6235% | 0.4157% | 0.9199% | 1.413% | 0.93%
174 | 8_30_21_W_1 |  | 6/9/2023 7:00:27 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_W_1.528 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 7.041% | 5.239% | 2.534% | 7.333% | 23.18% | 2.054% | 15.64% | 2.405% | 20.48% | 3.049% | 11.54% | 2.857% | 0.2741% | 3.915% | 30.67% | 6.219% | 4.346% | 2.036% | 5.668% | 4.917% | 17.37% | 1.562% | 3.034% | 24.23% | 228.2% | 77.13% | 66.99% | 560.6% | 124.6% | 2.615% | 1.174% | 1.916% | 1.18% | 2.608% | 2.37% | 1.791% | 1.311% | 1.401% | 19.74% | 20.14% | 4.853% | 4.683%
175 | LBLANK |  | 6/9/2023 7:06:16 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.529 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 11.94% | 4.41% | 0.04163% | 3.986% | 1.688% | 26.88% | 148.6% | 1.343e-06% | 27.43% | 161.3% | 39.65% | 1.181% | 2.072% | 6.074% | 22.92% | 69.02% | 24.54% | 8.059% | 1.963% | 128% | 2608% | 0.6816% | 101.3% | 21.17% | 50.09% | 33300% | 69.14% | 16.97% | 260.7% | 10.12% | 19.51% | 4.98% | 1.15% | 2.6% | 0.7163% | 1.393% | 1.442% | 2.139% | 92.32% | 13.78% | 25.49% | 1.845%
176 | CCV 1ppb TRACE |  | 6/9/2023 7:12:04 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.530 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.096% | 8.791% | 2.119% | 36.62% | 8.307% | 11.66% | 9.481% | 4.025% | 5.263% | 5.185% | 2.633% | 8.386% | 12.71% | 16.64% | 2.271% | 1.033% | 2.144% | 3.186% | 0.6976% | 6.556% | 18.16% | 5.99% | 10.09% | 4.892% | 5.898% | 15.73% | 14.73% | 4.445% | 13.02% | 2.287% | 10.71% | 1.685% | 8.847% | 0.06391% | 0.4677% | 1.674% | 3.025% | 5.982% | 2.145% | 1.546% | 1.159% | 0.5529%
177 | CCV 1ppb TRACE |  | 6/9/2023 7:18:22 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.531 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 6.238% | 2.367e-06% | 0.1594% | 0% | 0.235% | 3.066% | 140.4% | 1.343e-06% | 0% | 568.6% | 50.75% | 1.455% | 0.9682% | 2.287% | 49.17% | 82.52% | 13.37% | 1.243% | 0.3661% | 2.148% | 412335% | 0.08493% | 4.184% | 2.3% | 48.14% | 188.4% | 43.97% | 93.2% | 0% | 5.14% | 1.887% | 2.553% | 1.899% | 5.267% | 2.289% | 4.948% | 7.384% | 4.465% | 47.96% | 17.04% | 10.69% | 1.282%
178 | LFB |  | 6/9/2023 7:24:08 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.532 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.994% | 4.674% | 1.721% | 1.702% | 1.205% | 5.903% | 1.74% | 3.07% | 3.286% | 1.969% | 0.202% | 1.012% | 0.9918% | 3.176% | 1.946% | 2.801% | 4.394% | 1.548% | 1.591% | 2.928% | 1.102% | 3.37% | 2.385% | 2.264% | 3.129% | 0.4101% | 2.009% | 3.492% | 2.506% | 2.312% | 0.9785% | 3.142% | 1.134% | 2.348% | 1.089% | 0.4019% | 1.092% | 4.363% | 1.284% | 1.563% | 2.897% | 1.129%
179 | LFB |  | 6/9/2023 7:30:24 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.533 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4.066% | 1.718% | 0.3604% | 5.98% | 0.621% | 8.474% | 2.85% | 2.742% | 4.5% | 5.463% | 0.5772% | 0.1967% | 2.567% | 5.021% | 2.151% | 3.469% | 3.617% | 0.5835% | 0.4692% | 2.158% | 5.118% | 1.841% | 2.513% | 4.607% | 6.218% | 1.491% | 2.875% | 2.797% | 2.018% | 1.238% | 3.546% | 0.9668% | 1.803% | 2.092% | 0.8352% | 0.1633% | 0.9283% | 0.8912% | 2.159% | 0.2405% | 0.1738% | 0.828%
180 | CCV 10 ppb MAJORS |  | 6/9/2023 7:36:12 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.534 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4.478% | 4.13% | 0.8609% | 5.552% | 0.2169% | 6.816% | 4.667% | 6.945% | 6.073% | 3.249% | 1.391% | 2.069% | 1.191% | 5.693% | 4.701% | 5.062% | 4.557% | 1.038% | 1.771% | 2.275% | 8.046% | 2.05% | 0.4987% | 0.9237% | 4.923% | 1.252% | 1.099% | 4.086% | 6.472% | 1.093% | 0.2538% | 1.766% | 0.9397% | 3.918% | 1.11% | 1.656% | 0.3% | 2.833% | 0.9885% | 1.076% | 1.259% | 1.852%
181 | CCV 10 ppb MAJORS |  | 6/9/2023 7:42:30 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.535 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 5.292% | 4.558% | 0.3193% | 3.284% | 1.502% | 5.498% | 6.339% | 4.168% | 5.459% | 5.05% | 0.909% | 0.7093% | 1.981% | 7.816% | 8.37% | 3.609% | 6.844% | 1.137% | 0.9539% | 1.504% | 9.195% | 1.206% | 0.9484% | 2.495% | 5.745% | 1.847% | 1.837% | 3.558% | 2.106% | 0.9962% | 2.408% | 0.6312% | 1.746% | 4.298% | 1.738% | 4.305% | 1.668% | 2.034% | 1.246% | 1.013% | 0.6264% | 4.33%
182 | 8_30_21_W_6 |  | 6/9/2023 7:48:18 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_W_6.536 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 15.4% | 15.35% | 0.2711% | 13.09% | 4.659% | 18.57% | 13.55% | 12.53% | 13.93% | 25.43% | 6.992% | 2.639% | 1.809% | 11.75% | 14.98% | 12.15% | 8.183% | 3.278% | 5.822% | 9.032% | 20.8% | 1.333% | 0.7104% | 11% | 62.38% | 10.27% | 34.91% | 584.1% | 14.32% | 2.055% | 0.6355% | 1.034% | 0.6132% | 9.924% | 3.099% | 2.839% | 1.362% | 2.763% | 61.95% | 20.09% | 2.762% | 4.479%
183 | 8_30_21_W_6_LD |  | 6/9/2023 7:54:05 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_W_6_LD.537 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.388% | 6.065% | 2.4% | 9.644% | 2.462% | 10.68% | 15.7% | 2.72% | 16.93% | 11.3% | 7.361% | 3.711% | 6.746% | 0.7331% | 8.93% | 6.851% | 3.2% | 3.435% | 0.7379% | 6.462% | 16.72% | 5.604% | 4.446% | 28.32% | 2270% | 25.67% | 18.19% | 46.37% | 48.28% | 3.528% | 7.568% | 1.231% | 5.461% | 1.367% | 2.412% | 2.559% | 1.457% | 7.367% | 23.31% | 14.28% | 5.628% | 0.974%
184 | 8_30_21_W_6_LD |  | 6/9/2023 8:00:20 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_W_6_LD.538 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.048% | 12.47% | 0.5298% | 3.208% | 0.5623% | 3.689% | 19.76% | 7.6% | 19.01% | 43.68% | 12.35% | 1.446% | 1.171% | 5.582% | 9.424% | 16.28% | 6.742% | 1.272% | 0.8924% | 0.7245% | 6.062% | 2.682% | 2.67% | 32.23% | 163.5% | 46.24% | 38.05% | 55.28% | 69.25% | 1.686% | 1.426% | 1.037% | 0.2317% | 5.866% | 0.7221% | 11.92% | 1.205% | 0.8942% | 14.94% | 23.97% | 4.894% | 13.78%
185 | 8_30_21_W_6_DUP |  | 6/9/2023 8:06:08 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_W_6_DUP.539 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.452% | 3.872% | 2.28% | 6.516% | 2.098% | 6.057% | 3.876% | 6.373% | 19.05% | 36.26% | 9.129% | 1.83% | 1.908% | 3.913% | 19.89% | 11.19% | 0.6886% | 2.021% | 1.436% | 3.472% | 9.85% | 2.949% | 2.094% | 13.39% | 429.2% | 57.47% | 43.84% | 70.7% | 125.5% | 1.385% | 2.651% | 3.656% | 1.327% | 0.5923% | 1.53% | 1.609% | 1.009% | 4.07% | 22.82% | 12.67% | 0.9887% | 5.083%
186 | 8_30_21_W_6 |  | 6/9/2023 8:11:56 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_W_6.540 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 6.46% | 5.284% | 2.535% | 6.086% | 20.77% | 3.579% | 18.09% | 5.019% | 5.207% | 32.36% | 11.66% | 1.329% | 1.042% | 4.695% | 13.71% | 46.07% | 6.145% | 2.12% | 1.801% | 1.708% | 6.263% | 1.883% | 2.666% | 12.33% | 55.96% | 98.37% | 56.3% | 205.7% | 91.48% | 0.4099% | 1.687% | 1.139% | 1.143% | 2.507% | 0.671% | 29.4% | 8.17% | 3.673% | 12.68% | 30.63% | 4.821% | 30.82%
187 | 8_30_21_W_6_DUP_LSP |  | 6/9/2023 8:17:43 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_W_6_DUP_LSP.541 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.001% | 3.341% | 27.15% | 8.439% | 0.6343% | 6.243% | 5.854% | 4.977% | 4.712% | 6.375% | 1.388% | 0.7777% | 0.7105% | 4.254% | 5.191% | 4.995% | 4.63% | 0.7684% | 1.008% | 26.43% | 8.012% | 0.8958% | 0.3974% | 30.91% | 1.618% | 1.383% | 2.027% | 24.76% | 2.694% | 0.4598% | 1.386% | 23.01% | 0.3699% | 3.282% | 0.9642% | 1.085% | 2.727% | 0.9985% | 0.8698% | 0.6223% | 26.52% | 1.845%
188 | 8_30_21_W_6_DUP_LSP |  | 6/9/2023 8:24:01 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_W_6_DUP_LSP.542 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.492% | 3.523% | 3.049% | 6.238% | 0.9339% | 6.973% | 2.351% | 3.663% | 4.047% | 3.548% | 1.192% | 1.517% | 0.5419% | 3.566% | 4.574% | 3.429% | 4.801% | 0.8837% | 0.6958% | 0.859% | 4.563% | 2.481% | 1.286% | 0.503% | 0.7655% | 2.404% | 1.345% | 1.831% | 2.647% | 1.32% | 0.4473% | 0.3017% | 0.6686% | 3.76% | 0.3664% | 0.1975% | 1.051% | 1.714% | 1.804% | 0.9137% | 1.416% | 1.536%
189 | 8_30_21_W_8 |  | 6/9/2023 8:29:49 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_W_8.543 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 13.86% | 6.71% | 0.8871% | 12.22% | 8.516% | 13.8% | 12.48% | 12.29% | 2.526% | 9.115% | 11.13% | 2.942% | 0.7847% | 11.26% | 27.98% | 13.98% | 10.53% | 2.496% | 4.542% | 4.261% | 27% | 3.703% | 0.4845% | 8.316% | 163.2% | 20.94% | 25.97% | 32.11% | 52.82% | 2.549% | 1.609% | 0.3451% | 0.3964% | 8.808% | 3.149% | 0.07555% | 5.605% | 3.734% | 40.09% | 25.43% | 3.151% | 3.287%
190 | 8_30_21_W_8 |  | 6/9/2023 8:36:07 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_W_8.544 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.645% | 10.91% | 2.38% | 10.81% | 2.272% | 9.64% | 14.42% | 3.375% | 2.319% | 14.7% | 19.73% | 1.331% | 0.6009% | 1.482% | 31.81% | 26.82% | 5.02% | 0.7583% | 2.25% | 0.3909% | 20.32% | 3.863% | 1.451% | 22.46% | 249.6% | 13.33% | 34.17% | 59.16% | 102.3% | 3.203% | 1.551% | 0.8612% | 0.6009% | 1.655% | 1.862% | 1.852% | 1.763% | 2.611% | 45.95% | 16.14% | 3.158% | 3.331%
191 | 8_30_21_W_8_SP |  | 6/9/2023 8:41:55 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_W_8_SP.545 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 7.97% | 11.05% | 0.4091% | 6.002% | 0.9079% | 14.52% | 10.04% | 11.44% | 16.48% | 9.632% | 3.906% | 5.818% | 1.09% | 9.494% | 5.053% | 4.44% | 10.97% | 5.24% | 4.373% | 0.3944% | 8.807% | 2.24% | 1.966% | 6.148% | 15.78% | 7.02% | 6.004% | 14.07% | 11.94% | 4.442% | 0.3625% | 0.8077% | 0.8928% | 5.821% | 4.121% | 1.318% | 1.526% | 1.585% | 2.616% | 5.145% | 1.678% | 1.125%
192 | LBLANK |  | 6/9/2023 8:47:44 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.546 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 5.924% | 2.367e-06% | 0.09846% | 4.349% | 3.514% | 62.87% | 1288% | 1.343e-06% | 0% | 62.1% | 36.35% | 1.288% | 2.985% | 13.45% | 58.7% | 26.2% | 28.49% | 7.596% | 2.094% | 103.5% | 1111% | 1.283% | 191.8% | 10.59% | 58.73% | 365.5% | 137.1% | 556.1% | 0% | 7.389% | 3.013% | 0.3913% | 7.999% | 4.817% | 2.381% | 0.7951% | 2.896% | 7.852% | 107.3% | 54.43% | 15.37% | 3.776%
193 | CCV 1ppb TRACE |  | 6/9/2023 8:53:33 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.547 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 7.024% | 19.02% | 8.744% | 6.25% | 7.15% | 17.48% | 7.718% | 19.76% | 2.281% | 8.686% | 2.848% | 1.729% | 2.18% | 16.04% | 9.418% | 9.146% | 9.476% | 1.9% | 2.495% | 7.315% | 8.002% | 2.315% | 2.217% | 3.847% | 11.83% | 12.11% | 2.417% | 13.45% | 3.649% | 1.59% | 0.655% | 4.213% | 0.6761% | 6.604% | 0.6657% | 2.555% | 2.122% | 1.8% | 2.576% | 2.799% | 5.082% | 1.595%
194 | CCV 1ppb TRACE |  | 6/9/2023 8:59:50 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.548 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 6.051% | 9.14% | 4.178% | 30.5% | 4.176% | 19.8% | 3.414% | 2.695% | 5.665% | 4.836% | 3.641% | 1.406% | 0.6465% | 19.81% | 5.792% | 2.521% | 3.909% | 1.333% | 3.248% | 5.336% | 19.64% | 4.917% | 2.108% | 16.92% | 13.27% | 5.178% | 3.553% | 10.36% | 10.54% | 0.6238% | 2.016% | 0.51% | 0.9168% | 4.647% | 1.172% | 0.8966% | 0.5803% | 3.044% | 2.001% | 1.95% | 0.451% | 1.062%
195 | CCV 10 ppb MAJORS |  | 6/9/2023 9:05:39 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.549 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 7.743% | 7.759% | 1.75% | 6.668% | 1.424% | 3.139% | 5.494% | 7.472% | 6.522% | 4.361% | 0.3646% | 1.787% | 2.535% | 7.456% | 5.523% | 9.39% | 8.203% | 0.835% | 1.026% | 1.952% | 7.912% | 2.52% | 2.846% | 0.8718% | 2.828% | 4.348% | 3.244% | 3.961% | 2.009% | 0.318% | 2.221% | 2.923% | 2.596% | 5.936% | 1.276% | 0.3865% | 0.8899% | 2.217% | 1.541% | 0.9358% | 2.823% | 1.808%
196 | CCV 10 ppb MAJORS |  | 6/9/2023 9:11:56 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.550 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.254% | 3.334% | 1.767% | 7.041% | 1.681% | 2.113% | 3.32% | 3.161% | 0.2662% | 0.5944% | 1.361% | 0.8754% | 1.332% | 0.7173% | 0.4369% | 1.272% | 0.4594% | 1.424% | 0.8357% | 1.029% | 6.151% | 5.06% | 0.5289% | 1.175% | 5.416% | 3.064% | 2.069% | 4.155% | 8.438% | 0.5508% | 0.3467% | 0.8004% | 1.302% | 0.9883% | 1.996% | 1.524% | 1.028% | 4.687% | 0.7146% | 1.249% | 1.135% | 0.4573%
197 | 8_30_21_W_9 |  | 6/9/2023 9:17:44 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_W_9.551 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4.491% | 4.01% | 5.066% | 8.792% | 1.263% | 13.55% | 12.11% | 4.77% | 11.32% | 26.06% | 16.38% | 0.8198% | 2.05% | 1.364% | 29.47% | 40.14% | 6.521% | 0.9432% | 3.596% | 7.192% | 12.8% | 3.784% | 3.875% | 72.03% | 16720% | 105.4% | 52.56% | 139.3% | 98.29% | 0.1882% | 2.238% | 4.426% | 1.337% | 1.676% | 0.335% | 1.3% | 0.8568% | 1.434% | 27.62% | 33.66% | 7.996% | 4.129%
198 | 8_30_21_W_10 |  | 6/9/2023 9:23:32 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_W_10.552 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.813% | 3.422% | 2.599% | 10.72% | 5.757% | 5.83% | 11.31% | 4.95% | 15.69% | 13.85% | 8.034% | 4.437% | 3.401% | 4.093% | 19.69% | 27.25% | 7.728% | 2.373% | 7.806% | 7.543% | 23.01% | 1.171% | 2.086% | 23.91% | 143% | 16.18% | 34.35% | 22.89% | 73.94% | 2.751% | 3.288% | 3.153% | 1.325% | 3.178% | 2.383% | 0.4817% | 3.266% | 1.412% | 48.22% | 2.793% | 1.865% | 1.408%
199 | 8_30_21_W_11 |  | 6/9/2023 9:29:20 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_W_11.553 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4.389% | 0.6762% | 2.974% | 4.791% | 2.404% | 7.706% | 2.971% | 4.708% | 7.2% | 43.81% | 10.79% | 9.834% | 2.53% | 3.52% | 19.11% | 35.55% | 0.8942% | 7.438% | 13.47% | 6.193% | 15.71% | 1.237% | 1.547% | 29.99% | 182.3% | 33.44% | 31.37% | 73.53% | 37% | 8.359% | 2.674% | 1.405% | 0.9277% | 0.7839% | 6.11% | 3.35% | 1.79% | 0.9791% | 54.05% | 6.774% | 1.632% | 3.564%
200 | 8_30_21_F_1 |  | 6/9/2023 9:35:08 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_F_1.554 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.818% | 4.395% | 1.748% | 12.91% | 2.608% | 3.488% | 5.899% | 3.35% | 11.41% | 10.3% | 3.948% | 1.156% | 1.699% | 7.664% | 17.86% | 4.757% | 6.039% | 1.576% | 0.517% | 0.3426% | 4.485% | 3.034% | 0.5212% | 17.29% | 84.8% | 26.32% | 31.89% | 13.73% | 52.06% | 2.078% | 2.078% | 1.713% | 0.5226% | 2.909% | 0.521% | 0.8469% | 0.8928% | 3.183% | 16.56% | 10.51% | 4.856% | 1.446%
201 | 8_30_21_F_1 |  | 6/9/2023 9:41:25 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_F_1.555 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 7.53% | 8.627% | 3.193% | 8.484% | 3.606% | 9.825% | 11.26% | 10.82% | 15.33% | 10.75% | 5.088% | 3.175% | 0.335% | 8.748% | 4.998% | 8.996% | 7.026% | 0.5342% | 0.9663% | 1.259% | 16.98% | 1.365% | 1.645% | 20.9% | 18.29% | 30.86% | 33.6% | 21.6% | 23.53% | 1.141% | 0.4188% | 0.9547% | 0.4684% | 7.26% | 0.4881% | 0.6859% | 2.913% | 1.754% | 3.187% | 5.857% | 3.377% | 1.171%
202 | 8_30_21_F_1_DUP |  | 6/9/2023 9:47:13 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_F_1_DUP.556 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 6.548% | 5.553% | 17.84% | 13.89% | 2.308% | 5.251% | 5.929% | 4.311% | 5.597% | 2.617% | 1.93% | 0.6876% | 0.87% | 8.976% | 6.471% | 10.12% | 4.245% | 0.6914% | 0.883% | 14.17% | 10.45% | 1.903% | 1.672% | 54.13% | 74.29% | 39.4% | 9.34% | 25.91% | 70.1% | 0.1152% | 1.045% | 13.09% | 0.6088% | 4.189% | 0.942% | 1.752% | 1.458% | 1.68% | 25.82% | 21.32% | 13.01% | 1.424%
203 | 8_30_21_F_1_DUP |  | 6/9/2023 9:53:30 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_F_1_DUP.557 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 5.558% | 4.213% | 7.668% | 2.941% | 1.862% | 2.888% | 6.903% | 5.911% | 4.218% | 13.3% | 9.249% | 0.8739% | 1.148% | 2.167% | 29.69% | 7.366% | 2.644% | 0.5441% | 0.3573% | 6.312% | 4.046% | 1.089% | 1.206% | 24.15% | 52.57% | 13.15% | 51.78% | 7.264% | 34.23% | 1.222% | 0.8524% | 4.542% | 1.189% | 2.739% | 0.4843% | 0.9174% | 1.666% | 1.62% | 28.48% | 10.96% | 8.257% | 5.012%
204 | 8_30_21_F_6 |  | 6/9/2023 9:59:18 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_F_6.558 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.939% | 7.033% | 1.349% | 4.032% | 4.89% | 3.684% | 4.951% | 4.135% | 5.602% | 12.3% | 5.338% | 1.149% | 1.894% | 3.482% | 17.24% | 8.813% | 4.528% | 1.61% | 1.743% | 0.6164% | 6.319% | 5.037% | 1.5% | 10.73% | 83.09% | 21.13% | 21% | 9.295% | 23.68% | 1.589% | 1.746% | 0.9022% | 1.272% | 3.06% | 1.354% | 5.655% | 1.483% | 5.335% | 10.85% | 10.89% | 1.269% | 6.986%
205 | 8_30_21_F_6 |  | 6/9/2023 10:05:38 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_F_6.559 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10.18% | 12.02% | 1.289% | 16.7% | 5.852% | 14.83% | 14.09% | 12.7% | 15.22% | 22.56% | 2.635% | 1.275% | 1.952% | 18.47% | 10.2% | 9.566% | 13.7% | 1.628% | 0.384% | 2.068% | 6.79% | 5.309% | 1.288% | 11.12% | 94.1% | 19.27% | 9.018% | 20.71% | 0.8922% | 1.022% | 0.889% | 0.2916% | 0.8758% | 12.23% | 0.794% | 1.8% | 0.2391% | 5.492% | 12.07% | 12.45% | 2.974% | 1.463%
206 | 8_30_21_F_8 |  | 6/9/2023 10:11:26 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_F_8.560 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.255% | 3.768% | 3.273% | 1.158% | 2.254% | 2.003% | 5.937% | 2.208% | 10.5% | 27.42% | 4.147% | 1.184% | 8.02% | 6.344% | 23.99% | 19.65% | 4.337% | 0.9351% | 3.194% | 0.5447% | 10.08% | 7.39% | 8.555% | 35.17% | 72.19% | 34.21% | 15.73% | 5.41% | 60.17% | 1.083% | 8.542% | 1.055% | 6.394% | 3.857% | 0.8661% | 1.578% | 0.2328% | 5.7% | 13.17% | 3.299% | 3.91% | 1.675%
207 | 8_30_21_F_8 |  | 6/9/2023 10:17:43 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_F_8.561 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.693% | 4.847% | 2.374% | 28.26% | 3.888% | 3.363% | 2.204% | 1.808% | 2.666% | 5.608% | 1.908% | 1.509% | 5.159% | 2.629% | 5.586% | 13.8% | 0.2135% | 1.566% | 0.8485% | 3.101% | 6.826% | 3.873% | 5.116% | 7.968% | 17.03% | 20.58% | 25.93% | 14.28% | 4.72% | 0.502% | 5.52% | 0.171% | 4.617% | 0.9205% | 1.178% | 1.598% | 1.177% | 3.342% | 21.67% | 8.397% | 2.136% | 2.39%
208 | LBLANK |  | 6/9/2023 10:23:31 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.562 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 6.86% | 2.367e-06% | 0.3242% | 4.794% | 0.3724% | 43.76% | 6.051% | 1.343e-06% | 35.54% | 69.6% | 60.85% | 2.368% | 0.4645% | 13.4% | 0% | 19.88% | 25.96% | 10.17% | 3.571% | 100.6% | 1.756e-06% | 0.2731% | 58.49% | 17.36% | 0% | 2.889% | 62.51% | 579.3% | 102.6% | 5.045% | 19.51% | 2.063% | 1.995% | 0.6249% | 12.27% | 2.261% | 1.466% | 1.347% | 225.5% | 36.92% | 8.772% | 1.713%
209 | CCV 1ppb TRACE |  | 6/9/2023 10:29:20 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.563 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 8.107% | 8.048% | 3.358% | 40.7% | 2.384% | 5.548% | 5.101% | 4.262% | 2.089% | 4.863% | 0.8621% | 3.827% | 1.69% | 5.121% | 5.318% | 1.819% | 2.16% | 1.908% | 2.669% | 7.824% | 11.4% | 0.552% | 2.251% | 15.36% | 14.05% | 2.587% | 0.1879% | 10.17% | 14.63% | 1.099% | 1.698% | 0.9334% | 0.8896% | 1.613% | 1.208% | 0.4991% | 1.182% | 0.9958% | 1.905% | 0.2382% | 2.13% | 0.3499%
210 | CCV 1ppb TRACE |  | 6/9/2023 10:35:37 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.564 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.743% | 4.539% | 3.719% | 20.22% | 2.433% | 14.75% | 7.42% | 15.28% | 2.999% | 3.4% | 1.47% | 5.952% | 2.803% | 18.55% | 6.734% | 6.21% | 4.922% | 0.6822% | 1.862% | 4.413% | 18.82% | 3.318% | 2.154% | 1.126% | 10.59% | 5.083% | 5.158% | 4.663% | 3.446% | 0.5756% | 2.752% | 0.7658% | 1.496% | 5.835% | 0.5145% | 0.7479% | 1.299% | 2.912% | 0.4478% | 1.527% | 1.03% | 1.185%
211 | LFB |  | 6/9/2023 10:41:25 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.565 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.3722% | 0.7095% | 0.4092% | 2.302% | 1.635% | 2.961% | 2.288% | 2.799% | 2.072% | 0.7586% | 1.42% | 3.045% | 1.008% | 0.2743% | 0.8897% | 2.966% | 1.508% | 1.621% | 1.887% | 0.6751% | 0.5153% | 1.212% | 1.094% | 1.592% | 0.8952% | 2.448% | 3.068% | 3.443% | 3.982% | 1.435% | 1.842% | 0.6444% | 0.7507% | 0.3039% | 0.3871% | 0.5654% | 1.014% | 1.665% | 1.486% | 1.009% | 0.4028% | 1.691%
212 | LFB |  | 6/9/2023 10:47:41 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.566 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.363% | 0.9576% | 0.4975% | 10.58% | 3.418% | 4.138% | 1.906% | 4.119% | 3.203% | 0.586% | 0.8713% | 0.8031% | 2.468% | 1.838% | 0.9471% | 0.957% | 1.795% | 1.94% | 0.6845% | 1.999% | 6.535% | 4.748% | 1.459% | 1.278% | 1.138% | 2.432% | 2.577% | 0.8326% | 5.092% | 0.8677% | 1.182% | 1.263% | 1.28% | 1.553% | 1.24% | 0.9009% | 1.28% | 3.83% | 1.021% | 1.875% | 1.393% | 1.396%
213 | CCV 10 ppb MAJORS |  | 6/9/2023 10:53:35 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.567 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.116% | 1.419% | 1.892% | 11.33% | 3.979% | 1.882% | 2.378% | 2.882% | 3.499% | 1.75% | 0.9934% | 1.641% | 2.841% | 5.233% | 2.677% | 6.456% | 1.979% | 1.541% | 0.972% | 1.991% | 3.832% | 1.762% | 2.168% | 3.43% | 1.663% | 2.875% | 1.486% | 6.646% | 3.833% | 0.6272% | 1.671% | 0.9541% | 1.625% | 2.554% | 1.471% | 1.049% | 0.7704% | 1.192% | 1.844% | 1.14% | 0.7541% | 1.824%
214 | LBLANK |  | 6/9/2023 11:28:06 AM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.568 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth
215 | Cal blank |  | 6/9/2023 11:33:55 AM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\Cal blank.569 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth
216 | CAL BLANK |  | 6/9/2023 11:39:43 AM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CAL BLANK.570 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | [31 empty] | 0.6623% | 0.5934% | 0.8174% | 1.564% | 0.7852% | 1.523% | 0.8993%
217 | LOW 0.01 |  | 6/9/2023 11:45:32 AM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LOW 0.01.571 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | [31 empty] | 0.3673% | 1.84% | 0.9474% | 0.7389% | 1.658% | 0.8092% | 1.431%
218 | LOW 0.05 |  | 6/9/2023 11:51:20 AM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LOW 0.05.572 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | [8 empty] | 8.158% | 19.4% | 9.359% |  | 20.23% |  | 4.685% | 37.23% | 3.661% | 2.992% |  | 43.77% | 40.19% | 14.77% | 16.48% | [6 empty] | 14.47% | 1.201% | 1.424% | 1.024% | 0.635% | 0.6804% | 0.3827% | 1.29% | 1.942% | 4.04% | 5.517% | 2.236% | 1.877%
219 | LOW 0.1 |  | 6/9/2023 11:57:09 AM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LOW 0.1.573 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | [8 empty] | 14.52% | 5.52% | 3.847% |  | 1.591% |  | 5.302% | 3.323% | 6.302% | 4.49% |  | 18.92% | 28.53% | 5.948% | 4.728% | 13.89% | 13.94% | 42.43% | 5.763% | 27.66% | 36.73% | 5.669% | 0.8278% | 1.825% | 1.283% | 1.076% | 0.2006% | 1.706% | 1.319% | 1.875% | 4.93% | 0.8018% | 5.221% | 1.853%
220 | LOW 0.5 |  | 6/9/2023 12:02:57 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LOW 0.5.574 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth |  |  |  |  | 3.322% |  |  |  | 10.34% | 6.533% | 8.904% | 11.29% | 2.288% | 4.706% | 7.683% | 4.585% | 0.191% | 10.28% | 10.98% | 12.7% | 9.766% | 2.014% | 3.792% | 27.91% | 5.608% | 1.835% | 4.364% | 30.07% | 23.91% | 11.47% | 3.442% | 2.946% | 0.292% | 0.5819% | 7.365% | 14.01% | 0.7157% | 2.644% | 9.02% | 8.042% | 5.334% | 16.37%
221 | STD 1 |  | 6/9/2023 12:08:45 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 1.575 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.924% | 2.716% | 7.702% | 55.86% | 3.428% |  | 6.591% | 13.05% | 5.883% | 5.598% | 2.281% | 1.42% | 1.063% | 8.634% | 2.818% | 8.303% | 4.889% | 1.42% | 1.794% | 2.004% | 13.38% | 3.367% | 2.097% | 9.107% | 19.73% | 2.657% | 7.965% | 6.366% | 8.803% | 0.9724% | 0.6368% | 0.6116% | 0.7099% | 2.769% | 0.7788% | 1.58% | 2.355% | 2.789% | 2.468% | 0.8007% | 1.016% | 1.956%
222 | STD 5 |  | 6/9/2023 12:14:33 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 5.576 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.498% | 1.472% | 2.077% | 8.089% | 0.9462% | 5.991% | 3.129% | 2.419% | 0.8658% | 1.67% | 0.7557% | 3.465% | 1.997% | 3.909% | 1.084% | 1.334% | 0.8552% | 0.9415% | 1.091% | 1.891% | 13.44% | 2.882% | 2.072% | 3.484% | 6.277% | 4.369% | 5.434% | 5.056% | 2.557% | 0.9259% | 1.916% | 0.6203% | 2.11% | 1.185% | 0.5942% | 1.504% | 1.625% | 2.701% | 1.595% | 0.8763% | 1.427% | 1.644%
223 | STD 10 |  | 6/9/2023 12:20:21 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 10.577 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.1% | 2.517% | 1.957% | 10.46% | 2.287% | 9.367% | 4.132% | 2.674% | 3.772% | 4.212% | 1.472% | 1.384% | 1.066% | 4.805% | 2.225% | 4.068% | 3.547% | 0.4025% | 2.027% | 1.362% | 2.01% | 1.036% | 2.17% | 3.668% | 0.8594% | 0.1068% | 2.792% | 1.293% | 4.396% | 0.9713% | 1.559% | 0.9618% | 0.9593% | 3.411% | 0.9061% | 0.4659% | 0.1279% | 1.906% | 1.43% | 0.8215% | 0.6475% | 0.8374%
224 | STD 25 |  | 6/9/2023 12:26:09 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 25.578 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4.02% | 3.075% | 8.107% | 11.05% | 0.6036% | 5.732% | 5.581% | 5.775% | 2.015% | 3.802% | 0.5975% | 0.2141% | 0.9872% | 4.553% | 4.959% | 3.855% | 4.571% | 0.5216% | 0.5796% | 6.211% | 5.132% | 1.96% | 1.327% | 8.51% | 6.168% | 0.7182% | 1.338% | 6.46% | 0.1884% | 0.6202% | 0.8602% | 4.113% | 1.222% | 3.954% | 0.7926% | 1.09% | 1.664% | 2.126% | 1.297% | 0.2604% | 5.145% | 2.293%
225 | STD 50 |  | 6/9/2023 12:31:57 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 50.579 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.924% | 2.767% | 0.9789% | 3.799% | 2.465% | 1.341% | 3.821% | 1.875% | 1.943% | 1.457% | 2.304% | 2.225% | 0.909% | 2.548% | 1.949% | 2.43% | 2.054% | 2.115% | 2.199% | 0.433% | 3.287% | 2.918% | 1.018% | 3.148% | 0.5953% | 1.237% | 2.025% | 2.313% | 1.56% | 1.892% | 0.6722% | 0.4796% | 1.006% | 1.65% | 0.6768% | 6.87% | 1.442% | 2.755% | 2.178% | 2.295% | 0.4625% | 8.177%
226 | STD 100 |  | 6/9/2023 12:37:44 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 100.580 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9.915% | 9.293% | 15.6% | 7.434% | 0.6676% | 8.807% | 10.14% | 7.949% | 9.682% | 9.771% | 1.553% | 0.3685% | 0.6404% | 9.777% | 9.027% | 9.717% | 9.129% | 1.564% | 0.4478% | 15.53% | 8.463% | 1.675% | 0.7627% | 16.45% | 2.786% | 0.4332% | 1.536% | 14.58% | 0.8433% | 1.82% | 0.5626% | 13.82% | 0.7724% | 7.11% | 1.121% | 0.3927% | 0.5599% | 1.78% | 1.332% | 1.163% | 15.49% | 0.2586%
228 | LBLANK |  | 6/9/2023 12:43:32 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.581 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.619% | 14.47% | 1.949% | 8.149% | 1.371% | 40.6% | 26.65% | 0% | 84.43% | 12169% | 379.8% | 4.091% | 1.215% | 146.4% | 9.141% | 82.36% | 32.23% | 7.022% | 4.172% | 152.5% | 49.83% | 0.707% | 35.78% | 49.53% | 0% | 55.04% | 24.21% | 77.96% | 40.75% | 9.117% | 9.281% | 15.71% | 0.7355% | 0.8566% | 0.6528% | 0.7962% | 0.8173% | 2.215% | 57.51% | 140.8% | 19.02% | 9.897%
229 | LBLANK |  | 6/9/2023 1:13:17 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.583 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10.13% | 0% | 13.84% | 5.034% | 3.07% | 230.2% | 6.659% | 0% | 76.31% | 427% | 143% | 1.072% | 58.04% | 26.84% | 53.73% | 94.34% | 51.7% | 6.964% | 302.9% | 269.8% | 101.3% | 0.2534% | 35.98% | 199.1% | 41.25% | 54.8% | 42.81% | 138.8% | 0% | 1.318% | 16.81% | 15.24% | 0.4477% | 3.995% | 0.9617% | 0.8764% | 0.6516% | 3.854% | 31.2% | 32.41% | 11.29% | 6.864%
230 | LBLANK |  | 6/9/2023 1:19:06 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.584 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.68% | 0% | 0.3047% | 4.04% | 1.767% | 18.55% | 13.57% | 0% | 56.31% | 1.141% | 65.65% | 1.025% | 36.52% | 5.875% | 41.6% | 163.2% | 13.64% | 10.28% | 6.55% | 332.8% | 0% | 0.3052% | 65.1% | 269.5% | 0% | 55.34% | 21.37% | 848.8% | 0% | 11.14% | 6.872% | 1.166% | 0.3643% | 2.559% | 1.811% | 0.9652% | 2.079% | 0.8685% | 16.69% | 18.94% | 7.642% | 7.27%
231 | ICV PE |  | 6/9/2023 1:24:53 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\ICV PE.585 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.305% | 2.243% | 3.787% | 15.31% | 1.72% | 5.454% | 3.399% | 2.906% | 4.117% | 4.08% | 1.665% | 1.467% | 2.227% | 6.539% | 5.719% | 3.303% | 3.499% | 0.3588% | 0.7406% | 3.39% | 3.234% | 2.435% | 1.688% | 2.213% | 3.28% | 1.555% | 2.539% | 2.401% | 4.489% | 0.6035% | 0.4697% | 2.723% | 1.423% | 4.253% | 0.5821% | 0.7338% | 0.491% | 2.63% | 0.7522% | 0.9612% | 2.959% | 0.6399%
232 | ICV PE |  | 6/9/2023 1:31:10 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\ICV PE.586 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.145% | 4.137% | 1.581% | 4.997% | 0.7126% | 3.667% | 1.856% | 4.477% | 4.332% | 3.265% | 19.43% | 21.17% | 1.21% | 2.988% | 4.485% | 4.413% | 2.818% | 18.26% | 16.43% | 1.437% | 7.025% | 2.882% | 0.4041% | 2.636% | 2.559% | 0.4835% | 0.4682% | 0.5237% | 3.125% | 18.56% | 1.68% | 1.911% | 1.198% | 2.819% | 14.57% | 0.7403% | 0.8267% | 1.843% | 19.59% | 17.09% | 1.614% | 0.3482%
233 | CCV 1ppb TRACE |  | 6/9/2023 1:36:58 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.587 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9.766% | 18.05% | 1.649% | 13.76% | 24.16% | 11.52% | 10.41% | 7.88% | 11.87% | 9.978% | 1.724% | 3.643% | 1.99% | 46.73% | 11.43% | 13.97% | 14.91% | 1.68% | 3.377% | 4.979% | 12.89% | 0.8615% | 1.247% | 17.81% | 11.02% | 2.902% | 2.413% | 5.315% | 13.68% | 1.845% | 1.162% | 0.7307% | 0.5671% | 9.904% | 1.524% | 0.368% | 0.9798% | 1.813% | 0.6759% | 0.3892% | 1.774% | 2.334%
234 | CCV 1ppb TRACE |  | 6/9/2023 1:43:15 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.588 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.224% | 5.488% | 10.81% | 24.91% | 38.03% | 6.29% | 5.217% | 9.634% | 1.087% | 2.626% | 2.231% | 3.048% | 2.189% | 2.791% | 1.718% | 6.812% | 5.712% | 2.427% | 2.791% | 4.257% | 4.526% | 1.112% | 1.161% | 18.47% | 5.473% | 5.234% | 5.702% | 16.96% | 5.091% | 1.121% | 1.152% | 5.78% | 0.6314% | 2.097% | 1.011% | 0.4695% | 1.109% | 1.02% | 0.9868% | 0.1706% | 8.251% | 0.3223%
235 | LFB |  | 6/9/2023 1:49:03 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.589 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.971% | 2.693% | 2.087% | 2.464% | 2.013% | 1.746% | 2.139% | 3.888% | 2.24% | 2.933% | 0.566% | 2.028% | 1.221% | 3.181% | 0.8417% | 2.321% | 0.9363% | 0.3318% | 1.594% | 2.129% | 4.804% | 2.38% | 1.219% | 2.705% | 2.723% | 0.6988% | 1.754% | 3.182% | 5.108% | 0.2897% | 1.027% | 1.275% | 1.445% | 1.39% | 0.4243% | 1.086% | 1.012% | 2.508% | 0.623% | 0.2285% | 0.9035% | 1.009%
236 | LLOQ 0.01 |  | 6/9/2023 1:54:51 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LLOQ 0.01.590 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.2893% | 14.89% | 3.99% | 220.9% | 20.29% | 216.7% | 16.16% | 9.341% | 17.14% | 87.27% | 18.08% | 3.164% | 4.878% | 8.803% | 11.9% | 145.9% | 46.07% | 10.75% | 107.5% | 31.88% | 69.33% | 0.666% | 17.31% | 139.9% | 39.87% | 100.7% | 57.56% | 1442% | 0.8341% | 6.331% | 5.867% | 1.702% | 1.889% | 1.14% | 1.595% | 1.349% | 0.7058% | 4.215% | 7.97% | 15.3% | 2.014% | 4.699%
237 | CCV 10 ppb MAJORS |  | 6/9/2023 2:00:40 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.591 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.001% | 5.187% | 1.483% | 11.52% | 1.486% | 5.221% | 1.009% | 5.159% | 1.938% | 3.537% | 4.363% | 5.38% | 1.713% | 3.904% | 2.35% | 2.645% | 2.893% | 4.032% | 4.528% | 0.8566% | 5.814% | 1.457% | 1.392% | 2.487% | 8.134% | 3.491% | 1.345% | 3.213% | 6.148% | 3.165% | 0.9913% | 1.26% | 0.4545% | 3.227% | 3.636% | 1.393% | 1.728% | 2.451% | 5.345% | 4.442% | 2.354% | 1.365%
238 | CCV 10 ppb MAJORS |  | 6/9/2023 2:06:57 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.592 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.903% | 4.626% | 0.4587% | 9.105% | 2.421% | 3.949% | 3.686% | 1.546% | 1.112% | 0.4334% | 0.3089% | 0.6088% | 0.7884% | 3.248% | 2.253% | 0.4193% | 1.106% | 0.2633% | 1.198% | 0.3964% | 4.138% | 1.663% | 0.941% | 1.412% | 7.296% | 0.3244% | 1.709% | 1.227% | 2.412% | 0.3062% | 0.9473% | 0.2296% | 0.8037% | 1.232% | 0.336% | 0.8345% | 0.7409% | 1.433% | 1.181% | 0.304% | 0.8786% | 0.793%
239 | LLOQ 0.1 |  | 6/9/2023 2:12:45 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LLOQ 0.1.593 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10.65% | 2.291% | 1.396% | 7.909% | 9.729% | 25.1% | 9.767% | 22.46% | 5.976% | 7.188% | 2.736% | 2.304% | 3.603% | 177.7% | 6.558% | 21.27% | 4.665% | 4.967% | 44.59% | 17.17% | 58.26% | 0.5793% | 2.907% | 32.04% | 74.45% | 10.41% | 22.28% | 129.9% | 46.92% | 1.277% | 3.202% | 2.084% | 0.8149% | 1.134% | 0.3456% | 1.352% | 0.6818% | 1.147% | 5.171% | 1.98% | 1.465% | 1.867%
240 | 8_30_21_F_9 |  | 6/9/2023 2:19:27 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_F_9.594 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.285% | 3.063% | 3.453% | 11.04% | 2.263% | 4.453% | 5.46% | 3.663% | 4.256% | 19.64% | 5.764% | 2.603% | 1.777% | 0.8658% | 30.69% | 5.518% | 2.082% | 1.853% | 2.251% | 2.446% | 4.703% | 3.366% | 1.681% | 28.22% | 718.8% | 22.12% | 15.82% | 6.656% | 65.59% | 2.224% | 1.166% | 1.214% | 0.5334% | 0.7421% | 0.7146% | 0.5728% | 1.536% | 2.363% | 9.25% | 4.468% | 4.482% | 3.566%
241 | 8_30_21_F_10 |  | 6/9/2023 2:25:14 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_F_10.595 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.795% | 4.975% | 1.636% | 8.69% | 4.12% | 5.057% | 7.288% | 5.796% | 7.804% | 28.59% | 11.71% | 6.215% | 0.02903% | 9.897% | 21.09% | 20.13% | 5.482% | 6.433% | 5.322% | 1.46% | 4.761% | 1.922% | 1.079% | 30.63% | 59.86% | 33.73% | 19.55% | 16.14% | 65.38% | 6.337% | 0.9449% | 0.6525% | 0.3767% | 3.578% | 4.496% | 1.659% | 3.656% | 1.553% | 28.54% | 18.83% | 3.669% | 3.656%
242 | 8_30_21_F_11 |  | 6/9/2023 2:31:03 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_F_11.596 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.077% | 1.656% | 2.407% | 12.82% | 4.241% | 3.781% | 6.472% | 3.466% | 4.768% | 11.24% | 1.138% | 3.009% | 1.022% | 3.03% | 15.47% | 9.92% | 1.543% | 2.094% | 1.145% | 1.542% | 4.874% | 1.449% | 0.7656% | 9.815% | 12.89% | 17.49% | 12.03% | 27.05% | 53.31% | 1.178% | 1.023% | 0.5886% | 0.6527% | 1.832% | 1.149% | 9.273% | 1.648% | 2.532% | 7.363% | 10.7% | 2.005% | 11.26%
243 | 8_30_21_UF_1 |  | 6/9/2023 2:36:51 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_UF_1.597 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 5.877% | 9.434% | 3.663% | 25.94% | 0.9864% | 9.587% | 6.193% | 7.73% | 5.284% | 25.09% | 3.042% | 11.79% | 0.8765% | 23.83% | 16.43% | 8.003% | 6.271% | 1.487% | 1.044% | 1.191% | 7.214% | 3.476% | 1.051% | 4.477% | 70.28% | 18.15% | 8.93% | 21.53% | 637.7% | 1.527% | 0.6845% | 1.437% | 0.5175% | 5.774% | 1.353% | 1.215% | 1.494% | 4.192% | 14.04% | 8.23% | 5.444% | 2.637%
244 | 8_30_21_UF_6 |  | 6/9/2023 2:42:39 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_UF_6.598 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.2081% | 2.059% | 7.61% | 148.7% | 1.922% | 0.7858% | 2.4% | 2.179% | 3.276% | 20.36% | 20.94% | 63.08% | 11.71% | 16.83% | 17.24% | 13.26% | 0.163% | 13.87% | 16.73% | 1.526% | 9.814% | 13.19% | 11.39% | 13.19% | 138% | 18.93% | 6.889% | 13.37% | 75.09% | 14.54% | 10.97% | 0.5587% | 10.53% | 0.8394% | 11.8% | 0.9709% | 1.227% | 13.27% | 18.33% | 26.99% | 2.479% | 2.161%
245 | 8_30_21_UF_8 |  | 6/9/2023 2:48:27 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_UF_8.599 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.488% | 3.282% | 2.932% | 28.56% | 0.6545% | 3.572% | 4.818% | 1.357% | 6.672% | 12.31% | 9.112% | 21.61% | 0.8892% | 27.23% | 37.35% | 4.884% | 1.802% | 0.9575% | 2.719% | 0.8401% | 3.048% | 1.104% | 0.2311% | 16.23% | 40.87% | 21.36% | 15.33% | 38.75% | 88.12% | 1.116% | 0.2927% | 0.9796% | 0.3013% | 1.754% | 0.4663% | 0.7182% | 1.557% | 1.112% | 25.58% | 16.74% | 6.235% | 1.958%
246 | 8_30_21_UF_9 |  | 6/9/2023 2:54:15 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_UF_9.600 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4.436% | 4.695% | 24.54% | 80.78% | 0.7896% | 8.298% | 3.809% | 7.338% | 6.278% | 16.58% | 2.808% | 2.033% | 3.215% | 36.87% | 4.721% | 9.757% | 8.974% | 1.413% | 18.27% | 3.947% | 20.27% | 0.9898% | 2.037% | 41.75% | 121.7% | 14.84% | 8.583% | 8.618% | 41.63% | 0.953% | 3.687% | 3.237% | 2.818% | 5.445% | 0.6707% | 0.5695% | 1.097% | 2.003% | 3.5% | 17.81% | 4.097% | 1.672%
247 | 8_30_21_UF_9_LD |  | 6/9/2023 3:00:03 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_UF_9_LD.601 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.942% | 1.467% | 5.451% | 13.35% | 1.091% | 5.808% | 4.406% | 5.813% | 10.67% | 9.164% | 14.17% | 32.04% | 1.035% | 21.95% | 27.26% | 7.158% | 3.431% | 7.616% | 20.9% | 3.598% | 6.9% | 0.7477% | 0.2034% | 21.19% | 275.5% | 4.04% | 37.61% | 15.02% | 114.4% | 7.745% | 0.5259% | 2.52% | 0.3991% | 2.459% | 6.431% | 0.4269% | 1.181% | 1.131% | 28.44% | 17.16% | 5.594% | 0.8327%
248 | 8_30_21_UF_9_LD |  | 6/9/2023 3:06:19 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_UF_9_LD.602 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.159% | 6.412% | 4.087% | 29.27% | 1.008% | 3.548% | 7.561% | 3.022% | 5.694% | 5.169% | 5.54% | 11.26% | 1.071% | 35.63% | 22.9% | 8.509% | 2.432% | 1.417% | 6.241% | 1.809% | 12.4% | 1.409% | 1.44% | 35.35% | 137.5% | 27.28% | 12.94% | 5.695% | 28.07% | 1.092% | 0.7619% | 1.372% | 0.204% | 3.743% | 1.448% | 0.7068% | 0.9383% | 0.7132% | 3.997% | 15.09% | 5.28% | 4.105%
249 | 8_30_21_UF_10 |  | 6/9/2023 3:12:07 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_UF_10.603 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.49% | 1.927% | 0.854% | 25.56% | 0.763% | 2.407% | 4.58% | 1.33% | 4.445% | 12.85% | 9.464% | 8.995% | 1.499% | 1.487% | 15.46% | 2.984% | 4.536% | 0.9277% | 2.775% | 0.4165% | 8.078% | 3.435% | 2.566% | 17.48% | 193.1% | 24.36% | 21% | 9.376% | 52.68% | 1.614% | 1.318% | 0.2173% | 1.153% | 1.461% | 1.049% | 1.291% | 0.1052% | 2.285% | 14.11% | 3.285% | 3.309% | 3.745%
250 | LBLANK |  | 6/9/2023 3:17:56 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.604 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.735% | 14.6% | 1.643% | 3.323% | 2.409% | 129.3% | 11.3% | 5.373% | 2.3e-06% | 57.58% | 121.3% | 0.2917% | 21.99% | 12.92% | 59.63% | 198.3% | 12.08% | 10.37% | 2.043% | 119.4% | 34.95% | 0.408% | 6.333% | 87.52% | 42.87% | 385.2% | 78.58% | 260.8% | 20.27% | 16.7% | 14.89% | 8.876% | 0.5208% | 4.792% | 0.2839% | 0.8184% | 1.055% | 2.787% | 17.64% | 9.845% | 12.38% | 9.611%
251 | CCV 1ppb TRACE |  | 6/9/2023 3:23:44 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.605 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 6.505% | 7.889% | 2.362% | 28.11% | 12.94% | 42.12% | 3.487% | 18.55% | 10.46% | 10.46% | 1.185% | 8.8% | 2.098% | 23.36% | 6.939% | 4.603% | 4.263% | 1.951% | 2.556% | 2.145% | 13.62% | 0.7333% | 1.815% | 3.962% | 10.27% | 2.209% | 4.397% | 8.929% | 5.522% | 1.418% | 0.1164% | 1.358% | 0.5748% | 3.713% | 1.627% | 5.649% | 0.6799% | 1.078% | 2.144% | 2.981% | 0.8428% | 4.61%
252 | CCV 1ppb TRACE |  | 6/9/2023 3:30:02 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.606 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 6.06% | 7.246% | 1.295% | 31.44% | 31.92% | 25.69% | 4.52% | 6.678% | 6.098% | 4.76% | 4.061% | 4.611% | 1.78% | 8.45% | 1.371% | 4.599% | 4.775% | 2.718% | 4.127% | 2.565% | 3.991% | 1.063% | 1.593% | 9.065% | 0.2208% | 8.152% | 3.114% | 2.821% | 6.134% | 3.207% | 1.079% | 1.574% | 0.211% | 0.6129% | 1.745% | 0.8026% | 1.759% | 2.836% | 1.883% | 2.385% | 1.253% | 1.578%
253 | CCV 10 ppb MAJORS |  | 6/9/2023 3:35:51 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.607 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.135% | 0.9669% | 1.743% | 9.969% | 0.671% | 4.626% | 3.642% | 3.541% | 1.743% | 1.433% | 0.7051% | 4.496% | 11.34% | 0.5592% | 1.372% | 0.6122% | 1.445% | 2.079% | 0.7372% | 0.9167% | 5.1% | 11.44% | 10.66% | 5.454% | 7.3% | 10.92% | 8.991% | 2.097% | 8.8% | 1.454% | 10.58% | 0.8051% | 9.267% | 1.287% | 1.042% | 0.7017% | 2.35% | 10.17% | 0.8174% | 1.27% | 0.7153% | 0.748%
254 | CCV 10 ppb MAJORS |  | 6/9/2023 3:42:08 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.608 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.376% | 0.9306% | 14.27% | 8.328% | 24.56% | 3.595% | 2.187% | 3.918% | 2.521% | 0.413% | 0.4046% | 0.4177% | 0.4863% | 2.569% | 0.5052% | 0.6538% | 0.7631% | 1.024% | 0.4255% | 10.65% | 3.475% | 1.348% | 0.4422% | 16.48% | 3.246% | 2.911% | 1.412% | 11.8% | 5.809% | 0.9523% | 0.4916% | 11.06% | 0.7684% | 1.187% | 0.3132% | 1.553% | 18.64% | 1.259% | 1.853% | 0.256% | 12.55% | 2.723%
255 | 8_30_21_UF_10_LSP |  | 6/9/2023 3:47:56 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_UF_10_LSP.609 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.643% | 1.945% | 7.153% | 8.239% | 1.523% | 1.82% | 1.333% | 1.361% | 3.221% | 2.649% | 1.201% | 0.7062% | 3.436% | 1.975% | 1.858% | 0.7755% | 2.128% | 0.77% | 1.22% | 5.989% | 3.231% | 4.726% | 3.515% | 7.26% | 5.347% | 3.943% | 3.413% | 7.681% | 6.949% | 0.419% | 3.356% | 6.686% | 2.493% | 1.972% | 0.7269% | 4.064% | 0.2291% | 3.865% | 0.9556% | 0.7732% | 7.813% | 4.878%
256 | 8_30_21_UF_10_DUP_SP |  | 6/9/2023 3:53:45 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_UF_10_DUP_SP.610 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.7396% | 2.044% | 0.9192% | 5.56% | 1.271% | 3.713% | 2.907% | 1.572% | 3.029% | 2.674% | 2.764% | 3.703% | 0.05165% | 2.592% | 1.028% | 2.585% | 1.48% | 2.816% | 3.774% | 0.6878% | 5.87% | 1.176% | 0.1754% | 0.8635% | 2.606% | 2.053% | 0.8054% | 5.313% | 6.1% | 4.209% | 0.1832% | 0.5516% | 0.4731% | 1.335% | 2.533% | 1.251% | 0.691% | 1.585% | 2.763% | 2.608% | 0.5098% | 0.9616%
257 | 8_30_21_UF_11 |  | 6/9/2023 3:59:33 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_UF_11.611 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.3647% | 2.612% | 12.55% | 24.41% | 1.205% | 1.73% | 2.635% | 0.6732% | 10.71% | 7.372% | 10.54% | 2.511% | 0.5168% | 5.126% | 39.6% | 6.874% | 2.013% | 2.531% | 2.125% | 4.798% | 4.997% | 3.521% | 1.607% | 15.94% | 21.62% | 25.05% | 3.625% | 49.85% | 48.8% | 0.6065% | 1.467% | 2.354% | 0.2224% | 0.6777% | 0.3621% | 0.8093% | 0.6421% | 2.235% | 9.025% | 8.15% | 3.842% | 0.5777%
258 | 8_30_21_BB |  | 6/9/2023 4:05:21 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_BB.612 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.715% | 14.27% | 1.369% | 3.699% | 1179% | 31.68% | 13.05% | 9.722% | 37.04% | 36.9% | 8.674% | 22.09% | 12.59% | 1.917% | 94.63% | 119.2% | 4.948% | 13.52% | 401.1% | 51.91% | 0.4922% | 1.195% | 30.14% | 57.94% | 0% | 48.73% | 79.04% | 387.4% | 0% | 20.43% | 19.44% | 0.814% | 0.7483% | 0.7162% | 9.986% | 0.3897% | 0.8353% | 0.4636% | 16.44% | 31.7% | 1.199% | 0.8731%
259 | 8_30_21_FB |  | 6/9/2023 4:11:09 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_FB.613 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 15.06% | 34.55% | 1.233% | 1.755% | 7.47% | 26.56% | 24.35% | 14.29% | 76.76% | 9.411% | 7.669% | 9.262% | 3.33% | 30.78% | 42.34% | 0.7306% | 3.117% | 1.875% | 2.323% | 46.76% | 0% | 0.435% | 13.32% | 39.23% | 22.35% | 102.1% | 32.56% | 210.9% | 21.26% | 12.45% | 5.403% | 1.721% | 0.9475% | 0.1432% | 1.18% | 0.5146% | 3.437% | 2.342% | 8.646% | 63.85% | 5.174% | 2.656%
260 | 8_30_21_LAB_A |  | 6/9/2023 4:16:56 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_LAB_A.614 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 5.201% | 0% | 3.735% | 177.1% | 2.599% | 41.84% | 23.06% | 0% | 24.07% | 25.75% | 5.17% | 23.05% | 14.7% | 61.98% | 77.9% | 16.13% | 7.632% | 2.47% | 81.87% | 29.97% | 0% | 0.4176% | 17.32% | 12.75% | 0% | 233.9% | 45.59% | 3330% | 0% | 5.423% | 5.048% | 1.746% | 1.115% | 1.406% | 0.9107% | 2.348% | 2.376% | 1.211% | 16.36% | 22.6% | 1.579% | 7.27%
261 | 8_30_21_LAB_B |  | 6/9/2023 4:22:44 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_LAB_B.615 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 7.799% | 0% | 5.182% | 45.52% | 1.134% | 132% | 15.19% | 9.887% | 2.3e-06% | 41.8% | 8.397% | 45.37% | 17.1% | 43.16% | 60.78% | 9.056% | 8.078% | 8.164% | 49.74% | 85.82% | 74.62% | 0.4645% | 27.18% | 11.27% | 0% | 121.1% | 35.66% | 553.6% | 21.37% | 31.11% | 18.27% | 1.641% | 1.699% | 1.699% | 3.759% | 0.2761% | 0.878% | 3.94% | 0% | 10% | 1.982% | 0.5805%
262 | 9_13_21_W_1 |  | 6/9/2023 4:28:32 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_W_1.616 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.41% | 3.509% | 0.7826% | 10.33% | 1.307% | 8.982% | 9% | 2.832% | 13.99% | 41.53% | 10.76% | 1.024% | 0.4878% | 7.217% | 32.21% | 20.32% | 7.608% | 3.51% | 17.18% | 3.663% | 13.32% | 2.63% | 0.2564% | 38.16% | 38.87% | 53.45% | 54.81% | 121.2% | 87.79% | 1.333% | 0.5699% | 1.711% | 0.1605% | 4.198% | 1.216% | 0.3187% | 0.7964% | 3.785% | 36.13% | 76.85% | 5.268% | 1.904%
263 | 9_13_21_W_6 |  | 6/9/2023 4:34:20 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_W_6.617 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.282% | 6.001% | 2.025% | 5.199% | 7.729% | 4.488% | 17.75% | 3.542% | 8.557% | 16.22% | 9.102% | 1.166% | 0.6199% | 0.6468% | 17.36% | 4.256% | 4.078% | 2.291% | 2.396% | 5.485% | 24.36% | 2.377% | 1.123% | 23.49% | 158.7% | 1.08% | 8.62% | 29.29% | 185.8% | 0.1891% | 1.599% | 0.7042% | 1.049% | 0.39% | 0.4836% | 0.6744% | 2.204% | 2.677% | 33.52% | 19.32% | 0.8489% | 2.095%
264 | 9_13_21_W_8 |  | 6/9/2023 4:40:07 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_W_8.618 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.871% | 2.79% | 2.971% | 9.698% | 2.159% | 15.75% | 6.087% | 3.011% | 21.55% | 19.66% | 34.26% | 32.54% | 2.981% | 1.524% | 9.637% | 21.51% | 4.648% | 33.29% | 43.8% | 2.736% | 18.55% | 4.338% | 5.451% | 13.69% | 31.85% | 22.62% | 32.31% | 59.5% | 2210% | 31.2% | 2.55% | 2.482% | 2.347% | 1.263% | 23.81% | 0.8854% | 1.545% | 3.874% | 18.35% | 37.91% | 3.31% | 3.334%
265 | LBLANK |  | 6/9/2023 4:45:55 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.619 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 7.661% | 0% | 0.2981% | 8.49% | 0.948% | 168.9% | 9.937% | 0% | 59.22% | 139.2% | 55.77% | 1.537% | 11.7% | 4.046% | 93.25% | 326.7% | 10.61% | 3.152% | 0.8424% | 26.53% | 0% | 0.3606% | 57.73% | 45.92% | 0% | 360.2% | 52.39% | 282.8% | 0% | 20.9% | 14.45% | 0.8423% | 0.8979% | 2.629% | 1.447% | 0.8619% | 0.213% | 2.021% | 25.84% | 11.19% | 3.882% | 5.516%
266 | CCV 1ppb TRACE |  | 6/9/2023 4:51:44 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.620 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4.434% | 5.904% | 5.144% | 11.67% | 141.7% | 4.261% | 2.727% | 8.229% | 7.121% | 2.546% | 1.493% | 4.435% | 2.897% | 10.59% | 2.751% | 10.44% | 3.831% | 0.4711% | 2.079% | 11.59% | 17.99% | 1.713% | 1.471% | 7.628% | 9.975% | 8.215% | 9.362% | 25.84% | 8.189% | 1.47% | 1.554% | 1.05% | 0.6168% | 2.365% | 0.7681% | 1.512% | 0.7936% | 1.078% | 3.036% | 0.5124% | 1.343% | 0.2434%
267 | CCV 1ppb TRACE |  | 6/9/2023 4:58:02 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.621 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.279% | 4.477% | 1.658% | 29.66% | 43.41% | 16.98% | 2.931% | 6.399% | 1.304% | 0.5036% | 2.113% | 6.193% | 1.409% | 7.218% | 1.05% | 5.3% | 0.9493% | 1.64% | 1.356% | 2.044% | 15.82% | 1.141% | 1.274% | 17.95% | 5.149% | 3.635% | 4.712% | 11.34% | 16.91% | 1.953% | 1.563% | 1.26% | 0.9046% | 1.009% | 0.4907% | 0.8966% | 1.375% | 1.213% | 0.919% | 0.4004% | 0.6319% | 1.023%
268 | LFB |  | 6/9/2023 5:03:50 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.622 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.367% | 2.065% | 0.8945% | 3.709% | 1.459% | 2.052% | 4.04% | 6.098% | 0.2166% | 0.6783% | 1.158% | 1.358% | 0.7748% | 0.8552% | 3.031% | 1.757% | 0.8996% | 1.666% | 2.012% | 2.059% | 8.005% | 1.009% | 0.9985% | 4.012% | 4.183% | 1.46% | 0.7016% | 3.069% | 3.291% | 0.2506% | 1.327% | 2.218% | 0.7422% | 1.246% | 1.692% | 1.31% | 0.7621% | 1.379% | 0.4803% | 1.302% | 1.585% | 1.491%
269 | CCV 10 ppb MAJORS |  | 6/9/2023 5:09:38 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.623 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4.702% | 5.555% | 3.816% | 10.86% | 1.49% | 4.587% | 4.219% | 2.496% | 3.35% | 3.834% | 10.19% | 10.87% | 5.882% | 4.07% | 5.515% | 4.889% | 4.176% | 9.835% | 9.884% | 2.896% | 3.332% | 3.435% | 5.355% | 7.599% | 4.198% | 7.167% | 5.377% | 3.631% | 11.7% | 10.24% | 4.988% | 3.166% | 4.009% | 2.617% | 8.037% | 3.562% | 0.3541% | 2.135% | 8.805% | 9.728% | 4.397% | 4.964%
270 | CCV 10 ppb MAJORS |  | 6/9/2023 5:15:55 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.624 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.062% | 5.24% | 0.8341% | 3.531% | 1.401% | 3.368% | 3.37% | 2.221% | 1.938% | 0.8451% | 0.5243% | 0.7871% | 1.079% | 2.003% | 0.3734% | 0.8577% | 1.03% | 1.069% | 0.6643% | 1.334% | 1.874% | 1.058% | 1.303% | 1.51% | 3.037% | 2.027% | 2.174% | 3.022% | 1.007% | 0.5006% | 0.45% | 1.002% | 0.8197% | 0.786% | 0.3963% | 0.5005% | 1.964% | 1.09% | 0.965% | 1.004% | 1.073% | 1.296%
271 | 9_13_21_W_9 |  | 6/9/2023 5:21:43 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_W_9.625 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.882% | 2.292% | 18.42% | 17.61% | 3.194% | 9.084% | 2.493% | 4.734% | 19.45% | 25.38% | 7.747% | 1.277% | 1.781% | 1.511% | 43.8% | 48.22% | 1.076% | 2.177% | 6.591% | 10.04% | 18.59% | 1.424% | 2.934% | 55.11% | 64.24% | 47.69% | 16.34% | 157% | 93.95% | 0.5644% | 2.097% | 13.73% | 0.5977% | 1.193% | 0.7479% | 0.8537% | 0.6691% | 2.345% | 6.621% | 67.88% | 11.89% | 1.813%
272 | 9_13_21_W_10 |  | 6/9/2023 5:27:31 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_W_10.626 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.13% | 3.112% | 1.217% | 9.57% | 1.991% | 4.116% | 8.291% | 3.219% | 7.113% | 43.27% | 20.57% | 1.759% | 0.4019% | 1.543% | 15.31% | 14.63% | 1.568% | 2.136% | 0.4623% | 7.009% | 8.127% | 2.679% | 1.918% | 59.66% | 21.54% | 10.9% | 20.08% | 61.15% | 52.81% | 0.878% | 2.413% | 0.587% | 0.4249% | 0.1409% | 0.7192% | 1.239% | 0.4237% | 2.127% | 31.04% | 18.72% | 1.812% | 2.988%
273 | 9_13_21_W_10_DUP |  | 6/9/2023 5:33:20 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_W_10_DUP.627 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 6.148% | 2.055% | 3.116% | 12.41% | 2.268% | 7.145% | 3.071% | 4.861% | 5.992% | 59.5% | 16.39% | 2.044% | 1.176% | 1.443% | 42.56% | 14.06% | 2.22% | 0.9117% | 6.201% | 1.276% | 27.74% | 2.483% | 2.743% | 39.42% | 110% | 12.53% | 49.78% | 28.36% | 128.1% | 1.428% | 1.836% | 0.5151% | 1.36% | 1.166% | 0.4041% | 2.168% | 1.155% | 1.937% | 12.42% | 44.91% | 2.291% | 4.619%
274 | 9_13_21_W_11 |  | 6/9/2023 5:39:08 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_W_11.628 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.9873% | 9.306% | 9.027% | 1.197% | 1.472% | 10.85% | 4.692% | 2.071% | 11.55% | 21.81% | 5.952% | 0.6445% | 0.4468% | 0.6788% | 12.27% | 24.39% | 2.31% | 1.896% | 1.024% | 5.495% | 14.43% | 1.812% | 4.587% | 68.53% | 92.27% | 12.42% | 39.28% | 96.64% | 38.9% | 1.515% | 3.913% | 8.16% | 1.464% | 0.6951% | 0.5997% | 1.677% | 0.1948% | 1.709% | 39.43% | 14.5% | 5.961% | 1.217%
275 | 9_13_21_F_1 |  | 6/9/2023 5:44:56 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_F_1.629 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.2498% | 3.893% | 2.973% | 15.38% | 1.521% | 2.607% | 2.423% | 2.537% | 12.36% | 15.36% | 2.695% | 0.5833% | 1.459% | 4.799% | 16.42% | 4.407% | 2.848% | 1.151% | 2.608% | 4.917% | 7.814% | 1.763% | 1.354% | 11.07% | 692.3% | 26.19% | 22.38% | 23.85% | 80.98% | 1.879% | 0.7185% | 4.279% | 1.298% | 0.9488% | 0.6295% | 0.8144% | 1.132% | 2.898% | 29.69% | 13.32% | 0.7676% | 1.961%
276 | 9_13_21_F_6 |  | 6/9/2023 5:50:44 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_F_6.630 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 6.99% | 6.687% | 1.265% | 7.331% | 0.9935% | 8.797% | 4.68% | 8.27% | 7.241% | 16.25% | 7.015% | 2.851% | 2.167% | 9.635% | 5.313% | 13.41% | 10.25% | 2.037% | 1.93% | 0.3218% | 14.23% | 3.735% | 1.485% | 16.23% | 179.3% | 16.81% | 14.55% | 9.686% | 113.1% | 1.844% | 0.9211% | 0.2528% | 1.13% | 6.022% | 2.067% | 1.781% | 1.131% | 3.901% | 13.54% | 4.123% | 5.445% | 1.63%
277 | 9_13_21_F_8 |  | 6/9/2023 5:56:32 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_F_8.631 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 6.65% | 7.962% | 2.842% | 14.92% | 3.241% | 6.275% | 6.369% | 3.932% | 5.57% | 8.461% | 7.482% | 1.829% | 0.8426% | 7.098% | 31.69% | 10.05% | 6.306% | 1.173% | 0.263% | 0.4977% | 11.07% | 1.06% | 0.641% | 18.91% | 169.9% | 26.03% | 6.422% | 13.02% | 122.1% | 0.8639% | 1.055% | 1.081% | 0.1801% | 5.054% | 0.8072% | 0.9484% | 1.494% | 1.856% | 5.182% | 5.243% | 2.159% | 1.841%
278 | 9_13_21_F_9 |  | 6/9/2023 6:02:20 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_F_9.632 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 7.548% | 5.598% | 1.385% | 6.652% | 2.317% | 6.614% | 7% | 6.946% | 7.084% | 12.24% | 1.984% | 3.573% | 3.321% | 9.263% | 28.99% | 9.418% | 0.2457% | 2.959% | 3.067% | 0.6034% | 11.61% | 7.249% | 5.158% | 9.178% | 184.3% | 25.32% | 25.57% | 28.94% | 87.73% | 3.415% | 5.351% | 0.3624% | 3.666% | 5.978% | 1.724% | 0.8754% | 1.262% | 6.462% | 22.7% | 5.935% | 2.014% | 1.848%
279 | 9_13_21_F_10 |  | 6/9/2023 6:08:08 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_F_10.633 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.258% | 2.647% | 1.347% | 12.51% | 2.48% | 1.666% | 2.413% | 2.761% | 3.787% | 6.994% | 11.94% | 8.113% | 0.6897% | 3.271% | 39.96% | 6.598% | 2.017% | 9.647% | 10.95% | 2.159% | 5.523% | 0.4839% | 1.309% | 15.07% | 39.13% | 29.54% | 38.98% | 18.41% | 96.84% | 10.01% | 0.4173% | 0.383% | 0.8397% | 3.106% | 7.523% | 0.6527% | 0.4728% | 0.09927% | 11.19% | 37.65% | 5.445% | 2.102%
280 | 9_13_21_F_11 |  | 6/9/2023 6:13:56 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_F_11.634 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.686% | 3.898% | 1.099% | 8.776% | 2.706% | 5.155% | 7.845% | 3.63% | 1.42% | 19.84% | 11.13% | 5.054% | 0.8227% | 3.552% | 26.61% | 24.39% | 2.849% | 4.437% | 6.113% | 2.2% | 12.48% | 2.967% | 0.6122% | 19.65% | 182.4% | 6.586% | 23.99% | 12.6% | 81.94% | 4.49% | 0.3876% | 1.915% | 0.2078% | 3.012% | 3.366% | 2.789% | 1.788% | 3.673% | 19.83% | 8.484% | 1.916% | 5.078%
281 | 9_13_21_F_11_LD |  | 6/9/2023 6:19:44 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_F_11_LD.635 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 7.151% | 6.487% | 5.068% | 14.09% | 1.794% | 4.494% | 6.546% | 6.574% | 2.489% | 13.13% | 3.138% | 2.771% | 2.51% | 9.969% | 14.77% | 18.08% | 6.288% | 1.577% | 1.658% | 3.48% | 10.9% | 2.943% | 3.795% | 10.95% | 1122% | 20.4% | 24.21% | 8.477% | 74.64% | 2.294% | 3.166% | 2.377% | 2.108% | 5.02% | 0.06788% | 1.014% | 1.819% | 0.9668% | 3.662% | 7.039% | 5.716% | 3.36%
282 | 9_13_21_F_11_LD |  | 6/9/2023 6:26:00 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_F_11_LD.636 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.491% | 1.855% | 2.325% | 2.691% | 4.343% | 2.35% | 3.972% | 1.841% | 5.994% | 22.46% | 7.78% | 0.7892% | 1.296% | 1.663% | 24.66% | 10.24% | 1.424% | 0.994% | 1.099% | 0.8994% | 4.421% | 4.833% | 0.938% | 15.06% | 55.41% | 21.58% | 20.6% | 32.21% | 44.06% | 0.4689% | 1.415% | 1.663% | 0.1337% | 0.9135% | 1.239% | 1.175% | 1.141% | 4.819% | 22.06% | 2.766% | 2.383% | 0.1815%
283 | LBLANK |  | 6/9/2023 6:31:48 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.637 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 6.437% | 0% | 0.02527% | 16.06% | 4.386% | 347.5% | 9.807% | 0% | 59.47% | 45.11% | 130.6% | 1.16% | 21.46% | 2.413% | 95.79% | 57.45% | 17.61% | 13.32% | 1.31% | 10.62% | 0% | 0.1819% | 63.4% | 57.81% | 131% | 121.4% | 36.54% | 197.9% | 21.85% | 24.81% | 5.652% | 5.394% | 1.186% | 1.318% | 0.04205% | 1.383% | 1.301% | 3.179% | 18.61% | 24.1% | 3.288% | 5.721%
284 | CCV 1ppb TRACE |  | 6/9/2023 6:37:37 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.638 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.62% | 13.29% | 1.956% | 62.56% | 43.33% | 33.02% | 3.829% | 4.117% | 1.36% | 1.419% | 2.02% | 4.519% | 1.377% | 5.167% | 2.785% | 4.487% | 2.758% | 1.313% | 3.947% | 2.914% | 13.67% | 3.167% | 1.188% | 8.054% | 4.861% | 1.285% | 6.849% | 13.36% | 12.41% | 0.251% | 1.578% | 0.604% | 0.5305% | 1.872% | 0.5734% | 3.518% | 0.521% | 3.35% | 1.588% | 1.353% | 1.458% | 2.555%
285 | CCV 1ppb TRACE |  | 6/9/2023 6:43:54 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.639 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 7.872% | 12.55% | 7.387% | 114.1% | 74.97% | 8.745% | 11.4% | 7.128% | 6.538% | 4.728% | 2.081% | 2.985% | 1.487% | 7.576% | 3.121% | 10.37% | 3.595% | 1.941% | 2.402% | 8.055% | 20.94% | 3.634% | 3.24% | 4.328% | 4.382% | 7.166% | 6.165% | 10.19% | 5.211% | 1.072% | 0.8418% | 1.363% | 1.239% | 3.178% | 0.7482% | 14.15% | 0.2544% | 2.792% | 1.934% | 2.154% | 2.94% | 16.72%
286 | CCV 10 ppb MAJORS |  | 6/9/2023 6:49:43 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.640 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.741% | 5.015% | 14.26% | 5.126% | 1.099% | 1.669% | 1.255% | 1.909% | 0.9032% | 1.416% | 1.565% | 1.035% | 7.055% | 2.228% | 2.345% | 1.238% | 1.062% | 0.8639% | 1.541% | 11.02% | 4.703% | 8.674% | 7.185% | 11.69% | 5.974% | 9.169% | 4.742% | 10.36% | 7.422% | 0.9807% | 6.831% | 9.981% | 6.897% | 0.4346% | 0.357% | 0.6787% | 0.253% | 8.823% | 1.072% | 0.6601% | 12.15% | 1.204%
287 | CCV 10 ppb MAJORS |  | 6/9/2023 6:56:01 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.641 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.859% | 2.127% | 1.786% | 12.44% | 1.209% | 2.23% | 2.564% | 1.483% | 0.1942% | 0.5899% | 4.962% | 3.548% | 2.187% | 1.563% | 1.088% | 1.485% | 1.025% | 3.64% | 3.864% | 1.278% | 3.4% | 0.3443% | 1.051% | 0.5177% | 3.997% | 0.6025% | 2.794% | 4.309% | 1.816% | 3.564% | 0.8846% | 1.495% | 0.8839% | 1.35% | 1.917% | 10.99% | 0.7604% | 0.3437% | 4.109% | 3.562% | 0.4475% | 11.43%
288 | 9_13_21_F_11_DUP |  | 6/9/2023 7:01:49 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_F_11_DUP.642 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.105% | 4.259% | 17.66% | 2.815% | 1.843% | 2.952% | 6.858% | 3.386% | 5.575% | 17.1% | 3.951% | 1.548% | 0.6904% | 0.8537% | 11.27% | 3.456% | 4.599% | 0.4831% | 1.418% | 13.83% | 17.53% | 2.043% | 1.021% | 69.75% | 187% | 23.62% | 21.52% | 13.75% | 30.31% | 2.482% | 1.725% | 13.54% | 0.3303% | 0.9009% | 0.3685% | 0.5912% | 2.615% | 2.573% | 3.259% | 14.03% | 19.22% | 0.2891%
289 | 9_13_21_F_11_DUP_LSP |  | 6/9/2023 7:07:37 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_F_11_DUP_LSP.643 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.085% | 2.179% | 3.007% | 10.72% | 0.8862% | 2.183% | 3.632% | 0.8335% | 2.303% | 0.5411% | 0.4647% | 1.082% | 6.761% | 1.67% | 1.298% | 0.5869% | 1.855% | 1.908% | 1.396% | 1.567% | 2.552% | 3.736% | 6.886% | 1.664% | 10.94% | 6.688% | 8.191% | 1.285% | 5.444% | 1.256% | 6.104% | 1.282% | 6.093% | 1.678% | 1.42% | 2.223% | 0.2122% | 2.894% | 1.232% | 2.247% | 0.9985% | 1.468%
290 | 9_13_21_UF_1 |  | 6/9/2023 7:13:24 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_UF_1.644 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.332% | 5.058% | 21.29% | 48.8% | 1.86% | 5.156% | 6.385% | 3.009% | 5.287% | 34.25% | 7.084% | 7.904% | 1.314% | 5.327% | 8.529% | 4.757% | 9.223% | 0.7751% | 3.214% | 11.6% | 8.932% | 2.298% | 1.933% | 44.96% | 109.4% | 14.79% | 19.6% | 33.04% | 981.4% | 0.8515% | 1.421% | 9.958% | 1.597% | 2.334% | 0.3272% | 2.583% | 0.5097% | 2.582% | 10.36% | 26.61% | 16.19% | 5.831%
291 | 9_13_21_UF_6 |  | 6/9/2023 7:19:12 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_UF_6.645 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.781% | 4.19% | 40.28% | 243.6% | 2.748% | 1.708% | 1.666% | 1.238% | 5.529% | 9.813% | 8.195% | 41.24% | 3.904% | 13.24% | 13.45% | 7.83% | 0.7616% | 3.853% | 3.048% | 1.614% | 7.826% | 5.987% | 4.537% | 24.33% | 34.42% | 11.06% | 25.62% | 10.09% | 166.6% | 3.206% | 2.968% | 0.8797% | 2.995% | 0.8532% | 2.878% | 1.589% | 1.367% | 4.814% | 7.837% | 3.263% | 7.865% | 1.191%
292 | 9_13_21_UF_6_DUP |  | 6/9/2023 7:25:00 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_UF_6_DUP.646 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.733% | 1.046% | 63.08% | 337.2% | 1.898% | 2.665% | 4.102% | 0.4649% | 11.04% | 7.971% | 4.9% | 6.103% | 7.987% | 25.73% | 7.906% | 10.45% | 3.969% | 2.535% | 1.933% | 0.7288% | 8.386% | 4.987% | 9.388% | 30.06% | 140.3% | 28.91% | 14.38% | 17.25% | 106.6% | 1.768% | 8.53% | 0.7562% | 7.21% | 1.161% | 0.6748% | 1.753% | 0.9393% | 3.764% | 4.849% | 16.42% | 2.782% | 1.686%
293 | 9_13_21_UF_8 |  | 6/9/2023 7:30:49 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_UF_8.647 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.97% | 2.274% | 536% | 60.44% | 1.165% | 2.249% | 4.254% | 2.415% | 8.882% | 22.76% | 5.972% | 123% | 1.929% | 5.803% | 10.93% | 11.9% | 0.7911% | 0.9397% | 3.932% | 5.296% | 3.26% | 1.482% | 2.866% | 54.54% | 88.59% | 27.86% | 10.46% | 6.071% | 577.1% | 0.9256% | 1.777% | 3.854% | 1.74% | 1.457% | 0.741% | 0.6348% | 2.173% | 1.886% | 23.04% | 13.82% | 4.271% | 2.694%

## Unfactored Concentrations
| Sample Id | R | Acquisition Time | QC Status | Dataset File | Method File | Na KED 23/23 Helium KED_Low (ppb) | Mg KED 24/24 Helium KED_Low (ppb) | Al 27/27 (ppb) | Al KED 27/27 Helium KED_Low (ppb) | P Oshift 31/47 Oxygen DRC (ppb) | S (34) -1 34/34 Helium KED_Low (ppb) | K KED 39/39 Helium KED_Low (ppb) | Ca KED 43/43 Helium KED_Low (ppb) | V KED 51/51 Helium KED_Low (ppb) | Cr (52) 52/52 Helium KED_Low (ppb) | Cr (53) NH3 53/53 Ammonia DRC (ppb) | Fe (54) NH3 54/54 Ammonia DRC (ppb) | Mn Oshift 55/71 Oxygen DRC (ppb) | Fe (56) KED 56/56 Helium KED_Low (ppb) | Co KED 59/59 Helium KED_Low (ppb) | Ni KED 60/60 Helium KED_Low (ppb) | Cu (63) KED 63/63 Helium KED_Low (ppb) | Cu (65) NH3 65/65 Ammonia DRC (ppb) | Zn NH3 66/66 Ammonia DRC (ppb) | As 75/75 (ppb) | As KED 75/75 Helium KED_Low (ppb) | As RH Oshift 75/91 Oxygen DRC (ppb) | As Oshift 75/91 Oxygen DRC (ppb) | Se (77) 77/77 (ppb) | Se (77) Oshift 77/93 Oxygen DRC (ppb) | Se (78) Oshift 78/94 Oxygen DRC (ppb) | Se (80) Oshift 80/96 Oxygen DRC (ppb) | Se (82) 82/82 (ppb) | Se (82) Oshift 82/98 Oxygen DRC (ppb) | Mo (95) NH3 95/95 Ammonia DRC (ppb) | Mo (95) OShift 95/127 Oxygen DRC (ppb) | Rh (IS1) 103/103 (IS) | Rh (IS2) 103/103 (IS) Oxygen DRC | Rh (IS3) 103/103 (IS) Helium KED_Low | Rh (IS4) 103/103 (IS) Ammonia DRC | Rh (IS5) 103/103 (IS) Ammonia DRC | Rh (IS6) 103/103 (IS) Oxygen DRC | Rh OShift (IS1) 103/119 (IS) Oxygen DRC | Cd (111) NH3 111/111 Ammonia DRC (ppb) | Cd (114) NH3 114/114 Ammonia DRC (ppb) | Pb 208/208 (ppb) | Pb NH3 208/208 Ammonia DRC (ppb)
1 | Sac Blank |  | 06/08/2023 13:23:24 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\Sac Blank.356 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth
2 | sac blank cal |  | 06/08/2023 13:29:13 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\sac blank cal.357 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth
3 | CAL BLANK |  | 06/08/2023 13:35:02 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CAL BLANK.358 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth
4 | LOW 0.01 |  | 06/08/2023 13:40:50 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LOW 0.01.359 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth
5 | LOW 0.05 |  | 06/08/2023 13:46:39 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LOW 0.05.360 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth |  |  |  |  | 0.05 |  |  |  | 0.05 | 0.05 | 0.05 |  | 0.05 |  | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | [6 empty] | 0.05 | 0.05 | [7 empty] | 0.05 | 0.05 | 0.05 | 0.05
6 | LOW 0.1 |  | 06/08/2023 13:52:27 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LOW 0.1.361 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth |  |  |  |  | 0.08067 |  |  |  | 0.096063 | 0.090199 | 0.097504 |  | 0.12719 |  | 0.10617 | 0.094887 | 0.098589 | 0.097288 | 0.22882 | 0.077063 | 0.098368 | 0.10104 | 0.10364 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.11127 | 0.10952 | [7 empty] | 0.09857 | 0.097719 | 0.10179 | 0.10211
7 | LOW 0.5 |  | 06/08/2023 13:58:15 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LOW 0.5.362 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth |  |  |  |  | 0.15937 |  |  |  | 0.53897 | 0.49784 | 0.48804 | 0.5 | 0.60519 | 0.5 | 0.52415 | 0.51285 | 0.51696 | 0.49928 | 0.6686 | 0.48375 | 0.42073 | 0.49958 | 0.50119 | 0.46184 | 0.45275 | 0.49141 | 0.50515 | 0.6874 | 0.47689 | 0.55331 | 0.55418 | [7 empty] | 0.48389 | 0.49637 | 0.49775 | 0.508
8 | STD 1 |  | 06/08/2023 14:04:03 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 1.363 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1000 | 1000 | 1 | 1 | 0.19292 |  | 1000 | 1000 | 1.0447 | 0.98556 | 0.97875 | 1.0097 | 1.1676 | 0.96268 | 1.0039 | 1.0365 | 1.01 | 0.98232 | 1.1638 | 1.0379 | 0.89754 | 1.0048 | 0.99778 | 0.89607 | 0.98087 | 0.95947 | 0.97561 | 1.1989 | 0.90082 | 1.0975 | 1.0879 | [7 empty] | 0.96845 | 0.99495 | 0.98312 | 1.0259
9 | STD 5 |  | 06/08/2023 14:09:52 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 5.364 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4979.8464 | 4887.6514 | 4.8266 | 4.8375 | 0.96382 | 5000 | 5009.9906 | 5128.1752 | 4.7585 | 4.7674 | 4.8955 | 3.7294 | 5.5985 | 3.3639 | 4.8693 | 4.9649 | 4.6991 | 4.7033 | 5.4062 | 5.3697 | 4.1764 | 5.0013 | 4.9421 | 4.9 | 5.0152 | 4.8055 | 4.9715 | 5.7652 | 4.9769 | 5.327 | 5.3031 | [7 empty] | 4.833 | 4.9158 | 5.127 | 4.8222
10 | STD 10 |  | 06/08/2023 14:15:40 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 10.365 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9973.9677 | 9988.6715 | 9.4441 | 9.4057 | 2.1142 | 9955.8343 | 10200.131 | 10523.2366 | 10.0664 | 9.6701 | 10.0256 | 8.2264 | 10.9119 | 7.4698 | 10.1718 | 9.9669 | 9.4727 | 9.4746 | 10.5784 | 10.1302 | 9.3372 | 9.9509 | 9.8405 | 9.2726 | 9.8747 | 9.2901 | 9.7564 | 10.2538 | 9.4619 | 10.7162 | 10.3566 | [7 empty] | 9.8037 | 9.9269 | 9.4953 | 9.7702
11 | STD 25 |  | 06/08/2023 14:21:28 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 25.366 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 25315.1502 | 24630.9065 | 23.7289 | 24.0879 | 6.1982 | 26128.5984 | 25828.9422 | 26778.3022 | 25.9364 | 24.9548 | 25.6624 | 21.9204 | 28.0206 | 20.3561 | 25.5579 | 25.1387 | 24.2331 | 23.9183 | 26.678 | 26.3194 | 22.9203 | 25.2615 | 25.9924 | 23.6255 | 25.999 | 25.5927 | 25.9781 | 26.8225 | 25.666 | 27.2951 | 27.1937 | [7 empty] | 25.1116 | 25.1135 | 23.4982 | 22.1433
12 | STD 50 |  | 06/08/2023 14:27:16 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 50.367 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 49447.0918 | 48504.3884 | 51.7875 | 48.3518 | 13.535 | 50362.4194 | 51311.0486 | 52594.7474 | 52.3862 | 50.3191 | 51.2753 | 44.7757 | 55.9035 | 41.9217 | 50.6641 | 49.4133 | 47.4587 | 47.916 | 52.5765 | 52.2393 | 46.1585 | 50.1529 | 52.7285 | 47.5563 | 51.763 | 51.4843 | 52.3632 | 53.9298 | 53.109 | 55.2276 | 55.6385 | [7 empty] | 49.4126 | 49.4946 | 45.7608 | 47.8458
13 | STD 100 |  | 06/08/2023 14:33:03 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 100.368 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 100079.1431 | 93962.2249 | 103.1824 | 94.9518 | 28.971 | 101591.0065 | 103348.6199 | 106494.6693 | 107.8549 | 104.035 | 104.0079 | 90.2413 | 112.5456 | 86.3357 | 104.4566 | 98.7554 | 94.638 | 92.8021 | 100.6699 | 105.0888 | 91.5374 | 100.4066 | 103.1636 | 94.7812 | 105.5344 | 104.1968 | 104.8864 | 108.9352 | 105.6718 | 113.8257 | 113.1398 | [7 empty] | 97.3575 | 98.2282 | 88.3585 | 98.0191
14 | Calibration Curves |  | 06/08/2023 14:33:03
15 | LBLANK |  | 06/08/2023 14:38:52 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.369 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | -11.7416 | -41.0609 | -0.69918 | -0.76605 | -0.23825 | 228.7424 | 3.2182 | -30.7443 | 0.0068477 | 0.005656 | 0.0027179 | -0.26883 | -0.024393 | -0.17974 | 0.002683 | 0.022413 | -0.004841 | -0.0062042 | -0.11381 | -0.041553 | 0.0084722 | 0.16761 | 0.0085315 | 0.15939 | -0.018378 | 0.0039245 | 0.0059578 | 0.025924 | 0.00737 | 0.026908 | 8.4559e-05 | [7 empty] | 0.0015638 | 0.0006736 | -0.0017921 | -0.0013804
16 | LBLANK |  | 06/08/2023 14:57:55 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.370 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | -9.9038 | -42.7414 | -0.67285 | -0.78437 | -2.0623 | 141.26 | 2.8997 | -30.7443 | 0.0056732 | 0.0069391 | 0.0030874 | -0.51549 | -0.025186 | -0.55089 | 0.0021117 | 0.0043573 | -0.0080637 | -0.0072868 | -0.15372 | -0.021739 | 0.023137 | 0.15941 | -0.00034732 | 0.085389 | -0.010775 | 0.006613 | 0.00092024 | -0.022014 | 0.0047935 | -0.0059356 | -0.0081888 | [7 empty] | 0.00028119 | -0.00046423 | -0.0022749 | -0.0019186
17 | ICV PE |  | 06/08/2023 15:03:43 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\ICV PE.371 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9910.2106 | 9986.4089 | 9.9707 | 9.3487 | 19.0658 | 9589.0762 | 9988.7685 | 10156.8568 | 9.8933 | 10.0135 | 9.74 | 10.0444 | 10.4635 | 10.2 | 10.099 | 10.2574 | 9.7402 | 9.5326 | 9.5867 | 10.1386 | 8.2954 | 9.5977 | 10.0923 | 7.5385 | 10.1506 | 9.9938 | 10.0369 | 10.9266 | 10.1441 | 10.035 | 10.1813 | [7 empty] | 9.7246 | 9.8909 | 9.9805 | 9.4044
18 | ICV PE |  | 06/08/2023 15:09:59 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\ICV PE.372 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9801.7516 | 10242.1682 | 9.4473 | 9.6363 | 19.5113 | 9856.9094 | 9816.7945 | 10287.4771 | 10.0018 | 10.086 | 9.964 | 10.3213 | 10.5281 | 10.472 | 10.1863 | 10.3198 | 9.8091 | 9.7791 | 9.8334 | 9.7628 | 8.4124 | 9.1723 | 10.0533 | 7.3008 | 9.9368 | 9.6553 | 9.9805 | 10.0546 | 9.6627 | 10.0993 | 10.2349 | [7 empty] | 9.9102 | 10.114 | 9.601 | 9.8706
19 | CCV 1ppb TRACE |  | 06/08/2023 15:15:47 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.373 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 963.5918 | 970.3544 | 1.0188 | 0.95078 | -0.56547 | 1344.678 | 976.128 | 962.5777 | 1.0279 | 1.0071 | 0.99805 | 0.59771 | 1.0075 | 0.59076 | 1.0292 | 1.0562 | 1.0438 | 1.0291 | 0.85937 | 0.94406 | 0.80569 | 1.0664 | 0.99222 | 1.0858 | 0.96297 | 0.99192 | 1.03 | 1.0712 | 0.85187 | 1.0403 | 1.0379 | [7 empty] | 0.9723 | 1.0016 | 0.98748 | 1.0285
20 | CCV 1ppb TRACE |  | 06/08/2023 15:22:03 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.374 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1016.4433 | 1053.5242 | 1.0224 | 0.89479 | -0.98102 | 1072.081 | 986.0152 | 905.2676 | 0.98098 | 0.99507 | 0.99288 | 0.66355 | 1.0327 | 0.6398 | 1.0325 | 1.066 | 1.0819 | 1.0381 | 0.86327 | 1.0362 | 0.91812 | 1.0845 | 1.0114 | 1.0792 | 0.98803 | 0.94675 | 0.98258 | 1.0218 | 1.0772 | 1.0419 | 1.032 | [7 empty] | 1.0026 | 1.0096 | 0.98362 | 1.0027
21 | LFB |  | 06/08/2023 15:27:51 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.375 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9591.8017 | 9708.532 | 8.8195 | 9.0889 | 9.9079 | 10108.5395 | 9525.6146 | 9366.2479 | 9.761 | 9.5988 | 10.0447 | 10.0076 | 10.2607 | 9.7319 | 10.0643 | 10.1776 | 9.6645 | 9.8918 | 10.289 | 9.5474 | 9.0222 | 9.4115 | 10.1802 | 9.4066 | 9.9294 | 9.7458 | 9.9836 | 9.9387 | 9.7442 | 10.648 | 10.4375 | [7 empty] | 9.9167 | 10.1139 | 9.1342 | 9.5711
22 | LLOQ 0.01 |  | 06/08/2023 15:33:39 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LLOQ 0.01.376 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | -1.6627 | -30.1172 | -0.23747 | -0.2305 | -1.0393 | 358.4147 | 18.1369 | -16 | 0.01364 | 0.012039 | 0.010911 | -0.37113 | 0.0079147 | -0.38952 | 0.01076 | 0.01227 | 0.00791 | 0.011354 | -0.038124 | 0.033303 | 0.0045282 | 0.1668 | 0.0078158 | 0.013396 | 0.004558 | 0.015157 | 0.0084529 | -0.0978 | 0.017665 | 0.014279 | 0.012314 | [7 empty] | 0.010372 | 0.010054 | 0.0099926 | 0.0090555
23 | CCV 10 ppb MAJORS |  | 06/08/2023 15:39:28 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.377 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | -12.8383 | -41.1886 | -0.73032 | -0.80788 | -2.3731 | -1892.7639 | -3.8651 | -28.9428 | 0.0032821 | 0.0021771 | 0.0008629 | -0.53724 | -0.025739 | -1.1955 | 0.0040634 | -0.00086363 | -0.0091461 | -0.0092637 | -0.1815 | 0.53901 | -0.0057419 | 0.15706 | -0.0030757 | -1.1802 | -0.01521 | 0.0015557 | 0.00049202 | -0.049839 | 0.0041782 | -0.011053 | -0.013216 | [7 empty] | 0.00036437 | -0.00088094 | -0.0026241 | -0.0023583
24 | CCV 10 ppb MAJORS |  | 06/08/2023 15:45:42 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.378 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10570.1578 | 10848.3085 | 8.4876 | 9.8854 | 7.878 | 10722.0026 | 10934.0449 | 10683.3747 | 10.7684 | 10.625 | 9.9121 | 9.8704 | 10.3925 | 10.5423 | 10.752 | 10.8388 | 10.5791 | 9.6384 | 9.7375 | 9.9639 | 9.7083 | 9.3648 | 9.9952 | 9.2899 | 9.8888 | 9.6532 | 9.6943 | 10.0276 | 9.5757 | 10.4552 | 10.571 | [7 empty] | 9.7097 | 9.8231 | 9.3278 | 9.859
25 | LLOQ 0.1 |  | 06/08/2023 15:51:31 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LLOQ 0.1.379 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 99.835 | 75.8092 | 0.065718 | -0.051843 | -0.56644 | 279.1031 | 103.7612 | 103.8505 | 0.096542 | 0.11222 | 0.10337 | -0.098896 | 0.089628 | -0.1612 | 0.097228 | 0.11966 | 0.11388 | 0.10932 | 0.10192 | 0.15234 | 0.05636 | 0.24575 | 0.095374 | 0.14192 | 0.036329 | 0.11013 | 0.10514 | 0.1039 | 0.12918 | 0.09714 | 0.096442 | [7 empty] | 0.094455 | 0.09834 | 0.094652 | 0.099408
26 | STD 0.5 |  | 06/08/2023 15:57:19 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 0.5.380 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 501.3734 | 488.7631 | 0.46956 | 0.75969 | 0.31372 | 724.5439 | 471.3807 | 482.3687 | 0.51427 | 0.54837 | 0.56053 | 0.69189 | 0.51624 | 0.53141 | 0.51407 | 0.527 | 0.55145 | 0.60878 | 0.61458 | 0.53443 | 0.43089 | 0.61411 | 0.50895 | 0.60207 | 0.44687 | 0.53376 | 0.52509 | 0.45274 | 0.40338 | 0.57524 | 0.51412 | [7 empty] | 0.55709 | 0.55807 | 0.49293 | 0.51428
27 | 8_2_21_UF_1 |  | 06/08/2023 16:03:07 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_UF_1.381 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 5502.7123 | 5515.3958 | 1.5052 | 1.5284 | 5.452 | 12329.3522 | 1218.5089 | 28419.4625 | 0.396 | 0.062235 | 0.058805 | 0.32235 | 15.1885 | 0.62754 | 0.020633 | 0.24011 | 0.93314 | 0.90947 | 1.4522 | 4.3289 | 3.7252 | 4.3594 | 4.3801 | 0.32472 | 0.035145 | 0.049722 | 0.048206 | 0.30367 | 0.035435 | 5.2008 | 5.2956 | [7 empty] | 0.006901 | 0.0061183 | 0.013707 | 0.015818
28 | 8_2_21_UF_6 |  | 06/08/2023 16:08:55 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_UF_6.382 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10673.5634 | 6993.106 | 0.19846 | 0.11781 | 3.1604 | 12165.0275 | 2095.5035 | 34992.2024 | 1.0089 | 0.1001 | 0.089526 | 0.4944 | 13.7741 | 0.7065 | 0.022205 | 0.20966 | 2.1545 | 2.1857 | 2.2887 | 5.8023 | 4.8226 | 5.8373 | 6.1364 | 0.35773 | 0.090448 | 0.11082 | 0.11838 | 0.40007 | 0.11578 | 3.4543 | 3.4772 | [7 empty] | 0.018245 | 0.015568 | 0.032344 | 0.034105
29 | 8_2_21_UF_6_SP |  | 06/08/2023 16:14:43 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_UF_6_SP.383 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 16267.1399 | 12850.6642 | 5.5262 | 5.6011 | 10.7677 | 17182.7234 | 7761.035 | 40810.4893 | 6.4601 | 5.409 | 5.5981 | 6.5072 | 19.9571 | 6.7763 | 5.5375 | 5.6914 | 7.3306 | 7.3327 | 7.4577 | 10.9728 | 10.2813 | 10.6111 | 11.5188 | 5.2769 | 5.4806 | 5.3618 | 5.4247 | 5.9916 | 5.4688 | 9.0518 | 9.3038 | [7 empty] | 5.2923 | 5.283 | 5.1818 | 5.5548
30 | STD 1 |  | 06/08/2023 16:20:31 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 1.384 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1061.9021 | 1022.3951 | 1.1752 | 1.0968 | -0.058185 | 1483.1061 | 1098.78 | 1068.6497 | 1.0982 | 1.0578 | 0.99866 | 1.5526 | 1.0694 | 1.292 | 1.0498 | 1.0304 | 1.0501 | 1.0395 | 0.9972 | 1.0075 | 0.94017 | 1.1671 | 1.0899 | 1.1396 | 1.0234 | 0.97816 | 1.0129 | 1.068 | 0.92071 | 1.0446 | 1.0856 | [7 empty] | 1.0151 | 1.0189 | 1.0151 | 1.0442
31 | 8_2_21_UF_6_SP_LD |  | 06/08/2023 16:26:19 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_UF_6_SP_LD.385 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 16322.174 | 13028.9898 | 6.4476 | 6.0626 | 10.7892 | 17646.209 | 7587.6511 | 41803.4979 | 6.4397 | 5.4944 | 5.6204 | 5.9489 | 19.4195 | 6.1644 | 5.3805 | 5.673 | 7.4264 | 7.3524 | 7.5718 | 11.1056 | 10.0499 | 10.7755 | 11.521 | 5.2366 | 5.1493 | 5.2769 | 5.5188 | 6.3339 | 5.4734 | 9.1674 | 9.1538 | [7 empty] | 5.3426 | 5.34 | 5.1497 | 5.4843
32 | 8_2_21_UF_6_SP_LSP |  | 06/08/2023 16:32:06 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_UF_6_SP_LSP.386 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 73946.3549 | 66735.5321 | 20.0491 | 21.1067 | 26.6788 | 72722.0693 | 64321.6622 | 102925.7942 | 21.5038 | 19.8781 | 19.0814 | 20.7021 | 33.8219 | 22.1281 | 19.6579 | 19.4269 | 20.0057 | 19.2621 | 20.0408 | 23.9243 | 22.4755 | 22.9167 | 24.685 | 17.6732 | 18.1695 | 18.2202 | 18.7741 | 20.1524 | 18.5208 | 23.4714 | 23.1529 | [7 empty] | 17.3447 | 17.5556 | 16.0727 | 17.2485
33 | 8_2_21_UF_8 |  | 06/08/2023 16:37:54 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_UF_8.387 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 11514.0465 | 7057.0144 | 0.13691 | -0.039635 | 9.1591 | 10630.6795 | 2365.1317 | 32571.7492 | 1.365 | 0.17028 | 0.15347 | 1.1273 | 13.7759 | 1.3776 | 0.021529 | 0.1843 | 2.1273 | 2.1268 | 0.93688 | 6.4983 | 5.8167 | 6.3577 | 6.8399 | 0.41872 | 0.062266 | 0.1176 | 0.12602 | 0.53474 | 0.10292 | 2.8528 | 2.9017 | [7 empty] | 0.013118 | 0.011266 | 0.039558 | 0.042149
34 | 8_2_21_UF_9 |  | 06/08/2023 16:43:42 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_UF_9.388 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 8448.7078 | 6376.0207 | 0.13432 | 0.1801 | 7.3222 | 8589.4569 | 1837.3845 | 28046.7606 | 0.96084 | 0.088481 | 0.073535 | 0.1373 | 6.7035 | 0.37745 | 0.020649 | 0.1692 | 1.4687 | 1.4832 | 0.76957 | 4.6263 | 3.8894 | 4.5789 | 4.8122 | 0.34548 | 0.068057 | 0.089647 | 0.084482 | 0.32505 | 0.1046 | 1.9726 | 1.9953 | [7 empty] | 0.010172 | 0.0089667 | 0.044858 | 0.047201
35 | LBLANK |  | 06/08/2023 16:49:31 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.389 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | -10.0872 | -42.7414 | -0.6818 | -0.77198 | -2.1827 | 321.7818 | 2.5662 | -28.4866 | 0.0045481 | 0.0012304 | 0.0030117 | -0.50857 | -0.025009 | -0.61299 | 0.0020636 | 0.013225 | -0.0081338 | -0.0073653 | -0.16558 | 0.015361 | 0.021366 | 0.18901 | 0.032126 | 0.09901 | -0.00054305 | 0.014658 | 0.0074398 | 0.010049 | 0.012979 | -0.0097971 | -0.011526 | [7 empty] | 0.00024118 | -0.00038652 | -0.0023727 | -0.0017448
36 | CCV 1ppb TRACE |  | 06/08/2023 16:55:19 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.390 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 958.2467 | 954.1637 | 1.0925 | 1.0315 | -1.1519 | 1193.1936 | 960.3951 | 1029.761 | 0.99255 | 0.95997 | 1.0347 | 0.69564 | 1.0961 | 0.45499 | 0.99437 | 1.0008 | 1.0031 | 1.0593 | 0.92424 | 0.98373 | 0.86731 | 1.132 | 1.0571 | 1.0665 | 0.98164 | 0.94561 | 1.0226 | 0.9751 | 1.0735 | 1.0733 | 1.1139 | [7 empty] | 1.0537 | 1.0673 | 1.0081 | 1.077
37 | CCV 1ppb TRACE |  | 06/08/2023 17:01:37 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.391 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1012.0937 | 1074.9191 | 1.0866 | 0.89436 | -1.1615 | 1573.0823 | 1035.5289 | 1096.9143 | 1.0931 | 1.0022 | 1.0463 | 0.68502 | 1.0739 | 0.60055 | 1.054 | 1.1651 | 1.0659 | 1.0682 | 0.91587 | 0.96788 | 0.96679 | 1.1457 | 1.0544 | 1.0106 | 1.0155 | 1.0507 | 1.103 | 0.99177 | 1.1424 | 1.0737 | 1.0832 | [7 empty] | 1.0343 | 1.0425 | 0.99995 | 1.0728
38 | CCV 10 ppb MAJORS |  | 06/08/2023 17:07:25 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.392 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10449.6951 | 10827.3927 | 8.8754 | 9.385 | 8.2383 | 10773.6911 | 10424.1421 | 10695.1868 | 10.6484 | 10.3125 | 10.4227 | 10.1615 | 10.7974 | 10.6119 | 10.7221 | 10.7614 | 10.186 | 10.1361 | 10.2327 | 10.0147 | 9.7206 | 9.9877 | 10.6143 | 9.7898 | 10.0967 | 9.9586 | 10.3682 | 10.7447 | 10.154 | 10.8832 | 10.8953 | [7 empty] | 10.0934 | 10.2379 | 9.6027 | 10.3107
39 | CCV 10 ppb MAJORS |  | 06/08/2023 17:13:42 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.393 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10344.4684 | 10497.8461 | 8.8338 | 9.7847 | 8.2239 | 10880.8404 | 10208.6794 | 10646.9174 | 10.4707 | 10.391 | 10.2279 | 10.1228 | 10.7894 | 10.4285 | 10.7659 | 10.7974 | 10.1789 | 9.9797 | 10.0794 | 9.9816 | 9.4301 | 9.5176 | 10.4027 | 9.9573 | 9.9292 | 10.162 | 10.2351 | 10.7463 | 10.1786 | 10.762 | 10.8587 | [7 empty] | 9.9789 | 10.0871 | 9.5894 | 10.074
40 | 8_2_21_UF_10 |  | 06/08/2023 17:19:31 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_UF_10.394 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9598.4697 | 9886.1598 | 0.19407 | 0.19832 | 10.4002 | 16615.9048 | 2344.256 | 40010.3505 | 1.0446 | 0.07078 | 0.055195 | 0.2058 | 5.8855 | 0.62071 | 0.025608 | 0.21285 | 1.4485 | 1.3266 | 0.96015 | 5.5178 | 5.256 | 5.2426 | 5.626 | 0.44897 | 0.06676 | 0.1211 | 0.11179 | 0.44487 | 0.12025 | 2.0369 | 2.0181 | [7 empty] | 0.007992 | 0.0080177 | 0.062174 | 0.064752
41 | 8_2_21_UF_11 |  | 06/08/2023 17:25:19 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_UF_11.395 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9051.3898 | 10150.4829 | 0.10919 | 0.090119 | 9.0303 | 16970.7616 | 2386.5172 | 41748.6807 | 1.1614 | 0.90429 | 0.85462 | 7.1801 | 6.5406 | 8.2762 | 0.032921 | 0.2131 | 1.2657 | 1.2265 | 0.77614 | 5.4533 | 5.5926 | 5.3587 | 5.7281 | 0.44271 | 0.058287 | 0.095793 | 0.093982 | 0.40543 | 0.078401 | 1.9382 | 1.964 | [7 empty] | 0.00627 | 0.0052155 | 0.10511 | 0.11039
42 | 8_2_21_UF_11_DUP |  | 06/08/2023 17:31:09 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_UF_11_DUP.396 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 8241.29 | 9336.6655 | 0.078931 | -0.065985 | 8.828 | 14500.3266 | 2120.1682 | 36333.7762 | 0.97916 | 0.057679 | 0.053634 | -0.045502 | 6.4859 | 0.19076 | 0.021321 | 0.15092 | 1.1536 | 1.2326 | 0.81072 | 5.3027 | 4.4926 | 5.3103 | 5.5546 | 0.39162 | 0.087352 | 0.10057 | 0.098759 | 0.41687 | 0.14238 | 1.9005 | 1.8997 | [7 empty] | 0.0051917 | 0.0049684 | 0.10233 | 0.10706
43 | 8_2_21_F_1 |  | 06/08/2023 17:36:57 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_F_1.397 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4792.7967 | 6094.6747 | 2.7795 | 2.8282 | 3.7619 | 14307.4564 | 1236.5063 | 30800.8928 | 0.38222 | 0.05173 | 0.047766 | 6.5388 | 16.9092 | 6.9413 | 0.020599 | 0.29549 | 1.2505 | 1.2564 | 1.9154 | 4.5765 | 3.9524 | 4.8132 | 4.9396 | 0.28602 | 0.034128 | 0.0648 | 0.056557 | 0.32561 | 0.073003 | 5.9789 | 6.0645 | [7 empty] | 0.0075556 | 0.0077194 | 0.04365 | 0.045834
44 | 8_2_21_F_6 |  | 06/08/2023 17:44:00 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_F_6.398 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9870.3663 | 7479.4748 | 2.0716 | 1.8968 | 1.9888 | 13679.8545 | 2229.7006 | 37592.376 | 0.98798 | 0.41592 | 0.43471 | 10.283 | 13.7263 | 10.4716 | 0.02997 | 0.27013 | 4.0298 | 4.0867 | 3.4979 | 6.2147 | 5.5282 | 5.5969 | 5.9565 | 0.33265 | 0.11378 | 0.1182 | 0.099046 | 0.51878 | 0.1419 | 3.6802 | 3.4071 | [7 empty] | 0.01971 | 0.017708 | 0.095176 | 0.09544
45 | 8_2_21_F_8 |  | 06/08/2023 17:49:48 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_F_8.399 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 11499.5573 | 7495.2387 | 2.2742 | 2.8496 | 8.6639 | 11775.331 | 2639.3459 | 36017.8876 | 1.4132 | 0.12737 | 0.13059 | 6.1418 | 14.823 | 6.7343 | 0.025819 | 0.21305 | 4.3378 | 4.2282 | 2.2153 | 7.0102 | 6.3684 | 6.6055 | 6.9809 | 0.43361 | 0.065256 | 0.12849 | 0.096311 | 0.60259 | 0.097011 | 3.2096 | 3.1028 | [7 empty] | 0.011777 | 0.013197 | 0.09834 | 0.10214
46 | 8_2_21_F_8_DUP |  | 06/08/2023 17:55:36 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_F_8_DUP.400 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 13578.8846 | 8806.7453 | 2.264 | 2.6888 | 8.3294 | 14228.6245 | 2957.4 | 43547.8511 | 1.6198 | 0.062983 | 0.054222 | 5.78 | 15.1571 | 7.5621 | 0.034725 | 0.24438 | 5.1421 | 4.3268 | 2.4091 | 6.9238 | 6.9503 | 6.8477 | 7.2295 | 0.45331 | 0.12318 | 0.10519 | 0.13791 | 0.59814 | 0.14086 | 3.2218 | 3.126 | [7 empty] | 0.017314 | 0.01522 | 0.10086 | 0.10677
47 | 8_2_21_F_8_DUP_LD |  | 06/08/2023 18:01:24 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_F_8_DUP_LD.401 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 11009.9034 | 7560.4159 | 4.0612 | 4.4295 | 12.9022 | 11281.9472 | 2477.2675 | 35525.7403 | 1.4023 | 0.08326 | 0.076935 | 13.1077 | 15.4068 | 13.7768 | 0.028096 | 0.23385 | 4.3116 | 4.1097 | 2.5675 | 7.0105 | 5.7353 | 6.6985 | 7.103 | 0.44583 | 0.10884 | 0.11354 | 0.12107 | 0.60585 | 0.12467 | 3.05 | 3.2086 | [7 empty] | 0.013827 | 0.012919 | 0.10932 | 0.11128
48 | 8_2_21_F_8_DUP_LD |  | 06/08/2023 18:07:40 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_F_8_DUP_LD.402 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 11456.6737 | 7453.9698 | 4.1634 | 4.643 | 10.5713 | 11715.5566 | 2542.546 | 35658.55 | 1.4098 | 0.089804 | 0.077437 | 13.4773 | 15.5996 | 14.521 | 0.026216 | 0.22741 | 4.4075 | 4.2584 | 2.6589 | 6.951 | 5.5812 | 7.0372 | 7.2644 | 0.3823 | 0.12407 | 0.11039 | 0.11858 | 0.50839 | 0.077863 | 3.0866 | 3.1976 | [7 empty] | 0.014075 | 0.014319 | 0.10668 | 0.11434
49 | 8_2_21_F_8_DUP_LSP |  | 06/08/2023 18:13:28 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_F_8_DUP_LSP.403 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 70322.5868 | 62147.1141 | 18.0702 | 18.9009 | 27.2479 | 69112.0548 | 60451.4375 | 94052.3193 | 15.8954 | 14.0499 | 14.2974 | 21.5694 | 30.6798 | 21.9376 | 13.9673 | 13.88 | 17.0964 | 16.909 | 15.5205 | 20.4712 | 18.7385 | 19.3883 | 21.182 | 13.1075 | 13.9567 | 13.6072 | 13.7101 | 14.8969 | 13.5431 | 17.8698 | 17.8033 | [7 empty] | 12.8337 | 12.8337 | 11.7547 | 12.5433
50 | LBLANK |  | 06/08/2023 18:19:16 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.404 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | -6.8622 | -38.834 | -0.69722 | -0.77278 | -1.8744 | 551.2187 | 6.2902 | -30.7443 | 0.0034021 | 0.0031958 | 0.0038855 | -0.51479 | -0.024432 | -0.56598 | 0.0020661 | 0.017722 | -0.0067889 | -0.0060706 | -0.16573 | -0.022463 | 0.0055661 | 0.1642 | 0.0051691 | 0.17826 | -0.021858 | 0.0011072 | -0.00032446 | 0.010958 | 0.009361 | -0.0063377 | -0.0099329 | [7 empty] | 0.00022846 | -0.00016505 | -0.0018503 | -0.0016199
51 | CCV 1ppb TRACE |  | 06/08/2023 18:25:05 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.405 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1121.0672 | 1101.9409 | 1.1446 | 0.94174 | -1.0595 | 1850.787 | 1160.7013 | 1166.2788 | 1.1208 | 1.1122 | 1.0651 | 0.72696 | 1.147 | 0.83038 | 1.1214 | 1.1956 | 1.183 | 1.077 | 0.91968 | 1.0707 | 1.0168 | 1.1313 | 1.0815 | 1.0937 | 1.1362 | 1.0544 | 1.1351 | 1.1711 | 1.0063 | 1.1082 | 1.1224 | [7 empty] | 1.0277 | 1.0435 | 0.99686 | 1.0715
52 | CCV 1ppb TRACE |  | 06/08/2023 18:31:23 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.406 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1074.0423 | 1035.7773 | 1.2255 | 1.0572 | -1.1292 | 1627.1314 | 1053.138 | 1080.062 | 1.1531 | 1.0691 | 1.0641 | 0.70203 | 1.1187 | 0.78154 | 1.0968 | 1.1118 | 1.1253 | 1.093 | 0.90988 | 1.0676 | 0.83794 | 1.1339 | 1.0942 | 1.0592 | 1.0457 | 1.0296 | 1.0514 | 1.0642 | 0.96887 | 1.0689 | 1.1021 | [7 empty] | 1.0265 | 1.0467 | 1.0068 | 1.0562
53 | LFB |  | 06/08/2023 18:37:11 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.407 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10161.1608 | 10441.4137 | 9.3182 | 9.2826 | 11.0197 | 11171.1939 | 10037.7278 | 10717.8098 | 10.3496 | 10.2957 | 10.3642 | 10.4413 | 11.0351 | 10.3389 | 10.7057 | 10.6916 | 10.2054 | 10.2777 | 10.5841 | 10.1772 | 9.3958 | 9.9527 | 10.7524 | 10.1355 | 10.3593 | 10.1254 | 10.6399 | 11.1844 | 10.7475 | 10.7279 | 11.067 | [7 empty] | 10.0358 | 10.1841 | 9.3228 | 9.9476
54 | CCV 10 ppb MAJORS |  | 06/08/2023 18:42:59 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.408 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10458.6717 | 10458.0202 | 8.9889 | 9.0553 | 8.7825 | 10989.2516 | 10201.0749 | 10682.2509 | 10.4222 | 10.3563 | 10.7701 | 10.755 | 11.2439 | 10.4457 | 10.6393 | 10.8685 | 10.4075 | 10.5617 | 10.5997 | 10.1685 | 9.2452 | 9.8126 | 10.6672 | 9.9977 | 10.7093 | 10.3419 | 10.5523 | 10.5936 | 10.8071 | 11.173 | 11.1479 | [7 empty] | 10.2947 | 10.3651 | 9.4158 | 10.1256
55 | CCV 10 ppb MAJORS |  | 06/08/2023 18:49:21 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.409 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 11226.1341 | 11171.3778 | 8.0549 | 10.1827 | 8.6149 | 11890.9129 | 11159.7109 | 11518.3926 | 11.3565 | 11.4204 | 10.7401 | 10.5869 | 11.286 | 11.3798 | 11.5004 | 11.9248 | 11.1092 | 10.4923 | 10.4743 | 9.3896 | 10.3233 | 9.9616 | 10.7008 | 8.8304 | 10.6567 | 10.204 | 10.5193 | 9.8819 | 10.6587 | 11.0819 | 11.1202 | [7 empty] | 10.2097 | 10.2815 | 8.5632 | 10.1849
56 | LBLANK |  | 06/08/2023 19:26:31 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.410 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | -12.5329 | -42.7414 | -0.6932 | -0.76964 | -2.2011 | 88.4759 | 5.7847 | -30.7443 | 0.0034369 | 0.0049507 | 0.003945 | -0.5048 | -0.024582 | -0.52786 | 0.0018916 | 0.0051999 | -0.0063505 | -0.0070341 | -0.14708 | 0.080995 | 0.00023996 | 0.16015 | 0.00082783 | 0.18601 | -0.017423 | 0.0054974 | 0.0025319 | 0.064297 | -0.0014165 | -0.009219 | -0.010154 | [7 empty] | 0.00024033 | -0.00033309 | -0.0022167 | -0.0018331
57 | PE ICV |  | 06/08/2023 19:32:19 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\PE ICV.411 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10665.6053 | 10833.5347 | 9.3065 | 10.4375 | 20.2114 | 11365.3898 | 10528.3076 | 10918.0171 | 10.7558 | 10.6705 | 10.6362 | 10.865 | 11.4466 | 11.0672 | 10.9086 | 11.0381 | 10.6237 | 10.457 | 10.4061 | 9.9477 | 8.808 | 9.8602 | 10.7122 | 7.0581 | 10.1976 | 10.5756 | 10.7054 | 10.733 | 10.6431 | 10.7202 | 10.777 | [7 empty] | 10.2901 | 10.398 | 9.3585 | 10.0806
58 | 8_2_21_F_9 |  | 06/08/2023 19:38:07 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_F_9.412 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 7735.7764 | 6747.8197 | 1.9095 | 2.3241 | 6.9195 | 9781.4137 | 2025.2098 | 31186.209 | 1.0385 | 0.13673 | 0.11855 | 5.3315 | 7.8434 | 5.846 | 0.021986 | 0.19577 | 3.3724 | 3.1887 | 1.5009 | 5.034 | 4.5046 | 4.8741 | 5.2868 | 0.42041 | 0.039221 | 0.078945 | 0.08581 | 0.35005 | 0.085852 | 2.2429 | 2.2707 | [7 empty] | 0.012202 | 0.012582 | 0.21189 | 0.22323
59 | 8_2_21_F_10 |  | 06/08/2023 19:43:55 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_F_10.413 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 8124.2163 | 9838.9341 | 3.54 | 3.8248 | 9.6386 | 17244.5966 | 2387.7536 | 41618.5546 | 1.0745 | 0.065264 | 0.055426 | 4.3578 | 7.2034 | 5.2629 | 0.034649 | 0.23057 | 2.6651 | 2.4919 | 2.4205 | 5.6501 | 5.5331 | 5.5438 | 5.9772 | 0.54537 | 0.050007 | 0.084771 | 0.11108 | 0.49156 | 0.12348 | 2.1848 | 2.2009 | [7 empty] | 0.0097465 | 0.010742 | 0.13597 | 0.14464
60 | 8_2_21_F_10_DUP |  | 06/08/2023 19:49:43 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_F_10_DUP.414 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 7995.6742 | 9324.7709 | 3.539 | 3.4212 | 9.2696 | 16499.2578 | 2261.4864 | 39124.606 | 0.98698 | 0.26796 | 0.25495 | 5.5447 | 7.3102 | 5.939 | 0.027246 | 0.2237 | 2.5216 | 2.5323 | 2.3194 | 5.7496 | 5.1627 | 5.3895 | 5.9424 | 0.45504 | 0.073821 | 0.10134 | 0.10775 | 0.49512 | 0.13646 | 2.2017 | 2.226 | [7 empty] | 0.011615 | 0.0098388 | 0.13669 | 0.14047
61 | 8_2_21_F_11 |  | 06/08/2023 19:55:32 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_F_11.415 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 7542.4887 | 9632.7621 | 3.8746 | 3.9782 | 8.3791 | 16444.6123 | 2247.1013 | 39599.1672 | 1.0704 | 0.08026 | 0.06355 | 4.7243 | 8.0376 | 5.0257 | 0.027594 | 0.22973 | 2.7439 | 2.7966 | 1.5078 | 5.6662 | 4.7211 | 5.4801 | 5.9019 | 0.4609 | 0.12744 | 0.11729 | 0.10513 | 0.45718 | 0.09178 | 2.0434 | 2.0547 | [7 empty] | 0.0069367 | 0.0062619 | 0.39355 | 0.44953
62 | 8_2_21_F_11_SP |  | 06/08/2023 20:01:20 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_F_11_SP.416 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 13668.1948 | 15539.8554 | 9.4804 | 10.2143 | 14.5156 | 22293.1503 | 8195.0915 | 45968.7414 | 6.9717 | 5.976 | 6.1323 | 12.2207 | 14.2297 | 12.0601 | 5.7339 | 5.8872 | 8.0042 | 7.9533 | 6.9627 | 11.3835 | 9.9409 | 10.7033 | 11.5608 | 5.5694 | 5.6856 | 5.6063 | 5.5626 | 6.361 | 5.455 | 8.0623 | 8.1128 | [7 empty] | 5.5233 | 5.5237 | 5.7947 | 5.9916
63 | 8_2_21_W_1 |  | 06/08/2023 20:07:08 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_W_1.417 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 961.8842 | 1198.8293 | 6.7818 | 8.247 | 1.0777 | 3362.1385 | 278.7989 | 6094.6939 | 0.085682 | 0.036386 | 0.032159 | 7.9725 | 5.1767 | 8.2102 | 0.012373 | 0.098668 | 0.44846 | 0.46944 | 0.13885 | 1.0415 | 0.76934 | 1.1233 | 1.0536 | 0.20236 | 0.010864 | 0.015953 | 0.016704 | 0.036815 | 0.014058 | 1.1677 | 1.2008 | [7 empty] | 0.0025601 | 0.0024969 | 0.0357 | 0.035077
64 | 8_2_21_W_6 |  | 06/08/2023 20:12:56 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_W_6.418 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1956.6511 | 1441.6878 | 12.6143 | 14.3995 | 1.4443 | 3147.5132 | 499.7574 | 7569.36 | 0.23612 | 0.054883 | 0.050306 | 15.6104 | 6.2543 | 16.3252 | 0.015591 | 0.04863 | 1.5903 | 1.5115 | 1.009 | 1.3483 | 1.2493 | 1.4302 | 1.3798 | 0.27795 | 0.01459 | 0.024116 | 0.024613 | 0.10433 | 0.036505 | 0.73193 | 0.72366 | [7 empty] | 0.007499 | 0.0063291 | 0.12876 | 0.13115
65 | 8_2_21_W_6_DUP |  | 06/08/2023 20:18:44 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_W_6_DUP.419 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1946.0628 | 1473.4633 | 11.1201 | 12.242 | 1.3729 | 3088.0709 | 372.356 | 7293.0718 | 0.21385 | 0.044112 | 0.040668 | 13.4954 | 6.0153 | 14.4264 | 0.011063 | 0.080583 | 1.4585 | 1.4259 | 0.85122 | 1.3073 | 1.038 | 1.3848 | 1.3692 | 0.19814 | -0.012608 | 0.036488 | 0.023081 | 0.079564 | 0.0063549 | 0.71941 | 0.72878 | [7 empty] | 0.0074353 | 0.0056539 | 0.1184 | 0.11826
66 | LBLANK |  | 06/08/2023 20:24:33 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.420 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | -13.4076 | -42.7414 | -0.68761 | -0.75313 | -2.0219 | 429.4725 | 9.1767 | -30.7443 | 0.0040659 | 0.0044984 | 0.0028794 | -0.51172 | -0.025153 | -0.53747 | 0.0018928 | -0.0020792 | -0.0090217 | -0.0066424 | -0.16337 | 0.043694 | 0.00034057 | 0.159 | -0.00088352 | 0.15064 | -0.017448 | 0.0012374 | 0.0018458 | 0.016555 | -0.0014165 | -0.011277 | -0.013054 | [7 empty] | 0.00024654 | -0.00035722 | -0.0022487 | -0.0019217
67 | CCV 1ppb TRACE |  | 06/08/2023 20:30:21 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.421 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1103.0697 | 1020.7486 | 1.2577 | 1.2835 | -1.1208 | 1470.9034 | 1089.9895 | 1119.9024 | 1.1846 | 1.0814 | 1.0739 | 0.75535 | 1.1138 | 0.73685 | 1.1072 | 1.1212 | 1.1034 | 1.0981 | 0.93956 | 1.1154 | 1.0021 | 1.1325 | 1.0663 | 1.3082 | 0.90907 | 1.0418 | 1.0974 | 1.1586 | 1.0195 | 1.1305 | 1.0976 | [7 empty] | 1.0732 | 1.0758 | 1.0851 | 1.0828
68 | CCV 1ppb TRACE |  | 06/08/2023 20:36:39 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.422 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1055.8381 | 1030.8016 | 1.1158 | 1.1625 | -1.1697 | 1574.9501 | 1005.5786 | 1065.5611 | 1.1263 | 1.0592 | 1.0742 | 0.72877 | 1.1246 | 0.81247 | 1.1487 | 1.1242 | 1.1338 | 1.1124 | 0.91885 | 1.0537 | 1.0406 | 1.1137 | 1.086 | 1.1575 | 0.93309 | 0.99768 | 1.0916 | 1.0536 | 1.1759 | 1.0975 | 1.1172 | [7 empty] | 1.058 | 1.0696 | 1.0132 | 1.0869
69 | CCV 10 ppb MAJORS |  | 06/08/2023 20:42:27 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.423 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10394.4355 | 10123.1069 | 8.9606 | 10.4524 | 8.5693 | 10636.0102 | 9949.1138 | 10421.0072 | 10.3064 | 10.1256 | 10.7798 | 10.8039 | 11.3846 | 10.0072 | 10.3051 | 10.4742 | 10.0267 | 10.5788 | 10.747 | 10.0081 | 8.8278 | 9.7108 | 10.7809 | 9.6899 | 10.5624 | 10.2631 | 10.5892 | 10.53 | 10.8058 | 11.2381 | 11.3551 | [7 empty] | 10.4028 | 10.5105 | 9.7307 | 10.2385
70 | CCV 10 ppb MAJORS |  | 06/08/2023 20:48:44 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.424 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10712.17 | 10386.3941 | 8.9557 | 9.6579 | 8.7643 | 11081.0986 | 10424.5086 | 10876.6374 | 10.7028 | 10.5454 | 10.7862 | 10.679 | 11.3673 | 10.6017 | 11.0031 | 10.9919 | 10.4889 | 10.4813 | 10.6064 | 10.2466 | 9.8342 | 9.9471 | 10.8315 | 10.1333 | 10.4448 | 10.5841 | 10.8201 | 10.7146 | 10.7312 | 11.1392 | 11.438 | [7 empty] | 10.2411 | 10.4511 | 9.6987 | 10.0626
71 | 8_2_21_W_8 |  | 06/08/2023 20:54:33 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_W_8.425 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2176.9882 | 1423.0907 | 7.0443 | 7.4512 | 2.8816 | 2663.1656 | 452.5192 | 7023.7502 | 0.28368 | 0.075693 | 0.076502 | 11.0021 | 7.7723 | 10.6713 | 0.011914 | 0.057119 | 1.2964 | 1.3245 | 0.578 | 1.5218 | 1.2487 | 1.5534 | 1.5343 | 0.25968 | 0.00077636 | 0.037115 | 0.030831 | 0.097545 | 0.029071 | 0.63858 | 0.61072 | [7 empty] | 0.0049097 | 0.005197 | 0.07127 | 0.075417
72 | 8_2_21_W_9 |  | 06/08/2023 21:00:21 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_W_9.426 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1497.4462 | 1281.5104 | 6.1538 | 6.76 | 2.494 | 2142.3378 | 357.7206 | 6170.1572 | 0.19047 | 0.026515 | 0.029448 | 7.9639 | 4.8736 | 8.1376 | 0.009638 | 0.053089 | 0.83868 | 0.87619 | 0.29957 | 1.0537 | 0.87712 | 1.0987 | 1.0594 | 0.14221 | 0.0054482 | 0.022974 | 0.015819 | 0.0042289 | 0.010149 | 0.43297 | 0.43482 | [7 empty] | 0.004109 | 0.0031299 | 0.040506 | 0.044479
73 | 8_2_21_W_9_DUP |  | 06/08/2023 21:06:09 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_W_9_DUP.427 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1478.2644 | 1266.3299 | 5.7879 | 7.3462 | 1.9683 | 2086.8939 | 363.3827 | 5698.61 | 0.19705 | 0.021123 | 0.01551 | 7.0819 | 3.7761 | 7.186 | 0.0074554 | 0.05669 | 0.84133 | 0.82631 | 0.30295 | 1.0958 | 0.8746 | 1.1157 | 1.0425 | 0.36652 | -0.012923 | 0.025362 | 0.017463 | 0.032585 | 0.032311 | 0.43016 | 0.42212 | [7 empty] | 0.0032005 | 0.0029198 | 0.039574 | 0.041808
74 | 8_2_21_W_10 |  | 06/08/2023 21:11:57 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_W_10.428 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1513.6007 | 1788.0188 | 21.8921 | 24.7885 | 3.4727 | 3404.979 | 463.9478 | 7556.6481 | 0.21654 | 0.038975 | 0.038709 | 20.3872 | 7.0353 | 21.2392 | 0.019055 | 0.060138 | 1.0792 | 1.0593 | 1.07 | 1.1847 | 0.99959 | 1.2698 | 1.236 | 0.2779 | 0.01064 | 0.017367 | 0.029442 | 0.10327 | 0.014176 | 0.40585 | 0.41931 | [7 empty] | 0.0068326 | 0.0056922 | 0.1182 | 0.13465
75 | 8_2_21_W_11 |  | 06/08/2023 21:17:45 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_2_21_W_11.429 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1401.6033 | 1801.2425 | 37.3112 | 37.0185 | 4.1537 | 3478.566 | 395.2253 | 7375.8407 | 0.2625 | 0.051438 | 0.054972 | 32.0054 | 11.7222 | 31.78 | 0.022709 | 0.061571 | 1.3499 | 1.401 | 2.0755 | 1.2916 | 0.94269 | 1.2932 | 1.2748 | 0.33352 | -0.0077042 | 0.019186 | 0.027715 | 0.082626 | 0.041586 | 0.39542 | 0.39518 | [7 empty] | 0.011556 | 0.0086608 | 0.18973 | 0.19869
76 | 5_10_23_DB |  | 06/08/2023 21:23:33 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_10_23_DB.430 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | -11.218 | -42.7414 | -0.27088 | -0.32272 | 2.5741 | 426.4266 | 1.2341 | -30.7443 | 0.0028168 | 0.0222 | 0.013699 | -0.34176 | -0.0233 | -0.41047 | 0.0018855 | 0.0066639 | 0.035305 | 0.031233 | -0.16475 | 0.037983 | 0.00031299 | 0.15968 | 3.2582e-05 | 0.23609 | -0.017271 | 0.0013219 | 0.00062222 | 0.0048561 | 0.0102 | -0.012152 | -0.013247 | [7 empty] | 6.6371e-05 | -0.00077686 | 0.00060337 | 0.0011691
77 | 5_10_23_MFB |  | 06/08/2023 21:29:21 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_10_23_MFB.431 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 5365.9413 | 5422.2039 | 4.5214 | 5.1516 | 10.7006 | 5647.27 | 5228.4258 | 5402.1154 | 5.5095 | 5.5473 | 5.3196 | 5.3852 | 5.6264 | 5.6972 | 5.5512 | 5.4855 | 5.4827 | 5.218 | 5.1295 | 4.877 | 4.6087 | 4.8073 | 5.2235 | 4.9652 | 5.0612 | 5.2307 | 5.2507 | 5.146 | 5.2214 | 5.3179 | 5.4742 | [7 empty] | 5.1028 | 5.115 | 4.8536 | 5.247
78 | 5_10_23_T-239 |  | 06/08/2023 21:35:10 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_10_23_T-239.432 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 51632.5505 | 13752.5643 | 117.5695 | 122.9852 | 14.2376 | 1319.9589 | 4101.9059 | 54941.2253 | 2.0979 | 1.51 | 1.4545 | 45.1447 | 4.0369 | 47.3263 | 1.5975 | 1.8349 | 0.81993 | 0.84752 | 6.1572 | 2.4492 | 2.1755 | 2.4346 | 2.5447 | 4.809 | 2.0123 | 2.0016 | 2.071 | 2.3242 | 2.1223 | 1.1548 | 1.1845 | [7 empty] | 3.4745 | 3.3905 | 1.2032 | 1.2695
79 | 5_10_23_DB |  | 06/08/2023 21:40:58 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_10_23_DB.433 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | -9.9328 | -42.7414 | -0.32624 | -0.24404 | 3.028 | 465.0854 | 2.7504 | -23.6358 | 0.0046158 | 0.00478 | 0.0019313 | -0.44476 | -0.024751 | -0.49712 | 0.0018717 | 0.005688 | 0.017386 | 0.016541 | -0.17072 | 0.10607 | 0.0062134 | 0.15884 | -0.00081494 | 0.056404 | -0.021858 | -0.00020938 | 0.0030083 | -0.019409 | 0.0021361 | -0.011058 | -0.013027 | [7 empty] | 0.00032773 | -0.00079552 | -0.00090674 | 1.8169e-05
80 | 8_16_21_UF_1 |  | 06/08/2023 21:46:46 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_UF_1.434 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10571.0714 | 6203.0055 | 1.2856 | 1.6222 | 24.7266 | 26924.9868 | 2288.3182 | 40780.7317 | 0.49953 | 0.082009 | 0.080994 | 0.17123 | 14.5453 | 0.61986 | 0.033682 | 0.36497 | 0.74356 | 0.74384 | 0.80466 | 6.3777 | 5.7381 | 6.0383 | 6.713 | 0.58587 | 0.031463 | 0.061469 | 0.06928 | 0.41636 | 0.068041 | 19.8514 | 20.2361 | [7 empty] | 0.0084327 | 0.0079057 | 0.013877 | 0.013417
81 | LBLANK |  | 06/08/2023 21:52:35 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.435 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | -12.1461 | -42.7414 | -0.68933 | -0.75516 | -1.5152 | 289.9104 | 3.389 | -26.044 | 0.0040317 | 0.0038581 | 0.0037226 | -0.50852 | -0.025251 | -0.57173 | 0.0016624 | 0.0070594 | -0.008451 | -0.0066872 | -0.16251 | 0.057302 | 3.4913e-05 | 0.16914 | 0.010652 | 0.16521 | -0.012874 | 0.0054712 | -0.0001752 | -0.021405 | 0.0098861 | -0.0096086 | -0.010268 | [7 empty] | -2.3879e-05 | -0.00067103 | -0.0024194 | -0.0019407
82 | CCV 1ppb TRACE |  | 06/08/2023 21:58:24 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.436 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1083.8139 | 1074.0632 | 1.1396 | 1.6152 | -0.91991 | 1415.5344 | 1073.1178 | 1100.4318 | 1.1266 | 1.1065 | 1.0918 | 0.77772 | 1.1691 | 0.76269 | 1.1289 | 1.1687 | 1.1202 | 1.1199 | 0.93425 | 1.0681 | 0.96265 | 1.1286 | 1.0969 | 1.1809 | 1.0509 | 1.0511 | 1.053 | 1.1198 | 1.1705 | 1.118 | 1.1342 | [7 empty] | 1.0708 | 1.065 | 1.0249 | 1.0864
83 | CCV 1ppb TRACE |  | 06/08/2023 22:04:41 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.437 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1069.9421 | 1040.364 | 1.1282 | 1.1078 | -1.1396 | 1451.9629 | 1027.3233 | 1061.703 | 1.1249 | 1.0939 | 1.0874 | 0.73885 | 1.1369 | 0.78214 | 1.1261 | 1.0755 | 1.18 | 1.1194 | 0.95502 | 1.0204 | 0.95972 | 1.1353 | 1.0911 | 1.1983 | 1.1021 | 1.0175 | 1.0921 | 1.0745 | 1.0357 | 1.1005 | 1.1456 | [7 empty] | 1.0518 | 1.0477 | 1.0283 | 1.1594
84 | LFB |  | 06/08/2023 22:10:29 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.438 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10540.991 | 10387.9602 | 9.1304 | 10.6155 | 11.2237 | 11092.9045 | 9792.7414 | 10652.0255 | 10.5019 | 10.2952 | 10.547 | 10.6313 | 11.1506 | 10.2824 | 10.7343 | 10.6871 | 10.2373 | 10.3789 | 10.6509 | 10.1541 | 9.6477 | 9.7924 | 10.7244 | 9.98 | 10.4568 | 10.6225 | 10.7832 | 11.0253 | 10.9244 | 11.0404 | 11.11 | [7 empty] | 10.1415 | 10.2407 | 9.5039 | 10.0777
85 | CCV 10 ppb MAJORS |  | 06/08/2023 22:16:17 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.439 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10649.4991 | 10628.4209 | 8.8167 | 9.9724 | 9.1292 | 10967.8857 | 10420.3227 | 11135.6009 | 10.6916 | 10.8152 | 11.012 | 10.9275 | 11.4545 | 10.5824 | 11.0306 | 11.03 | 10.424 | 10.6278 | 10.662 | 10.0995 | 9.4107 | 9.7515 | 10.7147 | 10.0855 | 10.2857 | 10.1452 | 10.6641 | 10.4133 | 10.75 | 11.2823 | 11.3646 | [7 empty] | 10.2776 | 10.3963 | 9.593 | 10.2827
86 | CCV 10 ppb MAJORS |  | 06/08/2023 22:22:34 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.440 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10400.6221 | 10452.9318 | 8.8292 | 9.3935 | 8.7749 | 10752.0465 | 10239.9338 | 10814.7052 | 10.4035 | 10.4614 | 11.0016 | 10.8312 | 11.4635 | 10.3305 | 10.8005 | 10.5844 | 10.2011 | 10.6095 | 10.6924 | 10.2414 | 9.5224 | 9.7114 | 10.8677 | 10.2623 | 10.7228 | 10.7197 | 10.6732 | 10.8037 | 10.4807 | 11.253 | 11.5059 | [7 empty] | 10.2741 | 10.3278 | 9.6288 | 10.2904
87 | 8_16_21_UF_6 |  | 06/08/2023 22:28:23 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_UF_6.441 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 11079.104 | 6645.0905 | 0.42058 | 0.74148 | 13.1253 | 18395.6562 | 2034.406 | 37191.7853 | 0.86095 | 0.060879 | 0.054929 | -0.00851 | 12.7497 | 0.30446 | 0.024697 | 0.75079 | 1.6312 | 1.679 | 2.1443 | 5.7462 | 5.0336 | 5.5346 | 6.0317 | 0.65247 | 0.090114 | 0.081445 | 0.10752 | 0.54473 | 0.076008 | 9.8908 | 10.0246 | [7 empty] | 0.015344 | 0.013751 | 0.10338 | 0.11209
88 | 8_16_21_UF_8 |  | 06/08/2023 22:34:11 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_UF_8.442 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 11254.9186 | 7224.9877 | 0.41146 | 0.33161 | 17.9922 | 15457.1844 | 2225.5802 | 35532.3597 | 1.0786 | 0.099927 | 0.084717 | 0.24423 | 13.017 | 0.50941 | 0.022213 | 0.43559 | 1.8104 | 1.8638 | 1.2773 | 5.7752 | 5.4595 | 5.5413 | 5.9833 | 0.5607 | 0.060924 | 0.070799 | 0.082938 | 0.49311 | 0.10492 | 7.8815 | 7.9463 | [7 empty] | 0.012273 | 0.010406 | 0.045325 | 0.050077
89 | 8_16_21_UF_9 |  | 06/08/2023 22:39:59 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_UF_9.443 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9423.6986 | 6758.0814 | 0.81516 | 0.93307 | 17.9306 | 13195.6134 | 1737.2278 | 31442.2764 | 0.89715 | 0.062894 | 0.062216 | -0.08536 | 7.1378 | 0.10518 | 0.018465 | 0.79868 | 1.3202 | 1.3772 | 0.5968 | 4.4339 | 3.3166 | 4.3841 | 4.7403 | 0.53314 | 0.051313 | 0.08474 | 0.068754 | 0.34282 | 0.047698 | 5.1546 | 5.35 | [7 empty] | 0.006824 | 0.0066198 | 0.026635 | 0.027798
90 | 8_16_21_UF_10 |  | 06/08/2023 22:45:47 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_UF_10.444 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9344.4695 | 9408.4553 | 1.3491 | 1.347 | 19.8839 | 15807.3026 | 2031.4214 | 39179.9577 | 0.93195 | 0.097035 | 0.098869 | 0.48985 | 4.0983 | 0.77223 | 0.017635 | 0.31017 | 1.2177 | 1.2489 | 0.98653 | 4.7718 | 3.9598 | 4.6339 | 4.9976 | 0.56954 | 0.10879 | 0.10099 | 0.090313 | 0.39261 | 0.11255 | 2.5288 | 2.541 | [7 empty] | 0.008228 | 0.0076852 | 0.049217 | 0.05518
91 | 8_16_21_UF_11 |  | 06/08/2023 22:51:35 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_UF_11.445 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 8209.4572 | 9033.2918 | 0.24229 | 0.42865 | 17.8392 | 15456.9286 | 2007.2819 | 35032.4647 | 0.88307 | 0.048478 | 0.042249 | -0.22491 | 2.7931 | 0.030921 | 0.012948 | 0.14347 | 0.95225 | 1.0058 | 0.87123 | 4.6843 | 3.9403 | 4.4481 | 4.8207 | 0.57822 | 0.076297 | 0.093077 | 0.084714 | 0.32334 | 0.097566 | 2.0428 | 2.0711 | [7 empty] | 0.0031023 | 0.0030534 | 0.056075 | 0.05859
92 | 8_16_21_F_1 |  | 06/08/2023 22:57:23 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_F_1.446 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9353.5218 | 6316.9708 | 15.8961 | 16.1631 | 14.6921 | 25774.455 | 2228.6304 | 39524.5311 | 0.44209 | 0.069238 | 0.065122 | 5.6426 | 16.9444 | 5.6578 | 0.037541 | 0.36186 | 1.0085 | 1.0988 | 1.0366 | 6.6904 | 5.892 | 6.7298 | 7.3683 | 0.62193 | 0.044597 | 0.069379 | 0.077477 | 0.44857 | 0.080519 | 21.3321 | 22.3451 | [7 empty] | 0.0082081 | 0.006458 | 0.024809 | 0.02789
93 | 8_16_21_F_1_DUP |  | 06/08/2023 23:03:11 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_F_1_DUP.447 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9489.7858 | 6532.8029 | 1.6941 | 1.9541 | 13.0395 | 26862.9716 | 2258.6811 | 41132.2518 | 0.44466 | 0.073441 | 0.070815 | 5.5444 | 16.048 | 5.6965 | 0.035867 | 0.35031 | 1.155 | 1.1962 | 1.022 | 6.5484 | 5.8171 | 6.4902 | 6.9223 | 0.45405 | 0.041546 | 0.072526 | 0.076258 | 0.37299 | 0.06414 | 21.0106 | 20.9236 | [7 empty] | 0.0073884 | 0.006771 | 0.031263 | 0.034907
94 | 8_16_21_F_1_DUP_LD |  | 06/08/2023 23:08:58 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_F_1_DUP_LD.448 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 8935.6098 | 5978.774 | 3.0139 | 2.5672 | 14.828 | 25592.7418 | 2143.3792 | 38856.5913 | 0.40873 | 0.07541 | 0.077194 | 5.6818 | 16.0144 | 5.3947 | 0.041254 | 0.36392 | 1.1173 | 1.2298 | 1.103 | 6.8043 | 5.539 | 6.4043 | 7.0189 | 0.46777 | 0.032424 | 0.053156 | 0.061496 | 0.4353 | 0.081662 | 20.6859 | 21.3088 | [7 empty] | 0.0077573 | 0.0077656 | 0.031747 | 0.03639
95 | 8_16_21_F_1_DUP_LD |  | 06/08/2023 23:15:14 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_F_1_DUP_LD.449 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9368.8164 | 6397.8173 | 3.0205 | 2.9167 | 15.0735 | 26206.0947 | 2267.7375 | 40682.0768 | 0.46082 | 0.083791 | 0.077624 | 5.7046 | 16.0094 | 6.0862 | 0.03838 | 0.41682 | 1.1822 | 1.2214 | 1.0836 | 6.6369 | 5.9217 | 6.1786 | 6.8985 | 0.5778 | 0.031894 | 0.069253 | 0.061492 | 0.49295 | 0.10087 | 20.7664 | 21.104 | [7 empty] | 0.0098196 | 0.0078055 | 0.035166 | 0.03638
96 | 8_16_21_F_1_DUP_LSP |  | 06/08/2023 23:21:03 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_F_1_DUP_LSP.450 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 68464.0809 | 61915.0589 | 17.5294 | 18.0157 | 33.1685 | 83965.9752 | 59788.2915 | 101798.4975 | 15.5557 | 14.3962 | 14.9793 | 22.302 | 31.7038 | 21.6202 | 14.5808 | 14.407 | 14.3922 | 14.3345 | 14.5715 | 20.3666 | 18.6223 | 19.6715 | 21.4438 | 13.506 | 14.1598 | 13.8613 | 14.1908 | 15.1711 | 14.2148 | 37.0989 | 36.6893 | [7 empty] | 12.9117 | 13.0751 | 11.6411 | 11.4061
97 | LBLANK |  | 06/08/2023 23:26:52 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.451 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | -9.9378 | -41.697 | -0.69387 | -0.78747 | -1.8712 | 699.2796 | 7.3142 | -28.3815 | 0.00343 | 0.0018182 | 0.0030687 | -0.50681 | -0.024898 | -0.48754 | 0.0023115 | 0.012101 | -0.0051764 | -0.0076033 | -0.16429 | -0.0042435 | 0.00017371 | 0.1647 | 0.0060638 | 0.49126 | -0.013193 | 0.0038676 | 0.0043455 | 0.1072 | 0.0096492 | -0.00069744 | -0.0066249 | [7 empty] | 0.00014763 | -7.1062e-05 | -0.0019687 | -0.0019592
98 | CCV 1ppb TRACE |  | 06/08/2023 23:32:40 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.452 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1075.7356 | 1065.6145 | 1.1826 | 1.3138 | -1.0627 | 1548.7991 | 1080.3294 | 1122.7302 | 1.1198 | 1.0793 | 1.0998 | 0.81134 | 1.1546 | 0.75138 | 1.101 | 1.1429 | 1.1454 | 1.1363 | 0.96015 | 1.0722 | 0.95401 | 1.1202 | 1.0949 | 1.3668 | 1.0533 | 1.0713 | 1.0358 | 1.0991 | 1.1439 | 1.1506 | 1.1329 | [7 empty] | 1.0267 | 1.0695 | 1.0094 | 1.0642
99 | CCV 1ppb TRACE |  | 06/08/2023 23:38:58 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.453 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1075.3169 | 1055.3616 | 1.1189 | 1.2052 | -1.1385 | 1719.5689 | 1055.5793 | 1112.1303 | 1.1231 | 1.0601 | 1.0837 | 0.7447 | 1.1293 | 0.81715 | 1.1109 | 1.1305 | 1.1414 | 1.1071 | 0.93478 | 1.0399 | 1.019 | 1.104 | 1.0764 | 1.322 | 0.91512 | 1.0297 | 1.133 | 1.094 | 1.1194 | 1.117 | 1.1111 | [7 empty] | 1.0459 | 1.049 | 1.0043 | 1.0623
100 | CCV 10 ppb MAJORS |  | 06/08/2023 23:44:46 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.454 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10522.2205 | 10253.3821 | 8.8086 | 9.7325 | 8.9314 | 10993.611 | 9852.9594 | 10755.8497 | 10.6614 | 10.8242 | 10.9445 | 10.8515 | 11.4439 | 10.5267 | 11.0466 | 11.0528 | 10.5473 | 10.6169 | 10.7119 | 10.0794 | 9.3788 | 9.6111 | 10.6378 | 10.2046 | 10.3166 | 10.3514 | 10.6686 | 10.8022 | 10.4152 | 11.2476 | 11.2625 | [7 empty] | 10.2242 | 10.3082 | 9.5753 | 10.3607
101 | CCV 10 ppb MAJORS |  | 06/08/2023 23:51:04 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.455 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10195.537 | 10325.5385 | 8.7585 | 9.3852 | 8.878 | 10646.2834 | 10001.4775 | 10830.1385 | 10.6174 | 10.3862 | 11.0038 | 10.9492 | 11.5407 | 10.3497 | 10.9763 | 10.8296 | 10.279 | 10.5901 | 10.6735 | 10.2518 | 9.8277 | 9.6832 | 10.6885 | 10.3024 | 10.4745 | 10.3285 | 10.5242 | 10.828 | 10.9306 | 11.205 | 11.3154 | [7 empty] | 10.1536 | 10.3769 | 9.6745 | 10.1861
102 | 8_16_21_F_6 |  | 06/08/2023 23:56:52 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_F_6.456 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10046.722 | 6960.341 | 6.2659 | 5.388 | 11.3099 | 20296.4111 | 2204.5815 | 38446.2251 | 0.82544 | 0.084315 | 0.074965 | 7.9437 | 14.1443 | 7.9625 | 0.033056 | 0.97451 | 3.4583 | 3.4931 | 5.4122 | 6.563 | 5.2572 | 5.8898 | 6.5463 | 0.7746 | 0.080178 | 0.1091 | 0.098123 | 0.64363 | 0.072227 | 11.0302 | 11.1073 | [7 empty] | 0.016381 | 0.018191 | 7.3869 | 7.4684
103 | 8_16_21_F_6_SP |  | 06/09/2023 00:02:41 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_F_6_SP.457 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 16271.5237 | 12881.6732 | 10.3791 | 11.0115 | 19.4842 | 25868.5377 | 7584.8406 | 45633.0892 | 6.8078 | 5.783 | 5.8938 | 13.4726 | 19.6135 | 13.6655 | 5.9408 | 6.7165 | 8.8335 | 8.8356 | 11.0441 | 11.3394 | 10.4647 | 10.4667 | 11.7769 | 5.5791 | 5.5933 | 5.4287 | 5.4188 | 5.9908 | 5.766 | 16.653 | 16.5385 | [7 empty] | 5.3931 | 5.4459 | 5.457 | 5.8963
104 | 8_16_21_F_8 |  | 06/09/2023 00:08:29 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_F_8.458 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10845.9529 | 7778.4499 | 2.5567 | 3.1618 | 17.6744 | 18220.9661 | 2295.8651 | 39163.8822 | 1.2316 | 0.074592 | 0.071498 | 7.0254 | 14.6386 | 7.4707 | 0.028631 | 0.56907 | 4.0763 | 3.8964 | 2.2587 | 6.2272 | 5.6937 | 5.8262 | 6.5425 | 0.5391 | 0.036695 | 0.1213 | 0.10541 | 0.54698 | 0.076143 | 8.5618 | 8.7818 | [7 empty] | 0.0164 | 0.011852 | 0.14636 | 0.15587
105 | 8_16_21_F_8_DUP |  | 06/09/2023 00:14:18 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_F_8_DUP.459 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10608.858 | 7193.0931 | 2.6162 | 2.9868 | 17.2306 | 17236.1285 | 2205.7298 | 35965.26 | 1.0582 | 0.10975 | 0.098627 | 7.2888 | 14.1392 | 7.4583 | 0.03069 | 0.56016 | 3.8273 | 3.8112 | 2.4528 | 6.1559 | 5.1738 | 5.8211 | 6.4794 | 0.57169 | 0.091361 | 0.099505 | 0.091819 | 0.43962 | 0.11886 | 8.5976 | 8.688 | [7 empty] | 0.014038 | 0.01252 | 0.15159 | 0.1558
106 | 8_16_21_F_9 |  | 06/09/2023 00:20:06 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_F_9.460 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 8916.1791 | 6734.3524 | 4.8287 | 5.9473 | 18.5836 | 15603.3094 | 1828.1554 | 33564.631 | 0.95039 | 0.25638 | 0.22948 | 6.394 | 7.9853 | 6.4815 | 0.026678 | 1.2971 | 3.0096 | 2.9994 | 1.3806 | 5.0254 | 4.3433 | 4.6357 | 5.1686 | 0.58311 | 0.09904 | 0.085734 | 0.096962 | 0.36557 | 0.1002 | 5.8187 | 5.8929 | [7 empty] | 0.0084765 | 0.0091219 | 0.10136 | 0.10859
107 | 8_16_21_F_10 |  | 06/09/2023 00:25:54 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_F_10.461 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 8086.155 | 9531.6307 | 3.03 | 3.1353 | 17.296 | 17641.1863 | 2198.587 | 40772.3842 | 0.97611 | 0.051629 | 0.06129 | 3.9816 | 4.8606 | 4.355 | 0.022935 | 0.17813 | 2.3325 | 2.3191 | 1.5278 | 5.1293 | 4.2941 | 5.0224 | 5.5139 | 0.59498 | 0.047852 | 0.089891 | 0.12628 | 0.45934 | 0.094513 | 2.7255 | 2.8369 | [7 empty] | 0.01189 | 0.0093661 | 0.10068 | 0.10757
108 | 8_16_21_F_10_DUP |  | 06/09/2023 00:31:42 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_F_10_DUP.462 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 7704.4669 | 9102.0476 | 3.0858 | 3.2736 | 16.4146 | 16706.5181 | 2074.1772 | 38514.9004 | 0.90655 | 0.071571 | 0.07208 | 3.6301 | 4.4657 | 4.0366 | 0.020522 | 0.50282 | 2.3572 | 2.2581 | 1.3391 | 5.0827 | 4.1434 | 4.6025 | 5.0523 | 0.61589 | 0.10037 | 0.11553 | 0.082062 | 0.44922 | 0.078478 | 2.5753 | 2.5964 | [7 empty] | 0.0097216 | 0.0093537 | 0.10088 | 0.10362
109 | 8_16_21_F_11 |  | 06/09/2023 00:37:30 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_F_11.463 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 7740.7378 | 9437.2523 | 3.9765 | 4.8839 | 15.7305 | 17814.3672 | 2092.7175 | 38476.5404 | 0.97541 | 0.052122 | 0.062038 | 4.1038 | 3.5619 | 4.3378 | 0.024372 | 0.20454 | 2.2459 | 2.2737 | 1.5921 | 5.1555 | 4.1906 | 4.7951 | 5.3725 | 0.63921 | 0.10686 | 0.11108 | 0.11634 | 0.47189 | 0.11485 | 2.3272 | 2.3276 | [7 empty] | 0.0046543 | 0.0046445 | 0.12161 | 0.12237
110 | 8_16_21_W_1 |  | 06/09/2023 00:43:18 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_W_1.464 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1765.1783 | 1154.9932 | 3.3159 | 3.5896 | 6.0955 | 5981.2027 | 398.683 | 7850.108 | 0.084031 | 0.016073 | 0.018885 | 5.7488 | 4.8776 | 4.9137 | 0.013053 | 0.07904 | 0.29607 | 0.35193 | 0.089181 | 1.3938 | 1.1784 | 1.4489 | 1.4348 | 0.33019 | -0.012584 | 0.02188 | 0.008381 | 0.1005 | 0.021963 | 4.6313 | 4.2108 | [7 empty] | 0.0031552 | 0.0016466 | 0.019611 | 0.021345
111 | 8_16_21_W_6 |  | 06/09/2023 00:49:05 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_W_6.465 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2031.2873 | 1397.7537 | 8.4198 | 9.8428 | 1.7467 | 4389.6395 | 429.2917 | 8100.1365 | 0.18767 | 0.029422 | 0.024797 | 11.1882 | 4.9971 | 11.3048 | 0.0097791 | 0.065233 | 1.2216 | 1.1981 | 0.58714 | 1.2939 | 1.209 | 1.3557 | 1.3679 | 0.27836 | 0.015037 | 0.032131 | 0.018883 | 0.17089 | 0.025854 | 2.1625 | 2.1486 | [7 empty] | 0.0049499 | 0.0040669 | 0.09995 | 0.1038
112 | LBLANK |  | 06/09/2023 00:54:53 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.466 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | -12.6241 | -42.7414 | -0.70214 | -0.75557 | -1.8183 | 384.997 | 0.44467 | -30.7443 | 0.0028168 | 0.001162 | 0.0027566 | -0.50728 | -0.025033 | -0.58416 | 0.0016626 | -6.3058e-06 | -0.0079475 | -0.0067071 | -0.16464 | 0.037153 | -0.0057419 | 0.15754 | -0.002414 | 0.19771 | -0.021858 | 0.0012851 | -0.00013577 | -0.05994 | 0.010051 | -0.01102 | -0.011667 | [7 empty] | 6.7726e-05 | -0.00046338 | -0.0020734 | -0.0018346
113 | CCV 1ppb TRACE |  | 06/09/2023 01:00:42 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.467 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1059.0547 | 1023.5366 | 1.0866 | 1.2312 | -0.90312 | 1616.4827 | 1054.7718 | 1051.8822 | 1.1041 | 1.0617 | 1.0762 | 0.75827 | 1.1412 | 0.78446 | 1.1141 | 1.1303 | 1.1258 | 1.104 | 0.92732 | 0.99623 | 0.95554 | 1.1265 | 1.0652 | 1.2126 | 1.1471 | 1.0674 | 1.1117 | 0.99372 | 1.0214 | 1.1014 | 1.1344 | [7 empty] | 1.0313 | 1.0543 | 1.0218 | 1.0961
114 | CCV 1ppb TRACE |  | 06/09/2023 01:06:59 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.468 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1061.3316 | 1049.1345 | 1.0498 | 1.2418 | -1.0603 | 1495.5844 | 993.5035 | 1083.4123 | 1.0719 | 1.0829 | 1.0739 | 0.77724 | 1.1211 | 0.66232 | 1.1061 | 1.1197 | 1.1104 | 1.0858 | 0.90841 | 1.0095 | 0.94692 | 1.0991 | 1.057 | 1.2019 | 1.0128 | 1.0089 | 1.0729 | 1.0827 | 1.1624 | 1.0886 | 1.1345 | [7 empty] | 1.0284 | 1.0534 | 1.0357 | 1.1541
115 | LFB |  | 06/09/2023 01:12:47 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.469 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10411.443 | 10176.3145 | 8.551 | 10.5198 | 11.3376 | 11254.3572 | 10032.8819 | 10944.4045 | 10.4819 | 10.5276 | 10.5895 | 10.4133 | 11.249 | 10.4677 | 10.8796 | 10.6369 | 10.3729 | 10.3611 | 10.4729 | 9.8152 | 9.7479 | 9.3534 | 10.5721 | 9.653 | 10.4995 | 10.634 | 10.5591 | 10.5016 | 10.8548 | 10.8869 | 11.0634 | [7 empty] | 10.1906 | 10.2388 | 9.4088 | 10.2993
116 | CCV 10 ppb MAJORS |  | 06/09/2023 01:18:35 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.470 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10506.043 | 10123.676 | 8.4551 | 9.7246 | 8.9667 | 10850.9075 | 10202.4914 | 10915.2735 | 10.9189 | 10.5881 | 10.8122 | 10.7201 | 11.4696 | 10.6245 | 11.0236 | 10.9675 | 10.5965 | 10.5167 | 10.5185 | 10.0309 | 10.0877 | 9.4417 | 10.7086 | 10.1537 | 10.663 | 10.3829 | 10.4867 | 10.428 | 10.6185 | 11.2173 | 11.3052 | [7 empty] | 10.255 | 10.2996 | 9.7402 | 10.5455
117 | CCV 10 ppb MAJORS |  | 06/09/2023 01:24:53 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.471 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10686.6783 | 10343.3676 | 8.558 | 9.103 | 9.3316 | 11101.8856 | 10156.0171 | 10668.8401 | 10.7468 | 10.7033 | 10.9293 | 11.0091 | 11.5584 | 10.3537 | 11.069 | 11.1126 | 10.583 | 10.6431 | 10.6241 | 10.157 | 9.1867 | 9.286 | 10.754 | 10.1504 | 10.8812 | 10.5997 | 10.8801 | 10.7601 | 10.9323 | 11.2961 | 11.5315 | [7 empty] | 10.3244 | 10.4178 | 9.784 | 10.4368
118 | 8_16_21_W_6_DUP |  | 06/09/2023 01:30:41 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_W_6_DUP.472 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1869.1166 | 1291.9295 | 8.5166 | 10.4253 | 1.1172 | 3971.9093 | 358.78 | 7432.3982 | 0.17342 | 0.070123 | 0.069617 | 11.6022 | 4.945 | 11.7512 | 0.012874 | 0.055684 | 1.1641 | 1.1669 | 0.64832 | 1.232 | 1.1248 | 1.3465 | 1.3334 | 0.31151 | -0.017154 | 0.026298 | 0.023408 | 0.064823 | 0.022271 | 2.0557 | 2.1198 | [7 empty] | 0.0047591 | 0.004948 | 0.088961 | 0.092624
119 | 8_16_21_W_8 |  | 06/09/2023 01:36:29 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_W_8.473 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1878.916 | 1266.6595 | 8.7202 | 9.4988 | 3.2324 | 3119.5542 | 382.25 | 6749.4131 | 0.21912 | 0.019755 | 0.027086 | 11.2131 | 6.5776 | 10.6969 | 0.012795 | 0.047955 | 1.1116 | 1.1096 | 0.57888 | 1.2847 | 0.85359 | 1.3289 | 1.344 | 0.31885 | 0.025783 | 0.034648 | 0.023985 | 0.16236 | 0.018613 | 1.6847 | 1.7073 | [7 empty] | 0.0051431 | 0.004435 | 0.079979 | 0.082931
120 | 8_16_21_W_9 |  | 06/09/2023 01:42:18 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_W_9.474 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1667.4427 | 1295.857 | 9.4626 | 10.3281 | 2.3028 | 3325.6333 | 347.2434 | 6734.4457 | 0.20531 | 0.038542 | 0.041481 | 13.6975 | 5.3736 | 11.8755 | 0.010681 | 0.052368 | 0.90005 | 1.0718 | 0.47365 | 1.0637 | 0.9429 | 1.0893 | 1.0749 | 0.27703 | 0.020186 | 0.019257 | 0.021395 | 0.12087 | 0.018463 | 1.348 | 1.1603 | [7 empty] | 0.0054825 | 0.0032199 | 0.067566 | 0.071613
121 | 8_16_21_W_9_DUP |  | 06/09/2023 01:48:06 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_W_9_DUP.475 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1555.2133 | 1211.0954 | 10.4555 | 12.025 | 2.1509 | 2988.3617 | 332.3602 | 6131.1853 | 0.19029 | 0.031869 | 0.026556 | 12.8137 | 5.8979 | 12.7605 | 0.014257 | 0.053359 | 0.92282 | 0.92166 | 0.42129 | 1.0469 | 0.82014 | 1.1092 | 1.0602 | 0.2549 | -0.021858 | 0.031856 | 0.019689 | -0.0013693 | 0.026857 | 1.1181 | 1.1258 | [7 empty] | 0.0032262 | 0.0030247 | 0.074924 | 0.078887
122 | 8_16_21_W_10 |  | 06/09/2023 01:53:55 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_W_10.476 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1502.8483 | 1793.4033 | 11.6409 | 13.5315 | 1.9675 | 3485.8097 | 401.2559 | 7672.3893 | 0.19916 | 0.10517 | 0.10773 | 11.8556 | 3.8165 | 12.3045 | 0.0085364 | 0.048524 | 0.68656 | 0.7205 | 0.39574 | 1.0844 | 0.77951 | 1.0742 | 1.0892 | 0.41556 | 0.016727 | 0.012109 | 0.032909 | 0.12215 | 0.022971 | 0.52376 | 0.53315 | [7 empty] | 0.00392 | 0.0021605 | 0.072008 | 0.07607
123 | 8_16_21_W_11 |  | 06/09/2023 01:59:43 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_16_21_W_11.477 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1432.1919 | 1814.1799 | 5.5915 | 6.5667 | 1.0573 | 3683.2075 | 374.5935 | 7673.3255 | 0.17931 | 0.013845 | 0.016913 | 5.2821 | 1.8849 | 5.3094 | 0.0055511 | 0.039264 | 0.58629 | 0.59229 | 0.22478 | 1.0731 | 0.75443 | 1.0936 | 1.0763 | 0.28319 | 0.0024025 | 0.015336 | 0.028922 | 0.14949 | 0.027272 | 0.45219 | 0.45736 | [7 empty] | 0.0019914 | 0.0012061 | 0.041342 | 0.04724
124 | 5_10_23_DB |  | 06/09/2023 02:05:32 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_10_23_DB.478 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | -13.8492 | -42.7414 | -0.35374 | -0.37978 | 2.6292 | 297.698 | -0.52529 | -30.7443 | 0.0028168 | 0.0063856 | 0.0046394 | -0.41847 | -0.024294 | -0.50829 | 0.0018825 | 0.0063744 | 0.013461 | 0.017011 | -0.17112 | 0.021821 | 0.00015562 | 0.15649 | -0.0035484 | 0.22942 | -0.017172 | -0.00010356 | -0.00070609 | -0.17498 | 0.0064924 | -0.012048 | -0.013071 | [7 empty] | -2.3879e-05 | -0.00081166 | -0.00041208 | -0.00023984
125 | 5_10_23_MFB |  | 06/09/2023 02:11:19 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_10_23_MFB.479 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4939.388 | 4907.1418 | 4.2675 | 4.8362 | 9.5109 | 5401.6322 | 4689.4509 | 5154.9066 | 5.0984 | 5.1081 | 5.4579 | 5.5878 | 5.667 | 5.0288 | 5.2231 | 5.1693 | 5.0642 | 5.3721 | 5.2747 | 4.8911 | 4.3997 | 4.7282 | 5.214 | 5.1472 | 4.6424 | 5.1383 | 5.206 | 5.0991 | 5.2303 | 5.5275 | 5.4559 | [7 empty] | 5.2499 | 5.3153 | 4.9975 | 5.3953
126 | 5_10_23_T-239 |  | 06/09/2023 02:17:07 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_10_23_T-239.480 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 48745.8116 | 12605.4585 | 123.6386 | 110.6863 | 12.9154 | 1174.3707 | 3746.0799 | 54176.4977 | 2.0217 | 1.4535 | 1.4889 | 44.152 | 4.0816 | 45.9021 | 1.5274 | 1.8146 | 0.75729 | 0.8243 | 6.0527 | 2.6704 | 2.0411 | 2.2712 | 2.5167 | 5.3918 | 2.1303 | 1.908 | 2.1157 | 2.5691 | 2.2025 | 1.1549 | 1.1861 | [7 empty] | 3.4327 | 3.4631 | 1.3534 | 1.2989
127 | 5_10_23_DB |  | 06/09/2023 02:22:55 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_10_23_DB.481 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | -10.8239 | -41.6648 | -0.040188 | -0.11831 | 3.5216 | 477.3698 | 7.596 | -25.9111 | 0.0052826 | 0.017623 | 0.017607 | -0.26815 | -0.023836 | -0.27262 | 0.0016789 | -0.0028178 | 0.026063 | 0.030542 | -0.17011 | 0.02368 | -0.0057419 | 0.15837 | -0.0014099 | 0.24998 | -0.017364 | -0.0015974 | -0.0007842 | -0.059112 | 0.0062418 | -0.0099051 | -0.012984 | [7 empty] | 0.00015425 | -0.0008318 | 0.00018792 | 0.0012048
128 | LBLANK |  | 06/09/2023 02:28:44 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.482 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | -12.9698 | -42.7414 | -0.70369 | -0.80379 | -1.1278 | 455.5397 | 3.1372 | -25.9281 | 0.0034437 | 0.0050085 | 0.002168 | -0.51708 | -0.025263 | -0.53706 | 0.0014541 | -0.00027793 | -0.0085932 | -0.0077548 | -0.1653 | 0.034399 | -0.0057419 | 0.15719 | -0.0027479 | 0.27724 | -0.012592 | 0.0013548 | 0.0027863 | -0.037514 | -0.0014165 | -0.010916 | -0.012314 | [7 empty] | -2.3879e-05 | -0.00076703 | -0.0021525 | -0.0019502
129 | CCV 1ppb TRACE |  | 06/09/2023 02:34:32 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.483 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1052.5906 | 989.8255 | 1.0977 | 1.4093 | -0.93359 | 1502.2892 | 910.1179 | 939.8155 | 1.1138 | 1.1145 | 1.0856 | 0.73225 | 1.097 | 0.61284 | 1.0944 | 1.1226 | 1.0962 | 1.1169 | 0.9209 | 1.0637 | 0.9599 | 1.0416 | 1.0683 | 1.3095 | 0.90619 | 1.03 | 1.0523 | 1.1701 | 1.1188 | 1.109 | 1.0846 | [7 empty] | 1.0414 | 1.0653 | 1.0831 | 1.1025
130 | CCV 1ppb TRACE |  | 06/09/2023 02:40:49 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.484 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1041.0015 | 957.3049 | 0.98491 | 1.1163 | -1.0897 | 1367.7174 | 970.1029 | 995.2081 | 1.105 | 1.0515 | 1.0771 | 0.74848 | 1.1382 | 0.66903 | 1.1051 | 1.1126 | 1.108 | 1.1062 | 0.93832 | 1.0186 | 0.81931 | 1.0692 | 1.086 | 1.3417 | 1.1282 | 1.0149 | 1.1185 | 1.0496 | 1.1053 | 1.1056 | 1.1531 | [7 empty] | 1.0474 | 1.0586 | 1.0206 | 1.0894
131 | CCV 10 ppb MAJORS |  | 06/09/2023 02:46:38 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.485 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10356.5737 | 9983.1131 | 8.0889 | 9.4782 | 8.916 | 10990.5441 | 9940.5629 | 10699.272 | 10.8473 | 10.5968 | 10.6712 | 10.5622 | 11.5428 | 10.2954 | 10.9612 | 11.0111 | 10.5235 | 10.343 | 10.5023 | 9.9541 | 9.7851 | 9.3474 | 10.796 | 10.2314 | 10.7548 | 10.4601 | 10.8769 | 10.3907 | 10.8284 | 11.1769 | 11.5011 | [7 empty] | 10.1741 | 10.3961 | 9.7876 | 10.6141
132 | CCV 10 ppb MAJORS |  | 06/09/2023 02:52:55 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.486 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10271.5016 | 9936.3669 | 8.4039 | 9.3585 | 9.4053 | 10912.3798 | 9802.8815 | 11043.2201 | 10.6373 | 10.476 | 10.8694 | 10.6794 | 11.2887 | 10.4824 | 10.9061 | 10.7761 | 10.5068 | 10.5846 | 10.5224 | 10.1115 | 10.0138 | 9.4949 | 10.6146 | 10.4333 | 10.5058 | 10.1291 | 10.3115 | 10.9188 | 10.3008 | 11.1412 | 11.3613 | [7 empty] | 10.2073 | 10.2703 | 9.8385 | 10.4226
133 | MT 34 D |  | 06/09/2023 02:58:43 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\MT 34 D.487 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | -10.8277 | -42.7414 | 0.23901 | 0.43595 | 4.8407 | 543.6492 | 1.6478 | -14.5383 | 0.0058732 | 0.01294 | 0.014673 | -0.29073 | -0.022874 | -0.24256 | 0.0031352 | 0.033872 | 0.0080021 | 0.010263 | -0.10061 | 0.013508 | -0.0057419 | 0.15859 | -0.00097193 | 0.29121 | -0.021858 | 0.001239 | -0.0015406 | 0.011605 | -0.0014165 | 0.014307 | 0.01164 | [7 empty] | 6.4543e-05 | 0.00021047 | 0.025358 | 0.026957
134 | MT 35 D |  | 06/09/2023 03:04:31 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\MT 35 D.488 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | -13.012 | -42.7414 | 0.43198 | 0.41328 | 5.7319 | 258.3589 | -1.995 | -21.5825 | 0.0034198 | 0.008369 | 0.0095954 | -0.40778 | -0.02306 | -0.45272 | 0.0016593 | 0.015904 | 0.0074941 | 0.010364 | -0.12058 | 0.00090049 | 7.4596e-05 | 0.1579 | -0.0018912 | 0.1822 | -0.017295 | -0.0015974 | -0.00078806 | -0.056265 | 0.014027 | -0.0075421 | -0.012234 | [7 empty] | 6.7595e-05 | 5.5457e-05 | 0.026064 | 0.027221
135 | MT 36 D |  | 06/09/2023 03:10:18 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\MT 36 D.489 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | -11.711 | -42.7414 | 0.19463 | -0.058307 | 5.7752 | 520.1581 | 4.5843 | -21.1992 | 0.0028168 | 0.019254 | 0.015047 | -0.41859 | -0.022608 | -0.46609 | 0.0014541 | 0.0049225 | 0.0097397 | 0.012298 | -0.097188 | 0.035949 | -0.0057419 | 0.15833 | -0.0015026 | 0.13825 | -0.017301 | -0.00015644 | 0.00058315 | 0.12591 | 0.0024119 | -0.0071885 | -0.010263 | [7 empty] | 0.00033736 | -0.00028153 | 0.024544 | 0.026796
136 | MT 37 D |  | 06/09/2023 03:16:06 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\MT 37 D.490 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | -9.1713 | -41.6987 | 0.37209 | 0.25843 | 6.115 | 313.3414 | -0.073759 | -21.1912 | 0.0028168 | 0.0095054 | 0.0089949 | -0.4562 | -0.023182 | -0.4998 | 0.00231 | 0.0056653 | 0.0052343 | 0.008013 | -0.10296 | 0.038575 | -0.0057419 | 0.15712 | -0.0027621 | 0.22358 | -0.021858 | 0.0013298 | 0.0020793 | -0.050644 | -0.0014165 | -0.0078164 | -0.011081 | [7 empty] | 0.00033604 | -0.0003208 | 0.02506 | 0.027014
137 | MT 38 D |  | 06/09/2023 03:21:55 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\MT 38 D.491 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.9695 | -41.7114 | 0.73041 | 1.201 | 6.8098 | 269.772 | 11.3921 | -16.0414 | 0.0040534 | 0.012221 | 0.013329 | 3.6051 | -0.00028103 | 3.844 | 0.0029808 | 0.01975 | 0.094094 | 0.10123 | 0.28238 | 0.025252 | 0.012494 | 0.15797 | -0.0018506 | 0.24786 | -0.021858 | 0.007244 | 0.0006511 | 0.06173 | -0.0014165 | -0.0078569 | -0.0099949 | [7 empty] | 0.00034106 | -0.0002362 | 0.025678 | 0.027844
138 | MT 39 D |  | 06/09/2023 03:27:43 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\MT 39 D.492 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | -11.7463 | -42.7414 | 0.27651 | 0.26248 | 6.2095 | 336.2363 | 2.4481 | -23.7593 | 0.0052261 | 0.014992 | 0.01479 | 0.64182 | -0.017731 | 0.36372 | 0.0016667 | 0.003218 | 0.015908 | 0.018678 | -0.10496 | 0.0025863 | -0.0057419 | 0.15708 | -0.0028449 | 0.29666 | -0.021858 | 0.0013109 | -0.0021686 | -0.035541 | -0.0014165 | 0.0034843 | 0.0010705 | [7 empty] | 0.00077705 | 0.00028921 | 0.025475 | 0.027688
139 | MT 40 D |  | 06/09/2023 03:33:31 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\MT 40 D.493 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | -12.5533 | -42.7414 | 0.32587 | 0.37595 | 6.6491 | 451.7767 | 0.070073 | -23.5857 | 0.0028168 | 0.008623 | 0.0093602 | -0.40099 | -0.021118 | -0.41493 | 0.0033431 | 0.0086097 | 0.012825 | 0.013266 | -0.10572 | -0.0054084 | -0.0057419 | 0.15751 | -0.002378 | 0.20651 | -0.021858 | -0.00012784 | -6.3475e-05 | -0.032903 | -0.0014165 | -0.0096581 | -0.010228 | [7 empty] | 0.00052045 | 2.7819e-05 | 0.027373 | 0.029017
140 | MT 32 D |  | 06/09/2023 03:39:20 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\MT 32 D.494 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | -11.6123 | -42.7414 | 0.66214 | 0.84125 | 5.9762 | 406.3225 | 0.6661 | -18.4341 | 0.004075 | 0.011567 | 0.010195 | -0.3752 | -0.016008 | -0.3927 | 0.0014541 | 0.0049401 | 0.044793 | 0.043742 | -0.097942 | 0.04521 | 0.00032675 | 0.15862 | -0.0010317 | 0.15393 | -0.012499 | 0.0013672 | 1.1729e-05 | -0.036469 | -0.0014165 | -0.0096691 | -0.010303 | [7 empty] | 0.00015558 | -0.00013145 | 0.036823 | 0.040688
141 | MT 32 D LD |  | 06/09/2023 03:45:07 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\MT 32 D LD.495 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | -11.6557 | -42.7414 | 0.29154 | 0.27062 | 5.0144 | 402.165 | 5.8515 | -25.8977 | 0.0028168 | 0.016638 | 0.013089 | -0.42477 | -0.022634 | -0.47981 | 0.0016722 | 0.0052687 | 0.033957 | 0.033642 | -0.075663 | -0.0030021 | -0.0057419 | 0.1577 | -0.0021064 | 0.21335 | -0.017034 | -0.00011537 | -0.00070088 | -0.039275 | -0.0014165 | -0.0090254 | -0.0099686 | [7 empty] | -2.3879e-05 | 0.00027349 | 0.035452 | 0.040847
142 | MT 32 D LD |  | 06/09/2023 03:51:24 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\MT 32 D LD.496 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | -12.0832 | -41.6832 | 0.26833 | 0.36646 | 6.3501 | 571.4285 | 5.9722 | -25.8863 | 0.0034447 | 0.014069 | 0.012714 | -0.43123 | -0.022927 | -0.49223 | 0.0014541 | 0.0061521 | 0.030312 | 0.033012 | -0.075527 | 0.03637 | -0.0057419 | 0.15795 | -0.0018326 | 0.15998 | -0.01241 | -0.0001179 | 0.0016397 | 0.0083569 | 0.014351 | -0.0092743 | -0.010083 | [7 empty] | 0.00053082 | -0.00011403 | 0.035991 | 0.039001
143 | LBLANK |  | 06/09/2023 03:57:12 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.497 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | -12.5767 | -42.7414 | -0.70634 | -0.77097 | -1.131 | 313.9096 | 7.0578 | -28.3955 | 0.0034244 | 0.002147 | 0.0027951 | -0.51161 | -0.025165 | -0.56836 | 0.0014541 | 0.0024533 | -0.0027638 | -0.0070497 | -0.16616 | 0.051529 | -0.0057419 | 0.15721 | -0.0027403 | 0.18241 | -0.012245 | 0.0028582 | -0.00033899 | -0.03785 | -0.0014165 | -0.011642 | -0.013073 | [7 empty] | 0.0002271 | -0.00080048 | -0.0021612 | -0.0019083
144 | CCV 1ppb TRACE |  | 06/09/2023 04:03:01 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.498 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1036.5984 | 966.2173 | 0.95107 | 0.86853 | -0.84618 | 1444.5943 | 937.8407 | 992.7871 | 1.112 | 1.0465 | 1.0663 | 0.73551 | 1.1112 | 0.63826 | 1.0987 | 1.1617 | 1.0716 | 1.1164 | 0.92656 | 1.0232 | 0.89782 | 1.0393 | 1.0705 | 1.2681 | 1.2022 | 1.049 | 1.0599 | 1.0012 | 1.1514 | 1.1211 | 1.1208 | [7 empty] | 1.0339 | 1.0692 | 1.0381 | 1.1159
145 | CCV 1ppb TRACE |  | 06/09/2023 04:09:18 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.499 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 952.4867 | 920.3942 | 0.89949 | 1.3961 | -1.0523 | 1169.3073 | 894.8116 | 935.8916 | 1.0645 | 1.0183 | 1.0814 | 0.76452 | 1.116 | 0.58453 | 1.0508 | 1.0572 | 1.0662 | 1.1021 | 0.91888 | 0.99898 | 0.84647 | 1.0347 | 1.0648 | 1.1647 | 1.1172 | 1.0883 | 1.105 | 1.0828 | 0.94063 | 1.0975 | 1.14 | [7 empty] | 1.0635 | 1.0543 | 0.99776 | 1.1161
146 | LFB |  | 06/09/2023 04:15:06 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.500 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9655.3023 | 9422.9929 | 7.968 | 9.4554 | 11.2017 | 10083.7246 | 9111.2353 | 10266.8773 | 9.9671 | 10.052 | 10.5737 | 10.432 | 11.2055 | 9.8508 | 10.3153 | 10.4705 | 10.2236 | 10.2428 | 10.4287 | 9.8336 | 9.127 | 9.4018 | 10.7556 | 9.6608 | 10.8031 | 10.6016 | 10.9115 | 10.3717 | 10.5244 | 10.869 | 11.2639 | [7 empty] | 10.1629 | 10.2121 | 9.5371 | 10.3114
147 | LFB |  | 06/09/2023 04:21:22 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.501 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9656.6532 | 9273.0567 | 8.353 | 9.8519 | 11.0636 | 10003.9804 | 8980.1949 | 10216.0904 | 10.1707 | 9.9523 | 10.5156 | 10.364 | 11.1921 | 9.9731 | 10.4615 | 10.4314 | 10.0178 | 10.2409 | 10.5259 | 10.0852 | 9.556 | 9.1411 | 10.6072 | 10.1785 | 10.7257 | 10.5576 | 10.7993 | 10.7729 | 10.8282 | 10.8881 | 11.1021 | [7 empty] | 10.101 | 10.267 | 9.6012 | 10.2737
148 | CCV 10 ppb MAJORS |  | 06/09/2023 04:27:10 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.502 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9994.5089 | 9616.8679 | 7.8722 | 9.2965 | 8.7872 | 10492.0982 | 9483.7382 | 10467.4739 | 10.5612 | 10.6635 | 11.1277 | 11.1114 | 11.464 | 10.1943 | 10.712 | 10.8583 | 10.2157 | 10.7439 | 10.69 | 9.8211 | 9.5118 | 9.2813 | 10.6774 | 10.0933 | 10.4667 | 10.6094 | 10.6429 | 10.0384 | 10.721 | 11.5107 | 11.4835 | [7 empty] | 10.3355 | 10.4697 | 9.6241 | 10.5172
149 | CCV 10 ppb MAJORS |  | 06/09/2023 04:33:27 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.503 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10179.8295 | 9996.6224 | 8.1827 | 9.6363 | 8.8462 | 10714.2111 | 9654.1533 | 10923.3615 | 10.9967 | 10.6108 | 10.1528 | 9.9857 | 11.4208 | 10.6172 | 10.9366 | 10.8637 | 10.4136 | 9.9154 | 9.8559 | 10.0347 | 9.398 | 9.3666 | 10.6098 | 10.4625 | 10.2342 | 10.3869 | 10.5024 | 10.3776 | 10.6661 | 10.6211 | 11.3745 | [7 empty] | 9.5247 | 9.7712 | 9.7723 | 10.3593
150 | MT 33 D |  | 06/09/2023 04:39:15 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\MT 33 D.504 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | -9.4496 | -42.7414 | 0.31682 | 0.43273 | 4.979 | 382.9278 | 6.6417 | -28.3349 | 0.0046516 | 0.010729 | 0.0099189 | -0.42763 | -0.021513 | -0.45584 | 0.0023092 | 0.0052866 | 0.039574 | 0.036291 | -0.12176 | 0.033194 | -0.0057419 | 0.15945 | -5.0384e-05 | 0.44776 | -0.021858 | -0.00014835 | -0.00074856 | 0.055243 | 0.0063868 | 0.0074551 | -0.0085949 | [7 empty] | 0.00078518 | 0.00036522 | 0.029453 | 0.028925
151 | MT 33 D LSP |  | 06/09/2023 04:45:03 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\MT 33 D LSP.505 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 26044.1296 | 24339.7693 | 24.2298 | 27.2187 | 31.8764 | 25839.3101 | 25094.8771 | 27354.045 | 27.422 | 26.9878 | 28.0209 | 28.7198 | 29.5234 | 26.7482 | 27.2455 | 26.3297 | 25.5064 | 26.0889 | 25.668 | 25.809 | 21.7992 | 22.7301 | 26.1103 | 24.8809 | 25.3539 | 25.4079 | 25.9678 | 26.7067 | 25.5047 | 28.7165 | 29.4031 | [7 empty] | 24.8957 | 25.0185 | 25.0866 | 25.2813
152 | cal blank |  | 06/09/2023 04:50:51 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\cal blank.506 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | -11.63 | -42.7414 | -0.72278 | -0.82141 | -1.4135 | 390.0964 | -0.44732 | -30.7443 | 0.0040603 | -0.00030561 | 0.00015366 | -0.52038 | -0.025176 | -0.51838 | 0.0018945 | 0.006169 | -0.0115 | -0.0078486 | -0.17168 | 0.00068128 | 0.00047342 | 0.16085 | 0.0014075 | 0.31405 | -0.012608 | 0.0085479 | -0.0021793 | 0.020526 | 0.0063098 | 0.025833 | -0.0089063 | [7 empty] | 0.00043548 | -0.00062049 | -0.0018182 | -0.001494
153 | STD 0.01 |  | 06/09/2023 04:56:40 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 0.01.507 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | -0.88572 | -24.6378 | -0.21189 | -0.13153 | -0.85617 | 281.5926 | 9.01 | -18.3835 | 0.011137 | 0.018818 | 0.011412 | -0.36138 | 0.016297 | -0.36335 | 0.015899 | 0.010412 | 0.012683 | 0.01124 | -0.018936 | 0.028596 | -0.0057419 | 0.16948 | 0.01196 | 0.37887 | -0.012506 | 0.011739 | 0.0071279 | 0.028066 | 0.01034 | 0.032119 | 0.014964 | [7 empty] | 0.0083913 | 0.0090535 | 0.0094101 | 0.010552
154 | std 0.05 |  | 06/09/2023 05:02:29 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\std 0.05.508 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 42.3443 | 25.1399 | -0.15424 | 0.041214 | -0.14317 | 318.4276 | 55.8965 | 45.6529 | 0.058316 | 0.064247 | 0.058441 | -0.33816 | 0.033607 | -0.40509 | 0.066578 | 0.06618 | 0.062487 | 0.05813 | -0.056326 | 0.092106 | 0.061691 | 0.20528 | 0.052789 | 0.29763 | 0.029765 | 0.062492 | 0.065505 | 0.084585 | 0.050272 | 0.055842 | 0.044431 | [7 empty] | 0.055323 | 0.051647 | 0.049436 | 0.056223
155 | std 0.1 |  | 06/09/2023 05:08:17 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\std 0.1.509 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 90.5481 | 54.4037 | -0.44397 | -0.56456 | -1.1119 | 534.6699 | 91.0056 | 86.9093 | 0.11832 | 0.10389 | 0.10585 | -0.16717 | 0.093735 | -0.15083 | 0.11559 | 0.11577 | 0.12544 | 0.11139 | 0.031115 | 0.13689 | 0.10802 | 0.24482 | 0.10286 | 0.48144 | 0.10407 | 0.097214 | 0.11075 | 0.099131 | 0.10429 | 0.10617 | 0.10104 | [7 empty] | 0.10189 | 0.10235 | 0.099535 | 0.1057
156 | std 0.5 |  | 06/09/2023 05:14:05 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\std 0.5.510 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 505.7336 | 488.3675 | 0.42288 | 0.41398 | 0.27733 | 980.7375 | 503.4289 | 646.0592 | 0.52765 | 0.56339 | 0.54525 | 0.63737 | 0.55691 | 0.62369 | 0.57313 | 0.61677 | 0.60285 | 0.57109 | 0.5516 | 0.52234 | 0.42378 | 0.61217 | 0.53526 | 0.73029 | 0.5429 | 0.54155 | 0.52642 | 0.5162 | 0.52131 | 0.55013 | 0.55823 | [7 empty] | 0.51651 | 0.51024 | 0.51484 | 0.5379
157 | std 1 |  | 06/09/2023 05:19:54 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\std 1.511 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 993.727 | 948.1153 | 1.0216 | 1.1406 | 0.02795 | 1216.2986 | 921.8704 | 1088.4735 | 1.1009 | 1.0597 | 1.1338 | 1.1861 | 1.1349 | 1.0963 | 1.1082 | 1.1771 | 1.0809 | 1.1469 | 1.0834 | 0.9672 | 0.95137 | 1.0499 | 1.0523 | 1.1403 | 1.0078 | 1.0158 | 1.1041 | 0.92928 | 1.038 | 1.1582 | 1.1215 | [7 empty] | 1.059 | 1.0748 | 1.0071 | 1.0193
158 | std 5 |  | 06/09/2023 05:25:42 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\std 5.512 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4918.4856 | 4819.3986 | 4.1051 | 4.7761 | 5.8422 | 5449.0828 | 4829.0861 | 5105.3724 | 5.2605 | 5.2326 | 5.7527 | 5.6957 | 5.6315 | 5.0438 | 5.3103 | 5.4015 | 5.1114 | 5.6779 | 5.7051 | 4.7948 | 4.5134 | 4.7657 | 5.2924 | 5.0166 | 5.1047 | 5.2234 | 5.2969 | 5.0031 | 5.2336 | 5.9163 | 5.5273 | [7 empty] | 5.4796 | 5.5619 | 4.7133 | 5.2171
159 | std 10 |  | 06/09/2023 05:31:30 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\std 10.513 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10242.8307 | 9877.3074 | 8.4331 | 9.2918 | 10.2091 | 10556.8158 | 9404.0644 | 10529.487 | 10.6369 | 10.5505 | 10.83 | 10.776 | 11.3204 | 10.5035 | 10.9455 | 10.9796 | 10.4178 | 10.6069 | 10.5295 | 10.0338 | 9.5096 | 9.0234 | 10.4954 | 10.4257 | 10.3561 | 10.0568 | 10.3997 | 10.5136 | 10.327 | 11.2277 | 11.1581 | [7 empty] | 10.1995 | 10.3067 | 9.7309 | 10.2976
160 | LBLANK |  | 06/09/2023 05:37:19 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.514 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | -12.9681 | -41.683 | -0.70315 | -0.76901 | -1.8662 | 442.0872 | -1.2013 | -30.7443 | 0.0034565 | 0.0032836 | 0.0039677 | -0.50335 | -0.025041 | -0.51864 | 0.0025447 | 0.0029418 | -0.0078968 | -0.0070702 | -0.1673 | 0.042346 | 0.0066003 | 0.1583 | -0.0013996 | 0.35019 | -0.021858 | -0.00012043 | -0.0014366 | 0.00057204 | -0.0014165 | 0.0058947 | -0.010263 | [7 empty] | 0.00016021 | -0.00085873 | -0.0020506 | -0.0018641
161 | CCV 1ppb TRACE |  | 06/09/2023 05:43:07 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.515 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1095.1881 | 1052.2148 | 0.98874 | 1.4873 | -1.0254 | 1597.5271 | 1049.0373 | 1097.1047 | 1.2046 | 1.1554 | 1.0953 | 0.76683 | 1.1213 | 0.97591 | 1.2147 | 1.2524 | 1.1915 | 1.1161 | 0.91116 | 1.0394 | 1.2 | 1.0631 | 1.084 | 1.2623 | 0.98008 | 1.101 | 1.0509 | 1.0332 | 1.0577 | 1.1169 | 1.1325 | [7 empty] | 1.0408 | 1.0519 | 0.99979 | 1.0841
162 | CCV 1ppb TRACE |  | 06/09/2023 05:49:25 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.516 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1024.0303 | 931.1209 | 0.98854 | 1.3009 | -1.1378 | 1394.8765 | 938.2511 | 944.6971 | 1.135 | 1.0838 | 1.1303 | 0.79247 | 1.201 | 0.76478 | 1.0956 | 1.0554 | 1.0791 | 1.147 | 0.94372 | 1.0306 | 0.90568 | 1.1293 | 1.1229 | 1.2505 | 1.0986 | 1.116 | 1.1163 | 1.1167 | 1.0692 | 1.1405 | 1.2174 | [7 empty] | 1.0386 | 1.0706 | 1.0181 | 1.095
163 | CCV 10 ppb MAJORS |  | 06/09/2023 05:55:13 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.517 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10077.9948 | 9722.7396 | 8.0811 | 9.0427 | 8.8603 | 10644.8836 | 9633.7547 | 10401.7439 | 10.7421 | 10.5474 | 10.7787 | 10.5641 | 11.435 | 10.3795 | 10.753 | 10.4794 | 10.2912 | 10.486 | 10.3246 | 10.0781 | 9.8555 | 9.3768 | 10.6654 | 10.4149 | 9.7581 | 10.3503 | 10.4824 | 10.4773 | 10.3623 | 11.0465 | 11.3729 | [7 empty] | 10.0787 | 10.2221 | 9.812 | 13.6664
164 | CCV 10 ppb MAJORS |  | 06/09/2023 06:01:30 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.518 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10127.2038 | 9660.7643 | 8.3293 | 9.0625 | 8.9396 | 10355.8633 | 9514.7639 | 10478.8626 | 10.4444 | 10.3642 | 10.8827 | 10.6084 | 11.6028 | 10.0025 | 10.6399 | 10.6506 | 10.14 | 10.5524 | 10.5135 | 10.1763 | 9.8881 | 9.3709 | 10.7637 | 10.3398 | 10.5144 | 10.6341 | 10.6843 | 10.4701 | 10.8723 | 11.1997 | 11.5358 | [7 empty] | 10.1199 | 10.2084 | 9.744 | 10.3248
165 | std 25 |  | 06/09/2023 06:07:19 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\std 25.519 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 25398.1227 | 23629.8633 | 22.0588 | 23.81 | 28.0588 | 26197.1273 | 24586.3166 | 26664.5935 | 26.9583 | 26.2764 | 28.1751 | 28.7602 | 29.2513 | 25.9237 | 27.1549 | 26.2559 | 25.1535 | 26.6469 | 27.3665 | 26.3317 | 23.8956 | 24.1973 | 27.9375 | 25.7403 | 28.4751 | 28.1203 | 27.8609 | 27.8534 | 27.7892 | 29.1201 | 29.5809 | [7 empty] | 25.5387 | 25.9901 | 23.8259 | 25.3839
166 | std 50 |  | 06/09/2023 06:13:06 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\std 50.520 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 51651.8561 | 48467.7539 | 55.9609 | 50.8192 | 60.0335 | 53601.7771 | 50084.3154 | 56040.8476 | 56.0659 | 54.2931 | 58.1352 | 54.7715 | 63.5806 | 53.763 | 54.8729 | 52.9413 | 50.4784 | 52.6223 | 54.5789 | 53.1201 | 48.1276 | 52.5513 | 60.13 | 51.8461 | 59.4216 | 59.8176 | 60.1396 | 55.8621 | 60.35 | 59.4817 | 63.8111 | [7 empty] | 51.0029 | 51.2617 | 46.2789 | 49.1607
167 | std 50 |  | 06/09/2023 06:19:23 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\std 50.521 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 57514.8786 | 54098.408 | 57.5944 | 56.7123 | 62.1221 | 60725.4447 | 55251.255 | 61417.746 | 61.6791 | 59.963 | 58.3365 | 54.7018 | 60.2968 | 59.2017 | 59.5359 | 57.8086 | 55.5635 | 52.7725 | 54.6745 | 53.5853 | 52.5456 | 48.8933 | 56.3462 | 51.9037 | 55.7449 | 56.5226 | 57.2536 | 57.0671 | 57.2643 | 59.1213 | 60.6688 | [7 empty] | 50.2574 | 51.0328 | 46.1231 | 48.9931
168 | std 100 |  | 06/09/2023 06:25:11 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\std 100.522 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 112871.2702 | 104438.8781 | 119.8729 | 108.1934 | 122.6976 | 120006.7696 | 112264.9622 | 122218.2656 | 123.8248 | 118.1572 | 122.1673 | 111.7135 | 125.6833 | 113.8752 | 117.0495 | 111.5168 | 106.2146 | 103.9908 | 108.2579 | 109.2092 | 102.5304 | 99.6412 | 113.5297 | 104.3797 | 117.7598 | 114.707 | 116.6555 | 115.5131 | 116.6196 | 123.3314 | 125.0769 | [7 empty] | 100.9788 | 101.576 | 88.5866 | 95.0166
169 | std 100 |  | 06/09/2023 06:31:28 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\std 100.523 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 114581.4942 | 107924.3964 | 119.3748 | 109.6745 | 123.6732 | 124004.7376 | 113937.6549 | 126412.622 | 128.7455 | 122.1822 | 122.1227 | 113.51 | 124.9597 | 118.1089 | 119.4415 | 112.6024 | 109.9544 | 104.9428 | 107.6217 | 107.7039 | 105.1992 | 97.7652 | 113.9056 | 102.8396 | 116.2066 | 114.668 | 116.4887 | 114.8819 | 116.5623 | 123.9658 | 124.9704 | [7 empty] | 100.0453 | 100.768 | 88.0225 | 107.195
170 | LBLANK |  | 06/09/2023 06:37:16 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.524 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | -8.5472 | -40.5879 | -0.69949 | -0.7321 | -1.7305 | 632.9079 | 8.8247 | -30.7443 | 0.0067175 | 0.0043399 | 0.0038576 | -0.50047 | -0.023891 | -0.43979 | 0.0063974 | 0.014635 | -0.0041995 | -0.005282 | -0.16664 | 0.0017407 | 0.051217 | 0.1841 | 0.029384 | 0.33492 | -0.0076645 | 0.01045 | 0.0087335 | -0.016551 | -0.0014165 | 0.10219 | 0.0082823 | [7 empty] | 0.0014725 | 0.00047471 | -0.0020385 | -0.0012349
171 | PE ICV |  | 06/09/2023 06:43:04 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\PE ICV.525 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10784.0509 | 10806.6822 | 8.6196 | 10.2246 | 21.1985 | 11343.6038 | 10326.1352 | 11332.3659 | 11.0865 | 10.8954 | 11.1494 | 11.1034 | 11.7455 | 10.8905 | 11.1451 | 11.3409 | 10.8236 | 10.5072 | 10.28 | 10.0617 | 8.7484 | 9.1461 | 10.6438 | 7.6279 | 10.4446 | 10.6176 | 10.7366 | 10.6585 | 10.6997 | 10.8202 | 10.8647 | [7 empty] | 10.0225 | 10.1664 | 9.4283 | 10.0802
172 | 5ppb CCV |  | 06/09/2023 06:48:52 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5ppb CCV.526 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 5099.1538 | 5060.1377 | 4.5765 | 5.4187 | 5.2908 | 5757.1128 | 4796.6496 | 5266.0258 | 5.3398 | 5.292 | 5.5083 | 5.529 | 5.7169 | 5.0515 | 5.4146 | 5.3646 | 5.2757 | 5.393 | 5.3424 | 5.3082 | 4.6621 | 4.5705 | 5.3399 | 5.6864 | 5.4002 | 5.4184 | 5.4412 | 5.5335 | 5.1843 | 5.5856 | 5.6512 | [7 empty] | 4.9902 | 5.081 | 4.95 | 5.089
173 | 25ppb CCV |  | 06/09/2023 06:54:40 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\25ppb CCV.527 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 27052.5786 | 25739.4061 | 22.3597 | 26.7038 | 29.7251 | 28273.4832 | 25766.5741 | 28653.4629 | 28.9027 | 28.0587 | 28.8542 | 29.1463 | 29.5342 | 27.3924 | 28.4125 | 27.9263 | 27.1114 | 26.8724 | 27.6385 | 26.9164 | 25.8499 | 24.3267 | 28.1447 | 26.4734 | 27.7118 | 27.7448 | 28.189 | 28.8846 | 28.655 | 29.4376 | 29.4243 | [7 empty] | 25.8446 | 26.068 | 23.658 | 25.2795
174 | 8_30_21_W_1 |  | 06/09/2023 07:00:27 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_W_1.528 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1445.2147 | 1107.9637 | 20.3931 | 22.5768 | 4.6327 | 4780.9248 | 358.5827 | 7056.5057 | 0.080743 | 0.035841 | 0.034346 | 6.7471 | 3.5623 | 6.6475 | 0.011779 | 0.083671 | 0.34181 | 0.33565 | 0.28282 | 0.99561 | 0.90543 | 1.0697 | 1.0591 | 0.40392 | -0.0066747 | 0.014316 | 0.0086444 | 0.022783 | 0.015337 | 3.2238 | 3.3373 | [7 empty] | 0.0022111 | 0.001401 | 0.038595 | 0.041109
175 | LBLANK |  | 06/09/2023 07:06:16 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.529 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | -12.9562 | -41.6802 | -0.69969 | -0.75005 | -1.5046 | 495.6731 | 10.0863 | -30.7443 | 0.0041213 | 0.0038698 | 0.0020448 | -0.50349 | -0.024856 | -0.49308 | 0.0016759 | 0.0099529 | -0.0081047 | -0.0072256 | -0.16291 | -0.026293 | 0.00040846 | 0.16055 | 0.0012783 | 0.41103 | -0.016955 | -8.2658e-06 | 0.003197 | -0.074019 | 0.0028055 | 0.032649 | -0.008785 | [7 empty] | 0.00036211 | -0.0005511 | -0.0018278 | -0.0018018
176 | CCV 1ppb TRACE |  | 06/09/2023 07:12:04 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.530 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1051.1689 | 1020.0044 | 0.97899 | 1.2627 | -0.93036 | 1309.7646 | 998.1394 | 1088.2421 | 1.155 | 1.0781 | 1.1414 | 0.83223 | 1.0851 | 0.744 | 1.0979 | 1.1357 | 1.1135 | 1.1581 | 0.9574 | 0.96994 | 0.82821 | 1.0186 | 1.0262 | 1.2978 | 1.0821 | 0.99643 | 1.0561 | 1.0974 | 0.92919 | 1.1757 | 1.0745 | [7 empty] | 1.0438 | 1.0858 | 0.9751 | 1.0914
177 | CCV 1ppb TRACE |  | 06/09/2023 07:18:22 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.531 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | -12.226 | -42.7414 | -0.73088 | -0.82141 | -2.3965 | -1839.6045 | -6.5656 | -30.7443 | 0.0028168 | 0.00066113 | 0.0012662 | -0.53303 | -0.025781 | -1.2086 | 0.0032262 | -0.0040635 | -0.01051 | -0.0092415 | -0.18035 | 0.5533 | 2.4129e-06 | 0.15579 | -0.0043674 | -1.1983 | -0.017104 | 0.0013599 | -0.0028399 | -0.10368 | -0.0014165 | 0.024278 | -0.013457 | [7 empty] | 0.00032415 | -0.00095384 | -0.0025274 | -0.0022858
178 | LFB |  | 06/09/2023 07:24:08 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.532 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10254.4035 | 9776.8176 | 7.8853 | 9.9854 | 11.6517 | 10403.5171 | 9463.4849 | 10132.481 | 10.6011 | 10.5571 | 10.8327 | 10.9165 | 11.4834 | 10.0498 | 10.6216 | 10.4949 | 10.2395 | 10.4629 | 10.6167 | 9.6694 | 9.2012 | 9.2974 | 10.7981 | 9.6129 | 10.6855 | 10.7273 | 10.9476 | 10.3157 | 10.2766 | 10.9803 | 11.2306 | [7 empty] | 10.0459 | 10.144 | 8.9354 | 10.0646
179 | LFB |  | 06/09/2023 07:30:24 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.533 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10017.5401 | 9875.7031 | 8.3478 | 9.8424 | 11.7042 | 10429.6148 | 9441.887 | 10498.9096 | 10.5842 | 10.3864 | 11.0048 | 10.8546 | 11.5103 | 10.3028 | 10.6797 | 10.5309 | 10.309 | 10.4596 | 10.6408 | 10.151 | 9.3122 | 9.3152 | 10.8641 | 10.416 | 11.3736 | 10.5987 | 10.7255 | 10.4707 | 11.0097 | 11.1163 | 11.3147 | [7 empty] | 9.9849 | 10.0809 | 9.3547 | 9.9893
180 | CCV 10 ppb MAJORS |  | 06/09/2023 07:36:12 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.534 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10782.0523 | 10337.5341 | 7.9763 | 10.0551 | 9.3865 | 11819.8229 | 10201.5805 | 11595.8589 | 11.5809 | 11.3385 | 11.0865 | 10.9656 | 11.8503 | 11.1059 | 11.603 | 11.524 | 11.058 | 10.6293 | 10.6038 | 9.8505 | 10.1951 | 9.3939 | 10.8976 | 10.0099 | 10.7302 | 10.6745 | 10.6631 | 10.3899 | 11.0971 | 11.2673 | 11.6711 | [7 empty] | 10.2221 | 10.3652 | 9.41 | 10.2653
181 | CCV 10 ppb MAJORS |  | 06/09/2023 07:42:30 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.535 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10829.0372 | 10386.1016 | 8.1288 | 10.1892 | 9.1513 | 11204.1152 | 10091.0181 | 10646.2779 | 11.0517 | 10.9682 | 11.1525 | 11.1076 | 11.7863 | 10.9493 | 11.3019 | 11.151 | 10.829 | 10.7566 | 10.6965 | 10.0111 | 9.5612 | 9.4705 | 10.8046 | 10.4323 | 10.2157 | 10.3569 | 10.676 | 10.3559 | 10.4126 | 11.4757 | 11.5754 | [7 empty] | 10.309 | 10.3199 | 9.5899 | 11.0177
182 | 8_30_21_W_6 |  | 06/09/2023 07:48:18 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_W_6.536 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1713.3194 | 1384.8936 | 8.5795 | 10.302 | 1.9582 | 4070.5121 | 373.4556 | 7446.5251 | 0.16625 | 0.033713 | 0.034457 | 9.0289 | 3.6429 | 9.0657 | 0.0099182 | 0.10284 | 0.96963 | 0.94765 | 0.55169 | 0.96941 | 0.72783 | 1.0821 | 1.0943 | 0.38403 | 0.024257 | 0.029292 | 0.017423 | 0.023127 | 0.054745 | 1.8237 | 1.8439 | [7 empty] | 0.0042803 | 0.0038234 | 0.076649 | 0.088074
183 | 8_30_21_W_6_LD |  | 06/09/2023 07:54:05 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_W_6_LD.537 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1629.9337 | 1289.1539 | 11.4479 | 13.1858 | 2.9846 | 3923.8299 | 360.2321 | 7115.009 | 0.15889 | 0.024461 | 0.026272 | 8.9076 | 4.0091 | 8.7079 | 0.0090953 | 0.11234 | 0.91832 | 0.93419 | 0.54928 | 0.98709 | 0.6762 | 1.1936 | 1.1911 | 0.46986 | 0.00087091 | 0.032351 | 0.022948 | 0.19794 | 0.026533 | 1.7945 | 2.0344 | [7 empty] | 0.0074877 | 0.005747 | 0.081385 | 0.083486
184 | 8_30_21_W_6_LD |  | 06/09/2023 08:00:20 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_W_6_LD.538 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1730.1363 | 1408.2995 | 11.0373 | 13.895 | 3.3329 | 4336.153 | 392.4985 | 7976.5083 | 0.17034 | 0.027728 | 0.025805 | 8.9557 | 3.6647 | 9.6923 | 0.0098082 | 0.1094 | 1.008 | 0.9406 | 0.56106 | 0.97469 | 0.84022 | 1.0791 | 1.0993 | 0.36121 | 0.0096757 | 0.025031 | 0.022728 | 0.10577 | 0.038369 | 1.8232 | 1.8528 | [7 empty] | 0.0056183 | 0.0039239 | 0.083993 | 0.10159
185 | 8_30_21_W_6_DUP |  | 06/09/2023 08:06:08 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_W_6_DUP.539 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1644.1827 | 1291.7634 | 10.7325 | 12.8414 | 3.3864 | 3990.0734 | 362.3843 | 7128.1389 | 0.13476 | 0.02306 | 0.025681 | 9.7028 | 3.694 | 9.6211 | 0.0093403 | 0.096783 | 1.0186 | 1.0411 | 0.56784 | 0.95672 | 0.74399 | 1.0558 | 1.0624 | 0.33356 | -0.006285 | 0.010063 | 0.023626 | 0.13278 | 0.016409 | 1.821 | 1.8633 | [7 empty] | 0.0044532 | 0.0041624 | 0.095798 | 0.093852
186 | 8_30_21_W_6 |  | 06/09/2023 08:11:56 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_W_6.540 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1651.8639 | 1367.7319 | 8.1547 | 10.0046 | 2.8575 | 4201.5693 | 369.6612 | 7649.4613 | 0.14564 | 0.030538 | 0.03397 | 8.9265 | 3.6284 | 9.1088 | 0.0085592 | 0.11404 | 0.9474 | 0.92896 | 0.54314 | 0.93433 | 1.0003 | 1.0903 | 1.0954 | 0.26328 | -0.01652 | 0.01178 | 0.015682 | 0.034683 | 0.025238 | 1.8272 | 1.8592 | [7 empty] | 0.0054727 | 0.0049851 | 0.07855 | 0.08885
187 | 8_30_21_W_6_DUP_LSP |  | 06/09/2023 08:17:43 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_W_6_DUP_LSP.541 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 24410.2748 | 22811.342 | 29.2755 | 36.2223 | 26.8324 | 28535.2011 | 21352.6907 | 32565.2048 | 23.717 | 22.9373 | 22.4266 | 33.6525 | 27.1012 | 34.6269 | 23.114 | 23.4389 | 22.9904 | 21.925 | 21.8919 | 21.955 | 21.0684 | 19.1227 | 22.7403 | 20.8789 | 22.4305 | 21.4571 | 21.5996 | 21.79 | 22.1116 | 24.9539 | 25.2518 | [7 empty] | 20.5686 | 20.6769 | 19.5534 | 20.5707
188 | 8_30_21_W_6_DUP_LSP |  | 06/09/2023 08:24:01 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_W_6_DUP_LSP.542 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 23297.3304 | 22080.5706 | 29.1137 | 34.8303 | 27.1284 | 27411.4871 | 20719.8489 | 31220.0385 | 23.0276 | 22.4004 | 22.4249 | 33.4705 | 27.5932 | 33.3037 | 22.5586 | 22.6137 | 22.2724 | 22.1001 | 22.1271 | 21.7824 | 21.1742 | 19.9588 | 23.1001 | 20.637 | 21.8095 | 21.8018 | 22.021 | 21.6341 | 22.2679 | 24.9207 | 25.629 | [7 empty] | 20.3124 | 20.5536 | 18.8463 | 20.3084
189 | 8_30_21_W_8 |  | 06/09/2023 08:29:49 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_W_8.543 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1906.6865 | 1437.7037 | 8.7656 | 11.4516 | 5.2063 | 4353.9732 | 391.7948 | 7581.3015 | 0.17632 | 0.047681 | 0.039718 | 7.9169 | 4.0568 | 8.4035 | 0.011605 | 0.066095 | 1.0286 | 0.98613 | 0.35984 | 1.0637 | 0.95114 | 1.1036 | 1.1419 | 0.33572 | -0.011255 | 0.023672 | 0.023889 | 0.094173 | 0.025466 | 1.6545 | 1.6208 | [7 empty] | 0.0046769 | 0.0037236 | 0.061714 | 0.065929
190 | 8_30_21_W_8 |  | 06/09/2023 08:36:07 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_W_8.544 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1732.795 | 1236.7628 | 8.898 | 10.3313 | 3.4164 | 3866.3181 | 359.0262 | 7250.1434 | 0.2007 | 0.039833 | 0.038235 | 7.6022 | 4.0946 | 7.4649 | 0.0077961 | 0.046629 | 0.96687 | 0.95082 | 0.36748 | 1.0399 | 0.89146 | 1.1477 | 1.1739 | 0.36254 | 0.016704 | 0.022729 | 0.017255 | 0.088108 | 0.0078195 | 1.5942 | 1.6539 | [7 empty] | 0.0039508 | 0.0043502 | 0.061281 | 0.067404
191 | 8_30_21_W_8_SP |  | 06/09/2023 08:41:55 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_W_8_SP.545 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2886.0127 | 2390.8578 | 9.8296 | 12.4587 | 4.8817 | 4949.3681 | 1431.896 | 8732.8464 | 1.3548 | 1.1881 | 1.1646 | 8.2013 | 5.201 | 8.4444 | 1.1782 | 1.2798 | 2.0526 | 1.9756 | 1.4358 | 2.1199 | 2.015 | 2.1628 | 2.3512 | 1.3605 | 1.1318 | 1.179 | 1.1864 | 1.0825 | 1.2521 | 2.7271 | 2.8648 | [7 empty] | 1.0803 | 1.0862 | 1.1258 | 1.2355
192 | LBLANK |  | 06/09/2023 08:47:44 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.546 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | -13.3912 | -42.7414 | -0.70319 | -0.78213 | -1.3829 | 379.4154 | -0.11352 | -30.7443 | 0.0028168 | 0.0030021 | 0.0032259 | -0.50412 | -0.024715 | -0.6075 | 0.0021995 | 0.019936 | -0.0095497 | -0.0076463 | -0.1644 | -0.032562 | 0.0010604 | 0.15838 | -0.0012533 | 0.27624 | -0.016323 | 0.0014387 | 0.0019614 | 0.023643 | -0.0014165 | -0.0080142 | -0.011205 | [7 empty] | 0.00026628 | -0.00070797 | -0.0021271 | -0.001969
193 | CCV 1ppb TRACE |  | 06/09/2023 08:53:33 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.547 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1031.8537 | 953.3088 | 1.0689 | 1.3093 | -0.91718 | 1623.6707 | 1033.0956 | 1057.154 | 1.1153 | 1.1272 | 1.1152 | 0.76352 | 1.1597 | 0.63389 | 1.104 | 1.0999 | 1.0772 | 1.1377 | 0.92757 | 1.0033 | 0.96884 | 1.0672 | 1.0893 | 1.3203 | 1.0705 | 1.1046 | 1.14 | 1.0993 | 1.0261 | 1.1307 | 1.1511 | [7 empty] | 1.0569 | 1.0482 | 1.0727 | 1.1084
194 | CCV 1ppb TRACE |  | 06/09/2023 08:59:50 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.548 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1079.761 | 984.0317 | 0.97784 | 1.0862 | -1.0429 | 1784.1959 | 987.648 | 1085.102 | 1.1324 | 1.0799 | 1.1164 | 0.80172 | 1.1584 | 0.7542 | 1.0848 | 1.1162 | 1.1407 | 1.134 | 0.94972 | 1.0102 | 0.90142 | 1.0622 | 1.0651 | 1.2202 | 1.0165 | 1.1738 | 1.1152 | 1.1001 | 1.1936 | 1.1413 | 1.1734 | [7 empty] | 1.0182 | 1.0627 | 1.0167 | 1.103
195 | CCV 10 ppb MAJORS |  | 06/09/2023 09:05:39 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.549 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10992.7877 | 10545.2027 | 7.9606 | 11.007 | 9.3178 | 11988.486 | 10257.7245 | 11514.3664 | 11.4202 | 11.2655 | 11.0759 | 11.0225 | 11.7624 | 11.1772 | 11.4556 | 11.429 | 11.0224 | 10.6617 | 10.6764 | 10.1097 | 10.0917 | 9.1025 | 10.8117 | 10.5986 | 11.059 | 10.5726 | 10.6964 | 10.5923 | 10.839 | 11.4291 | 11.6291 | [7 empty] | 10.2806 | 10.4932 | 9.6686 | 10.4674
196 | CCV 10 ppb MAJORS |  | 06/09/2023 09:11:56 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.550 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10389.8526 | 9967.1841 | 7.8848 | 9.9334 | 9.1764 | 11508.9672 | 9789.0791 | 10913.4873 | 10.8624 | 10.6877 | 11.1521 | 11.1438 | 11.8195 | 10.6551 | 10.862 | 10.8946 | 10.5392 | 10.7941 | 10.7343 | 10.1943 | 9.6135 | 9.5865 | 10.9978 | 10.3312 | 10.5074 | 10.5744 | 10.8123 | 10.8018 | 10.7438 | 11.5463 | 11.7391 | [7 empty] | 10.354 | 10.5002 | 9.5143 | 10.3809
197 | 8_30_21_W_9 |  | 06/09/2023 09:17:44 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_W_9.551 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1519.6953 | 1264.3598 | 13.9836 | 17.0116 | 3.0075 | 3130.2946 | 343.3971 | 6979.7717 | 0.18591 | 0.034953 | 0.034332 | 15.494 | 5.5517 | 15.7576 | 0.014522 | 0.044203 | 1.0424 | 1.017 | 0.61058 | 0.86314 | 0.81785 | 0.98912 | 0.97815 | 0.3089 | 0.00011383 | 0.015647 | 0.027173 | 0.071097 | 0.021757 | 1.149 | 1.1669 | [7 empty] | 0.0033906 | 0.0038938 | 0.091138 | 0.09389
198 | 8_30_21_W_10 |  | 06/09/2023 09:23:32 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_W_10.552 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1508.8286 | 1728.0829 | 12.4715 | 14.8945 | 4.0738 | 4020.6105 | 388.0208 | 8256.4271 | 0.19598 | 0.026693 | 0.031073 | 11.7787 | 3.588 | 12.3134 | 0.012576 | 0.045773 | 0.80954 | 0.76206 | 0.56606 | 0.91022 | 0.85357 | 1.0244 | 1.0021 | 0.19143 | 0.011741 | 0.030336 | 0.023721 | 0.10261 | 0.022207 | 0.72576 | 0.7619 | [7 empty] | 0.0045009 | 0.0040684 | 0.079107 | 0.083173
199 | 8_30_21_W_11 |  | 06/09/2023 09:29:20 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_W_11.553 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1480.3533 | 1579.9002 | 14.717 | 16.97 | 4.2731 | 3884.0715 | 359.3443 | 7233.5574 | 0.21692 | 0.026761 | 0.027798 | 13.7296 | 4.0971 | 13.2398 | 0.010209 | 0.055881 | 0.779 | 0.82247 | 0.64369 | 0.9376 | 0.71423 | 1.0074 | 1.014 | 0.21525 | -0.010649 | 0.018059 | 0.024725 | 0.090709 | 0.022367 | 0.6933 | 0.69912 | [7 empty] | 0.0046013 | 0.0058572 | 0.24568 | 0.27797
200 | 8_30_21_F_1 |  | 06/09/2023 09:35:08 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_F_1.554 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 7546.9993 | 5893.4038 | 4.1234 | 5.4336 | 8.2326 | 23277.9562 | 1918.9465 | 39001.853 | 0.36695 | 0.13047 | 0.13971 | 3.8402 | 12.858 | 4.2528 | 0.035297 | 0.4696 | 0.91798 | 0.92522 | 5.6603 | 4.7334 | 4.1202 | 4.211 | 5.0723 | 0.58023 | 0.043061 | 0.04526 | 0.051136 | 0.40318 | 0.028399 | 16.5593 | 17.217 | [7 empty] | 0.0076911 | 0.006764 | 0.094584 | 0.1036
201 | 8_30_21_F_1 |  | 06/09/2023 09:41:25 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_F_1.555 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 7238.8248 | 5347.5665 | 4.3052 | 4.9601 | 7.5795 | 22524.8113 | 1795.6325 | 36153.42 | 0.36185 | 0.13422 | 0.14649 | 3.9721 | 12.7933 | 3.9015 | 0.034114 | 0.43107 | 0.86434 | 0.94563 | 5.7101 | 4.8098 | 4.1496 | 4.4161 | 5.0594 | 0.4978 | 0.053205 | 0.031363 | 0.044577 | 0.33395 | 0.095837 | 16.8223 | 17.1193 | [7 empty] | 0.0057911 | 0.0074352 | 0.097379 | 0.10333
202 | 8_30_21_F_1_DUP |  | 06/09/2023 09:47:13 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_F_1_DUP.556 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 7752.4812 | 5762.7149 | 4.9108 | 5.3855 | 7.619 | 23318.8905 | 1795.3641 | 37666.5611 | 0.38171 | 0.13172 | 0.12302 | 3.8841 | 12.5475 | 4.2022 | 0.03405 | 0.47364 | 0.91571 | 0.94654 | 5.6642 | 5.1486 | 4.0809 | 4.4115 | 5.0008 | 0.72635 | 0.035867 | 0.056912 | 0.04277 | 0.4554 | 0.051839 | 16.6007 | 17.0399 | [7 empty] | 0.0075863 | 0.006655 | 0.10496 | 0.10493
203 | 8_30_21_F_1_DUP |  | 06/09/2023 09:53:30 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_F_1_DUP.557 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 7650.2474 | 5679.3873 | 4.371 | 5.7819 | 7.4682 | 22870.8844 | 1842.7719 | 36766.0404 | 0.37471 | 0.12309 | 0.12712 | 3.8458 | 12.7895 | 3.9592 | 0.03163 | 0.45206 | 0.87886 | 0.94584 | 5.6941 | 4.6671 | 4.0617 | 4.4278 | 5.057 | 0.49434 | 0.018815 | 0.042613 | 0.04766 | 0.39072 | 0.067201 | 16.5282 | 17.0674 | [7 empty] | 0.0080843 | 0.0059629 | 0.095829 | 0.10651
204 | 8_30_21_F_6 |  | 06/09/2023 09:59:18 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_F_6.558 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 8437.8915 | 6658.332 | 2.5772 | 2.7705 | 0.89808 | 19039.1386 | 1830.7598 | 36241.4473 | 0.72013 | 0.049118 | 0.063106 | 6.3566 | 12.4064 | 6.0683 | 0.028888 | 0.25293 | 2.7501 | 2.9508 | 3.3973 | 4.889 | 4.3966 | 4.629 | 5.2435 | 0.51356 | 0.053797 | 0.10099 | 0.083519 | 0.36134 | 0.13043 | 9.4547 | 9.5542 | [7 empty] | 0.02032 | 0.016284 | 0.12611 | 0.16512
205 | 8_30_21_F_6 |  | 06/09/2023 10:05:38 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_F_6.559 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 8363.874 | 6860.6339 | 2.5257 | 2.7755 | 0.78046 | 19766.9928 | 1784.9782 | 38114.3138 | 0.71425 | 0.06332 | 0.064272 | 6.3954 | 12.7851 | 6.2713 | 0.027418 | 0.24759 | 2.9144 | 2.9855 | 3.4426 | 4.8534 | 4.1816 | 4.6165 | 5.2947 | 0.59542 | 0.077346 | 0.10724 | 0.092745 | 0.37676 | 0.10154 | 9.5752 | 9.6967 | [7 empty] | 0.020889 | 0.0171 | 0.12263 | 0.13493
206 | 8_30_21_F_8 |  | 06/09/2023 10:11:26 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_F_8.560 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9010.4641 | 6602.8872 | 1.5135 | 1.8563 | 2.452 | 18134.7389 | 1923.611 | 36632.3517 | 0.90541 | 0.068773 | 0.062577 | 6.8386 | 15.1166 | 6.8628 | 0.026937 | 0.2484 | 3.1665 | 3.3355 | 1.6868 | 5.28 | 4.4256 | 5.3196 | 6.1959 | 0.5456 | 0.041655 | 0.099006 | 0.090011 | 0.42065 | 0.090699 | 8.5001 | 9.4844 | [7 empty] | 0.017546 | 0.013855 | 0.10288 | 0.11192
207 | 8_30_21_F_8 |  | 06/09/2023 10:17:43 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_F_8.561 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9078.2803 | 6883.7996 | 1.5155 | 1.2477 | 2.3067 | 18594.5226 | 1935.3788 | 38345.1762 | 0.90178 | 0.067133 | 0.063359 | 6.8157 | 13.4181 | 7.1124 | 0.028799 | 0.22869 | 3.2746 | 3.2554 | 1.6816 | 5.3109 | 4.9143 | 4.8249 | 5.4584 | 0.59737 | 0.087629 | 0.09677 | 0.092591 | 0.45284 | 0.085654 | 8.4679 | 8.398 | [7 empty] | 0.015406 | 0.014741 | 0.10661 | 0.11109
208 | LBLANK |  | 06/09/2023 10:23:31 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.562 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | -12.8327 | -42.7414 | -0.70053 | -0.72136 | -2.2825 | 603.3681 | 1.615 | -30.7443 | 0.003544 | 0.001236 | 0.0036236 | -0.49989 | -0.024707 | -0.50489 | 0.0014541 | 0.0085312 | -0.01003 | -0.0063536 | -0.16236 | -0.033331 | -0.0057419 | 0.16006 | 0.00081604 | 0.24527 | -0.021858 | 0.0035984 | 0.0038821 | -0.010501 | 0.0076898 | -0.008223 | -0.009355 | [7 empty] | 7.9106e-05 | -0.00056883 | -0.0024448 | -0.0018272
209 | CCV 1ppb TRACE |  | 06/09/2023 10:29:20 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.563 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1084.9647 | 993.13 | 0.98655 | 1.2115 | -1.1015 | 1898.0804 | 976.3604 | 1229.5396 | 1.1575 | 1.1118 | 1.1371 | 0.81718 | 1.1555 | 0.81617 | 1.1495 | 1.1762 | 1.1615 | 1.1289 | 0.97656 | 0.98685 | 0.99232 | 1.0786 | 1.0923 | 1.1737 | 1.1536 | 1.0609 | 1.1133 | 1.0824 | 1.0107 | 1.1457 | 1.1698 | [7 empty] | 1.0658 | 1.0737 | 0.9912 | 1.0887
210 | CCV 1ppb TRACE |  | 06/09/2023 10:35:37 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.564 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1091.3034 | 1077.1134 | 0.98195 | 1.379 | -1.1796 | 1868.8095 | 966.2532 | 1166.725 | 1.1809 | 1.1451 | 1.1464 | 0.81924 | 1.1395 | 0.86577 | 1.1539 | 1.1773 | 1.165 | 1.133 | 0.95846 | 0.95168 | 0.98756 | 1.0578 | 1.086 | 1.247 | 1.0386 | 1.1102 | 1.0989 | 1.1179 | 0.99256 | 1.1434 | 1.1772 | [7 empty] | 1.0656 | 1.0671 | 0.9878 | 1.0953
211 | LFB |  | 06/09/2023 10:41:25 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.565 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9832.9683 | 9512.4877 | 7.8929 | 10.3227 | 12.0467 | 10952.736 | 9277.6341 | 10700.7818 | 10.3999 | 10.2808 | 10.8458 | 10.6165 | 11.4097 | 9.9622 | 10.3816 | 10.4985 | 10.1137 | 10.2807 | 10.6458 | 9.8879 | 9.0591 | 8.9786 | 10.8905 | 9.7347 | 10.7941 | 10.8602 | 10.8534 | 10.7314 | 11.0785 | 11.0507 | 11.4922 | [7 empty] | 10.05 | 10.2693 | 9.0812 | 10.0593
212 | LFB |  | 06/09/2023 10:47:41 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.566 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9938.389 | 9253.4913 | 8.1182 | 9.945 | 12.5979 | 10929.352 | 9350.0958 | 10403.04 | 10.348 | 10.2445 | 10.8123 | 10.6193 | 11.4597 | 9.8958 | 10.5201 | 10.4879 | 10.1432 | 10.3584 | 10.6537 | 10.2057 | 9.4876 | 9.351 | 10.8188 | 10.1654 | 10.7933 | 10.7603 | 11.044 | 10.8664 | 10.9967 | 11.1718 | 11.446 | [7 empty] | 10.1133 | 10.2828 | 9.1207 | 10.1033
213 | CCV 10 ppb MAJORS |  | 06/09/2023 10:53:35 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.567 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10573.7087 | 10097.63 | 7.9675 | 9.7495 | 9.4746 | 11710.2334 | 9917.0711 | 11235.9814 | 11.097 | 10.8697 | 11.2383 | 11.0506 | 12.0634 | 10.9045 | 11.0704 | 10.927 | 10.5331 | 10.6238 | 10.6343 | 10.1742 | 9.6743 | 9.2865 | 11.1364 | 10.4692 | 11.1464 | 10.938 | 11.154 | 10.8519 | 10.7872 | 11.5057 | 11.9297 | [7 empty] | 10.2002 | 10.353 | 9.3466 | 10.2881
214 | LBLANK |  | 06/09/2023 11:28:06 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.568 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth
215 | Cal blank |  | 06/09/2023 11:33:55 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\Cal blank.569 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth
216 | CAL BLANK |  | 06/09/2023 11:39:43 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CAL BLANK.570 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth
217 | LOW 0.01 |  | 06/09/2023 11:45:32 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LOW 0.01.571 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth
218 | LOW 0.05 |  | 06/09/2023 11:51:20 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LOW 0.05.572 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | [8 empty] | 0.05 | 0.05 | 0.05 |  | 0.05 |  | 0.05 | 0.05 | 0.05 | 0.05 |  | 0.05 | 0.05 | 0.05 | 0.05 | [6 empty] | 0.05 | 0.05 | [7 empty] | 0.05 | 0.05 | 0.05 | 0.05
219 | LOW 0.1 |  | 06/09/2023 11:57:09 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LOW 0.1.573 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | [8 empty] | 0.087104 | 0.10707 | 0.099403 |  | 0.12409 |  | 0.09896 | 0.091533 | 0.098328 | 0.098929 |  | 0.084465 | 0.083518 | 0.098642 | 0.096277 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.10618 | 0.11419 | [7 empty] | 0.10143 | 0.099011 | 0.10268 | 0.097276
220 | LOW 0.5 |  | 06/09/2023 12:02:57 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LOW 0.5.574 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth |  |  |  |  | 0.5 |  |  |  | 0.49413 | 0.47486 | 0.48139 | 0.5 | 0.65836 | 0.5 | 0.48102 | 0.47385 | 0.50812 | 0.51569 | 0.5 | 0.37164 | 0.56696 | 0.50036 | 0.49609 | 0.36485 | 0.41213 | 0.5072 | 0.50701 | 0.46012 | 0.56548 | 0.55952 | 0.57619 | [7 empty] | 0.49547 | 0.49443 | 0.50215 | 0.55518
221 | STD 1 |  | 06/09/2023 12:08:45 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 1.575 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1000 | 1000 | 1 | 1 | 0.55522 |  | 1000 | 1000 | 1.0279 | 1.0008 | 0.99475 | 0.91334 | 1.2405 | 0.88241 | 0.99038 | 0.92553 | 0.99065 | 0.98827 | 0.96184 | 0.81413 | 1.0616 | 1.0095 | 1.0216 | 0.77729 | 0.85579 | 0.9918 | 0.99198 | 0.97996 | 1.0996 | 1.0939 | 1.1454 | [7 empty] | 1.0051 | 0.98941 | 0.98213 | 0.95279
222 | STD 5 |  | 06/09/2023 12:14:33 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 5.576 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4856.1427 | 4918.7394 | 4.7459 | 4.8287 | 3.53 | 5000 | 5077.9857 | 5012.7903 | 4.9915 | 4.8182 | 4.9916 | 3.8132 | 5.9735 | 3.7391 | 4.7624 | 4.7094 | 4.8572 | 4.7822 | 4.8327 | 4.228 | 4.9347 | 5.0035 | 5.1086 | 4.7663 | 4.4998 | 4.8919 | 4.9788 | 5.1468 | 5.2046 | 5.3354 | 5.4776 | [7 empty] | 4.9252 | 4.9921 | 4.8438 | 4.675
223 | STD 10 |  | 06/09/2023 12:20:21 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 10.577 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10101.1051 | 10091.8913 | 9.7017 | 9.4484 | 6.9961 | 10349.7893 | 10447.8693 | 10249.2277 | 10.3488 | 10.0408 | 10.1648 | 8.3378 | 11.6951 | 8.0378 | 10.0127 | 10.1035 | 9.8306 | 9.7489 | 9.8438 | 8.8053 | 10.1776 | 10.0077 | 10.2766 | 9.5276 | 9.6872 | 9.763 | 9.8211 | 10.0907 | 10.3835 | 10.8496 | 10.9932 | [7 empty] | 9.8264 | 10.0065 | 9.6466 | 9.4106
224 | STD 25 |  | 06/09/2023 12:26:09 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 25.578 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 25070.0782 | 24991.2149 | 27.3334 | 23.5337 | 19.5029 | 26155.6288 | 25647.6427 | 25044.7638 | 26.1792 | 25.2582 | 25.8833 | 22.2877 | 28.7751 | 21.3861 | 25.1612 | 24.6289 | 24.4845 | 24.5423 | 25.3559 | 25.6451 | 27.0333 | 25.1241 | 26.6622 | 27.4735 | 24.4891 | 26.0348 | 25.6789 | 28.557 | 27.3585 | 27.3519 | 27.7447 | [7 empty] | 25.1122 | 25.2793 | 26.0154 | 22.9519
225 | STD 50 |  | 06/09/2023 12:31:57 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 50.579 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 51529.8733 | 50332.3498 | 59.2987 | 49.0069 | 41.4208 | 52095.888 | 52273.9954 | 51171.8421 | 53.8646 | 51.6358 | 52.4368 | 42.433 | 57.0428 | 45.2206 | 51.0892 | 50.3715 | 49.6855 | 48.2667 | 49.7924 | 45.9809 | 53.5519 | 50.614 | 52.7338 | 48.0343 | 50.456 | 51.9228 | 51.5557 | 50.7775 | 53.0673 | 54.2528 | 55.2244 | [7 empty] | 48.8322 | 48.9903 | 44.7382 | 47.6419
226 | STD 100 |  | 06/09/2023 12:37:44 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 100.580 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 108498.323 | 105277.2281 | 121.9483 | 102.2174 | 84.9644 | 109393.8023 | 111350.031 | 107409.5804 | 117.1948 | 112.5209 | 113.0715 | 90.6799 | 120.057 | 95.0887 | 108.7277 | 104.4821 | 102.5525 | 98.3322 | 100.4032 | 95.2398 | 111.6938 | 100.136 | 107.2627 | 100.3927 | 104.0236 | 107.5178 | 106.6472 | 106.6085 | 109.4809 | 115.3563 | 114.7526 | [7 empty] | 98.2079 | 98.9217 | 86.6002 | 84.4994
227 | Calibration Curves |  | 06/09/2023 12:37:44
228 | LBLANK |  | 06/09/2023 12:43:32 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.581 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 28.8843 | 13.3084 | -0.54623 | -0.71016 | -0.48328 | 469.8215 | 33.7033 | 65.8758 | 0.0062431 | 3.8361e-05 | -0.00024657 | -0.19936 | -0.02705 | -0.046952 | 0.0037233 | 0.0082987 | -0.0047314 | -0.0066264 | -0.13116 | -0.05553 | 0.023883 | 0.26669 | 0.0063026 | 0.29045 | 0.001725 | 0.008059 | 0.0085619 | 0.042051 | -0.029558 | 0.088502 | -0.0059357 | [7 empty] | 0.0015747 | -0.0003043 | -0.0029503 | -0.0012646
229 | LBLANK |  | 06/09/2023 13:13:17 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.583 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 33.848 | 12.1963 | -0.63327 | -0.58786 | -0.67693 | -305.5873 | 35.2432 | 65.8758 | 0.0039552 | -0.00098948 | -0.00052821 | -0.51706 | 0.0012718 | -0.53421 | 0.0019316 | 0.0042903 | -0.0068075 | -0.0065981 | 0.032302 | -0.024094 | 0.024445 | 0.26285 | 0.0018791 | 0.18619 | -0.042869 | 0.0084746 | 0.0029837 | 0.030816 | -0.041654 | 0.049407 | 0.0070623 | [7 empty] | 0.0011205 | -0.0010036 | -0.0033749 | -0.0018772
230 | LBLANK |  | 06/09/2023 13:19:06 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.584 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 28.0894 | 12.1963 | -0.72088 | -0.7574 | -0.84198 | 423.1132 | 35.3658 | 65.8758 | 0.0019891 | -0.0060389 | -0.0016394 | -0.53294 | 0.0010633 | -0.50355 | 0.0014738 | 0.0041766 | -0.009049 | -0.0076148 | -0.14069 | -0.0026811 | 0.010141 | 0.26237 | 0.0013302 | 0.044187 | -0.053078 | 0.0085013 | 0.0059206 | -0.0062396 | -0.041654 | 0.044702 | 0.003428 | [7 empty] | 0.0010166 | -0.0011082 | -0.0033389 | -0.0020406
231 | ICV PE |  | 06/09/2023 13:24:53 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\ICV PE.585 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10402.4606 | 10790.0138 | 8.6359 | 9.6861 | 18.1854 | 10365.3137 | 10205.5983 | 10292.2368 | 10.733 | 10.5487 | 9.9666 | 10.139 | 10.057 | 10.548 | 10.6406 | 10.687 | 10.49 | 9.6931 | 9.4901 | 9.6037 | 9.5931 | 9.4309 | 10.1099 | 6.8086 | 10.1598 | 9.9229 | 9.7291 | 9.6263 | 10.1502 | 9.5231 | 9.4949 | [7 empty] | 9.6305 | 9.7447 | 9.216 | 9.4524
232 | ICV PE |  | 06/09/2023 13:31:10 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\ICV PE.586 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10122.2283 | 10855.4755 | 8.9054 | 9.4452 | 18.1516 | 9683.2897 | 9953.6138 | 10338.8355 | 10.4088 | 10.3483 | 11.5413 | 11.7797 | 10.0963 | 10.3264 | 10.2743 | 10.5963 | 10.4382 | 11.2558 | 11.1597 | 9.7487 | 9.4015 | 9.4857 | 10.1739 | 7.0021 | 10.1279 | 9.7158 | 9.8288 | 9.6421 | 10.1201 | 10.9642 | 9.5224 | [7 empty] | 11.1516 | 11.2557 | 9.3969 | 9.2748
233 | CCV 1ppb TRACE |  | 06/09/2023 13:36:58 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.587 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1012.0115 | 1031.7646 | 1.1072 | 1.1547 | 0.21933 | 1080.0053 | 1042.3484 | 1008.9618 | 1.0311 | 0.98852 | 1.0257 | 0.64356 | 1.0076 | 0.62252 | 0.97966 | 0.98319 | 1.0464 | 1.0593 | 0.89225 | 0.92828 | 1.0605 | 1.1684 | 1.015 | 1.1284 | 0.98134 | 0.92326 | 0.99144 | 0.90329 | 1.0864 | 1.0278 | 0.99404 | [7 empty] | 1.01 | 1.0126 | 0.98687 | 1.0111
234 | CCV 1ppb TRACE |  | 06/09/2023 13:43:15 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.588 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1043.7191 | 1063.8988 | 1.0537 | 1.2045 | 0.072191 | 1221.5046 | 1066.8359 | 1057.3282 | 1.079 | 1.0358 | 1.0503 | 0.67741 | 1.0355 | 0.78205 | 1.0177 | 1.0107 | 1.0884 | 1.0503 | 0.88945 | 1.1188 | 1.1159 | 1.1659 | 1.0081 | 1.0799 | 0.97475 | 1.0336 | 1.0393 | 1.0669 | 0.95948 | 1.0445 | 1.0203 | [7 empty] | 1.0321 | 1.027 | 0.96723 | 0.99754
235 | LFB |  | 06/09/2023 13:49:03 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.589 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9526.0582 | 9724.8269 | 8.4123 | 8.6757 | 10.5802 | 9700.8945 | 9346.9596 | 9322.9533 | 9.6882 | 9.7567 | 9.794 | 9.6263 | 9.7864 | 9.3537 | 9.5731 | 9.715 | 9.8361 | 9.7629 | 9.8083 | 9.8132 | 10.5522 | 9.1845 | 10.0564 | 9.7069 | 10.5715 | 9.8014 | 9.8874 | 9.8269 | 9.9559 | 9.7918 | 9.6215 | [7 empty] | 9.6619 | 9.7692 | 8.9503 | 9.0958
236 | LLOQ 0.01 |  | 06/09/2023 13:54:51 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LLOQ 0.01.590 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 34.7293 | 24.036 | -0.18318 | -0.13095 | 0.33499 | -45.9485 | 49.555 | 76.3561 | 0.021211 | 0.0062971 | 0.0080325 | -0.36542 | 0.037102 | -0.41536 | 0.0093336 | 0.011693 | 0.0076755 | 0.010278 | 0.025539 | 0.061947 | 0.016909 | 0.27198 | 0.011849 | 0.11901 | -0.037815 | 0.014684 | 0.0096409 | 0.0045345 | -0.028951 | 0.043077 | 0.023431 | [7 empty] | 0.012204 | 0.0080342 | 0.0083319 | 0.0096463
237 | CCV 10 ppb MAJORS |  | 06/09/2023 14:00:40 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.591 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10350.2821 | 10395.9406 | 8.2917 | 8.2791 | 8.8623 | 10539.3843 | 10188.7254 | 9940.4101 | 10.6015 | 10.5051 | 9.6156 | 9.4123 | 10.2109 | 10.2352 | 10.4859 | 10.5562 | 10.4632 | 9.3715 | 9.2352 | 10.1257 | 11.4212 | 9.7152 | 10.2558 | 9.7888 | 10.1078 | 9.6431 | 9.8179 | 9.8217 | 9.8025 | 9.5089 | 10.0476 | [7 empty] | 9.3278 | 9.4768 | 9.4024 | 9.5241
238 | CCV 10 ppb MAJORS |  | 06/09/2023 14:06:57 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.592 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9896.7305 | 9948.2934 | 8.4312 | 8.5052 | 8.6271 | 9698.5241 | 9916.7378 | 9574.7797 | 10.2203 | 10.0061 | 10.3564 | 10.2811 | 10.3055 | 9.8619 | 10.1985 | 10.1822 | 10.0987 | 10.0408 | 9.7984 | 10.1363 | 10.6039 | 9.5365 | 10.2556 | 10.169 | 10.0833 | 10.0715 | 9.7977 | 9.7178 | 10.01 | 10.2348 | 10.1442 | [7 empty] | 9.9873 | 10.0255 | 9.4366 | 9.4457
239 | LLOQ 0.1 |  | 06/09/2023 14:12:45 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LLOQ 0.1.593 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 133.0755 | 130.3552 | -0.42096 | -0.35051 | -0.37528 | 202.177 | 139.043 | 145.5491 | 0.10403 | 0.10336 | 0.10204 | -0.17154 | 0.10846 | -0.051706 | 0.11566 | 0.10179 | 0.11814 | 0.10638 | 0.45672 | 0.2031 | 0.15196 | 0.35565 | 0.10433 | 0.2092 | 0.094324 | 0.091263 | 0.1066 | 0.10935 | 0.072271 | 0.11944 | 0.10119 | [7 empty] | 0.09904 | 0.097686 | 0.095693 | 0.095319
240 | 8_30_21_F_9 |  | 06/09/2023 14:19:27 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_F_9.594 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 7372.8488 | 6724.842 | 1.1053 | 1.5575 | 3.444 | 13230.5621 | 1643.2144 | 30115.7601 | 0.72611 | 0.057552 | 0.058717 | 4.0314 | 6.5002 | 4.3067 | 0.016903 | 0.16078 | 2.6419 | 2.6617 | 0.85072 | 4.3507 | 3.9003 | 4.2577 | 4.3082 | 0.38355 | -0.0023603 | 0.073499 | 0.083641 | 0.43748 | 0.056395 | 5.1663 | 5.1573 | [7 empty] | 0.012089 | 0.010099 | 0.089304 | 0.091094
241 | 8_30_21_F_10 |  | 06/09/2023 14:25:14 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_F_10.595 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 7605.4317 | 8562.5889 | 2.4702 | 2.7483 | 5.653 | 15614.9794 | 2027.2265 | 35400.3667 | 0.823 | 0.071903 | 0.071788 | 3.8707 | 4.7868 | 4.3119 | 0.016704 | 0.16923 | 2.2124 | 2.1133 | 2.2426 | 4.6293 | 4.2277 | 4.4619 | 4.6144 | 0.32003 | 0.048495 | 0.085906 | 0.093583 | 0.50439 | 0.033279 | 3.3088 | 3.4063 | [7 empty] | 0.010856 | 0.0087869 | 0.099545 | 0.10543
242 | 8_30_21_F_11 |  | 06/09/2023 14:31:03 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_F_11.596 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 7528.8781 | 9191.9043 | 2.542 | 2.5494 | 3.7502 | 15500.1986 | 1970.5247 | 34033.8756 | 0.90535 | 0.073176 | 0.069375 | 3.3053 | 3.7728 | 3.8449 | 0.018276 | 0.18032 | 2.1842 | 2.116 | 1.4155 | 4.7482 | 4.4507 | 4.4323 | 4.6714 | 0.54324 | 0.07571 | 0.085125 | 0.11545 | 0.51325 | 0.069863 | 3.1972 | 3.2026 | [7 empty] | 0.0093834 | 0.0071903 | 0.10305 | 0.096277
243 | 8_30_21_UF_1 |  | 06/09/2023 14:36:51 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_UF_1.597 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 8320.4135 | 5831.4832 | 1.9205 | 2.0267 | 19.7294 | 19352.791 | 1924.7639 | 33123.7424 | 0.38764 | 0.058208 | 0.061593 | -0.081697 | 10.4533 | 0.32625 | 0.025743 | 0.37545 | 0.63925 | 0.63523 | 2.4048 | 4.672 | 4.4787 | 4.4706 | 4.5227 | 0.58761 | -0.013792 | 0.050874 | 0.03954 | 0.3523 | -0.0045232 | 13.6187 | 13.7337 | [7 empty] | 0.0081768 | 0.0054316 | 0.076785 | 0.072067
244 | 8_30_21_UF_6 |  | 06/09/2023 14:42:39 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_UF_6.598 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9523.0812 | 6935.5368 | 0.85983 | 0.59981 | 15.3217 | 16554.712 | 1911.6149 | 34065.3098 | 0.65319 | 0.058027 | 0.056616 | 0.22156 | 10.8045 | 0.65174 | 0.023888 | 0.25852 | 1.7779 | 1.6438 | 1.3796 | 4.8385 | 4.3707 | 4.8383 | 4.9559 | 0.56287 | 0.041442 | 0.05797 | 0.10526 | 0.43462 | 0.025084 | 7.8126 | 8.2503 | [7 empty] | 0.014868 | 0.014218 | 0.080954 | 0.087163
245 | 8_30_21_UF_8 |  | 06/09/2023 14:48:27 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_UF_8.599 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9534.3572 | 7103.6571 | 1.4054 | 1.5788 | 16.8898 | 15586.1169 | 2014.0234 | 32621.0707 | 0.83508 | 0.072469 | 0.052878 | -0.12639 | 11.2526 | 0.23476 | 0.020179 | 0.19625 | 1.8802 | 1.8273 | 0.85176 | 4.9927 | 4.8737 | 4.875 | 5.0245 | 0.57168 | 0.081949 | 0.099428 | 0.078658 | 0.47057 | 0.032948 | 6.9325 | 6.9145 | [7 empty] | 0.015126 | 0.011403 | 0.06056 | 0.059831
246 | 8_30_21_UF_9 |  | 06/09/2023 14:54:15 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_UF_9.600 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 8404.8755 | 6524.6541 | 0.24392 | 0.23937 | 15.5168 | 12499.1699 | 1647.4951 | 29454.6348 | 0.71135 | 0.05764 | 0.052054 | -0.17014 | 6.1432 | 0.24216 | 0.019327 | 0.18918 | 1.7099 | 1.6538 | 3.1745 | 4.0082 | 4.4781 | 3.956 | 4.107 | 0.30994 | 0.015163 | 0.070336 | 0.072725 | 0.41079 | 0.01944 | 4.795 | 4.8001 | [7 empty] | 0.010693 | 0.0074593 | 0.059537 | 0.06337
247 | 8_30_21_UF_9_LD |  | 06/09/2023 15:00:03 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_UF_9_LD.601 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 8031.3725 | 6515.491 | 1.0267 | 1.335 | 18.9922 | 12539.6398 | 1624.7286 | 29269.6609 | 0.70477 | 0.048263 | 0.053716 | -0.064952 | 5.9675 | 0.19273 | 0.017213 | 0.16256 | 1.5967 | 1.7289 | 1.5396 | 3.9674 | 3.57 | 3.9774 | 4.0085 | 0.30317 | 0.025474 | 0.068014 | 0.093703 | 0.32015 | 0.01881 | 5.0348 | 4.6685 | [7 empty] | 0.011324 | 0.0079519 | 0.065401 | 0.067274
248 | 8_30_21_UF_9_LD |  | 06/09/2023 15:06:19 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_UF_9_LD.602 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 7965.0959 | 6546.9333 | 1.0362 | 1.2345 | 17.4724 | 12303.4553 | 1641.874 | 28653.0401 | 0.72786 | 0.058323 | 0.056416 | -0.10597 | 6.0293 | 0.18329 | 0.011504 | 0.18564 | 1.6287 | 1.6307 | 0.9205 | 3.9099 | 3.9799 | 3.8943 | 4.0846 | 0.4118 | 0.043376 | 0.079304 | 0.077703 | 0.29052 | 0.07594 | 4.7215 | 4.738 | [7 empty] | 0.0096816 | 0.0089778 | 0.063337 | 0.065462
249 | 8_30_21_UF_10 |  | 06/09/2023 15:12:07 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_UF_10.603 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 8472.3085 | 8484.3442 | 1.9179 | 2.1762 | 17.9355 | 14571.7212 | 1948.3658 | 34763.217 | 0.81316 | 0.068043 | 0.06421 | 0.83048 | 4.2891 | 1.1461 | 0.020972 | 0.15756 | 1.3915 | 1.402 | 1.3438 | 4.394 | 3.9521 | 4.2927 | 4.4186 | 0.40296 | 0.027 | 0.066 | 0.10395 | 0.4738 | 0.067368 | 3.2265 | 3.2417 | [7 empty] | 0.011232 | 0.0082623 | 0.07988 | 0.080989
250 | LBLANK |  | 06/09/2023 15:17:56 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.604 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 26.2709 | 13.319 | -0.69936 | -0.76018 | -0.37202 | 130.629 | 27.3721 | 70.2176 | 0.0013424 | -0.0047257 | -0.0006886 | -0.5246 | 0.0011045 | -0.54598 | 0.00060981 | 0.0032985 | -0.010553 | -0.0080581 | -0.15021 | 0.089651 | 0.038237 | 0.27621 | 0.016887 | 0.23529 | -0.042548 | 0.00072501 | 0.0060475 | 0.028391 | -0.037289 | 0.0026916 | 0.0037586 | [7 empty] | 0.0010229 | -0.0013405 | -0.0033056 | -0.0018326
251 | CCV 1ppb TRACE |  | 06/09/2023 15:23:44 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.605 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1030.3002 | 1107.7589 | 1.0969 | 1.1722 | 0.19496 | 1257.8902 | 1038.9043 | 1095.2024 | 1.0234 | 1.069 | 1.0357 | 0.69805 | 1.0109 | 0.6622 | 1.0115 | 1.0752 | 1.0758 | 1.0485 | 0.86927 | 1.0899 | 1.1851 | 1.166 | 1.0262 | 1.1277 | 0.96804 | 0.92377 | 0.97121 | 1.1089 | 0.96313 | 1.0169 | 1.0125 | [7 empty] | 1.0137 | 1.0271 | 1.0182 | 1.1293
252 | CCV 1ppb TRACE |  | 06/09/2023 15:30:02 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.606 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1048.8007 | 1147.2892 | 1.0941 | 0.72079 | 0.088427 | 1163.7187 | 1048.0881 | 1103.3704 | 1.0533 | 1.0081 | 1.0417 | 0.74341 | 1.0271 | 0.64882 | 1.0202 | 1.0339 | 1.0599 | 1.0761 | 0.89906 | 1.0757 | 1.0736 | 1.2161 | 1.0325 | 1.039 | 1.1258 | 1.0488 | 1.0353 | 0.93568 | 0.82707 | 1.0314 | 1.0313 | [7 empty] | 1.0504 | 1.0656 | 1.0235 | 1.0292
253 | CCV 10 ppb MAJORS |  | 06/09/2023 15:35:51 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.607 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10079.2403 | 10296.5081 | 8.5051 | 9.6642 | 8.4516 | 10013.7953 | 9964.8395 | 9910.8312 | 10.4844 | 10.0899 | 10.2684 | 10.1233 | 11.6735 | 9.6451 | 9.9282 | 10.0368 | 10.0718 | 9.9945 | 9.8268 | 10.1057 | 10.6171 | 10.7958 | 11.7603 | 9.9018 | 11.5121 | 11.2319 | 11.3624 | 9.8211 | 11.813 | 10.2358 | 11.5954 | [7 empty] | 10.1339 | 10.274 | 9.7834 | 9.9058
254 | CCV 10 ppb MAJORS |  | 06/09/2023 15:42:08 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.608 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10060.8796 | 10229.1112 | 8.2553 | 9.5989 | 5.9915 | 10299.688 | 9854.0791 | 10178.4009 | 10.3675 | 10.232 | 10.1164 | 10.1097 | 10.3263 | 9.9755 | 10.0747 | 10.3532 | 10.1758 | 9.8394 | 9.6371 | 9.7508 | 10.6018 | 9.6584 | 10.23 | 9.8167 | 10.3005 | 9.8925 | 9.9642 | 9.3935 | 9.948 | 10.0776 | 10.1505 | [7 empty] | 9.9543 | 10.0537 | 9.5638 | 9.8555
255 | 8_30_21_UF_10_LSP |  | 06/09/2023 15:47:56 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_UF_10_LSP.609 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 17365.4278 | 17581.9066 | 11.1992 | 11.476 | 23.5128 | 23197.1074 | 11782.5598 | 40956.2313 | 10.9144 | 9.9621 | 10.2421 | 11.2112 | 13.1278 | 11.3792 | 9.8013 | 9.9271 | 10.8114 | 10.8244 | 10.8601 | 13.827 | 14.4124 | 12.3795 | 13.252 | 10.0911 | 9.5057 | 9.438 | 9.3648 | 10.1008 | 9.7307 | 12.9794 | 12.1851 | [7 empty] | 9.8159 | 9.8213 | 9.3499 | 9.3586
256 | 8_30_21_UF_10_DUP_SP |  | 06/09/2023 15:53:45 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_UF_10_DUP_SP.610 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 14108.5777 | 14084.5612 | 5.2743 | 5.4299 | 24.5372 | 20751.6749 | 7570.4687 | 42227.5235 | 6.4444 | 5.5419 | 5.8325 | 5.7491 | 9.909 | 5.7674 | 5.4167 | 5.6421 | 6.7151 | 6.8213 | 6.7085 | 9.8112 | 10.3554 | 9.2842 | 9.9296 | 5.8174 | 5.5191 | 5.3775 | 5.356 | 5.5505 | 5.6788 | 9.0124 | 8.636 | [7 empty] | 5.4407 | 5.5574 | 5.1485 | 5.3223
257 | 8_30_21_UF_11 |  | 06/09/2023 15:59:33 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_UF_11.611 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 8065.7075 | 8504.7753 | 0.61339 | 0.65688 | 17.1705 | 14908.8046 | 1987.4744 | 32827.2518 | 0.83344 | 0.067081 | 0.057486 | -0.10642 | 3.381 | 0.26452 | 0.015242 | 0.19224 | 1.439 | 1.4216 | 0.83114 | 4.545 | 4.5633 | 4.4215 | 4.5489 | 0.43427 | 0.045713 | 0.08479 | 0.10218 | 0.32428 | 0.059543 | 3.0612 | 3.0904 | [7 empty] | 0.0099806 | 0.0049391 | 0.05923 | 0.062604
258 | 8_30_21_BB |  | 06/09/2023 16:05:21 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_BB.612 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 29.1599 | 14.6033 | 3.0502 | 2.9447 | 0.0016889 | 329.1188 | 47.4722 | 72.9402 | 0.0053734 | 0.024807 | 0.021117 | 0.73705 | 0.012399 | 0.95052 | 0.001103 | 0.0068628 | 0.10484 | 0.094086 | 0.0061283 | 0.074308 | 0.032691 | 0.27196 | 0.0117 | 0.086407 | -0.053078 | 0.015527 | 0.0061205 | 0.014641 | -0.041654 | 0.0041907 | 0.0047277 | [7 empty] | 0.0011339 | -0.00085688 | 0.033817 | 0.035714
259 | 8_30_21_FB |  | 06/09/2023 16:11:09 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_FB.613 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 55.394 | 24.3595 | 1.427 | 1.9395 | -0.28418 | 303.1345 | 38.5668 | 103.9039 | 0.0040616 | 0.15976 | 0.1519 | 0.17986 | 0.13601 | 0.17232 | 0.007061 | 9.4736 | 0.12651 | 0.1175 | 5.5104 | 0.071241 | 0.010141 | 0.27043 | 0.010284 | 0.1428 | -0.042193 | 0.0074574 | 0.0085552 | -0.0093558 | -0.0371 | 0.0044411 | 0.004955 | [7 empty] | 0.0023058 | -0.0001493 | 0.063451 | 0.065462
260 | 8_30_21_LAB_A |  | 06/09/2023 16:16:56 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_LAB_A.614 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 27.2854 | 12.1963 | -0.25959 | -0.11942 | -0.66346 | 253.0791 | 33.439 | 65.8758 | 0.0047203 | 0.070653 | 0.080143 | 0.092313 | 0.0088334 | 0.10053 | 0.00087106 | 0.28522 | 0.099523 | 0.096049 | 0.013109 | 0.072263 | 0.010141 | 0.267 | 0.006365 | 0.20081 | -0.053078 | 0.0025344 | 0.0030486 | 0.0012292 | -0.041654 | 0.023982 | 0.025113 | [7 empty] | 0.0011326 | -0.0011926 | 0.032982 | 0.032868
261 | 8_30_21_LAB_B |  | 06/09/2023 16:22:44 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\8_30_21_LAB_B.615 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 28.2395 | 12.1963 | -0.22676 | -0.31875 | -0.6428 | 148.4438 | 37.4518 | 73.0302 | 0.0013424 | 0.021579 | 0.028696 | -0.091964 | 0.0051528 | -0.16986 | 0.00064181 | 0.26659 | 0.075653 | 0.092292 | -0.020739 | 0.057848 | 0.017818 | 0.26559 | 0.0047472 | 0.14761 | -0.053078 | 0.0041921 | 0.0038174 | -0.0094023 | -0.037078 | 0.0017026 | 0.0028624 | [7 empty] | 0.0009187 | -0.0013025 | 0.032653 | 0.03178
262 | 9_13_21_W_1 |  | 06/09/2023 16:28:32 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_W_1.616 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1431.059 | 1177.3561 | 6.4934 | 8.0822 | 2.6273 | 4178.8156 | 453.1148 | 6784.2511 | 0.058839 | 0.028678 | 0.035264 | 7.2768 | 4.381 | 7.1902 | 0.0093741 | 0.068688 | 0.33552 | 0.32986 | 0.07969 | 0.88049 | 0.75282 | 1.0125 | 0.83515 | 0.141 | -0.02497 | 0.01121 | 0.01363 | 0.05912 | -0.027642 | 3.1849 | 3.1184 | [7 empty] | 0.0032873 | 0.00089727 | 0.034703 | 0.034128
263 | 9_13_21_W_6 |  | 06/09/2023 16:34:20 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_W_6.617 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1482.1859 | 1322.6815 | 5.8551 | 6.3051 | 1.231 | 3068.5802 | 335.1227 | 6205.8169 | 0.12314 | 0.032214 | 0.029852 | 7.776 | 3.1643 | 7.7823 | 0.0059579 | 0.05734 | 0.8517 | 0.82494 | 0.49123 | 0.90553 | 0.90412 | 1.07 | 0.87993 | 0.22008 | -0.018736 | 0.025334 | 0.017873 | 0.15201 | -0.0087112 | 1.1615 | 1.157 | [7 empty] | 0.0059896 | 0.0028704 | 0.060291 | 0.062624
264 | 9_13_21_W_8 |  | 06/09/2023 16:40:07 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_W_8.618 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1627.3536 | 1373.6259 | 5.9716 | 7.1458 | 1.6691 | 2720.0914 | 400.5724 | 6467.0618 | 0.17124 | 0.060711 | 0.076185 | 9.5979 | 2.8293 | 8.1304 | 0.0074319 | 0.1123 | 0.69324 | 0.83719 | 0.33101 | 1.0229 | 0.79213 | 1.0746 | 0.90938 | 0.12963 | -0.030413 | 0.025321 | 0.024499 | 0.10167 | 0.00061869 | 1.1729 | 0.98398 | [7 empty] | 0.00515 | 0.0027356 | 0.048866 | 0.050529
265 | LBLANK |  | 06/09/2023 16:45:55 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.619 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 27.7253 | 12.1963 | -0.73915 | -0.75231 | -0.75297 | 113.8578 | 32.3785 | 65.8758 | 0.0020398 | -0.0028941 | -0.00087695 | -0.55435 | 0.00083728 | -0.55519 | 0.00041452 | -0.0014944 | -0.010466 | -0.008599 | -0.15804 | 0.10023 | 0.010141 | 0.26283 | 0.0018582 | 0.12282 | -0.053078 | 0.00082217 | 0.0046871 | -0.036165 | -0.041654 | 0.0019808 | 0.0031888 | [7 empty] | 0.0012397 | -0.0011874 | -0.0037063 | -0.0022137
266 | CCV 1ppb TRACE |  | 06/09/2023 16:51:44 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.620 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 993.1619 | 1041.3404 | 0.48807 | 0.54156 | -0.019812 | 957.7847 | 999.2865 | 1065.9513 | 0.97358 | 0.96627 | 0.99705 | 0.77491 | 0.95789 | 0.78497 | 0.97097 | 0.96372 | 0.97835 | 0.98649 | 0.8434 | 0.99404 | 1.079 | 1.1465 | 0.99948 | 0.97419 | 1.0197 | 0.92347 | 1.0124 | 0.83096 | 0.9417 | 0.98114 | 0.98494 | [7 empty] | 0.97995 | 0.99255 | 0.99272 | 1.0217
267 | CCV 1ppb TRACE |  | 06/09/2023 16:58:02 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.621 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 966.7869 | 971.3645 | 0.46476 | 0.64628 | -0.065615 | 1056.7536 | 949.6856 | 1108.4414 | 1.0172 | 0.98682 | 0.9931 | 0.79067 | 0.94275 | 0.74872 | 0.96006 | 0.9526 | 1.0304 | 0.98146 | 0.84948 | 1.0453 | 0.88978 | 1.1391 | 0.97712 | 1.0156 | 0.896 | 0.90304 | 1.0256 | 0.93563 | 0.97357 | 0.97889 | 0.98462 | [7 empty] | 1.0051 | 1.0004 | 0.98972 | 1.0196
268 | LFB |  | 06/09/2023 17:03:50 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.622 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9262.4846 | 9471.6401 | 8.1244 | 9.0627 | 10.738 | 9258.761 | 9226.0767 | 9558.0387 | 9.5624 | 9.4227 | 9.5924 | 9.4899 | 9.6537 | 9.1613 | 9.3395 | 9.3215 | 9.4256 | 9.4542 | 9.5007 | 9.5848 | 10.4672 | 9.226 | 9.914 | 9.6342 | 10.2524 | 9.8394 | 9.8034 | 9.6658 | 10.2339 | 9.6754 | 9.7665 | [7 empty] | 9.7837 | 9.8395 | 9.4397 | 9.6013
269 | CCV 10 ppb MAJORS |  | 06/09/2023 17:09:38 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.623 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9677.6664 | 9970.5728 | 8.16 | 8.6489 | 8.2377 | 10400.4483 | 9814.0069 | 9921.448 | 10.2958 | 9.8427 | 10.7959 | 10.7359 | 9.9488 | 9.6644 | 9.8859 | 10.0682 | 9.9351 | 10.4581 | 10.7515 | 9.9187 | 10.7892 | 9.1103 | 9.7958 | 10.2309 | 10.1506 | 9.8266 | 9.5771 | 9.5753 | 9.9265 | 10.6589 | 9.8607 | [7 empty] | 10.6465 | 10.7446 | 9.7305 | 9.8364
270 | CCV 10 ppb MAJORS |  | 06/09/2023 17:15:55 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.624 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9642.6703 | 9692.0215 | 7.9438 | 8.8318 | 8.3094 | 9318.3186 | 9668.4219 | 10123.8807 | 9.798 | 9.5417 | 10.0094 | 9.9519 | 10.1888 | 9.5323 | 9.6938 | 9.8957 | 9.7993 | 9.7489 | 10.0579 | 9.7141 | 10.4829 | 9.2205 | 10.0749 | 9.5894 | 10.3581 | 9.6943 | 9.7155 | 9.4027 | 10.0792 | 9.9916 | 10.0249 | [7 empty] | 9.8276 | 9.9611 | 9.4716 | 9.6015
271 | 9_13_21_W_9 |  | 06/09/2023 17:21:43 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_W_9.625 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1452.5202 | 1355.911 | 4.5469 | 4.6595 | 1.2511 | 2646.2283 | 331.7143 | 5866.5146 | 0.14078 | 0.018604 | 0.019793 | 4.5486 | 1.7309 | 4.569 | 0.0056605 | 0.035378 | 0.58871 | 0.57913 | 0.079389 | 0.90221 | 0.79963 | 0.97404 | 0.79185 | 0.37823 | -0.030474 | 0.016464 | 0.026782 | 0.064454 | -0.0088088 | 0.76106 | 0.75983 | [7 empty] | 0.0031368 | 0.0011386 | 0.033869 | 0.033499
272 | 9_13_21_W_10 |  | 06/09/2023 17:27:31 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_W_10.626 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1353.6894 | 1691.1696 | 11.2151 | 12.1638 | 1.9732 | 2706.3437 | 362.0248 | 6470.5883 | 0.16784 | 0.013093 | 0.017975 | 10.2469 | 2.8566 | 9.9885 | 0.0079445 | 0.044731 | 0.60759 | 0.62425 | 0.38713 | 0.92803 | 0.75668 | 1.0908 | 0.92338 | 0.20922 | -0.047208 | 0.029727 | 0.025942 | 0.10217 | -0.031922 | 0.4947 | 0.48757 | [7 empty] | 0.0038632 | 0.0020711 | 0.067946 | 0.065375
273 | 9_13_21_W_10_DUP |  | 06/09/2023 17:33:20 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_W_10_DUP.627 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1305.2801 | 1722.5311 | 9.5939 | 9.7557 | 2.0356 | 3004.2621 | 390.6448 | 6378.6575 | 0.15885 | 0.020603 | 0.027796 | 8.5626 | 2.5026 | 8.7277 | 0.0084008 | 0.035213 | 0.50826 | 0.55308 | 0.33577 | 0.92111 | 0.75867 | 1.0603 | 0.89266 | 0.21189 | -0.024034 | 0.02236 | 0.023347 | 0.14582 | -0.017446 | 0.48412 | 0.48843 | [7 empty] | 0.0042131 | 0.0017923 | 0.079493 | 0.081176
274 | 9_13_21_W_11 |  | 06/09/2023 17:39:08 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_W_11.628 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1262.2545 | 1648.2715 | 22.8041 | 26.0731 | 2.0672 | 2849.4449 | 350.9965 | 6326.2494 | 0.1627 | 0.034828 | 0.040507 | 18.1507 | 3.584 | 18.1442 | 0.0093079 | 0.048495 | 0.84507 | 0.82269 | 0.80306 | 0.86672 | 0.78054 | 1.0454 | 0.86199 | 0.23744 | -0.029495 | 0.024452 | 0.025147 | 0.05197 | -0.022082 | 0.44564 | 0.44327 | [7 empty] | 0.0056532 | 0.0033254 | 0.092055 | 0.090943
275 | 9_13_21_F_1 |  | 06/09/2023 17:44:56 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_F_1.629 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 7265.8993 | 5740.635 | 3.3542 | 3.4301 | 7.8391 | 20453.4331 | 1892.4248 | 34467.0894 | 0.31926 | 0.17214 | 0.18037 | 4.5784 | 13.2675 | 4.8002 | 0.033831 | 0.48595 | 0.82159 | 0.89081 | 1.5823 | 4.2432 | 3.6625 | 3.9073 | 4.0164 | 0.62471 | -0.0053425 | 0.054099 | 0.047152 | 0.32966 | 0.017664 | 15.9339 | 16.0864 | [7 empty] | 0.0097698 | 0.0065376 | 0.12043 | 0.11513
276 | 9_13_21_F_6 |  | 06/09/2023 17:50:44 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_F_6.630 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 8090.0289 | 7439.234 | 2.5007 | 2.4975 | 1.4491 | 15734.8164 | 1885.3472 | 36122.8731 | 0.6942 | 0.10499 | 0.089059 | 4.3423 | 12.1021 | 5.0803 | 0.028459 | 0.40665 | 2.6074 | 2.4808 | 2.6743 | 4.1111 | 3.96 | 3.8671 | 4.1292 | 0.49143 | 0.036575 | 0.081188 | 0.098457 | 0.47141 | 0.046711 | 6.0134 | 6.0815 | [7 empty] | 0.01602 | 0.015804 | 0.13682 | 0.14381
277 | 9_13_21_F_8 |  | 06/09/2023 17:56:32 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_F_8.631 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 8700.1091 | 7057.6205 | 2.8348 | 3.0194 | 2.7141 | 14805.0086 | 1876.6058 | 34430.3418 | 0.88873 | 0.062953 | 0.068192 | 5.221 | 10.8373 | 5.9127 | 0.023603 | 0.26911 | 2.7927 | 2.6821 | 2.0266 | 4.5302 | 4.663 | 4.2861 | 4.497 | 0.56704 | 0.011983 | 0.073642 | 0.080893 | 0.50217 | 0.031963 | 5.1192 | 5.0869 | [7 empty] | 0.013257 | 0.011262 | 0.19062 | 0.19259
278 | 9_13_21_F_9 |  | 06/09/2023 18:02:20 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_F_9.632 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 7694.3349 | 7066.9083 | 6.851 | 7.2079 | 2.3701 | 12644.5093 | 1735.8204 | 30955.1119 | 0.71985 | 0.067375 | 0.067746 | 5.4635 | 3.9336 | 5.8622 | 0.018383 | 0.18979 | 2.354 | 2.2648 | 5.1581 | 3.8069 | 3.3942 | 3.5428 | 3.6982 | 0.50296 | 0.03041 | 0.07367 | 0.090924 | 0.38532 | 0.037234 | 3.7604 | 3.7511 | [7 empty] | 0.0087501 | 0.005871 | 0.11027 | 0.11734
279 | 9_13_21_F_10 |  | 06/09/2023 18:08:08 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_F_10.633 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 6720.4679 | 8466.093 | 3.0675 | 3.5375 | 3.3289 | 13953.7113 | 1839.8809 | 33615.5471 | 0.76553 | 0.062824 | 0.065684 | 3.7798 | 4.6232 | 3.8619 | 0.01702 | 0.1833 | 1.783 | 1.8893 | 1.8002 | 4.2404 | 4.1209 | 4.1177 | 4.2793 | 0.44616 | 0.026058 | 0.092804 | 0.083112 | 0.3748 | 0.039088 | 2.5619 | 2.4679 | [7 empty] | 0.0097031 | 0.0075681 | 0.11264 | 0.1188
280 | 9_13_21_F_11 |  | 06/09/2023 18:13:56 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_F_11.634 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 6546.1441 | 8488.2297 | 9.7718 | 10.8019 | 2.353 | 14529.1266 | 1821.4119 | 33490.9011 | 0.77635 | 0.064325 | 0.065044 | 4.3432 | 4.9411 | 4.2834 | 0.015147 | 0.17821 | 2.1329 | 2.2587 | 3.8006 | 4.1032 | 3.8567 | 4.0333 | 4.1773 | 0.45241 | 0.020084 | 0.083356 | 0.083279 | 0.44844 | 0.059348 | 2.4099 | 2.2889 | [7 empty] | 0.011052 | 0.0079887 | 0.10648 | 0.11068
281 | 9_13_21_F_11_LD |  | 06/09/2023 18:19:44 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_F_11_LD.635 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 6606.4935 | 8650.7375 | 8.6731 | 8.9113 | 3.3946 | 14347.0358 | 1856.2064 | 33978.0754 | 0.72166 | 0.060198 | 0.059493 | 4.032 | 5.0291 | 4.2061 | 0.018694 | 0.17477 | 2.168 | 2.1982 | 3.5672 | 4.2156 | 3.7024 | 4.0482 | 4.2137 | 0.48908 | -0.0038044 | 0.088635 | 0.089946 | 0.33682 | 0.030011 | 2.3118 | 2.3086 | [7 empty] | 0.010098 | 0.008941 | 0.10821 | 0.10923
282 | 9_13_21_F_11_LD |  | 06/09/2023 18:26:00 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_F_11_LD.636 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 6419.967 | 8408.2134 | 8.4553 | 9.3243 | 3.4651 | 13677.5987 | 1830.7418 | 32686.1297 | 0.71774 | 0.064445 | 0.055345 | 3.9319 | 4.9914 | 4.0516 | 0.017413 | 0.19412 | 2.1034 | 2.1427 | 3.4963 | 4.1084 | 3.8159 | 4.0614 | 4.134 | 0.4565 | 0.050707 | 0.090979 | 0.089569 | 0.42839 | 0.034225 | 2.2537 | 2.2884 | [7 empty] | 0.011467 | 0.0067253 | 0.10273 | 0.10759
283 | LBLANK |  | 06/09/2023 18:31:48 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.637 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 26.8542 | 12.1963 | -0.73585 | -0.73226 | -0.78764 | 45.8085 | 40.4007 | 65.8758 | 0.0020443 | -0.005768 | -0.0012446 | -0.54655 | 0.00078471 | -0.5523 | 0.00042811 | 0.0049746 | -0.0099498 | -0.0077734 | -0.1623 | 0.04554 | 0.010141 | 0.262 | 0.00090522 | 0.17269 | -0.030217 | 0.0043397 | 0.0038854 | 0.057765 | -0.036988 | 0.0015626 | 0.0031974 | [7 empty] | 0.0010293 | -0.0012204 | -0.0035345 | -0.0022193
284 | CCV 1ppb TRACE |  | 06/09/2023 18:37:37 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.638 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 990.5272 | 1004.2552 | 0.51466 | 0.59794 | -0.024345 | 1087.9378 | 1024.5979 | 998.689 | 1.0132 | 0.99961 | 0.99533 | 0.77654 | 0.97632 | 0.7694 | 0.98053 | 0.99964 | 0.96853 | 0.98302 | 0.81642 | 0.98861 | 0.99486 | 1.1792 | 0.98661 | 1.1434 | 0.97474 | 0.996 | 0.95995 | 0.95554 | 0.99607 | 0.97894 | 0.9928 | [7 empty] | 0.97334 | 0.97579 | 1.0095 | 1.1754
285 | CCV 1ppb TRACE |  | 06/09/2023 18:43:54 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.639 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 978.5032 | 984.4313 | 0.50882 | 0.41605 | -0.075403 | 1341.7701 | 1067.8231 | 1044.0416 | 1.0319 | 1.0151 | 0.96791 | 0.79594 | 0.97529 | 0.87173 | 0.98077 | 0.99058 | 1.081 | 0.98512 | 0.84038 | 1.0323 | 1.0578 | 1.1271 | 0.98683 | 1.1375 | 0.93332 | 1.0285 | 0.9997 | 0.93953 | 1.0488 | 0.96639 | 0.97562 | [7 empty] | 0.98266 | 0.98763 | 1.0123 | 1.1862
286 | CCV 10 ppb MAJORS |  | 06/09/2023 18:49:43 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.640 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9357.2943 | 9934.5988 | 8.0297 | 7.555 | 8.1385 | 9656.7771 | 9593.9796 | 9754.7959 | 9.9263 | 9.6521 | 9.8255 | 9.8189 | 10.8173 | 9.2958 | 9.6017 | 9.4213 | 9.6505 | 9.5815 | 9.7832 | 9.743 | 10.1818 | 10.2667 | 10.9586 | 9.6974 | 10.5995 | 10.6169 | 10.7044 | 9.5458 | 10.7835 | 9.7857 | 10.8307 | [7 empty] | 9.7434 | 9.9244 | 9.7802 | 9.9968
287 | CCV 10 ppb MAJORS |  | 06/09/2023 18:56:01 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.641 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9359.2825 | 9758.1549 | 8.2043 | 8.6707 | 8.0438 | 9660.2451 | 9328.1657 | 9747.3858 | 9.7615 | 9.5959 | 10.1773 | 10.1279 | 10.1694 | 9.3479 | 9.6939 | 9.4181 | 9.5992 | 9.934 | 10.1681 | 10.0508 | 9.7274 | 9.4649 | 10.1579 | 9.8794 | 9.8433 | 9.6168 | 9.9831 | 9.6275 | 9.9185 | 10.1765 | 10.0221 | [7 empty] | 9.986 | 10.1948 | 9.7992 | 9.0283
288 | 9_13_21_F_11_DUP |  | 06/09/2023 19:01:49 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_F_11_DUP.642 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 6319.7421 | 8200.8959 | 7.0696 | 7.8935 | 2.3763 | 14023.5136 | 1757.5997 | 33340.1601 | 0.70129 | 0.094072 | 0.11062 | 4.3052 | 4.9288 | 4.6348 | 0.019315 | 0.17221 | 2.1572 | 2.1897 | 3.3516 | 4.0512 | 3.7302 | 4.0078 | 4.1217 | 0.41564 | 0.015127 | 0.077222 | 0.091013 | 0.47286 | 0.076565 | 2.2531 | 2.2715 | [7 empty] | 0.011208 | 0.0084817 | 0.11267 | 0.12116
289 | 9_13_21_F_11_DUP_LSP |  | 06/09/2023 19:07:37 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_F_11_DUP_LSP.643 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 15282.9974 | 17161.2892 | 15.887 | 17.6304 | 12.2822 | 21879.0557 | 11122.8162 | 39866.3127 | 10.731 | 9.8767 | 10.0676 | 14.3913 | 15.5426 | 14.3204 | 9.7356 | 9.7202 | 11.4251 | 11.2439 | 12.4523 | 13.1433 | 13.9219 | 13.2762 | 14.7963 | 9.7463 | 10.5485 | 10.3849 | 10.3646 | 9.5374 | 10.1288 | 11.9488 | 12.9806 | [7 empty] | 9.5699 | 9.7408 | 9.4744 | 9.8775
290 | 9_13_21_UF_1 |  | 06/09/2023 19:13:24 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_UF_1.644 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 8406.4411 | 5776.4857 | 1.295 | 1.0836 | 6.5518 | 20677.0751 | 1914.2748 | 33226.2221 | 0.31017 | 0.052862 | 0.060253 | 0.44958 | 12.9073 | 0.74231 | 0.030787 | 0.48929 | 0.56333 | 0.62138 | 1.1934 | 4.3231 | 3.62 | 3.8744 | 3.9622 | 0.691 | -0.0097598 | 0.040885 | 0.047865 | 0.38443 | 0.0041033 | 15.4972 | 15.4994 | [7 empty] | 0.0093707 | 0.0068681 | 0.051488 | 0.049953
291 | 9_13_21_UF_6 |  | 06/09/2023 19:19:12 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_UF_6.645 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 8590.7769 | 6678.1146 | 0.062114 | -0.063557 | 2.1955 | 14559.0137 | 1710.2953 | 33109.0591 | 0.65611 | 0.066997 | 0.059011 | 0.11675 | 11.413 | 0.44434 | 0.019112 | 0.22479 | 1.6633 | 1.7502 | 1.6765 | 4.0614 | 3.8797 | 3.8915 | 3.9651 | 0.51524 | 0.030526 | 0.085183 | 0.074196 | 0.35931 | 0.033624 | 5.855 | 5.6403 | [7 empty] | 0.016644 | 0.014431 | 0.063749 | 0.068039
292 | 9_13_21_UF_6_DUP |  | 06/09/2023 19:25:00 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_UF_6_DUP.646 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 8727.8469 | 6832.6696 | 0.034985 | -0.045113 | 2.2106 | 14646.7774 | 1777.8245 | 32536.5065 | 0.6123 | 0.088386 | 0.081406 | 0.21757 | 13.4677 | 0.57943 | 0.021634 | 0.23609 | 1.7106 | 1.714 | 1.7324 | 4.0114 | 3.5668 | 4.5439 | 4.7133 | 0.57374 | 0.056426 | 0.10855 | 0.098715 | 0.38333 | 0.040242 | 5.7429 | 6.7444 | [7 empty] | 0.016624 | 0.012808 | 0.064614 | 0.069286
293 | 9_13_21_UF_8 |  | 06/09/2023 19:30:49 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_UF_8.647 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9186.9545 | 7155.0816 | 0.0098375 | 0.42558 | 2.6685 | 13504.8571 | 1878.8146 | 32664.3532 | 0.82355 | 0.053444 | 0.049778 | 0.017128 | 10.3329 | 0.47986 | 0.020961 | 0.21961 | 1.5915 | 1.5671 | 0.95528 | 4.217 | 4.2853 | 4.3014 | 4.4073 | 0.44348 | 0.032998 | 0.072723 | 0.089949 | 0.35168 | 0.0091943 | 4.9401 | 4.9173 | [7 empty] | 0.011493 | 0.0079384 | 0.050025 | 0.056634

## Internal Standards
Row Index Number | Rh (IS1) 103/103 | Rh (IS2) 103/103 | Rh (IS3) 103/103 | Rh (IS4) 103/103 | Rh (IS5) 103/103 | Rh (IS6) 103/103 | Rh OShift (IS1) 103/119
1
2
3 | 100% | 100% | 100% | 100% | 100% | 100% | 100%
4 | 99.46% | 99.5% | 98.99% | 98.96% | 99.02% | 97.28% | 98.84%
5 | 96.71% | 97.54% | 96.66% | 96.06% | 99.78% | 97.35% | 96.72%
6 | 98.13% | 92.78% | 98.04% | 98.58% | 99.13% | 96.3% | 92.28%
7 | 98.47% | 95.89% | 93.43% | 97.42% | 97.81% | 93.98% | 95.72%
8 | 97.03% | 95.25% | 95.49% | 95.17% | 95.56% | 93.62% | 92.09%
9 | 87.4% | 91.43% | 94.1% | 90.52% | 93.86% | 91.41% | 86.68%
10 | 91.65% | 91.28% | 90.53% | 89.4% | 91.31% | 89.1% | 88.07%
11 | 87.69% | 84.4% | 86.12% | 83.56% | 92.98% | 82.26% | 80.14%
12 | 85.02% | 79.93% | 83.63% | 80.44% | 80.75% | 79.72% | 77.01%
13 | 81.4% | 74.98% | 78.53% | 76.12% | 76.6% | 73.82% | 70.48%
15 | 102.9% | 100.2% | 103.7% | 100.8% | 92.28% | 98.09% | 99.95%
16 | 95.56% | 94.59% | 89.02% | 95.99% | 98.16% | 92.08% | 96.41%
17 | 88.2% | 87.61% | 90.95% | 88.43% | 95.41% | 86.78% | 85.58%
18 | 91.96% | 87.46% | 90.05% | 87.41% | 90.23% | 85.93% | 89.14%
19 | 96.74% | 94.33% | 94.85% | 95.68% | 96.16% | 90.36% | 94.25%
20 | 94.29% | 90.06% | 94.61% | 91.67% | 94.59% | 88.07% | 90.02%
21 | 89.3% | 83.99% | 88.41% | 83.1% | 87.21% | 84.01% | 84.41%
22 | 99.38% | 92.53% | 94.36% | 94% | 97.31% | 90.98% | 93.77%
23 | 107.1% | 103.8% | 108.8% | 112.7% | 111.4% | 98.11% | 101.1%
24 | 88.57% | 83.78% | 83.11% | 85.07% | 86.08% | 82.86% | 83.09%
25 | 94.78% | 90.18% | 96.15% | 92.76% | 94.03% | 88.56% | 91.39%
26 | 91.97% | 88.13% | 92.05% | 82.4% | 92.73% | 87.25% | 90.2%
27 | 83.67% | 79.94% | 84.3% | 81.16% | 83.78% | 80.21% | 76.14%
28 | 81.81% | 77.88% | 82% | 77.37% | 81.77% | 79.76% | 76.9%
29 | 81.37% | 76.81% | 80.53% | 77.39% | 79.56% | 69.11% | 77.34%
30 | 88.61% | 83% | 87.74% | 85.13% | 89.22% | 81.6% | 82.16%
31 | 81.42% | 75.68% | 80.92% | 76.64% | 79.86% | 75.82% | 75.03%
32 | 75.9% | 69.82% | 69.7% | 69.47% | 73.23% | 68.88% | 69.17%
33 | 83.22% | 79.02% | 83.93% | 79.34% | 81.91% | 79.23% | 79.69%
34 | 82.4% | 77.74% | 82.35% | 78.66% | 82.28% | 79.25% | 77.31%
35 | 86.27% | 81.95% | 89.6% | 83.17% | 85.6% | 80.24% | 82.36%
36 | 85.61% | 80.25% | 90.93% | 82.14% | 84.6% | 78.3% | 79.73%
37 | 85.88% | 81.63% | 86.3% | 82.83% | 84.84% | 79.32% | 79.74%
38 | 81.72% | 75.72% | 80.3% | 77.33% | 80.42% | 75.71% | 74.68%
39 | 83.52% | 77.58% | 81.94% | 78.91% | 81.81% | 76.17% | 78.78%
40 | 78.69% | 75.06% | 75.54% | 74.57% | 78.54% | 72.26% | 75.88%
41 | 77.13% | 73.95% | 74.45% | 74.16% | 77.13% | 72.86% | 74.44%
42 | 78.39% | 73.7% | 79.34% | 74.04% | 78.14% | 72.97% | 72.62%
43 | 79.4% | 75.04% | 80.44% | 75.97% | 78.95% | 74.3% | 72.83%
44 | 77.16% | 80.91% | 79.25% | 73.9% | 77.36% | 72% | 80.85%
45 | 78.35% | 77.11% | 77.21% | 74.18% | 78.67% | 72.72% | 76.21%
46 | 79.75% | 74.63% | 69.41% | 73.19% | 78.86% | 74.67% | 73.69%
47 | 77.71% | 75.01% | 80.04% | 77.67% | 78.93% | 63.97% | 74.41%
48 | 79.23% | 74.27% | 78.96% | 76% | 78.71% | 72.6% | 71.65%
49 | 72.34% | 65.82% | 70.29% | 66.15% | 69.13% | 64.67% | 66.24%
50 | 87.41% | 81.96% | 87.76% | 86.09% | 86.87% | 79.05% | 85.09%
51 | 86.4% | 80.37% | 84.44% | 83.57% | 85.82% | 76.89% | 81.88%
52 | 86.52% | 80.71% | 87.14% | 83.11% | 85.91% | 79.77% | 82.79%
53 | 82.41% | 75.56% | 83.18% | 77.18% | 80.41% | 74.51% | 75.8%
54 | 83.63% | 76.25% | 84.05% | 77.71% | 80.9% | 73.25% | 77%
55 | 91.58% | 76.22% | 78.94% | 78.44% | 80.23% | 73.93% | 76%
56 | 85.03% | 77.73% | 82.56% | 81.98% | 85.25% | 77.23% | 84.3%
57 | 83.75% | 74.18% | 80.37% | 76.29% | 80.84% | 72.14% | 74.96%
58 | 81.13% | 74.33% | 79.28% | 75.2% | 78.6% | 72.59% | 76.21%
59 | 79.7% | 73.11% | 74.76% | 73.77% | 79.08% | 71.89% | 74.13%
60 | 78.69% | 72.95% | 77.93% | 73.86% | 78.88% | 72.47% | 75.74%
61 | 79.33% | 72.73% | 78.09% | 73.39% | 71.6% | 71.19% | 73.69%
62 | 75.99% | 71.74% | 76.7% | 72.21% | 75.83% | 71.41% | 71.83%
63 | 83.52% | 75.72% | 82.01% | 79.01% | 83.77% | 75.26% | 75.6%
64 | 83.19% | 77.23% | 80.86% | 78.56% | 81.83% | 74.98% | 76.63%
65 | 82.95% | 75.64% | 80.32% | 78.17% | 81.33% | 74.33% | 77.22%
66 | 83.07% | 78.1% | 83.15% | 79.97% | 83.25% | 76.59% | 78.61%
67 | 80.52% | 77.33% | 82.86% | 78.91% | 83.14% | 75.74% | 77.4%
68 | 84.95% | 77.28% | 84.06% | 80.48% | 82.62% | 75.24% | 80.37%
69 | 82.58% | 72.99% | 83.29% | 74.95% | 79.58% | 71.87% | 75.24%
70 | 82.38% | 73.71% | 81.22% | 75.78% | 80.19% | 71.8% | 74.5%
71 | 83.85% | 77.21% | 83.89% | 75.63% | 81.35% | 77.24% | 77.69%
72 | 85.28% | 76.76% | 83.16% | 77.62% | 82.42% | 74.59% | 79.09%
73 | 84.11% | 78.05% | 83.55% | 79.17% | 82.5% | 75.05% | 77.96%
74 | 85.45% | 75.63% | 81.43% | 77.61% | 78.2% | 74.45% | 76.88%
75 | 83.93% | 75.13% | 81.73% | 76.88% | 81.44% | 74.15% | 77.17%
76 | 82.22% | 76.04% | 82.79% | 79.24% | 82.54% | 73.76% | 78.84%
77 | 84.38% | 73.56% | 77.52% | 76.48% | 80.43% | 71.93% | 75.51%
78 | 79.23% | 68.64% | 74.57% | 72.51% | 75.09% | 66.64% | 70.15%
79 | 96.34% | 80.96% | 84.07% | 82.88% | 86.37% | 77.22% | 86.13%
80 | 81.28% | 72.09% | 77.54% | 74.13% | 77.55% | 71.28% | 75.18%
81 | 85.67% | 78.08% | 85.19% | 80.56% | 84.13% | 76.74% | 80.29%
82 | 85.43% | 76.63% | 83.13% | 79.22% | 83.49% | 75.77% | 79.2%
83 | 87.62% | 76.9% | 84.26% | 79.71% | 79.71% | 81.03% | 78.6%
84 | 82.63% | 73.1% | 79.36% | 76.15% | 78.63% | 72.02% | 74.32%
85 | 83.51% | 74.06% | 81.09% | 75.51% | 79.25% | 72.62% | 75.56%
86 | 83.26% | 73.61% | 82.67% | 76.59% | 79.12% | 72.31% | 76.52%
87 | 80.98% | 72.1% | 79.33% | 73.46% | 76.4% | 72.25% | 73.91%
88 | 80.31% | 71.87% | 76.72% | 72.9% | 75.73% | 72.02% | 72.98%
89 | 81.13% | 71.85% | 78.27% | 73.29% | 75.67% | 71.38% | 73.63%
90 | 80.51% | 72.32% | 76.2% | 72.48% | 74.62% | 71.52% | 73.79%
91 | 78.84% | 71.54% | 78.4% | 72.24% | 74.68% | 70.71% | 73.49%
92 | 79.74% | 68.56% | 77.83% | 71.85% | 75.65% | 66.94% | 70.36%
93 | 79.68% | 71.76% | 75.48% | 71.8% | 75.62% | 72.01% | 71.69%
94 | 78.88% | 70.91% | 78.83% | 72.16% | 75.25% | 71.24% | 72.79%
95 | 79.27% | 71.76% | 76.26% | 71.95% | 76.41% | 70.31% | 75.11%
96 | 72.73% | 63.69% | 67.84% | 64.36% | 74.27% | 62.38% | 63.93%
97 | 84.29% | 80.77% | 83.44% | 84.6% | 86.43% | 79.26% | 86.39%
98 | 88.12% | 79.46% | 84.84% | 81.42% | 85.03% | 77.07% | 82.75%
99 | 88.22% | 79.87% | 85.61% | 81.74% | 85.38% | 77.36% | 83.14%
100 | 84.9% | 75.42% | 81.01% | 76.37% | 79.11% | 72.69% | 77.56%
101 | 84.93% | 74.76% | 82.61% | 76.88% | 80.24% | 74.11% | 76.62%
102 | 78.09% | 72.02% | 78.26% | 72.93% | 76.54% | 72.31% | 75.18%
103 | 81.58% | 71.29% | 74.12% | 71.12% | 75.25% | 69.57% | 74.4%
104 | 82.36% | 72.11% | 75.84% | 74.64% | 76.83% | 72.39% | 76.02%
105 | 81.15% | 71.17% | 75.96% | 72.61% | 77.88% | 70.33% | 74.4%
106 | 81.25% | 72.34% | 76.62% | 72.92% | 76.78% | 68.86% | 76.33%
107 | 81.33% | 70.37% | 75.95% | 72.86% | 75.47% | 70.21% | 72.91%
108 | 81% | 73.9% | 77.04% | 74.87% | 77.69% | 72.62% | 76.75%
109 | 80.32% | 70.76% | 76.67% | 71.7% | 76.04% | 71.62% | 74.89%
110 | 84.24% | 75.52% | 83.36% | 70.63% | 80% | 71.32% | 76.77%
111 | 88.3% | 75.67% | 80.6% | 77.78% | 80.89% | 74.44% | 79.05%
112 | 85.85% | 77.29% | 86.75% | 79.27% | 80.4% | 75.99% | 79.6%
113 | 88.8% | 74.55% | 81.81% | 79.24% | 82.36% | 72.64% | 75.09%
114 | 86.13% | 75.69% | 82.02% | 79.8% | 78.62% | 73.63% | 77.78%
115 | 84.86% | 70.85% | 76.43% | 73.71% | 75.98% | 70.25% | 74.42%
116 | 84.69% | 72.31% | 77.51% | 74.78% | 77.45% | 72.3% | 76.2%
117 | 84.3% | 71.72% | 78.96% | 75.01% | 77.43% | 70.24% | 77.18%
118 | 87.28% | 74.94% | 83.28% | 77.88% | 80.88% | 72.97% | 76.92%
119 | 85.32% | 73.26% | 84.08% | 76.14% | 79.17% | 72.49% | 76.97%
120 | 84.54% | 74.3% | 79.62% | 67.2% | 79.12% | 72.3% | 78.43%
121 | 83.97% | 72.98% | 81.03% | 76.43% | 78.02% | 71.69% | 74.45%
122 | 83.97% | 72.83% | 78.44% | 75.76% | 78.4% | 71.66% | 79.22%
123 | 83.91% | 71.96% | 78.49% | 75.99% | 74.61% | 71.94% | 75.75%
124 | 85.06% | 74.36% | 81.66% | 78.32% | 78.76% | 71.89% | 79.7%
125 | 84.14% | 71.44% | 79.83% | 73.09% | 76.77% | 70.81% | 74.5%
126 | 74.93% | 65.18% | 72.38% | 69.51% | 71.63% | 64.34% | 70.95%
127 | 91.66% | 76.59% | 83.05% | 82.23% | 81.78% | 73.63% | 81.57%
128 | 88.57% | 75.15% | 82.06% | 84.53% | 82.5% | 73.8% | 79.73%
129 | 84.34% | 75.66% | 82.42% | 78.23% | 80.5% | 72.19% | 83.72%
130 | 87.33% | 73.82% | 83.14% | 79.11% | 80.11% | 73.23% | 80.6%
131 | 84.44% | 69.31% | 77.46% | 73.94% | 75.36% | 69.45% | 74.44%
132 | 83.93% | 71.58% | 77.4% | 74.06% | 76.78% | 66.89% | 74.34%
133 | 89.67% | 78.2% | 84.79% | 83.48% | 85.23% | 75.7% | 86.38%
134 | 89.64% | 76.49% | 86.53% | 81.07% | 83.04% | 73.75% | 82.34%
135 | 89.66% | 76.79% | 83.93% | 80.58% | 83.29% | 73.4% | 80.04%
136 | 88.62% | 75.73% | 84.74% | 80.68% | 84.47% | 73.19% | 83.17%
137 | 88.81% | 75.6% | 81.73% | 79.93% | 82.04% | 69.9% | 79.86%
138 | 87.7% | 74.73% | 85.5% | 76.68% | 82.31% | 72.3% | 80.9%
139 | 89.42% | 76.09% | 84.63% | 80.47% | 79.06% | 73.59% | 80.9%
140 | 89.35% | 75.31% | 79.95% | 80.86% | 83.34% | 74.13% | 81.4%
141 | 87.18% | 74.06% | 82.2% | 81.97% | 77.46% | 83.34% | 80.6%
142 | 86.79% | 72.71% | 81.65% | 79.3% | 79.96% | 70.65% | 78.06%
143 | 87.21% | 68.2% | 84.68% | 82.43% | 81.76% | 73.2% | 71.73%
144 | 85.82% | 72.58% | 80.86% | 76.85% | 78.52% | 71.26% | 80.68%
145 | 88.6% | 72.51% | 83.52% | 76.55% | 79.08% | 70.34% | 80.6%
146 | 82.89% | 67.17% | 76.28% | 71.85% | 73.86% | 66.55% | 71.39%
147 | 82.78% | 69.07% | 77% | 72.89% | 74.25% | 69.35% | 74.53%
148 | 85.76% | 69.23% | 77.78% | 72.55% | 74.68% | 68.82% | 74.03%
149 | 84.56% | 70.17% | 77.06% | 78.69% | 77.03% | 68.88% | 73.87%
150 | 87.27% | 75.8% | 82.51% | 81.02% | 81.4% | 73.34% | 81.8%
151 | 77.97% | 66.55% | 74.66% | 70.16% | 73.05% | 69.62% | 70.38%
152 | 86.93% | 76.41% | 81.63% | 79.06% | 79.54% | 72.72% | 79.91%
153 | 87.51% | 74.66% | 80.67% | 78.8% | 80.79% | 70.79% | 81.32%
154 | 86.64% | 74.58% | 80.85% | 77.86% | 78.57% | 72.79% | 78.63%
155 | 87.7% | 75.13% | 82.86% | 79.61% | 80.97% | 72.96% | 82.72%
156 | 87.63% | 74.97% | 81.12% | 79.84% | 79.99% | 72.76% | 80.96%
157 | 86.74% | 73.35% | 81.35% | 74.85% | 82.21% | 71.78% | 79.16%
158 | 87.19% | 71.03% | 78.44% | 69.75% | 76.22% | 64.33% | 74.54%
159 | 82.68% | 69.68% | 75.99% | 72.84% | 74.97% | 67.03% | 75.37%
160 | 85.94% | 74.52% | 81.59% | 79.6% | 77.08% | 71.74% | 80.86%
161 | 88.39% | 73.37% | 77.01% | 78.45% | 79.2% | 72.65% | 80.37%
162 | 86.4% | 70.84% | 81.22% | 75.38% | 78.45% | 71.86% | 74.84%
163 | 81.81% | 68.72% | 76.71% | 72.71% | 60.91% | 68.12% | 72.7%
164 | 81.89% | 68.26% | 76.6% | 72.22% | 74.51% | 67.44% | 72.86%
165 | 78.83% | 64.82% | 73.23% | 68.12% | 68.83% | 64.83% | 68.8%
166 | 75.09% | 58.15% | 69.78% | 63.71% | 65.7% | 60.31% | 60.94%
167 | 74.73% | 61.18% | 64.44% | 63.91% | 65.02% | 60.17% | 64.65%
168 | 67.03% | 55.16% | 61.17% | 56.92% | 59.11% | 54.45% | 57.49%
169 | 68.64% | 55.86% | 60.32% | 57.32% | 59.31% | 54.65% | 59.54%
170 | 87.23% | 73.36% | 79.98% | 77.65% | 79.28% | 72.23% | 80.04%
171 | 83.11% | 68.83% | 73.01% | 71.8% | 74.61% | 67.94% | 74.45%
172 | 81.57% | 70.52% | 78.34% | 73.51% | 74.88% | 68.99% | 77.98%
173 | 78.52% | 64.29% | 70.55% | 66.76% | 68.52% | 63.11% | 68.37%
174 | 85.99% | 69.67% | 77.6% | 75.02% | 76.37% | 67.46% | 74.03%
175 | 80.87% | 70.58% | 79.6% | 75.58% | 77.05% | 69.25% | 76.5%
176 | 86.32% | 73.62% | 78.62% | 72.31% | 75.2% | 66.22% | 80.06%
177 | 91.81% | 74.5% | 89.85% | 84.11% | 85.48% | 67.7% | 79.78%
178 | 84.03% | 65.53% | 73.58% | 69.33% | 70.98% | 64.61% | 70.77%
179 | 80.51% | 65.68% | 73.13% | 68.88% | 71.45% | 62.84% | 71.16%
180 | 82.83% | 66.02% | 71.09% | 70.1% | 71.7% | 64.48% | 71.18%
181 | 81.6% | 66.23% | 72.4% | 69.62% | 67.66% | 65.16% | 70.19%
182 | 82.96% | 68.19% | 74.83% | 71.99% | 70.48% | 66.45% | 73.88%
183 | 83.63% | 62.19% | 75.01% | 71.73% | 75.86% | 66.25% | 65.53%
184 | 80.61% | 66.61% | 70.56% | 69.79% | 63.66% | 64.8% | 72.73%
185 | 81.82% | 66.57% | 74.88% | 70.18% | 74.24% | 65.38% | 72.11%
186 | 83.72% | 66.38% | 70.6% | 68.93% | 71.36% | 68.32% | 71.36%
187 | 76.7% | 59.98% | 62.48% | 63.08% | 64.6% | 58.64% | 65.69%
188 | 76.04% | 60.61% | 66.26% | 64.13% | 65.73% | 59.9% | 64.6%
189 | 81.34% | 65.75% | 70.13% | 67.02% | 71.55% | 52.74% | 72.69%
190 | 79.07% | 63.78% | 73.13% | 67.84% | 70.28% | 62.36% | 69.25%
191 | 77.51% | 62.51% | 69.44% | 68.22% | 68.52% | 61.24% | 67.01%
192 | 81.11% | 66.86% | 74.61% | 74.78% | 72.67% | 61.11% | 73.79%
193 | 77.24% | 65.05% | 73.54% | 68.94% | 69.99% | 62.75% | 71.3%
194 | 79.95% | 64.03% | 72.83% | 67.95% | 69.95% | 62.57% | 69.08%
195 | 76.5% | 60.91% | 66.04% | 65.35% | 66.32% | 59.03% | 67.24%
196 | 77.86% | 61.65% | 70.2% | 65.34% | 67.53% | 60.04% | 65.78%
197 | 78.85% | 63.81% | 72.88% | 68.37% | 71.61% | 63.73% | 68.72%
198 | 79.54% | 62.38% | 69.42% | 67.23% | 68.5% | 62.1% | 66%
199 | 78.56% | 61.93% | 71.69% | 66% | 66.2% | 60.27% | 67.63%
200 | 75.23% | 59.18% | 65.74% | 63.28% | 65.05% | 59.99% | 67.7%
201 | 74.65% | 60.57% | 69.78% | 62.69% | 64.91% | 60.02% | 65.75%
202 | 69.54% | 60.69% | 67.16% | 63.29% | 66.22% | 59.43% | 65.19%
203 | 74.41% | 60.17% | 67.53% | 62.93% | 65.38% | 59.07% | 65.11%
204 | 73.67% | 60.41% | 68.62% | 62.3% | 55.02% | 59.39% | 64.83%
205 | 74.5% | 60.06% | 68.36% | 61.39% | 65.33% | 60.46% | 65.28%
206 | 73.74% | 55.67% | 68.63% | 61.82% | 64.27% | 59.12% | 61%
207 | 73.75% | 60.95% | 66.15% | 62.27% | 65.57% | 58.39% | 65.08%
208 | 78.28% | 64% | 71.19% | 64.15% | 68.39% | 61.56% | 70.86%
209 | 78.7% | 63.42% | 70.25% | 67.2% | 69.34% | 60.44% | 68.81%
210 | 79.66% | 63.97% | 70.52% | 67.27% | 69.7% | 61.8% | 70.66%
211 | 75.43% | 58.71% | 68.28% | 62.56% | 64.34% | 58.13% | 66.21%
212 | 75.3% | 59.84% | 69.47% | 63.35% | 65.19% | 56.98% | 64.4%
213 | 78.2% | 59.96% | 69.85% | 65.14% | 66.94% | 59.47% | 66.8%
214
215
216 | 100% | 100% | 100% | 100% | 100% | 100% | 100%
217 | 99.73% | 99.78% | 101.4% | 101% | 100.1% | 101% | 99.07%
218 | 100.4% | 98.87% | 98.01% | 100.3% | 100.1% | 101% | 100.4%
219 | 96.34% | 97.39% | 98.29% | 98.56% | 98.37% | 96.57% | 100.6%
220 | 97.53% | 98.18% | 99.78% | 98.22% | 84.07% | 97.44% | 100.8%
221 | 98.97% | 95.61% | 98.46% | 96.89% | 97.01% | 95.76% | 96.52%
222 | 95.92% | 92.23% | 95.64% | 92.68% | 92.41% | 94.3% | 92.07%
223 | 92.95% | 90.12% | 91.28% | 90.49% | 89.93% | 91.84% | 90.63%
224 | 82.57% | 86.35% | 89.71% | 86.35% | 86.35% | 88.01% | 88.08%
225 | 87.36% | 83.91% | 84.54% | 82.75% | 76.52% | 83.74% | 80.6%
226 | 80.09% | 76.37% | 77.75% | 73.38% | 75.62% | 78.48% | 75.21%
228 | 93.48% | 100.1% | 100.8% | 102.9% | 102.2% | 105.6% | 104.3%
229 | 94.94% | 95.75% | 97.84% | 94.91% | 97.26% | 94.72% | 99.05%
230 | 99.61% | 95.68% | 97.3% | 97.23% | 98.89% | 84.65% | 97.67%
231 | 96.09% | 91.25% | 87.83% | 90.87% | 91.34% | 91.23% | 92.26%
232 | 94.95% | 90.59% | 90.59% | 82.69% | 94.66% | 91.91% | 91.64%
233 | 99.4% | 96.96% | 102.2% | 96.39% | 96.78% | 96.31% | 99.47%
234 | 97.93% | 94.74% | 98.44% | 94.27% | 95.5% | 95.27% | 96.78%
235 | 94.3% | 89.18% | 92.98% | 89.97% | 90.47% | 92.42% | 92.17%
236 | 97.94% | 95.36% | 103.2% | 97.37% | 96.99% | 96.06% | 95.85%
237 | 93.59% | 88.01% | 88.92% | 93.93% | 88.94% | 90.2% | 87.55%
238 | 93.76% | 88.24% | 92.31% | 89.01% | 90.21% | 90.74% | 89.46%
239 | 97.83% | 95.77% | 97.96% | 95.98% | 97.76% | 97.33% | 96.95%
240 | 89.62% | 86.42% | 91.44% | 86.23% | 87.98% | 89.2% | 85.45%
241 | 88.04% | 86.25% | 88.37% | 86.77% | 84.62% | 89.9% | 86.88%
242 | 88.9% | 86.97% | 88.18% | 85.33% | 92.91% | 87.37% | 89.32%
243 | 88.05% | 86.76% | 87.28% | 86.34% | 89.97% | 90.92% | 85.55%
244 | 89.67% | 85.56% | 87.57% | 89.03% | 86.24% | 85.8% | 85.31%
245 | 89.77% | 86.53% | 88.41% | 85.44% | 87.93% | 89.35% | 86.4%
246 | 90.49% | 85.92% | 86.44% | 84.72% | 86.58% | 92.66% | 87.54%
247 | 89.92% | 86.85% | 87.73% | 80.71% | 85.95% | 84.16% | 85.89%
248 | 89.66% | 85.8% | 88.38% | 84.62% | 85.51% | 87.34% | 88.44%
249 | 90.53% | 85.15% | 87.69% | 84.04% | 86.18% | 87.68% | 85.59%
250 | 90.53% | 92.37% | 99.93% | 92.75% | 93.67% | 91.76% | 95.18%
251 | 95.42% | 91.58% | 92.72% | 90.03% | 85.05% | 91.66% | 95.23%
252 | 93.91% | 89.22% | 92.56% | 88.15% | 90.05% | 90.55% | 88.47%
253 | 91.38% | 76.63% | 87.42% | 84.92% | 86.81% | 86.51% | 78.54%
254 | 94.68% | 86.93% | 88.77% | 88.2% | 87.97% | 113.8% | 86.77%
255 | 90.66% | 88.21% | 86.65% | 82.57% | 86.44% | 85.9% | 88.5%
256 | 88.1% | 83.28% | 85.68% | 79.91% | 82.81% | 85.74% | 84.04%
257 | 88.03% | 83.79% | 87.56% | 82.88% | 84.79% | 85.67% | 84.04%
258 | 93.75% | 91.17% | 92.24% | 94.94% | 91.52% | 91.96% | 90.71%
259 | 92.57% | 89.42% | 91.45% | 90.34% | 89.75% | 88.11% | 91.03%
260 | 90.71% | 88.13% | 91.64% | 90.05% | 90.09% | 88.47% | 88.58%
261 | 92.41% | 89.2% | 92.1% | 84.07% | 89.48% | 88.32% | 88.17%
262 | 92.9% | 86.59% | 90.66% | 85.72% | 90.19% | 85.6% | 88.33%
263 | 90.47% | 85.59% | 90.77% | 85.35% | 86.25% | 86.27% | 85.41%
264 | 90.37% | 85.84% | 86.79% | 76.61% | 88.09% | 84.38% | 88.02%
265 | 90.68% | 86.81% | 91.48% | 89.17% | 85.63% | 87.13% | 89.26%
266 | 90.64% | 85.77% | 90.1% | 86.62% | 86.48% | 86.07% | 88.78%
267 | 91.82% | 86.73% | 91.51% | 86.44% | 87.87% | 85.33% | 88.5%
268 | 90.98% | 81.99% | 86.51% | 83.61% | 85.07% | 81.96% | 83.14%
269 | 91.13% | 84.78% | 87.87% | 80.8% | 85.92% | 84.74% | 85.97%
270 | 91.84% | 84.24% | 90.44% | 85.32% | 86.31% | 86.15% | 86.85%
271 | 84.77% | 86.09% | 92.1% | 86.87% | 88.04% | 87.15% | 87.68%
272 | 90.6% | 83.1% | 89.15% | 85.04% | 86.4% | 86.21% | 84.83%
273 | 90.39% | 83.58% | 89.09% | 84.78% | 85.69% | 84.4% | 85.61%
274 | 90.47% | 82.63% | 87.35% | 83.39% | 85.32% | 83.21% | 83.23%
275 | 82.52% | 81.53% | 86.01% | 82.69% | 84.22% | 84.03% | 82.36%
276 | 89.26% | 81.93% | 82.25% | 83.37% | 83.75% | 84.75% | 86.09%
277 | 87.67% | 82.3% | 82.29% | 81.97% | 83.03% | 84.27% | 84.32%
278 | 87.62% | 82.68% | 82.95% | 81.38% | 82.13% | 85.86% | 85.61%
279 | 85.79% | 79.97% | 84.07% | 77.78% | 81.22% | 81.88% | 81.35%
280 | 86.75% | 79.91% | 83% | 76.35% | 80.45% | 79.93% | 81.19%
281 | 84.47% | 78.61% | 82.92% | 78.68% | 81.52% | 80.36% | 80.17%
282 | 86.37% | 79.78% | 85.25% | 79.44% | 81.22% | 81.93% | 79.7%
283 | 89.92% | 86.31% | 88.49% | 87.12% | 86.62% | 84.91% | 85.72%
284 | 90.89% | 84.8% | 87.41% | 85.84% | 76.51% | 85.46% | 83.62%
285 | 90.78% | 84.41% | 85.74% | 86.39% | 77.73% | 85.08% | 88.25%
286 | 88.8% | 74.18% | 84.58% | 81.52% | 82.34% | 83.03% | 74.64%
287 | 88.17% | 80.45% | 87.35% | 80.9% | 89.66% | 85.46% | 81.4%
288 | 87.35% | 78.5% | 83.2% | 78.62% | 79.99% | 80.23% | 79.2%
289 | 85.73% | 72.76% | 80.51% | 77.71% | 77.58% | 77.79% | 75.68%
290 | 77.6% | 78.69% | 82.25% | 78.62% | 80.14% | 86.09% | 79.13%
291 | 84.93% | 81.58% | 82.97% | 79.1% | 80.34% | 81.69% | 81.78%
292 | 84.83% | 70.21% | 82.99% | 79.69% | 79.75% | 81.71% | 70.61%
293 | 87.77% | 79.43% | 79.61% | 78.4% | 80.48% | 81.03% | 79.42%

## QC
QC Standards
QC STD 1
|  | Sample Id | R | Acquisition Time | Na KED 23/23 | Mg KED 24/24 | Al 27/27 | Al KED 27/27 | P Oshift 31/47 | S (34) -1 34/34 | K KED 39/39 | Ca KED 43/43 | V KED 51/51 | Cr (52) 52/52 | Cr (53) NH3 53/53 | Fe (54) NH3 54/54 | Mn Oshift 55/71 | Fe (56) KED 56/56 | Co KED 59/59 | Ni KED 60/60 | Cu (63) KED 63/63 | Cu (65) NH3 65/65 | Zn NH3 66/66 | As 75/75 | As KED 75/75 | As RH Oshift 75/91 | As Oshift 75/91 | Se (77) 77/77 | Se (77) Oshift 77/93 | Se (78) Oshift 78/94 | Se (80) Oshift 80/96 | Se (82) 82/82 | Se (82) Oshift 82/98 | Mo (95) NH3 95/95 | Mo (95) OShift 95/127 | Rh (IS1) 103/103 (IS) | Rh (IS2) 103/103 (IS) | Rh (IS3) 103/103 (IS) | Rh (IS4) 103/103 (IS) | Rh (IS5) 103/103 (IS) | Rh (IS6) 103/103 (IS) | Rh OShift (IS1) 103/119 (IS) | Cd (111) NH3 111/111 | Cd (114) NH3 114/114 | Pb 208/208 | Pb NH3 208/208
| 15 | LBLANK |  | 06/08/2023 14:38:52 | -11.7416 | -41.0609 | -0.69918 | -0.76605 | -0.23825 | 228.7424 | 3.2182 | -30.7443 | 0.0068477 | 0.005656 | 0.0027179 | -0.26883 | -0.024393 | -0.17974 | 0.002683 | 0.022413 | -0.004841 | -0.0062042 | -0.11381 | -0.041553 | 0.0084722 | 0.16761 | 0.0085315 | 0.15939 | -0.018378 | 0.0039245 | 0.0059578 | 0.025924 | 0.00737 | 0.026908 | 8.4559e-05 | [7 empty] | 0.0015638 | 0.0006736 | -0.0017921 | -0.0013804
|  | RSD |  |  | 9.396% | 3.545% | 0.03034% | 3.325% | 1.943% | 60.37% | 245.2% | 1.343e-06% | 47.25% | 21.69% | 52.88% | 2.043% | 5.178% | 17.3% | 46.33% | 21.09% | 81.76% | 3.692% | 1.725% | 70.82% | 164.6% | 0.6165% | 18.78% | 58.66% | 32.79% | 48.2% | 52.67% | 20.93% | 119.2% | 7.571% | 1299% | [7 empty] | 39.5% | 22.63% | 0.539% | 1.52%
|  | % Recovery | [33 empty] | 102.9% | 100.2% | 103.7% | 100.8% | 92.28% | 98.09% | 99.95%
| 35 | LBLANK |  | 06/08/2023 16:49:31 | -10.0872 | -42.7414 | -0.6818 | -0.77198 | -2.1827 | 321.7818 | 2.5662 | -28.4866 | 0.0045481 | 0.0012304 | 0.0030117 | -0.50857 | -0.025009 | -0.61299 | 0.0020636 | 0.013225 | -0.0081338 | -0.0073653 | -0.16558 | 0.015361 | 0.021366 | 0.18901 | 0.032126 | 0.09901 | -0.00054305 | 0.014658 | 0.0074398 | 0.010049 | 0.012979 | -0.0097971 | -0.011526 | [7 empty] | 0.00024118 | -0.00038652 | -0.0023727 | -0.0017448
|  | RSD |  |  | 26.69% | 2.367e-06% | 0.4326% | 6.284% | 1.48% | 36.04% | 468.7% | 13.73% | 2.97% | 278.5% | 40.48% | 1.069% | 1.059% | 13.36% | 51.16% | 63.16% | 21.47% | 16.77% | 4.424% | 204.8% | 119.5% | 1.71% | 9.078% | 88.62% | 1342% | 48.76% | 39.8% | 406.4% | 48.76% | 9.044% | 7.349% | [7 empty] | 110.6% | 34.45% | 3.21% | 6.163%
|  | % Recovery | [33 empty] | 86.27% | 81.95% | 89.6% | 83.17% | 85.6% | 80.24% | 82.36%
| 50 | LBLANK |  | 06/08/2023 18:19:16 | -6.8622 | -38.834 | -0.69722 | -0.77278 | -1.8744 | 551.2187 | 6.2902 | -30.7443 | 0.0034021 | 0.0031958 | 0.0038855 | -0.51479 | -0.024432 | -0.56598 | 0.0020661 | 0.017722 | -0.0067889 | -0.0060706 | -0.16573 | -0.022463 | 0.0055661 | 0.1642 | 0.0051691 | 0.17826 | -0.021858 | 0.0011072 | -0.00032446 | 0.010958 | 0.009361 | -0.0063377 | -0.0099329 | [7 empty] | 0.00022846 | -0.00016505 | -0.0018503 | -0.0016199
|  | RSD |  |  | 37.58% | 4.333% | 0.08046% | 6.288% | 1.602% | 21.77% | 185.4% | 1.343e-06% | 29.8% | 23.12% | 46.7% | 0.9067% | 2.148% | 7.656% | 29.65% | 47.61% | 30.89% | 10.38% | 1.264% | 115.8% | 175.9% | 0.6536% | 21.33% | 60.7% | 0% | 211.6% | 690.9% | 254.8% | 1.054% | 17.55% | 3.078% | [7 empty] | 109.5% | 239% | 4.751% | 6.632%
|  | % Recovery | [33 empty] | 87.41% | 81.96% | 87.76% | 86.09% | 86.87% | 79.05% | 85.09%
| 66 | LBLANK |  | 06/08/2023 20:24:33 | -13.4076 | -42.7414 | -0.68761 | -0.75313 | -2.0219 | 429.4725 | 9.1767 | -30.7443 | 0.0040659 | 0.0044984 | 0.0028794 | -0.51172 | -0.025153 | -0.53747 | 0.0018928 | -0.0020792 | -0.0090217 | -0.0066424 | -0.16337 | 0.043694 | 0.00034057 | 0.159 | -0.00088352 | 0.15064 | -0.017448 | 0.0012374 | 0.0018458 | 0.016555 | -0.0014165 | -0.011277 | -0.013054 | [7 empty] | 0.00024654 | -0.00035722 | -0.0022487 | -0.0019217
|  | RSD |  |  | 5.704% | 2.367e-06% | 0.8827% | 7.858% | 0.5153% | 33.21% | 66.34% | 1.343e-06% | 53.21% | 37.82% | 7.727% | 1.973% | 1.455% | 13.45% | 40.14% | 275.4% | 41.37% | 9.844% | 1.936% | 18.75% | 3093% | 0.7996% | 149.1% | 22.16% | 43.78% | 198.5% | 165% | 374.1% | 0% | 6.011% | 3.778% | [7 empty] | 108.8% | 77.68% | 13.71% | 2.259%
|  | % Recovery | [33 empty] | 83.07% | 78.1% | 83.15% | 79.97% | 83.25% | 76.59% | 78.61%
| 81 | LBLANK |  | 06/08/2023 21:52:35 | -12.1461 | -42.7414 | -0.68933 | -0.75516 | -1.5152 | 289.9104 | 3.389 | -26.044 | 0.0040317 | 0.0038581 | 0.0037226 | -0.50852 | -0.025251 | -0.57173 | 0.0016624 | 0.0070594 | -0.008451 | -0.0066872 | -0.16251 | 0.057302 | 3.4913e-05 | 0.16914 | 0.010652 | 0.16521 | -0.012874 | 0.0054712 | -0.0001752 | -0.021405 | 0.0098861 | -0.0096086 | -0.010268 | [7 empty] | -2.3879e-05 | -0.00067103 | -0.0024194 | -0.0019407
|  | RSD |  |  | 12.15% | 2.367e-06% | 0.435% | 10.06% | 0.9292% | 45.56% | 84.01% | 15.64% | 26.11% | 89.41% | 34.88% | 0.3115% | 1.193% | 3.925% | 21.71% | 23.36% | 6.993% | 10.28% | 2.406% | 8.043% | 28659% | 0.4781% | 11.38% | 35.14% | 60.43% | 118.3% | 1772% | 533% | 113.9% | 2.637% | 1.585% | [7 empty] | 0% | 25.44% | 4.191% | 2.895%
|  | % Recovery | [33 empty] | 85.67% | 78.08% | 85.19% | 80.56% | 84.13% | 76.74% | 80.29%
| 97 | LBLANK |  | 06/08/2023 23:26:52 | -9.9378 | -41.697 | -0.69387 | -0.78747 | -1.8712 | 699.2796 | 7.3142 | -28.3815 | 0.00343 | 0.0018182 | 0.0030687 | -0.50681 | -0.024898 | -0.48754 | 0.0023115 | 0.012101 | -0.0051764 | -0.0076033 | -0.16429 | -0.0042435 | 0.00017371 | 0.1647 | 0.0060638 | 0.49126 | -0.013193 | 0.0038676 | 0.0043455 | 0.1072 | 0.0096492 | -0.00069744 | -0.0066249 | [7 empty] | 0.00014763 | -7.1062e-05 | -0.0019687 | -0.0019592
|  | RSD |  |  | 26.67% | 4.338% | 1.185% | 7.465% | 2.553% | 22.63% | 32.99% | 14.42% | 30.97% | 154% | 37.32% | 0.4176% | 1.601% | 6.832% | 15.84% | 48.63% | 72.29% | 3.381% | 1.444% | 994.4% | 5898% | 0.5689% | 19.1% | 37.94% | 56.9% | 58.77% | 48.55% | 87.95% | 115.4% | 254.5% | 10.65% | [7 empty] | 201.2% | 438.8% | 12.21% | 6.509%
|  | % Recovery | [33 empty] | 84.29% | 80.77% | 83.44% | 84.6% | 86.43% | 79.26% | 86.39%
| 112 | LBLANK |  | 06/09/2023 00:54:53 | -12.6241 | -42.7414 | -0.70214 | -0.75557 | -1.8183 | 384.997 | 0.44467 | -30.7443 | 0.0028168 | 0.001162 | 0.0027566 | -0.50728 | -0.025033 | -0.58416 | 0.0016626 | -6.3058e-06 | -0.0079475 | -0.0067071 | -0.16464 | 0.037153 | -0.0057419 | 0.15754 | -0.002414 | 0.19771 | -0.021858 | 0.0012851 | -0.00013577 | -0.05994 | 0.010051 | -0.01102 | -0.011667 | [7 empty] | 6.7726e-05 | -0.00046338 | -0.0020734 | -0.0018346
|  | RSD |  |  | 16.81% | 2.367e-06% | 0.2235% | 3.845% | 0.6843% | 47.1% | 1605% | 1.343e-06% | 0% | 110.8% | 18.08% | 1.911% | 0.9897% | 5.113% | 21.72% | 76495% | 33.17% | 4.954% | 3.209% | 63.76% | 1.756e-06% | 0.4787% | 36.09% | 90.11% | 0% | 194.3% | 3160% | 87.22% | 113.7% | 2.687% | 4.589% | [7 empty] | 234.3% | 53.46% | 13.47% | 6.586%
|  | % Recovery | [33 empty] | 85.85% | 77.29% | 86.75% | 79.27% | 80.4% | 75.99% | 79.6%
| 128 | LBLANK |  | 06/09/2023 02:28:44 | -12.9698 | -42.7414 | -0.70369 | -0.80379 | -1.1278 | 455.5397 | 3.1372 | -25.9281 | 0.0034437 | 0.0050085 | 0.002168 | -0.51708 | -0.025263 | -0.53706 | 0.0014541 | -0.00027793 | -0.0085932 | -0.0077548 | -0.1653 | 0.034399 | -0.0057419 | 0.15719 | -0.0027479 | 0.27724 | -0.012592 | 0.0013548 | 0.0027863 | -0.037514 | -0.0014165 | -0.010916 | -0.012314 | [7 empty] | -2.3879e-05 | -0.00076703 | -0.0021525 | -0.0019502
|  | RSD |  |  | 5.874% | 2.367e-06% | 0.004685% | 3.797% | 2.125% | 43.53% | 215.9% | 16.09% | 31.53% | 53.22% | 49.81% | 1.053% | 2.341% | 3.732% | 0% | 895.1% | 39.14% | 9.129% | 4.35% | 66.47% | 1.756e-06% | 0.3225% | 23.35% | 54.61% | 127.5% | 377.4% | 75% | 217.9% | 0% | 3.237% | 0.8123% | [7 empty] | 0% | 38.34% | 7.585% | 5.859%
|  | % Recovery | [33 empty] | 88.57% | 75.15% | 82.06% | 84.53% | 82.5% | 73.8% | 79.73%
| 143 | LBLANK |  | 06/09/2023 03:57:12 | -12.5767 | -42.7414 | -0.70634 | -0.77097 | -1.131 | 313.9096 | 7.0578 | -28.3955 | 0.0034244 | 0.002147 | 0.0027951 | -0.51161 | -0.025165 | -0.56836 | 0.0014541 | 0.0024533 | -0.0027638 | -0.0070497 | -0.16616 | 0.051529 | -0.0057419 | 0.15721 | -0.0027403 | 0.18241 | -0.012245 | 0.0028582 | -0.00033899 | -0.03785 | -0.0014165 | -0.011642 | -0.013073 | [7 empty] | 0.0002271 | -0.00080048 | -0.0021612 | -0.0019083
|  | RSD |  |  | 10.03% | 2.367e-06% | 0.1306% | 11.33% | 2.465% | 34.92% | 68.5% | 14.33% | 30.74% | 32.06% | 22.02% | 1.893% | 0.5321% | 13.12% | 0% | 268.9% | 100.7% | 6.673% | 0.1776% | 54.93% | 1.756e-06% | 0.243% | 16.03% | 54.31% | 68.05% | 270% | 839.9% | 143.6% | 0% | 2.56% | 0.2545% | [7 empty] | 191.4% | 26.57% | 7.004% | 9.276%
|  | % Recovery | [33 empty] | 87.21% | 68.2% | 84.68% | 82.43% | 81.76% | 73.2% | 71.73%
| 160 | LBLANK |  | 06/09/2023 05:37:19 | -12.9681 | -41.683 | -0.70315 | -0.76901 | -1.8662 | 442.0872 | -1.2013 | -30.7443 | 0.0034565 | 0.0032836 | 0.0039677 | -0.50335 | -0.025041 | -0.51864 | 0.0025447 | 0.0029418 | -0.0078968 | -0.0070702 | -0.1673 | 0.042346 | 0.0066003 | 0.1583 | -0.0013996 | 0.35019 | -0.021858 | -0.00012043 | -0.0014366 | 0.00057204 | -0.0014165 | 0.0058947 | -0.010263 | [7 empty] | 0.00016021 | -0.00085873 | -0.0020506 | -0.0018641
|  | RSD |  |  | 5.887% | 4.398% | 0.08874% | 0.1209% | 1.321% | 27.98% | 808.3% | 1.343e-06% | 32.06% | 63.44% | 39.93% | 0.883% | 1.067% | 3.433% | 38.95% | 166.1% | 24.88% | 16.34% | 1.32% | 83.98% | 323.9% | 0.6162% | 79.5% | 30.72% | 0% | 2124% | 148% | 8275% | 0% | 17.64% | 7.384% | [7 empty] | 199% | 13.69% | 18.21% | 8.247%
|  | % Recovery | [33 empty] | 85.94% | 74.52% | 81.59% | 79.6% | 77.08% | 71.74% | 80.86%
| 175 | LBLANK |  | 06/09/2023 07:06:16 | -12.9562 | -41.6802 | -0.69969 | -0.75005 | -1.5046 | 495.6731 | 10.0863 | -30.7443 | 0.0041213 | 0.0038698 | 0.0020448 | -0.50349 | -0.024856 | -0.49308 | 0.0016759 | 0.0099529 | -0.0081047 | -0.0072256 | -0.16291 | -0.026293 | 0.00040846 | 0.16055 | 0.0012783 | 0.41103 | -0.016955 | -8.2658e-06 | 0.003197 | -0.074019 | 0.0028055 | 0.032649 | -0.008785 | [7 empty] | 0.00036211 | -0.0005511 | -0.0018278 | -0.0018018
|  | RSD |  |  | 11.94% | 4.41% | 0.04163% | 3.986% | 1.688% | 26.88% | 148.6% | 1.343e-06% | 27.43% | 161.3% | 39.65% | 1.181% | 2.072% | 6.074% | 22.92% | 69.02% | 24.54% | 8.059% | 1.963% | 128% | 2608% | 0.6816% | 101.3% | 21.17% | 50.09% | 33300% | 69.14% | 16.97% | 260.7% | 10.12% | 19.51% | [7 empty] | 92.32% | 13.78% | 25.49% | 1.845%
|  | % Recovery | [33 empty] | 80.87% | 70.58% | 79.6% | 75.58% | 77.05% | 69.25% | 76.5%
| 192 | LBLANK |  | 06/09/2023 08:47:44 | -13.3912 | -42.7414 | -0.70319 | -0.78213 | -1.3829 | 379.4154 | -0.11352 | -30.7443 | 0.0028168 | 0.0030021 | 0.0032259 | -0.50412 | -0.024715 | -0.6075 | 0.0021995 | 0.019936 | -0.0095497 | -0.0076463 | -0.1644 | -0.032562 | 0.0010604 | 0.15838 | -0.0012533 | 0.27624 | -0.016323 | 0.0014387 | 0.0019614 | 0.023643 | -0.0014165 | -0.0080142 | -0.011205 | [7 empty] | 0.00026628 | -0.00070797 | -0.0021271 | -0.001969
|  | RSD |  |  | 5.924% | 2.367e-06% | 0.09846% | 4.349% | 3.514% | 62.87% | 1288% | 1.343e-06% | 0% | 62.1% | 36.35% | 1.288% | 2.985% | 13.45% | 58.7% | 26.2% | 28.49% | 7.596% | 2.094% | 103.5% | 1111% | 1.283% | 191.8% | 10.59% | 58.73% | 365.5% | 137.1% | 556.1% | 0% | 7.389% | 3.013% | [7 empty] | 107.3% | 54.43% | 15.37% | 3.776%
|  | % Recovery | [33 empty] | 81.11% | 66.86% | 74.61% | 74.78% | 72.67% | 61.11% | 73.79%
| 208 | LBLANK |  | 06/09/2023 10:23:31 | -12.8327 | -42.7414 | -0.70053 | -0.72136 | -2.2825 | 603.3681 | 1.615 | -30.7443 | 0.003544 | 0.001236 | 0.0036236 | -0.49989 | -0.024707 | -0.50489 | 0.0014541 | 0.0085312 | -0.01003 | -0.0063536 | -0.16236 | -0.033331 | -0.0057419 | 0.16006 | 0.00081604 | 0.24527 | -0.021858 | 0.0035984 | 0.0038821 | -0.010501 | 0.0076898 | -0.008223 | -0.009355 | [7 empty] | 7.9106e-05 | -0.00056883 | -0.0024448 | -0.0018272
|  | RSD |  |  | 6.86% | 2.367e-06% | 0.3242% | 4.794% | 0.3724% | 43.76% | 6.051% | 1.343e-06% | 35.54% | 69.6% | 60.85% | 2.368% | 0.4645% | 13.4% | 0% | 19.88% | 25.96% | 10.17% | 3.571% | 100.6% | 1.756e-06% | 0.2731% | 58.49% | 17.36% | 0% | 2.889% | 62.51% | 579.3% | 102.6% | 5.045% | 19.51% | [7 empty] | 225.5% | 36.92% | 8.772% | 1.713%
|  | % Recovery | [33 empty] | 78.28% | 64% | 71.19% | 64.15% | 68.39% | 61.56% | 70.86%
| 228 | LBLANK |  | 06/09/2023 12:43:32 | 28.8843 | 13.3084 | -0.54623 | -0.71016 | -0.48328 | 469.8215 | 33.7033 | 65.8758 | 0.0062431 | 3.8361e-05 | -0.00024657 | -0.19936 | -0.02705 | -0.046952 | 0.0037233 | 0.0082987 | -0.0047314 | -0.0066264 | -0.13116 | -0.05553 | 0.023883 | 0.26669 | 0.0063026 | 0.29045 | 0.001725 | 0.008059 | 0.0085619 | 0.042051 | -0.029558 | 0.088502 | -0.0059357 | [7 empty] | 0.0015747 | -0.0003043 | -0.0029503 | -0.0012646
|  | RSD |  |  | 2.619% | 14.47% | 1.949% | 8.149% | 1.371% | 40.6% | 26.65% | 0% | 84.43% | 12169% | 379.8% | 4.091% | 1.215% | 146.4% | 9.141% | 82.36% | 32.23% | 7.022% | 4.172% | 152.5% | 49.83% | 0.707% | 35.78% | 49.53% | 0% | 55.04% | 24.21% | 77.96% | 40.75% | 9.117% | 9.281% | [7 empty] | 57.51% | 140.8% | 19.02% | 9.897%
|  | % Recovery | [33 empty] | 93.48% | 100.1% | 100.8% | 102.9% | 102.2% | 105.6% | 104.3%
| 230 | LBLANK |  | 06/09/2023 13:19:06 | 28.0894 | 12.1963 | -0.72088 | -0.7574 | -0.84198 | 423.1132 | 35.3658 | 65.8758 | 0.0019891 | -0.0060389 | -0.0016394 | -0.53294 | 0.0010633 | -0.50355 | 0.0014738 | 0.0041766 | -0.009049 | -0.0076148 | -0.14069 | -0.0026811 | 0.010141 | 0.26237 | 0.0013302 | 0.044187 | -0.053078 | 0.0085013 | 0.0059206 | -0.0062396 | -0.041654 | 0.044702 | 0.003428 | [7 empty] | 0.0010166 | -0.0011082 | -0.0033389 | -0.0020406
|  | RSD |  |  | 2.68% | 0% | 0.3047% | 4.04% | 1.767% | 18.55% | 13.57% | 0% | 56.31% | 1.141% | 65.65% | 1.025% | 36.52% | 5.875% | 41.6% | 163.2% | 13.64% | 10.28% | 6.55% | 332.8% | 0% | 0.3052% | 65.1% | 269.5% | 0% | 55.34% | 21.37% | 848.8% | 0% | 11.14% | 6.872% | [7 empty] | 16.69% | 18.94% | 7.642% | 7.27%
|  | % Recovery | [33 empty] | 99.61% | 95.68% | 97.3% | 97.23% | 98.89% | 84.65% | 97.67%
| 250 | LBLANK |  | 06/09/2023 15:17:56 | 26.2709 | 13.319 | -0.69936 | -0.76018 | -0.37202 | 130.629 | 27.3721 | 70.2176 | 0.0013424 | -0.0047257 | -0.0006886 | -0.5246 | 0.0011045 | -0.54598 | 0.00060981 | 0.0032985 | -0.010553 | -0.0080581 | -0.15021 | 0.089651 | 0.038237 | 0.27621 | 0.016887 | 0.23529 | -0.042548 | 0.00072501 | 0.0060475 | 0.028391 | -0.037289 | 0.0026916 | 0.0037586 | [7 empty] | 0.0010229 | -0.0013405 | -0.0033056 | -0.0018326
|  | RSD |  |  | 2.735% | 14.6% | 1.643% | 3.323% | 2.409% | 129.3% | 11.3% | 5.373% | 2.3e-06% | 57.58% | 121.3% | 0.2917% | 21.99% | 12.92% | 59.63% | 198.3% | 12.08% | 10.37% | 2.043% | 119.4% | 34.95% | 0.408% | 6.333% | 87.52% | 42.87% | 385.2% | 78.58% | 260.8% | 20.27% | 16.7% | 14.89% | [7 empty] | 17.64% | 9.845% | 12.38% | 9.611%
|  | % Recovery | [33 empty] | 90.53% | 92.37% | 99.93% | 92.75% | 93.67% | 91.76% | 95.18%
| 265 | LBLANK |  | 06/09/2023 16:45:55 | 27.7253 | 12.1963 | -0.73915 | -0.75231 | -0.75297 | 113.8578 | 32.3785 | 65.8758 | 0.0020398 | -0.0028941 | -0.00087695 | -0.55435 | 0.00083728 | -0.55519 | 0.00041452 | -0.0014944 | -0.010466 | -0.008599 | -0.15804 | 0.10023 | 0.010141 | 0.26283 | 0.0018582 | 0.12282 | -0.053078 | 0.00082217 | 0.0046871 | -0.036165 | -0.041654 | 0.0019808 | 0.0031888 | [7 empty] | 0.0012397 | -0.0011874 | -0.0037063 | -0.0022137
|  | RSD |  |  | 7.661% | 0% | 0.2981% | 8.49% | 0.948% | 168.9% | 9.937% | 0% | 59.22% | 139.2% | 55.77% | 1.537% | 11.7% | 4.046% | 93.25% | 326.7% | 10.61% | 3.152% | 0.8424% | 26.53% | 0% | 0.3606% | 57.73% | 45.92% | 0% | 360.2% | 52.39% | 282.8% | 0% | 20.9% | 14.45% | [7 empty] | 25.84% | 11.19% | 3.882% | 5.516%
|  | % Recovery | [33 empty] | 90.68% | 86.81% | 91.48% | 89.17% | 85.63% | 87.13% | 89.26%
| 283 | LBLANK |  | 06/09/2023 18:31:48 | 26.8542 | 12.1963 | -0.73585 | -0.73226 | -0.78764 | 45.8085 | 40.4007 | 65.8758 | 0.0020443 | -0.005768 | -0.0012446 | -0.54655 | 0.00078471 | -0.5523 | 0.00042811 | 0.0049746 | -0.0099498 | -0.0077734 | -0.1623 | 0.04554 | 0.010141 | 0.262 | 0.00090522 | 0.17269 | -0.030217 | 0.0043397 | 0.0038854 | 0.057765 | -0.036988 | 0.0015626 | 0.0031974 | [7 empty] | 0.0010293 | -0.0012204 | -0.0035345 | -0.0022193
|  | RSD |  |  | 6.437% | 0% | 0.02527% | 16.06% | 4.386% | 347.5% | 9.807% | 0% | 59.47% | 45.11% | 130.6% | 1.16% | 21.46% | 2.413% | 95.79% | 57.45% | 17.61% | 13.32% | 1.31% | 10.62% | 0% | 0.1819% | 63.4% | 57.81% | 131% | 121.4% | 36.54% | 197.9% | 21.85% | 24.81% | 5.652% | [7 empty] | 18.61% | 24.1% | 3.288% | 5.721%
|  | % Recovery | [33 empty] | 89.92% | 86.31% | 88.49% | 87.12% | 86.62% | 84.91% | 85.72%
QC STD 2
|  | Sample Id | R | Acquisition Time | Na KED 23/23 | Mg KED 24/24 | Al 27/27 | Al KED 27/27 | P Oshift 31/47 | S (34) -1 34/34 | K KED 39/39 | Ca KED 43/43 | V KED 51/51 | Cr (52) 52/52 | Cr (53) NH3 53/53 | Fe (54) NH3 54/54 | Mn Oshift 55/71 | Fe (56) KED 56/56 | Co KED 59/59 | Ni KED 60/60 | Cu (63) KED 63/63 | Cu (65) NH3 65/65 | Zn NH3 66/66 | As 75/75 | As KED 75/75 | As RH Oshift 75/91 | As Oshift 75/91 | Se (77) 77/77 | Se (77) Oshift 77/93 | Se (78) Oshift 78/94 | Se (80) Oshift 80/96 | Se (82) 82/82 | Se (82) Oshift 82/98 | Mo (95) NH3 95/95 | Mo (95) OShift 95/127 | Rh (IS1) 103/103 (IS) | Rh (IS2) 103/103 (IS) | Rh (IS3) 103/103 (IS) | Rh (IS4) 103/103 (IS) | Rh (IS5) 103/103 (IS) | Rh (IS6) 103/103 (IS) | Rh OShift (IS1) 103/119 (IS) | Cd (111) NH3 111/111 | Cd (114) NH3 114/114 | Pb 208/208 | Pb NH3 208/208
| 17 | ICV PE |  | 06/08/2023 15:03:43 | 9910.2106 | 9986.4089 | 9.9707 | 9.3487 | 19.0658 | 9589.0762 | 9988.7685 | 10156.8568 | 9.8933 | 10.0135 | 9.74 | 10.0444 | 10.4635 | 10.2 | 10.099 | 10.2574 | 9.7402 | 9.5326 | 9.5867 | 10.1386 | 8.2954 | 9.5977 | 10.0923 | 7.5385 | 10.1506 | 9.9938 | 10.0369 | 10.9266 | 10.1441 | 10.035 | 10.1813 | [7 empty] | 9.7246 | 9.8909 | 9.9805 | 9.4044
|  | RSD |  |  | 2.223% | 0.8283% | 2.429% | 7.278% | 0.9562% | 4.827% | 2.313% | 3.424% | 1.68% | 1.887% | 1.548% | 0.1172% | 0.6104% | 1.271% | 1.646% | 3.179% | 3.409% | 1.286% | 1.249% | 3.058% | 5.003% | 2.139% | 1.177% | 1.791% | 2.544% | 2.087% | 0.355% | 2.413% | 5.166% | 0.6731% | 0.961% | [7 empty] | 0.7057% | 1.047% | 1.497% | 6.894%
|  | % Recovery |  |  | 99.1% | 99.86% |  | 93.49% | 95.33% | 95.89% | 99.89% | 101.6% | 98.93% | 100.1% | 97.4% | 100.4% | 104.6% | 102% | 101% | 102.6% | 97.4% | 95.33% | 95.87% | 101.4% | 82.95% |  | 100.9% | 75.38% | 101.5% | 99.94% | 100.4% | 109.3% | 101.4% | 100.3% | 101.8% | 88.2% | 87.61% | 90.95% | 88.43% | 95.41% | 86.78% | 85.58% | 97.25% | 98.91% | 99.81% | 94.04%
| 18 | ICV PE |  | 06/08/2023 15:09:59 | 9801.7516 | 10242.1682 | 9.4473 | 9.6363 | 19.5113 | 9856.9094 | 9816.7945 | 10287.4771 | 10.0018 | 10.086 | 9.964 | 10.3213 | 10.5281 | 10.472 | 10.1863 | 10.3198 | 9.8091 | 9.7791 | 9.8334 | 9.7628 | 8.4124 | 9.1723 | 10.0533 | 7.3008 | 9.9368 | 9.6553 | 9.9805 | 10.0546 | 9.6627 | 10.0993 | 10.2349 | [7 empty] | 9.9102 | 10.114 | 9.601 | 9.8706
|  | RSD |  |  | 1.154% | 2.763% | 1.016% | 13.98% | 1.589% | 2.888% | 3.431% | 4.709% | 5.656% | 3.959% | 1.046% | 3.558% | 0.5038% | 4.834% | 2.037% | 2.63% | 2.991% | 1.795% | 0.7392% | 1.846% | 4.95% | 2.678% | 1.04% | 1.701% | 3.344% | 0.5396% | 2.422% | 0.7387% | 4.983% | 1.248% | 0.943% | [7 empty] | 2.22% | 1.873% | 0.02433% | 0.7608%
|  | % Recovery |  |  | 98.02% | 102.4% |  | 96.36% | 97.56% | 98.57% | 98.17% | 102.9% | 100% | 100.9% | 99.64% | 103.2% | 105.3% | 104.7% | 101.9% | 103.2% | 98.09% | 97.79% | 98.33% | 97.63% | 84.12% |  | 100.5% | 73.01% | 99.37% | 96.55% | 99.8% | 100.5% | 96.63% | 101% | 102.3% | 91.96% | 87.46% | 90.05% | 87.41% | 90.23% | 85.93% | 89.14% | 99.1% | 101.1% | 96.01% | 98.71%
| 231 | ICV PE |  | 06/09/2023 13:24:53 | 10402.4606 | 10790.0138 | 8.6359 | 9.6861 | 18.1854 | 10365.3137 | 10205.5983 | 10292.2368 | 10.733 | 10.5487 | 9.9666 | 10.139 | 10.057 | 10.548 | 10.6406 | 10.687 | 10.49 | 9.6931 | 9.4901 | 9.6037 | 9.5931 | 9.4309 | 10.1099 | 6.8086 | 10.1598 | 9.9229 | 9.7291 | 9.6263 | 10.1502 | 9.5231 | 9.4949 | [7 empty] | 9.6305 | 9.7447 | 9.216 | 9.4524
|  | RSD |  |  | 2.305% | 2.243% | 3.787% | 15.31% | 1.72% | 5.454% | 3.399% | 2.906% | 4.117% | 4.08% | 1.665% | 1.467% | 2.227% | 6.539% | 5.719% | 3.303% | 3.499% | 0.3588% | 0.7406% | 3.39% | 3.234% | 2.435% | 1.688% | 2.213% | 3.28% | 1.555% | 2.539% | 2.401% | 4.489% | 0.6035% | 0.4697% | [7 empty] | 0.7522% | 0.9612% | 2.959% | 0.6399%
|  | % Recovery |  |  | 104% | 107.9% |  | 96.86% | 90.93% | 103.7% | 102.1% | 102.9% | 107.3% | 105.5% | 99.67% | 101.4% | 100.6% | 105.5% | 106.4% | 106.9% | 104.9% | 96.93% | 94.9% | 96.04% | 95.93% |  | 101.1% | 68.09% | 101.6% | 99.23% | 97.29% | 96.26% | 101.5% | 95.23% | 94.95% | 96.09% | 91.25% | 87.83% | 90.87% | 91.34% | 91.23% | 92.26% | 96.31% | 97.45% | 92.16% | 94.52%
| 232 | ICV PE |  | 06/09/2023 13:31:10 | 10122.2283 | 10855.4755 | 8.9054 | 9.4452 | 18.1516 | 9683.2897 | 9953.6138 | 10338.8355 | 10.4088 | 10.3483 | 11.5413 | 11.7797 | 10.0963 | 10.3264 | 10.2743 | 10.5963 | 10.4382 | 11.2558 | 11.1597 | 9.7487 | 9.4015 | 9.4857 | 10.1739 | 7.0021 | 10.1279 | 9.7158 | 9.8288 | 9.6421 | 10.1201 | 10.9642 | 9.5224 | [7 empty] | 11.1516 | 11.2557 | 9.3969 | 9.2748
|  | RSD |  |  | 2.145% | 4.137% | 1.581% | 4.997% | 0.7126% | 3.667% | 1.856% | 4.477% | 4.332% | 3.265% | 19.43% | 21.17% | 1.21% | 2.988% | 4.485% | 4.413% | 2.818% | 18.26% | 16.43% | 1.437% | 7.025% | 2.882% | 0.4041% | 2.636% | 2.559% | 0.4835% | 0.4682% | 0.5237% | 3.125% | 18.56% | 1.68% | [7 empty] | 19.59% | 17.09% | 1.614% | 0.3482%
|  | % Recovery |  |  | 101.2% | 108.6% |  | 94.45% | 90.76% | 96.83% | 99.54% | 103.4% | 104.1% | 103.5% | 115.4% | 117.8% | 101% | 103.3% | 102.7% | 106% | 104.4% | 112.6% | 111.6% | 97.49% | 94.02% |  | 101.7% | 70.02% | 101.3% | 97.16% | 98.29% | 96.42% | 101.2% | 109.6% | 95.22% | 94.95% | 90.59% | 90.59% | 82.69% | 94.66% | 91.91% | 91.64% | 111.5% | 112.6% | 93.97% | 92.75%
QC STD 3
|  | Sample Id | R | Acquisition Time | Na KED 23/23 | Mg KED 24/24 | Al 27/27 | Al KED 27/27 | P Oshift 31/47 | S (34) -1 34/34 | K KED 39/39 | Ca KED 43/43 | V KED 51/51 | Cr (52) 52/52 | Cr (53) NH3 53/53 | Fe (54) NH3 54/54 | Mn Oshift 55/71 | Fe (56) KED 56/56 | Co KED 59/59 | Ni KED 60/60 | Cu (63) KED 63/63 | Cu (65) NH3 65/65 | Zn NH3 66/66 | As 75/75 | As KED 75/75 | As RH Oshift 75/91 | As Oshift 75/91 | Se (77) 77/77 | Se (77) Oshift 77/93 | Se (78) Oshift 78/94 | Se (80) Oshift 80/96 | Se (82) 82/82 | Se (82) Oshift 82/98 | Mo (95) NH3 95/95 | Mo (95) OShift 95/127 | Rh (IS1) 103/103 (IS) | Rh (IS2) 103/103 (IS) | Rh (IS3) 103/103 (IS) | Rh (IS4) 103/103 (IS) | Rh (IS5) 103/103 (IS) | Rh (IS6) 103/103 (IS) | Rh OShift (IS1) 103/119 (IS) | Cd (111) NH3 111/111 | Cd (114) NH3 114/114 | Pb 208/208 | Pb NH3 208/208
| 19 | CCV 1ppb TRACE |  | 06/08/2023 15:15:47 | 963.5918 | 970.3544 | 1.0188 | 0.95078 | -0.56547 | 1344.678 | 976.128 | 962.5777 | 1.0279 | 1.0071 | 0.99805 | 0.59771 | 1.0075 | 0.59076 | 1.0292 | 1.0562 | 1.0438 | 1.0291 | 0.85937 | 0.94406 | 0.80569 | 1.0664 | 0.99222 | 1.0858 | 0.96297 | 0.99192 | 1.03 | 1.0712 | 0.85187 | 1.0403 | 1.0379 | [7 empty] | 0.9723 | 1.0016 | 0.98748 | 1.0285
|  | RSD |  |  | 1.162% | 5.001% | 2.309% | 19.08% | 10.67% | 6.306% | 2.165% | 13.27% | 1.727% | 4.504% | 1.897% | 5.401% | 0.6934% | 8.199% | 3.01% | 2.485% | 2.425% | 1.376% | 2.029% | 3.459% | 12.09% | 0.4992% | 3.414% | 5.736% | 14.98% | 8.319% | 3.773% | 8.458% | 9.872% | 3.394% | 1.22% | [7 empty] | 3.668% | 2.098% | 1.251% | 0.6195%
|  | % Recovery |  |  | 96.36% | 97.04% |  | 95.08% | -56.55% | 134.5% | 97.61% | 96.26% | 102.8% | 100.7% | 99.8% | 59.77% | 100.8% | 59.08% | 102.9% | 105.6% | 104.4% | 102.9% | 85.94% | 94.41% | 80.57% | 106.6% | 99.22% | 108.6% | 96.3% | 99.19% | 103% | 107.1% | 85.19% | 104% | 103.8% | 96.74% | 94.33% | 94.85% | 95.68% | 96.16% | 90.36% | 94.25% | 97.23% | 100.2% | 98.75% | 102.8%
| 20 | CCV 1ppb TRACE |  | 06/08/2023 15:22:03 | 1016.4433 | 1053.5242 | 1.0224 | 0.89479 | -0.98102 | 1072.081 | 986.0152 | 905.2676 | 0.98098 | 0.99507 | 0.99288 | 0.66355 | 1.0327 | 0.6398 | 1.0325 | 1.066 | 1.0819 | 1.0381 | 0.86327 | 1.0362 | 0.91812 | 1.0845 | 1.0114 | 1.0792 | 0.98803 | 0.94675 | 0.98258 | 1.0218 | 1.0772 | 1.0419 | 1.032 | [7 empty] | 1.0026 | 1.0096 | 0.98362 | 1.0027
|  | RSD |  |  | 0.7425% | 2.604% | 2.975% | 75.11% | 2.806% | 15.97% | 6% | 10.71% | 6.019% | 0.4856% | 3.595% | 2.118% | 2.585% | 13.66% | 1.704% | 7.099% | 3.202% | 2.404% | 3.488% | 4.056% | 15.75% | 3.187% | 1.595% | 7.981% | 8.32% | 6.159% | 3.616% | 2.032% | 13.45% | 3.233% | 1.371% | [7 empty] | 1.588% | 3.928% | 1.893% | 1.754%
|  | % Recovery |  |  | 101.6% | 105.4% |  | 89.48% | -98.1% | 107.2% | 98.6% | 90.53% | 98.1% | 99.51% | 99.29% | 66.36% | 103.3% | 63.98% | 103.2% | 106.6% | 108.2% | 103.8% | 86.33% | 103.6% | 91.81% | 108.4% | 101.1% | 107.9% | 98.8% | 94.68% | 98.26% | 102.2% | 107.7% | 104.2% | 103.2% | 94.29% | 90.06% | 94.61% | 91.67% | 94.59% | 88.07% | 90.02% | 100.3% | 101% | 98.36% | 100.3%
| 36 | CCV 1ppb TRACE |  | 06/08/2023 16:55:19 | 958.2467 | 954.1637 | 1.0925 | 1.0315 | -1.1519 | 1193.1936 | 960.3951 | 1029.761 | 0.99255 | 0.95997 | 1.0347 | 0.69564 | 1.0961 | 0.45499 | 0.99437 | 1.0008 | 1.0031 | 1.0593 | 0.92424 | 0.98373 | 0.86731 | 1.132 | 1.0571 | 1.0665 | 0.98164 | 0.94561 | 1.0226 | 0.9751 | 1.0735 | 1.0733 | 1.1139 | [7 empty] | 1.0537 | 1.0673 | 1.0081 | 1.077
|  | RSD |  |  | 2.481% | 8.255% | 5.117% | 24.65% | 1.263% | 22.47% | 3.243% | 18.88% | 2.906% | 6.405% | 3.065% | 0.9744% | 1.944% | 20.8% | 4.163% | 6.897% | 4.492% | 0.8384% | 1.579% | 3.061% | 10.91% | 2.83% | 1.747% | 6.434% | 20.95% | 1.344% | 3.271% | 10.45% | 18.29% | 1.312% | 0.9194% | [7 empty] | 2.765% | 1.843% | 1.96% | 1.377%
|  | % Recovery |  |  | 95.82% | 95.42% |  | 103.1% | -115.2% | 119.3% | 96.04% | 103% | 99.25% | 96% | 103.5% | 69.56% | 109.6% | 45.5% | 99.44% | 100.1% | 100.3% | 105.9% | 92.42% | 98.37% | 86.73% | 113.2% | 105.7% | 106.6% | 98.16% | 94.56% | 102.3% | 97.51% | 107.3% | 107.3% | 111.4% | 85.61% | 80.25% | 90.93% | 82.14% | 84.6% | 78.3% | 79.73% | 105.4% | 106.7% | 100.8% | 107.7%
| 37 | CCV 1ppb TRACE |  | 06/08/2023 17:01:37 | 1012.0937 | 1074.9191 | 1.0866 | 0.89436 | -1.1615 | 1573.0823 | 1035.5289 | 1096.9143 | 1.0931 | 1.0022 | 1.0463 | 0.68502 | 1.0739 | 0.60055 | 1.054 | 1.1651 | 1.0659 | 1.0682 | 0.91587 | 0.96788 | 0.96679 | 1.1457 | 1.0544 | 1.0106 | 1.0155 | 1.0507 | 1.103 | 0.99177 | 1.1424 | 1.0737 | 1.0832 | [7 empty] | 1.0343 | 1.0425 | 0.99995 | 1.0728
|  | RSD |  |  | 2.657% | 5.919% | 2.587% | 24.3% | 4.86% | 12.09% | 9.893% | 4.535% | 4.429% | 5.222% | 0.9794% | 8.135% | 0.6665% | 8.366% | 1.646% | 3.511% | 1.756% | 2.095% | 1.308% | 2.774% | 8.875% | 1.071% | 2.712% | 2.807% | 5.17% | 6.982% | 3.569% | 11.4% | 12.91% | 2.18% | 1.317% | [7 empty] | 1.776% | 1.291% | 0.8901% | 0.335%
|  | % Recovery |  |  | 101.2% | 107.5% |  | 89.44% | -116.1% | 157.3% | 103.6% | 109.7% | 109.3% | 100.2% | 104.6% | 68.5% | 107.4% | 60.05% | 105.4% | 116.5% | 106.6% | 106.8% | 91.59% | 96.79% | 96.68% | 114.6% | 105.4% | 101.1% | 101.5% | 105.1% | 110.3% | 99.18% | 114.2% | 107.4% | 108.3% | 85.88% | 81.63% | 86.3% | 82.83% | 84.84% | 79.32% | 79.74% | 103.4% | 104.2% | 100% | 107.3%
| 51 | CCV 1ppb TRACE |  | 06/08/2023 18:25:05 | 1121.0672 | 1101.9409 | 1.1446 | 0.94174 | -1.0595 | 1850.787 | 1160.7013 | 1166.2788 | 1.1208 | 1.1122 | 1.0651 | 0.72696 | 1.147 | 0.83038 | 1.1214 | 1.1956 | 1.183 | 1.077 | 0.91968 | 1.0707 | 1.0168 | 1.1313 | 1.0815 | 1.0937 | 1.1362 | 1.0544 | 1.1351 | 1.1711 | 1.0063 | 1.1082 | 1.1224 | [7 empty] | 1.0277 | 1.0435 | 0.99686 | 1.0715
|  | RSD |  |  | 2.331% | 8.691% | 1.699% | 43.67% | 2.929% | 6.333% | 6.957% | 10.58% | 9.599% | 2.52% | 1.772% | 2.354% | 3.355% | 4.001% | 5.016% | 7.016% | 2.914% | 1.615% | 3.697% | 4.285% | 24.01% | 6.136% | 2.489% | 8.225% | 5.311% | 3.005% | 1.337% | 14.44% | 5.332% | 1.43% | 0.5322% | [7 empty] | 1.55% | 3.184% | 0.9004% | 2.085%
|  | % Recovery |  |  | 112.1% | 110.2% |  | 94.17% | -106% | 185.1% | 116.1% | 116.6% | 112.1% | 111.2% | 106.5% | 72.7% | 114.7% | 83.04% | 112.1% | 119.6% | 118.3% | 107.7% | 91.97% | 107.1% | 101.7% | 113.1% | 108.1% | 109.4% | 113.6% | 105.4% | 113.5% | 117.1% | 100.6% | 110.8% | 112.2% | 86.4% | 80.37% | 84.44% | 83.57% | 85.82% | 76.89% | 81.88% | 102.8% | 104.3% | 99.69% | 107.1%
| 52 | CCV 1ppb TRACE |  | 06/08/2023 18:31:23 | 1074.0423 | 1035.7773 | 1.2255 | 1.0572 | -1.1292 | 1627.1314 | 1053.138 | 1080.062 | 1.1531 | 1.0691 | 1.0641 | 0.70203 | 1.1187 | 0.78154 | 1.0968 | 1.1118 | 1.1253 | 1.093 | 0.90988 | 1.0676 | 0.83794 | 1.1339 | 1.0942 | 1.0592 | 1.0457 | 1.0296 | 1.0514 | 1.0642 | 0.96887 | 1.0689 | 1.1021 | [7 empty] | 1.0265 | 1.0467 | 1.0068 | 1.0562
|  | RSD |  |  | 3.46% | 5.725% | 15.48% | 21.47% | 5.102% | 18.01% | 5.245% | 13.64% | 4.191% | 3.896% | 2.844% | 3.08% | 1.599% | 7.866% | 3.734% | 3.788% | 1.877% | 1.785% | 3.312% | 11.78% | 5.957% | 1.98% | 0.9404% | 0.2673% | 9.309% | 0.2283% | 9.394% | 13.28% | 1.191% | 1.693% | 0.8009% | [7 empty] | 1.274% | 2.246% | 6.379% | 2.392%
|  | % Recovery |  |  | 107.4% | 103.6% |  | 105.7% | -112.9% | 162.7% | 105.3% | 108% | 115.3% | 106.9% | 106.4% | 70.2% | 111.9% | 78.15% | 109.7% | 111.2% | 112.5% | 109.3% | 90.99% | 106.8% | 83.79% | 113.4% | 109.4% | 105.9% | 104.6% | 103% | 105.1% | 106.4% | 96.89% | 106.9% | 110.2% | 86.52% | 80.71% | 87.14% | 83.11% | 85.91% | 79.77% | 82.79% | 102.7% | 104.7% | 100.7% | 105.6%
| 67 | CCV 1ppb TRACE |  | 06/08/2023 20:30:21 | 1103.0697 | 1020.7486 | 1.2577 | 1.2835 | -1.1208 | 1470.9034 | 1089.9895 | 1119.9024 | 1.1846 | 1.0814 | 1.0739 | 0.75535 | 1.1138 | 0.73685 | 1.1072 | 1.1212 | 1.1034 | 1.0981 | 0.93956 | 1.1154 | 1.0021 | 1.1325 | 1.0663 | 1.3082 | 0.90907 | 1.0418 | 1.0974 | 1.1586 | 1.0195 | 1.1305 | 1.0976 | [7 empty] | 1.0732 | 1.0758 | 1.0851 | 1.0828
|  | RSD |  |  | 4.784% | 2.119% | 8.358% | 47.74% | 1.377% | 6.556% | 4.448% | 9.7% | 6.646% | 3.352% | 0.7191% | 2.422% | 3.199% | 9.266% | 1.855% | 0.2962% | 6.259% | 1.079% | 3.591% | 4.944% | 4.509% | 0.525% | 1.33% | 16.38% | 11.62% | 7.884% | 8.88% | 7.251% | 4.437% | 1.609% | 1.764% | [7 empty] | 4.182% | 0.888% | 6.799% | 0.4381%
|  | % Recovery |  |  | 110.3% | 102.1% |  | 128.4% | -112.1% | 147.1% | 109% | 112% | 118.5% | 108.1% | 107.4% | 75.54% | 111.4% | 73.69% | 110.7% | 112.1% | 110.3% | 109.8% | 93.96% | 111.5% | 100.2% | 113.2% | 106.6% | 130.8% | 90.91% | 104.2% | 109.7% | 115.9% | 101.9% | 113% | 109.8% | 80.52% | 77.33% | 82.86% | 78.91% | 83.14% | 75.74% | 77.4% | 107.3% | 107.6% | 108.5% | 108.3%
| 68 | CCV 1ppb TRACE |  | 06/08/2023 20:36:39 | 1055.8381 | 1030.8016 | 1.1158 | 1.1625 | -1.1697 | 1574.9501 | 1005.5786 | 1065.5611 | 1.1263 | 1.0592 | 1.0742 | 0.72877 | 1.1246 | 0.81247 | 1.1487 | 1.1242 | 1.1338 | 1.1124 | 0.91885 | 1.0537 | 1.0406 | 1.1137 | 1.086 | 1.1575 | 0.93309 | 0.99768 | 1.0916 | 1.0536 | 1.1759 | 1.0975 | 1.1172 | [7 empty] | 1.058 | 1.0696 | 1.0132 | 1.0869
|  | RSD |  |  | 3.278% | 1.157% | 1.659% | 16.88% | 3.089% | 17.18% | 5.519% | 6.396% | 1.058% | 4.257% | 3.77% | 1.954% | 1.575% | 7.836% | 4.597% | 4.591% | 0.5582% | 1.714% | 0.4569% | 4.39% | 12.42% | 1.947% | 1.352% | 12.07% | 8.007% | 3.385% | 3.045% | 6.142% | 2.949% | 2.42% | 1.228% | [7 empty] | 2.5% | 1.356% | 1.968% | 0.6185%
|  | % Recovery |  |  | 105.6% | 103.1% |  | 116.2% | -117% | 157.5% | 100.6% | 106.6% | 112.6% | 105.9% | 107.4% | 72.88% | 112.5% | 81.25% | 114.9% | 112.4% | 113.4% | 111.2% | 91.89% | 105.4% | 104.1% | 111.4% | 108.6% | 115.8% | 93.31% | 99.77% | 109.2% | 105.4% | 117.6% | 109.8% | 111.7% | 84.95% | 77.28% | 84.06% | 80.48% | 82.62% | 75.24% | 80.37% | 105.8% | 107% | 101.3% | 108.7%
| 82 | CCV 1ppb TRACE |  | 06/08/2023 21:58:24 | 1083.8139 | 1074.0632 | 1.1396 | 1.6152 | -0.91991 | 1415.5344 | 1073.1178 | 1100.4318 | 1.1266 | 1.1065 | 1.0918 | 0.77772 | 1.1691 | 0.76269 | 1.1289 | 1.1687 | 1.1202 | 1.1199 | 0.93425 | 1.0681 | 0.96265 | 1.1286 | 1.0969 | 1.1809 | 1.0509 | 1.0511 | 1.053 | 1.1198 | 1.1705 | 1.118 | 1.1342 | [7 empty] | 1.0708 | 1.065 | 1.0249 | 1.0864
|  | RSD |  |  | 6.005% | 7.083% | 1.648% | 20.41% | 8.354% | 12.66% | 8.62% | 7.982% | 4.33% | 4.201% | 0.9479% | 2.161% | 0.2774% | 4.848% | 0.5987% | 4.91% | 7.602% | 1.707% | 2.708% | 5.788% | 12.78% | 2.012% | 3.66% | 7.08% | 12.66% | 6.97% | 4.557% | 16.63% | 4.772% | 0.3417% | 3.779% | [7 empty] | 2.849% | 0.9939% | 1.55% | 0.5666%
|  | % Recovery |  |  | 108.4% | 107.4% |  | 161.5% | -91.99% | 141.6% | 107.3% | 110% | 112.7% | 110.7% | 109.2% | 77.77% | 116.9% | 76.27% | 112.9% | 116.9% | 112% | 112% | 93.42% | 106.8% | 96.26% | 112.9% | 109.7% | 118.1% | 105.1% | 105.1% | 105.3% | 112% | 117.1% | 111.8% | 113.4% | 85.43% | 76.63% | 83.13% | 79.22% | 83.49% | 75.77% | 79.2% | 107.1% | 106.5% | 102.5% | 108.6%
| 83 | CCV 1ppb TRACE |  | 06/08/2023 22:04:41 | 1069.9421 | 1040.364 | 1.1282 | 1.1078 | -1.1396 | 1451.9629 | 1027.3233 | 1061.703 | 1.1249 | 1.0939 | 1.0874 | 0.73885 | 1.1369 | 0.78214 | 1.1261 | 1.0755 | 1.18 | 1.1194 | 0.95502 | 1.0204 | 0.95972 | 1.1353 | 1.0911 | 1.1983 | 1.1021 | 1.0175 | 1.0921 | 1.0745 | 1.0357 | 1.1005 | 1.1456 | [7 empty] | 1.0518 | 1.0477 | 1.0283 | 1.1594
|  | RSD |  |  | 4.588% | 1.511% | 4.976% | 22.13% | 3.95% | 23.19% | 15.54% | 8.873% | 0.4676% | 7.781% | 1.598% | 3.509% | 1.138% | 12.74% | 1.893% | 8.555% | 1.807% | 2.074% | 1.596% | 2.021% | 12.21% | 3.622% | 2.633% | 12.29% | 4.591% | 15.26% | 4.963% | 7.591% | 21.5% | 0.942% | 1.888% | [7 empty] | 1.934% | 1.949% | 1.194% | 10.74%
|  | % Recovery |  |  | 107% | 104% |  | 110.8% | -114% | 145.2% | 102.7% | 106.2% | 112.5% | 109.4% | 108.7% | 73.88% | 113.7% | 78.21% | 112.6% | 107.5% | 118% | 111.9% | 95.5% | 102% | 95.97% | 113.5% | 109.1% | 119.8% | 110.2% | 101.7% | 109.2% | 107.5% | 103.6% | 110.1% | 114.6% | 87.62% | 76.9% | 84.26% | 79.71% | 79.71% | 81.03% | 78.6% | 105.2% | 104.8% | 102.8% | 115.9%
| 98 | CCV 1ppb TRACE |  | 06/08/2023 23:32:40 | 1075.7356 | 1065.6145 | 1.1826 | 1.3138 | -1.0627 | 1548.7991 | 1080.3294 | 1122.7302 | 1.1198 | 1.0793 | 1.0998 | 0.81134 | 1.1546 | 0.75138 | 1.101 | 1.1429 | 1.1454 | 1.1363 | 0.96015 | 1.0722 | 0.95401 | 1.1202 | 1.0949 | 1.3668 | 1.0533 | 1.0713 | 1.0358 | 1.0991 | 1.1439 | 1.1506 | 1.1329 | [7 empty] | 1.0267 | 1.0695 | 1.0094 | 1.0642
|  | RSD |  |  | 4.296% | 4.885% | 0.4031% | 47.52% | 1.08% | 7.823% | 4.816% | 10.16% | 1.514% | 6.874% | 2.262% | 3.731% | 1.667% | 25.96% | 3.567% | 8.211% | 1.17% | 3.462% | 3.128% | 4.291% | 11.71% | 2.282% | 0.5215% | 6.129% | 24.56% | 8.058% | 7.707% | 11.04% | 8.011% | 2.212% | 0.8737% | [7 empty] | 2.793% | 2.241% | 0.6873% | 0.3072%
|  | % Recovery |  |  | 107.6% | 106.6% |  | 131.4% | -106.3% | 154.9% | 108% | 112.3% | 112% | 107.9% | 110% | 81.13% | 115.5% | 75.14% | 110.1% | 114.3% | 114.5% | 113.6% | 96.01% | 107.2% | 95.4% | 112% | 109.5% | 136.7% | 105.3% | 107.1% | 103.6% | 109.9% | 114.4% | 115.1% | 113.3% | 88.12% | 79.46% | 84.84% | 81.42% | 85.03% | 77.07% | 82.75% | 102.7% | 107% | 100.9% | 106.4%
| 99 | CCV 1ppb TRACE |  | 06/08/2023 23:38:58 | 1075.3169 | 1055.3616 | 1.1189 | 1.2052 | -1.1385 | 1719.5689 | 1055.5793 | 1112.1303 | 1.1231 | 1.0601 | 1.0837 | 0.7447 | 1.1293 | 0.81715 | 1.1109 | 1.1305 | 1.1414 | 1.1071 | 0.93478 | 1.0399 | 1.019 | 1.104 | 1.0764 | 1.322 | 0.91512 | 1.0297 | 1.133 | 1.094 | 1.1194 | 1.117 | 1.1111 | [7 empty] | 1.0459 | 1.049 | 1.0043 | 1.0623
|  | RSD |  |  | 2.571% | 1.871% | 6.885% | 29.6% | 2.997% | 15.1% | 1.643% | 11.37% | 2.76% | 1.443% | 2.357% | 3.356% | 1.882% | 4.034% | 2.848% | 2.701% | 2.004% | 1.214% | 2.521% | 5.127% | 19.81% | 2.442% | 2.268% | 9.896% | 8.968% | 5.18% | 7.126% | 2.661% | 2.642% | 2.407% | 0.7582% | [7 empty] | 1.193% | 1.476% | 1.204% | 1.45%
|  | % Recovery |  |  | 107.5% | 105.5% |  | 120.5% | -113.9% | 172% | 105.6% | 111.2% | 112.3% | 106% | 108.4% | 74.47% | 112.9% | 81.72% | 111.1% | 113.1% | 114.1% | 110.7% | 93.48% | 104% | 101.9% | 110.4% | 107.6% | 132.2% | 91.51% | 103% | 113.3% | 109.4% | 111.9% | 111.7% | 111.1% | 88.22% | 79.87% | 85.61% | 81.74% | 85.38% | 77.36% | 83.14% | 104.6% | 104.9% | 100.4% | 106.2%
| 113 | CCV 1ppb TRACE |  | 06/09/2023 01:00:42 | 1059.0547 | 1023.5366 | 1.0866 | 1.2312 | -0.90312 | 1616.4827 | 1054.7718 | 1051.8822 | 1.1041 | 1.0617 | 1.0762 | 0.75827 | 1.1412 | 0.78446 | 1.1141 | 1.1303 | 1.1258 | 1.104 | 0.92732 | 0.99623 | 0.95554 | 1.1265 | 1.0652 | 1.2126 | 1.1471 | 1.0674 | 1.1117 | 0.99372 | 1.0214 | 1.1014 | 1.1344 | [7 empty] | 1.0313 | 1.0543 | 1.0218 | 1.0961
|  | RSD |  |  | 4.805% | 13.67% | 3.046% | 33.43% | 3.128% | 25.46% | 3.528% | 11.2% | 4.656% | 3.537% | 1.945% | 4.309% | 2.502% | 11.06% | 2.29% | 7.407% | 3.296% | 0.57% | 2.292% | 2.848% | 10.5% | 3.644% | 1.014% | 7.623% | 7.357% | 6.521% | 5.882% | 13.68% | 0.9607% | 1.228% | 0.5477% | [7 empty] | 3.054% | 0.8178% | 1.033% | 0.9342%
|  | % Recovery |  |  | 105.9% | 102.4% |  | 123.1% | -90.31% | 161.6% | 105.5% | 105.2% | 110.4% | 106.2% | 107.6% | 75.83% | 114.1% | 78.45% | 111.4% | 113% | 112.6% | 110.4% | 92.73% | 99.62% | 95.55% | 112.6% | 106.5% | 121.3% | 114.7% | 106.7% | 111.2% | 99.37% | 102.1% | 110.1% | 113.4% | 88.8% | 74.55% | 81.81% | 79.24% | 82.36% | 72.64% | 75.09% | 103.1% | 105.4% | 102.2% | 109.6%
| 114 | CCV 1ppb TRACE |  | 06/09/2023 01:06:59 | 1061.3316 | 1049.1345 | 1.0498 | 1.2418 | -1.0603 | 1495.5844 | 993.5035 | 1083.4123 | 1.0719 | 1.0829 | 1.0739 | 0.77724 | 1.1211 | 0.66232 | 1.1061 | 1.1197 | 1.1104 | 1.0858 | 0.90841 | 1.0095 | 0.94692 | 1.0991 | 1.057 | 1.2019 | 1.0128 | 1.0089 | 1.0729 | 1.0827 | 1.1624 | 1.0886 | 1.1345 | [7 empty] | 1.0284 | 1.0534 | 1.0357 | 1.1541
|  | RSD |  |  | 8.124% | 10.6% | 0.429% | 18.34% | 1.398% | 12.22% | 10.04% | 13.28% | 5.883% | 6.251% | 3.886% | 8.608% | 1.598% | 16.69% | 1.887% | 1.816% | 4.449% | 3.643% | 2.84% | 5.003% | 10.5% | 1.24% | 2.078% | 4.191% | 6.115% | 3.185% | 4.766% | 8.825% | 6.321% | 3.763% | 3.626% | [7 empty] | 2.484% | 4.304% | 0.8858% | 0.1672%
|  | % Recovery |  |  | 106.1% | 104.9% |  | 124.2% | -106% | 149.6% | 99.35% | 108.3% | 107.2% | 108.3% | 107.4% | 77.72% | 112.1% | 66.23% | 110.6% | 112% | 111% | 108.6% | 90.84% | 100.9% | 94.69% | 109.9% | 105.7% | 120.2% | 101.3% | 100.9% | 107.3% | 108.3% | 116.2% | 108.9% | 113.5% | 86.13% | 75.69% | 82.02% | 79.8% | 78.62% | 73.63% | 77.78% | 102.8% | 105.3% | 103.6% | 115.4%
| 129 | CCV 1ppb TRACE |  | 06/09/2023 02:34:32 | 1052.5906 | 989.8255 | 1.0977 | 1.4093 | -0.93359 | 1502.2892 | 910.1179 | 939.8155 | 1.1138 | 1.1145 | 1.0856 | 0.73225 | 1.097 | 0.61284 | 1.0944 | 1.1226 | 1.0962 | 1.1169 | 0.9209 | 1.0637 | 0.9599 | 1.0416 | 1.0683 | 1.3095 | 0.90619 | 1.03 | 1.0523 | 1.1701 | 1.1188 | 1.109 | 1.0846 | [7 empty] | 1.0414 | 1.0653 | 1.0831 | 1.1025
|  | RSD |  |  | 6.406% | 3.733% | 17.19% | 7.225% | 2.413% | 10.93% | 3.829% | 7.783% | 6.758% | 2.08% | 1.39% | 6.3% | 2.164% | 7.075% | 1.084% | 0.6213% | 2.351% | 2.787% | 3.637% | 11.53% | 10.85% | 3.414% | 1.176% | 19.05% | 14.16% | 4.028% | 5.142% | 26.03% | 7.477% | 1.641% | 1.87% | [7 empty] | 1.806% | 3.353% | 10.89% | 1.05%
|  | % Recovery |  |  | 105.3% | 98.98% |  | 140.9% | -93.36% | 150.2% | 91.01% | 93.98% | 111.4% | 111.5% | 108.6% | 73.22% | 109.7% | 61.28% | 109.4% | 112.3% | 109.6% | 111.7% | 92.09% | 106.4% | 95.99% | 104.2% | 106.8% | 130.9% | 90.62% | 103% | 105.2% | 117% | 111.9% | 110.9% | 108.5% | 84.34% | 75.66% | 82.42% | 78.23% | 80.5% | 72.19% | 83.72% | 104.1% | 106.5% | 108.3% | 110.2%
| 130 | CCV 1ppb TRACE |  | 06/09/2023 02:40:49 | 1041.0015 | 957.3049 | 0.98491 | 1.1163 | -1.0897 | 1367.7174 | 970.1029 | 995.2081 | 1.105 | 1.0515 | 1.0771 | 0.74848 | 1.1382 | 0.66903 | 1.1051 | 1.1126 | 1.108 | 1.1062 | 0.93832 | 1.0186 | 0.81931 | 1.0692 | 1.086 | 1.3417 | 1.1282 | 1.0149 | 1.1185 | 1.0496 | 1.1053 | 1.1056 | 1.1531 | [7 empty] | 1.0474 | 1.0586 | 1.0206 | 1.0894
|  | RSD |  |  | 0.5636% | 4.067% | 1.493% | 13.46% | 1.843% | 19.4% | 4.469% | 14.11% | 9.084% | 1.688% | 3.437% | 3.464% | 3.53% | 8.405% | 0.09538% | 9.096% | 2.027% | 2.825% | 2.471% | 1.586% | 8.737% | 4.012% | 1.336% | 4.678% | 20.19% | 12.51% | 7.133% | 7.205% | 12.7% | 2.599% | 2.032% | [7 empty] | 1.852% | 0.9524% | 1.04% | 1.213%
|  | % Recovery |  |  | 104.1% | 95.73% |  | 111.6% | -109% | 136.8% | 97.01% | 99.52% | 110.5% | 105.1% | 107.7% | 74.85% | 113.8% | 66.9% | 110.5% | 111.3% | 110.8% | 110.6% | 93.83% | 101.9% | 81.93% | 106.9% | 108.6% | 134.2% | 112.8% | 101.5% | 111.8% | 105% | 110.5% | 110.6% | 115.3% | 87.33% | 73.82% | 83.14% | 79.11% | 80.11% | 73.23% | 80.6% | 104.7% | 105.9% | 102.1% | 108.9%
| 144 | CCV 1ppb TRACE |  | 06/09/2023 04:03:01 | 1036.5984 | 966.2173 | 0.95107 | 0.86853 | -0.84618 | 1444.5943 | 937.8407 | 992.7871 | 1.112 | 1.0465 | 1.0663 | 0.73551 | 1.1112 | 0.63826 | 1.0987 | 1.1617 | 1.0716 | 1.1164 | 0.92656 | 1.0232 | 0.89782 | 1.0393 | 1.0705 | 1.2681 | 1.2022 | 1.049 | 1.0599 | 1.0012 | 1.1514 | 1.1211 | 1.1208 | [7 empty] | 1.0339 | 1.0692 | 1.0381 | 1.1159
|  | RSD |  |  | 1.993% | 2.923% | 3.241% | 25.58% | 6.551% | 5.061% | 2.358% | 5.915% | 6.012% | 3.912% | 1.052% | 6.607% | 1.309% | 12.62% | 2.29% | 1.846% | 2.385% | 0.7465% | 2.126% | 4.449% | 13.11% | 3.834% | 2.868% | 10.61% | 5.143% | 3.461% | 1.915% | 19.97% | 17.34% | 1.25% | 1.462% | [7 empty] | 1.115% | 1.361% | 1.296% | 0.5146%
|  | % Recovery |  |  | 103.7% | 96.62% |  | 86.85% | -84.62% | 144.5% | 93.78% | 99.28% | 111.2% | 104.7% | 106.6% | 73.55% | 111.1% | 63.83% | 109.9% | 116.2% | 107.2% | 111.6% | 92.66% | 102.3% | 89.78% | 103.9% | 107% | 126.8% | 120.2% | 104.9% | 106% | 100.1% | 115.1% | 112.1% | 112.1% | 85.82% | 72.58% | 80.86% | 76.85% | 78.52% | 71.26% | 80.68% | 103.4% | 106.9% | 103.8% | 111.6%
| 145 | CCV 1ppb TRACE |  | 06/09/2023 04:09:18 | 952.4867 | 920.3942 | 0.89949 | 1.3961 | -1.0523 | 1169.3073 | 894.8116 | 935.8916 | 1.0645 | 1.0183 | 1.0814 | 0.76452 | 1.116 | 0.58453 | 1.0508 | 1.0572 | 1.0662 | 1.1021 | 0.91888 | 0.99898 | 0.84647 | 1.0347 | 1.0648 | 1.1647 | 1.1172 | 1.0883 | 1.105 | 1.0828 | 0.94063 | 1.0975 | 1.14 | [7 empty] | 1.0635 | 1.0543 | 0.99776 | 1.1161
|  | RSD |  |  | 0.3913% | 3.283% | 1.843% | 30.98% | 1.79% | 14.12% | 4.768% | 13.67% | 3.244% | 2.024% | 2.863% | 3.58% | 1.618% | 4.056% | 2.515% | 3.253% | 2.682% | 1.958% | 3.53% | 3.705% | 14.27% | 2.147% | 1.335% | 10.06% | 13.51% | 6.882% | 10.74% | 23.34% | 2.324% | 2.218% | 2.33% | [7 empty] | 1.826% | 0.6011% | 3.413% | 0.7153%
|  | % Recovery |  |  | 95.25% | 92.04% |  | 139.6% | -105.2% | 116.9% | 89.48% | 93.59% | 106.5% | 101.8% | 108.1% | 76.45% | 111.6% | 58.45% | 105.1% | 105.7% | 106.6% | 110.2% | 91.89% | 99.9% | 84.65% | 103.5% | 106.5% | 116.5% | 111.7% | 108.8% | 110.5% | 108.3% | 94.06% | 109.7% | 114% | 88.6% | 72.51% | 83.52% | 76.55% | 79.08% | 70.34% | 80.6% | 106.3% | 105.4% | 99.78% | 111.6%
| 161 | CCV 1ppb TRACE |  | 06/09/2023 05:43:07 | 1095.1881 | 1052.2148 | 0.98874 | 1.4873 | -1.0254 | 1597.5271 | 1049.0373 | 1097.1047 | 1.2046 | 1.1554 | 1.0953 | 0.76683 | 1.1213 | 0.97591 | 1.2147 | 1.2524 | 1.1915 | 1.1161 | 0.91116 | 1.0394 | 1.2 | 1.0631 | 1.084 | 1.2623 | 0.98008 | 1.101 | 1.0509 | 1.0332 | 1.0577 | 1.1169 | 1.1325 | [7 empty] | 1.0408 | 1.0519 | 0.99979 | 1.0841
|  | RSD |  |  | 10.24% | 9.453% | 1.165% | 56.84% | 7.465% | 31.08% | 9.255% | 14.56% | 7.835% | 5.327% | 1.703% | 0.8516% | 0.9503% | 23.83% | 6.137% | 6.553% | 11.67% | 2.695% | 0.04525% | 2.35% | 14% | 1.092% | 1.434% | 5.952% | 18.21% | 2.681% | 5.53% | 13.23% | 8.371% | 1.608% | 3.128% | [7 empty] | 2.413% | 0.9292% | 1.607% | 1.668%
|  | % Recovery |  |  | 109.5% | 105.2% |  | 148.7% | -102.5% | 159.8% | 104.9% | 109.7% | 120.5% | 115.5% | 109.5% | 76.68% | 112.1% | 97.59% | 121.5% | 125.2% | 119.1% | 111.6% | 91.12% | 103.9% | 120% | 106.3% | 108.4% | 126.2% | 98.01% | 110.1% | 105.1% | 103.3% | 105.8% | 111.7% | 113.3% | 88.39% | 73.37% | 77.01% | 78.45% | 79.2% | 72.65% | 80.37% | 104.1% | 105.2% | 99.98% | 108.4%
| 162 | CCV 1ppb TRACE |  | 06/09/2023 05:49:25 | 1024.0303 | 931.1209 | 0.98854 | 1.3009 | -1.1378 | 1394.8765 | 938.2511 | 944.6971 | 1.135 | 1.0838 | 1.1303 | 0.79247 | 1.201 | 0.76478 | 1.0956 | 1.0554 | 1.0791 | 1.147 | 0.94372 | 1.0306 | 0.90568 | 1.1293 | 1.1229 | 1.2505 | 1.0986 | 1.116 | 1.1163 | 1.1167 | 1.0692 | 1.1405 | 1.2174 | [7 empty] | 1.0386 | 1.0706 | 1.0181 | 1.095
|  | RSD |  |  | 4.816% | 1.523% | 2.935% | 7.275% | 2.997% | 3.66% | 4.012% | 6.026% | 9.329% | 3.856% | 2.958% | 5.47% | 6.54% | 5.144% | 4.622% | 2.074% | 2.654% | 1.956% | 3.751% | 3.749% | 12.37% | 4.069% | 6.078% | 12.37% | 15.24% | 3.65% | 7.184% | 10.47% | 18.33% | 0.7193% | 5.747% | [7 empty] | 2.438% | 0.4244% | 2.513% | 1.264%
|  | % Recovery |  |  | 102.4% | 93.11% |  | 130.1% | -113.8% | 139.5% | 93.83% | 94.47% | 113.5% | 108.4% | 113% | 79.25% | 120.1% | 76.48% | 109.6% | 105.5% | 107.9% | 114.7% | 94.37% | 103.1% | 90.57% | 112.9% | 112.3% | 125% | 109.9% | 111.6% | 111.6% | 111.7% | 106.9% | 114% | 121.7% | 86.4% | 70.84% | 81.22% | 75.38% | 78.45% | 71.86% | 74.84% | 103.9% | 107.1% | 101.8% | 109.5%
| 176 | CCV 1ppb TRACE |  | 06/09/2023 07:12:04 | 1051.1689 | 1020.0044 | 0.97899 | 1.2627 | -0.93036 | 1309.7646 | 998.1394 | 1088.2421 | 1.155 | 1.0781 | 1.1414 | 0.83223 | 1.0851 | 0.744 | 1.0979 | 1.1357 | 1.1135 | 1.1581 | 0.9574 | 0.96994 | 0.82821 | 1.0186 | 1.0262 | 1.2978 | 1.0821 | 0.99643 | 1.0561 | 1.0974 | 0.92919 | 1.1757 | 1.0745 | [7 empty] | 1.0438 | 1.0858 | 0.9751 | 1.0914
|  | RSD |  |  | 2.096% | 8.791% | 2.119% | 36.62% | 8.307% | 11.66% | 9.481% | 4.025% | 5.263% | 5.185% | 2.633% | 8.386% | 12.71% | 16.64% | 2.271% | 1.033% | 2.144% | 3.186% | 0.6976% | 6.556% | 18.16% | 5.99% | 10.09% | 4.892% | 5.898% | 15.73% | 14.73% | 4.445% | 13.02% | 2.287% | 10.71% | [7 empty] | 2.145% | 1.546% | 1.159% | 0.5529%
|  | % Recovery |  |  | 105.1% | 102% |  | 126.3% | -93.04% | 131% | 99.81% | 108.8% | 115.5% | 107.8% | 114.1% | 83.22% | 108.5% | 74.4% | 109.8% | 113.6% | 111.4% | 115.8% | 95.74% | 96.99% | 82.82% | 101.9% | 102.6% | 129.8% | 108.2% | 99.64% | 105.6% | 109.7% | 92.92% | 117.6% | 107.5% | 86.32% | 73.62% | 78.62% | 72.31% | 75.2% | 66.22% | 80.06% | 104.4% | 108.6% | 97.51% | 109.1%
| 177 | CCV 1ppb TRACE |  | 06/09/2023 07:18:22 | -12.226 | -42.7414 | -0.73088 | -0.82141 | -2.3965 | -1839.6045 | -6.5656 | -30.7443 | 0.0028168 | 0.00066113 | 0.0012662 | -0.53303 | -0.025781 | -1.2086 | 0.0032262 | -0.0040635 | -0.01051 | -0.0092415 | -0.18035 | 0.5533 | 2.4129e-06 | 0.15579 | -0.0043674 | -1.1983 | -0.017104 | 0.0013599 | -0.0028399 | -0.10368 | -0.0014165 | 0.024278 | -0.013457 | [7 empty] | 0.00032415 | -0.00095384 | -0.0025274 | -0.0022858
|  | RSD |  |  | 6.238% | 2.367e-06% | 0.1594% | 0% | 0.235% | 3.066% | 140.4% | 1.343e-06% | 0% | 568.6% | 50.75% | 1.455% | 0.9682% | 2.287% | 49.17% | 82.52% | 13.37% | 1.243% | 0.3661% | 2.148% | 412335% | 0.08493% | 4.184% | 2.3% | 48.14% | 188.4% | 43.97% | 93.2% | 0% | 5.14% | 1.887% | [7 empty] | 47.96% | 17.04% | 10.69% | 1.282%
|  | % Recovery |  |  | -1.223% | -4.274% |  | -82.14% | -239.7% | -184% | -0.6566% | -3.074% | 0.2817% | 0.06611% | 0.1266% | -53.3% | -2.578% | -120.9% | 0.3226% | -0.4064% | -1.051% | -0.9241% | -18.04% | 55.33% | 0.0002413% | 15.58% | -0.4367% | -119.8% | -1.71% | 0.136% | -0.284% | -10.37% | -0.1417% | 2.428% | -1.346% | 91.81% | 74.5% | 89.85% | 84.11% | 85.48% | 67.7% | 79.78% | 0.03241% | -0.09538% | -0.2527% | -0.2286%
| 193 | CCV 1ppb TRACE |  | 06/09/2023 08:53:33 | 1031.8537 | 953.3088 | 1.0689 | 1.3093 | -0.91718 | 1623.6707 | 1033.0956 | 1057.154 | 1.1153 | 1.1272 | 1.1152 | 0.76352 | 1.1597 | 0.63389 | 1.104 | 1.0999 | 1.0772 | 1.1377 | 0.92757 | 1.0033 | 0.96884 | 1.0672 | 1.0893 | 1.3203 | 1.0705 | 1.1046 | 1.14 | 1.0993 | 1.0261 | 1.1307 | 1.1511 | [7 empty] | 1.0569 | 1.0482 | 1.0727 | 1.1084
|  | RSD |  |  | 7.024% | 19.02% | 8.744% | 6.25% | 7.15% | 17.48% | 7.718% | 19.76% | 2.281% | 8.686% | 2.848% | 1.729% | 2.18% | 16.04% | 9.418% | 9.146% | 9.476% | 1.9% | 2.495% | 7.315% | 8.002% | 2.315% | 2.217% | 3.847% | 11.83% | 12.11% | 2.417% | 13.45% | 3.649% | 1.59% | 0.655% | [7 empty] | 2.576% | 2.799% | 5.082% | 1.595%
|  | % Recovery |  |  | 103.2% | 95.33% |  | 130.9% | -91.72% | 162.4% | 103.3% | 105.7% | 111.5% | 112.7% | 111.5% | 76.35% | 116% | 63.39% | 110.4% | 110% | 107.7% | 113.8% | 92.76% | 100.3% | 96.88% | 106.7% | 108.9% | 132% | 107% | 110.5% | 114% | 109.9% | 102.6% | 113.1% | 115.1% | 77.24% | 65.05% | 73.54% | 68.94% | 69.99% | 62.75% | 71.3% | 105.7% | 104.8% | 107.3% | 110.8%
| 194 | CCV 1ppb TRACE |  | 06/09/2023 08:59:50 | 1079.761 | 984.0317 | 0.97784 | 1.0862 | -1.0429 | 1784.1959 | 987.648 | 1085.102 | 1.1324 | 1.0799 | 1.1164 | 0.80172 | 1.1584 | 0.7542 | 1.0848 | 1.1162 | 1.1407 | 1.134 | 0.94972 | 1.0102 | 0.90142 | 1.0622 | 1.0651 | 1.2202 | 1.0165 | 1.1738 | 1.1152 | 1.1001 | 1.1936 | 1.1413 | 1.1734 | [7 empty] | 1.0182 | 1.0627 | 1.0167 | 1.103
|  | RSD |  |  | 6.051% | 9.14% | 4.178% | 30.5% | 4.176% | 19.8% | 3.414% | 2.695% | 5.665% | 4.836% | 3.641% | 1.406% | 0.6465% | 19.81% | 5.792% | 2.521% | 3.909% | 1.333% | 3.248% | 5.336% | 19.64% | 4.917% | 2.108% | 16.92% | 13.27% | 5.178% | 3.553% | 10.36% | 10.54% | 0.6238% | 2.016% | [7 empty] | 2.001% | 1.95% | 0.451% | 1.062%
|  | % Recovery |  |  | 108% | 98.4% |  | 108.6% | -104.3% | 178.4% | 98.76% | 108.5% | 113.2% | 108% | 111.6% | 80.17% | 115.8% | 75.42% | 108.5% | 111.6% | 114.1% | 113.4% | 94.97% | 101% | 90.14% | 106.2% | 106.5% | 122% | 101.7% | 117.4% | 111.5% | 110% | 119.4% | 114.1% | 117.3% | 79.95% | 64.03% | 72.83% | 67.95% | 69.95% | 62.57% | 69.08% | 101.8% | 106.3% | 101.7% | 110.3%
| 209 | CCV 1ppb TRACE |  | 06/09/2023 10:29:20 | 1084.9647 | 993.13 | 0.98655 | 1.2115 | -1.1015 | 1898.0804 | 976.3604 | 1229.5396 | 1.1575 | 1.1118 | 1.1371 | 0.81718 | 1.1555 | 0.81617 | 1.1495 | 1.1762 | 1.1615 | 1.1289 | 0.97656 | 0.98685 | 0.99232 | 1.0786 | 1.0923 | 1.1737 | 1.1536 | 1.0609 | 1.1133 | 1.0824 | 1.0107 | 1.1457 | 1.1698 | [7 empty] | 1.0658 | 1.0737 | 0.9912 | 1.0887
|  | RSD |  |  | 8.107% | 8.048% | 3.358% | 40.7% | 2.384% | 5.548% | 5.101% | 4.262% | 2.089% | 4.863% | 0.8621% | 3.827% | 1.69% | 5.121% | 5.318% | 1.819% | 2.16% | 1.908% | 2.669% | 7.824% | 11.4% | 0.552% | 2.251% | 15.36% | 14.05% | 2.587% | 0.1879% | 10.17% | 14.63% | 1.099% | 1.698% | [7 empty] | 1.905% | 0.2382% | 2.13% | 0.3499%
|  | % Recovery |  |  | 108.5% | 99.31% |  | 121.1% | -110.2% | 189.8% | 97.64% | 123% | 115.7% | 111.2% | 113.7% | 81.72% | 115.5% | 81.62% | 114.9% | 117.6% | 116.2% | 112.9% | 97.66% | 98.69% | 99.23% | 107.9% | 109.2% | 117.4% | 115.4% | 106.1% | 111.3% | 108.2% | 101.1% | 114.6% | 117% | 78.7% | 63.42% | 70.25% | 67.2% | 69.34% | 60.44% | 68.81% | 106.6% | 107.4% | 99.12% | 108.9%
| 210 | CCV 1ppb TRACE |  | 06/09/2023 10:35:37 | 1091.3034 | 1077.1134 | 0.98195 | 1.379 | -1.1796 | 1868.8095 | 966.2532 | 1166.725 | 1.1809 | 1.1451 | 1.1464 | 0.81924 | 1.1395 | 0.86577 | 1.1539 | 1.1773 | 1.165 | 1.133 | 0.95846 | 0.95168 | 0.98756 | 1.0578 | 1.086 | 1.247 | 1.0386 | 1.1102 | 1.0989 | 1.1179 | 0.99256 | 1.1434 | 1.1772 | [7 empty] | 1.0656 | 1.0671 | 0.9878 | 1.0953
|  | RSD |  |  | 2.743% | 4.539% | 3.719% | 20.22% | 2.433% | 14.75% | 7.42% | 15.28% | 2.999% | 3.4% | 1.47% | 5.952% | 2.803% | 18.55% | 6.734% | 6.21% | 4.922% | 0.6822% | 1.862% | 4.413% | 18.82% | 3.318% | 2.154% | 1.126% | 10.59% | 5.083% | 5.158% | 4.663% | 3.446% | 0.5756% | 2.752% | [7 empty] | 0.4478% | 1.527% | 1.03% | 1.185%
|  | % Recovery |  |  | 109.1% | 107.7% |  | 137.9% | -118% | 186.9% | 96.63% | 116.7% | 118.1% | 114.5% | 114.6% | 81.92% | 114% | 86.58% | 115.4% | 117.7% | 116.5% | 113.3% | 95.85% | 95.17% | 98.76% | 105.8% | 108.6% | 124.7% | 103.9% | 111% | 109.9% | 111.8% | 99.26% | 114.3% | 117.7% | 79.66% | 63.97% | 70.52% | 67.27% | 69.7% | 61.8% | 70.66% | 106.6% | 106.7% | 98.78% | 109.5%
| 233 | CCV 1ppb TRACE |  | 06/09/2023 13:36:58 | 1012.0115 | 1031.7646 | 1.1072 | 1.1547 | 0.21933 | 1080.0053 | 1042.3484 | 1008.9618 | 1.0311 | 0.98852 | 1.0257 | 0.64356 | 1.0076 | 0.62252 | 0.97966 | 0.98319 | 1.0464 | 1.0593 | 0.89225 | 0.92828 | 1.0605 | 1.1684 | 1.015 | 1.1284 | 0.98134 | 0.92326 | 0.99144 | 0.90329 | 1.0864 | 1.0278 | 0.99404 | [7 empty] | 1.01 | 1.0126 | 0.98687 | 1.0111
|  | RSD |  |  | 9.766% | 18.05% | 1.649% | 13.76% | 24.16% | 11.52% | 10.41% | 7.88% | 11.87% | 9.978% | 1.724% | 3.643% | 1.99% | 46.73% | 11.43% | 13.97% | 14.91% | 1.68% | 3.377% | 4.979% | 12.89% | 0.8615% | 1.247% | 17.81% | 11.02% | 2.902% | 2.413% | 5.315% | 13.68% | 1.845% | 1.162% | [7 empty] | 0.6759% | 0.3892% | 1.774% | 2.334%
|  | % Recovery |  |  | 101.2% | 103.2% |  | 115.5% | 21.93% | 108% | 104.2% | 100.9% | 103.1% | 98.85% | 102.6% | 64.36% | 100.8% | 62.25% | 97.97% | 98.32% | 104.6% | 105.9% | 89.23% | 92.83% | 106.1% | 116.8% | 101.5% | 112.8% | 98.13% | 92.33% | 99.14% | 90.33% | 108.6% | 102.8% | 99.4% | 99.4% | 96.96% | 102.2% | 96.39% | 96.78% | 96.31% | 99.47% | 101% | 101.3% | 98.69% | 101.1%
| 234 | CCV 1ppb TRACE |  | 06/09/2023 13:43:15 | 1043.7191 | 1063.8988 | 1.0537 | 1.2045 | 0.072191 | 1221.5046 | 1066.8359 | 1057.3282 | 1.079 | 1.0358 | 1.0503 | 0.67741 | 1.0355 | 0.78205 | 1.0177 | 1.0107 | 1.0884 | 1.0503 | 0.88945 | 1.1188 | 1.1159 | 1.1659 | 1.0081 | 1.0799 | 0.97475 | 1.0336 | 1.0393 | 1.0669 | 0.95948 | 1.0445 | 1.0203 | [7 empty] | 1.0321 | 1.027 | 0.96723 | 0.99754
|  | RSD |  |  | 3.224% | 5.488% | 10.81% | 24.91% | 38.03% | 6.29% | 5.217% | 9.634% | 1.087% | 2.626% | 2.231% | 3.048% | 2.189% | 2.791% | 1.718% | 6.812% | 5.712% | 2.427% | 2.791% | 4.257% | 4.526% | 1.112% | 1.161% | 18.47% | 5.473% | 5.234% | 5.702% | 16.96% | 5.091% | 1.121% | 1.152% | [7 empty] | 0.9868% | 0.1706% | 8.251% | 0.3223%
|  | % Recovery |  |  | 104.4% | 106.4% |  | 120.5% | 7.219% | 122.2% | 106.7% | 105.7% | 107.9% | 103.6% | 105% | 67.74% | 103.6% | 78.2% | 101.8% | 101.1% | 108.8% | 105% | 88.94% | 111.9% | 111.6% | 116.6% | 100.8% | 108% | 97.47% | 103.4% | 103.9% | 106.7% | 95.95% | 104.4% | 102% | 97.93% | 94.74% | 98.44% | 94.27% | 95.5% | 95.27% | 96.78% | 103.2% | 102.7% | 96.72% | 99.75%
| 251 | CCV 1ppb TRACE |  | 06/09/2023 15:23:44 | 1030.3002 | 1107.7589 | 1.0969 | 1.1722 | 0.19496 | 1257.8902 | 1038.9043 | 1095.2024 | 1.0234 | 1.069 | 1.0357 | 0.69805 | 1.0109 | 0.6622 | 1.0115 | 1.0752 | 1.0758 | 1.0485 | 0.86927 | 1.0899 | 1.1851 | 1.166 | 1.0262 | 1.1277 | 0.96804 | 0.92377 | 0.97121 | 1.1089 | 0.96313 | 1.0169 | 1.0125 | [7 empty] | 1.0137 | 1.0271 | 1.0182 | 1.1293
|  | RSD |  |  | 6.505% | 7.889% | 2.362% | 28.11% | 12.94% | 42.12% | 3.487% | 18.55% | 10.46% | 10.46% | 1.185% | 8.8% | 2.098% | 23.36% | 6.939% | 4.603% | 4.263% | 1.951% | 2.556% | 2.145% | 13.62% | 0.7333% | 1.815% | 3.962% | 10.27% | 2.209% | 4.397% | 8.929% | 5.522% | 1.418% | 0.1164% | [7 empty] | 2.144% | 2.981% | 0.8428% | 4.61%
|  | % Recovery |  |  | 103% | 110.8% |  | 117.2% | 19.5% | 125.8% | 103.9% | 109.5% | 102.3% | 106.9% | 103.6% | 69.8% | 101.1% | 66.22% | 101.1% | 107.5% | 107.6% | 104.8% | 86.93% | 109% | 118.5% | 116.6% | 102.6% | 112.8% | 96.8% | 92.38% | 97.12% | 110.9% | 96.31% | 101.7% | 101.3% | 95.42% | 91.58% | 92.72% | 90.03% | 85.05% | 91.66% | 95.23% | 101.4% | 102.7% | 101.8% | 112.9%
| 252 | CCV 1ppb TRACE |  | 06/09/2023 15:30:02 | 1048.8007 | 1147.2892 | 1.0941 | 0.72079 | 0.088427 | 1163.7187 | 1048.0881 | 1103.3704 | 1.0533 | 1.0081 | 1.0417 | 0.74341 | 1.0271 | 0.64882 | 1.0202 | 1.0339 | 1.0599 | 1.0761 | 0.89906 | 1.0757 | 1.0736 | 1.2161 | 1.0325 | 1.039 | 1.1258 | 1.0488 | 1.0353 | 0.93568 | 0.82707 | 1.0314 | 1.0313 | [7 empty] | 1.0504 | 1.0656 | 1.0235 | 1.0292
|  | RSD |  |  | 6.06% | 7.246% | 1.295% | 31.44% | 31.92% | 25.69% | 4.52% | 6.678% | 6.098% | 4.76% | 4.061% | 4.611% | 1.78% | 8.45% | 1.371% | 4.599% | 4.775% | 2.718% | 4.127% | 2.565% | 3.991% | 1.063% | 1.593% | 9.065% | 0.2208% | 8.152% | 3.114% | 2.821% | 6.134% | 3.207% | 1.079% | [7 empty] | 1.883% | 2.385% | 1.253% | 1.578%
|  | % Recovery |  |  | 104.9% | 114.7% |  | 72.08% | 8.843% | 116.4% | 104.8% | 110.3% | 105.3% | 100.8% | 104.2% | 74.34% | 102.7% | 64.88% | 102% | 103.4% | 106% | 107.6% | 89.91% | 107.6% | 107.4% | 121.6% | 103.2% | 103.9% | 112.6% | 104.9% | 103.5% | 93.57% | 82.71% | 103.1% | 103.1% | 93.91% | 89.22% | 92.56% | 88.15% | 90.05% | 90.55% | 88.47% | 105% | 106.6% | 102.4% | 102.9%
| 266 | CCV 1ppb TRACE |  | 06/09/2023 16:51:44 | 993.1619 | 1041.3404 | 0.48807 | 0.54156 | -0.019812 | 957.7847 | 999.2865 | 1065.9513 | 0.97358 | 0.96627 | 0.99705 | 0.77491 | 0.95789 | 0.78497 | 0.97097 | 0.96372 | 0.97835 | 0.98649 | 0.8434 | 0.99404 | 1.079 | 1.1465 | 0.99948 | 0.97419 | 1.0197 | 0.92347 | 1.0124 | 0.83096 | 0.9417 | 0.98114 | 0.98494 | [7 empty] | 0.97995 | 0.99255 | 0.99272 | 1.0217
|  | RSD |  |  | 4.434% | 5.904% | 5.144% | 11.67% | 141.7% | 4.261% | 2.727% | 8.229% | 7.121% | 2.546% | 1.493% | 4.435% | 2.897% | 10.59% | 2.751% | 10.44% | 3.831% | 0.4711% | 2.079% | 11.59% | 17.99% | 1.713% | 1.471% | 7.628% | 9.975% | 8.215% | 9.362% | 25.84% | 8.189% | 1.47% | 1.554% | [7 empty] | 3.036% | 0.5124% | 1.343% | 0.2434%
|  | % Recovery |  |  | 99.32% | 104.1% |  | 54.16% | -1.981% | 95.78% | 99.93% | 106.6% | 97.36% | 96.63% | 99.71% | 77.49% | 95.79% | 78.5% | 97.1% | 96.37% | 97.83% | 98.65% | 84.34% | 99.4% | 107.9% | 114.7% | 99.95% | 97.42% | 102% | 92.35% | 101.2% | 83.1% | 94.17% | 98.11% | 98.49% | 90.64% | 85.77% | 90.1% | 86.62% | 86.48% | 86.07% | 88.78% | 98% | 99.26% | 99.27% | 102.2%
| 267 | CCV 1ppb TRACE |  | 06/09/2023 16:58:02 | 966.7869 | 971.3645 | 0.46476 | 0.64628 | -0.065615 | 1056.7536 | 949.6856 | 1108.4414 | 1.0172 | 0.98682 | 0.9931 | 0.79067 | 0.94275 | 0.74872 | 0.96006 | 0.9526 | 1.0304 | 0.98146 | 0.84948 | 1.0453 | 0.88978 | 1.1391 | 0.97712 | 1.0156 | 0.896 | 0.90304 | 1.0256 | 0.93563 | 0.97357 | 0.97889 | 0.98462 | [7 empty] | 1.0051 | 1.0004 | 0.98972 | 1.0196
|  | RSD |  |  | 2.279% | 4.477% | 1.658% | 29.66% | 43.41% | 16.98% | 2.931% | 6.399% | 1.304% | 0.5036% | 2.113% | 6.193% | 1.409% | 7.218% | 1.05% | 5.3% | 0.9493% | 1.64% | 1.356% | 2.044% | 15.82% | 1.141% | 1.274% | 17.95% | 5.149% | 3.635% | 4.712% | 11.34% | 16.91% | 1.953% | 1.563% | [7 empty] | 0.919% | 0.4004% | 0.6319% | 1.023%
|  | % Recovery |  |  | 96.68% | 97.14% |  | 64.63% | -6.562% | 105.7% | 94.97% | 110.8% | 101.7% | 98.68% | 99.31% | 79.07% | 94.27% | 74.87% | 96.01% | 95.26% | 103% | 98.15% | 84.95% | 104.5% | 88.98% | 113.9% | 97.71% | 101.6% | 89.6% | 90.3% | 102.6% | 93.56% | 97.36% | 97.89% | 98.46% | 91.82% | 86.73% | 91.51% | 86.44% | 87.87% | 85.33% | 88.5% | 100.5% | 100% | 98.97% | 102%
| 284 | CCV 1ppb TRACE |  | 06/09/2023 18:37:37 | 990.5272 | 1004.2552 | 0.51466 | 0.59794 | -0.024345 | 1087.9378 | 1024.5979 | 998.689 | 1.0132 | 0.99961 | 0.99533 | 0.77654 | 0.97632 | 0.7694 | 0.98053 | 0.99964 | 0.96853 | 0.98302 | 0.81642 | 0.98861 | 0.99486 | 1.1792 | 0.98661 | 1.1434 | 0.97474 | 0.996 | 0.95995 | 0.95554 | 0.99607 | 0.97894 | 0.9928 | [7 empty] | 0.97334 | 0.97579 | 1.0095 | 1.1754
|  | RSD |  |  | 2.62% | 13.29% | 1.956% | 62.56% | 43.33% | 33.02% | 3.829% | 4.117% | 1.36% | 1.419% | 2.02% | 4.519% | 1.377% | 5.167% | 2.785% | 4.487% | 2.758% | 1.313% | 3.947% | 2.914% | 13.67% | 3.167% | 1.188% | 8.054% | 4.861% | 1.285% | 6.849% | 13.36% | 12.41% | 0.251% | 1.578% | [7 empty] | 1.588% | 1.353% | 1.458% | 2.555%
|  | % Recovery |  |  | 99.05% | 100.4% |  | 59.79% | -2.434% | 108.8% | 102.5% | 99.87% | 101.3% | 99.96% | 99.53% | 77.65% | 97.63% | 76.94% | 98.05% | 99.96% | 96.85% | 98.3% | 81.64% | 98.86% | 99.49% | 117.9% | 98.66% | 114.3% | 97.47% | 99.6% | 95.99% | 95.55% | 99.61% | 97.89% | 99.28% | 90.89% | 84.8% | 87.41% | 85.84% | 76.51% | 85.46% | 83.62% | 97.33% | 97.58% | 100.9% | 117.5%
| 285 | CCV 1ppb TRACE |  | 06/09/2023 18:43:54 | 978.5032 | 984.4313 | 0.50882 | 0.41605 | -0.075403 | 1341.7701 | 1067.8231 | 1044.0416 | 1.0319 | 1.0151 | 0.96791 | 0.79594 | 0.97529 | 0.87173 | 0.98077 | 0.99058 | 1.081 | 0.98512 | 0.84038 | 1.0323 | 1.0578 | 1.1271 | 0.98683 | 1.1375 | 0.93332 | 1.0285 | 0.9997 | 0.93953 | 1.0488 | 0.96639 | 0.97562 | [7 empty] | 0.98266 | 0.98763 | 1.0123 | 1.1862
|  | RSD |  |  | 7.872% | 12.55% | 7.387% | 114.1% | 74.97% | 8.745% | 11.4% | 7.128% | 6.538% | 4.728% | 2.081% | 2.985% | 1.487% | 7.576% | 3.121% | 10.37% | 3.595% | 1.941% | 2.402% | 8.055% | 20.94% | 3.634% | 3.24% | 4.328% | 4.382% | 7.166% | 6.165% | 10.19% | 5.211% | 1.072% | 0.8418% | [7 empty] | 1.934% | 2.154% | 2.94% | 16.72%
|  | % Recovery |  |  | 97.85% | 98.44% |  | 41.61% | -7.54% | 134.2% | 106.8% | 104.4% | 103.2% | 101.5% | 96.79% | 79.59% | 97.53% | 87.17% | 98.08% | 99.06% | 108.1% | 98.51% | 84.04% | 103.2% | 105.8% | 112.7% | 98.68% | 113.8% | 93.33% | 102.9% | 99.97% | 93.95% | 104.9% | 96.64% | 97.56% | 90.78% | 84.41% | 85.74% | 86.39% | 77.73% | 85.08% | 88.25% | 98.27% | 98.76% | 101.2% | 118.6%
QC STD 4
|  | Sample Id | R | Acquisition Time | Na KED 23/23 | Mg KED 24/24 | Al 27/27 | Al KED 27/27 | P Oshift 31/47 | S (34) -1 34/34 | K KED 39/39 | Ca KED 43/43 | V KED 51/51 | Cr (52) 52/52 | Cr (53) NH3 53/53 | Fe (54) NH3 54/54 | Mn Oshift 55/71 | Fe (56) KED 56/56 | Co KED 59/59 | Ni KED 60/60 | Cu (63) KED 63/63 | Cu (65) NH3 65/65 | Zn NH3 66/66 | As 75/75 | As KED 75/75 | As RH Oshift 75/91 | As Oshift 75/91 | Se (77) 77/77 | Se (77) Oshift 77/93 | Se (78) Oshift 78/94 | Se (80) Oshift 80/96 | Se (82) 82/82 | Se (82) Oshift 82/98 | Mo (95) NH3 95/95 | Mo (95) OShift 95/127 | Rh (IS1) 103/103 (IS) | Rh (IS2) 103/103 (IS) | Rh (IS3) 103/103 (IS) | Rh (IS4) 103/103 (IS) | Rh (IS5) 103/103 (IS) | Rh (IS6) 103/103 (IS) | Rh OShift (IS1) 103/119 (IS) | Cd (111) NH3 111/111 | Cd (114) NH3 114/114 | Pb 208/208 | Pb NH3 208/208
| 21 | LFB |  | 06/08/2023 15:27:51 | 9591.8017 | 9708.532 | 8.8195 | 9.0889 | 9.9079 | 10108.5395 | 9525.6146 | 9366.2479 | 9.761 | 9.5988 | 10.0447 | 10.0076 | 10.2607 | 9.7319 | 10.0643 | 10.1776 | 9.6645 | 9.8918 | 10.289 | 9.5474 | 9.0222 | 9.4115 | 10.1802 | 9.4066 | 9.9294 | 9.7458 | 9.9836 | 9.9387 | 9.7442 | 10.648 | 10.4375 | [7 empty] | 9.9167 | 10.1139 | 9.1342 | 9.5711
|  | RSD |  |  | 0.4299% | 4.166% | 1.49% | 3.203% | 2.599% | 4.081% | 0.39% | 2.886% | 2.349% | 2.248% | 3.542% | 4.888% | 1.18% | 2.608% | 2.204% | 3.981% | 0.5029% | 4.352% | 5.151% | 1.552% | 3.605% | 2.171% | 2.16% | 4.458% | 1.342% | 1.226% | 2.018% | 2.886% | 2.605% | 5.182% | 1.785% | [7 empty] | 5.601% | 4.302% | 1.217% | 1.099%
|  | % Recovery |  |  | 95.92% | 97.09% | 88.2% | 90.89% | 99.08% | 101.1% | 95.26% | 93.66% | 97.61% | 95.99% | 100.4% | 100.1% | 102.6% | 97.32% | 100.6% | 101.8% | 96.65% | 98.92% | 102.9% | 95.47% | 90.22% | 94.12% | 101.8% | 94.07% | 99.29% | 97.46% | 99.84% | 99.39% | 97.44% | 106.5% | 104.4% | 89.3% | 83.99% | 88.41% | 83.1% | 87.21% | 84.01% | 84.41% | 99.17% | 101.1% | 91.34% | 95.71%
| 53 | LFB |  | 06/08/2023 18:37:11 | 10161.1608 | 10441.4137 | 9.3182 | 9.2826 | 11.0197 | 11171.1939 | 10037.7278 | 10717.8098 | 10.3496 | 10.2957 | 10.3642 | 10.4413 | 11.0351 | 10.3389 | 10.7057 | 10.6916 | 10.2054 | 10.2777 | 10.5841 | 10.1772 | 9.3958 | 9.9527 | 10.7524 | 10.1355 | 10.3593 | 10.1254 | 10.6399 | 11.1844 | 10.7475 | 10.7279 | 11.067 | [7 empty] | 10.0358 | 10.1841 | 9.3228 | 9.9476
|  | RSD |  |  | 1.853% | 0.4268% | 1.647% | 5.207% | 0.9278% | 4.228% | 3.691% | 3.677% | 1.671% | 2.105% | 0.9434% | 1.748% | 2.922% | 1.424% | 0.8571% | 1.151% | 2.268% | 1.077% | 1.741% | 1.479% | 5.035% | 3.296% | 0.8654% | 1.933% | 5.543% | 4.132% | 4.198% | 0.6829% | 2.537% | 0.6451% | 1.835% | [7 empty] | 0.3681% | 0.9523% | 1.54% | 1.922%
|  | % Recovery |  |  | 101.6% | 104.4% | 93.18% | 92.83% | 110.2% | 111.7% | 100.4% | 107.2% | 103.5% | 103% | 103.6% | 104.4% | 110.4% | 103.4% | 107.1% | 106.9% | 102.1% | 102.8% | 105.8% | 101.8% | 93.96% | 99.53% | 107.5% | 101.4% | 103.6% | 101.3% | 106.4% | 111.8% | 107.5% | 107.3% | 110.7% | 82.41% | 75.56% | 83.18% | 77.18% | 80.41% | 74.51% | 75.8% | 100.4% | 101.8% | 93.23% | 99.48%
| 84 | LFB |  | 06/08/2023 22:10:29 | 10540.991 | 10387.9602 | 9.1304 | 10.6155 | 11.2237 | 11092.9045 | 9792.7414 | 10652.0255 | 10.5019 | 10.2952 | 10.547 | 10.6313 | 11.1506 | 10.2824 | 10.7343 | 10.6871 | 10.2373 | 10.3789 | 10.6509 | 10.1541 | 9.6477 | 9.7924 | 10.7244 | 9.98 | 10.4568 | 10.6225 | 10.7832 | 11.0253 | 10.9244 | 11.0404 | 11.11 | [7 empty] | 10.1415 | 10.2407 | 9.5039 | 10.0777
|  | RSD |  |  | 1.421% | 1.357% | 3.028% | 3.766% | 2.196% | 4.277% | 4.188% | 1.714% | 0.8205% | 0.985% | 0.2087% | 2.029% | 1.604% | 1.396% | 2.036% | 1.402% | 0.9228% | 1.134% | 0.197% | 1.735% | 4.899% | 1.59% | 1.929% | 5.72% | 2.765% | 3.396% | 1.75% | 2.753% | 3.042% | 1.7% | 2.259% | [7 empty] | 0.2889% | 1.944% | 1.933% | 1.276%
|  | % Recovery |  |  | 105.4% | 103.9% | 91.3% | 106.2% | 112.2% | 110.9% | 97.93% | 106.5% | 105% | 103% | 105.5% | 106.3% | 111.5% | 102.8% | 107.3% | 106.9% | 102.4% | 103.8% | 106.5% | 101.5% | 96.48% | 97.92% | 107.2% | 99.8% | 104.6% | 106.2% | 107.8% | 110.3% | 109.2% | 110.4% | 111.1% | 82.63% | 73.1% | 79.36% | 76.15% | 78.63% | 72.02% | 74.32% | 101.4% | 102.4% | 95.04% | 100.8%
| 115 | LFB |  | 06/09/2023 01:12:47 | 10411.443 | 10176.3145 | 8.551 | 10.5198 | 11.3376 | 11254.3572 | 10032.8819 | 10944.4045 | 10.4819 | 10.5276 | 10.5895 | 10.4133 | 11.249 | 10.4677 | 10.8796 | 10.6369 | 10.3729 | 10.3611 | 10.4729 | 9.8152 | 9.7479 | 9.3534 | 10.5721 | 9.653 | 10.4995 | 10.634 | 10.5591 | 10.5016 | 10.8548 | 10.8869 | 11.0634 | [7 empty] | 10.1906 | 10.2388 | 9.4088 | 10.2993
|  | RSD |  |  | 1.901% | 3.397% | 2.067% | 12.01% | 1.126% | 3.011% | 1.445% | 4.269% | 3.19% | 2.563% | 0.8825% | 1.437% | 0.7557% | 3.366% | 3.377% | 2.75% | 3.374% | 0.6168% | 1.154% | 0.6864% | 6.076% | 2.81% | 0.8805% | 2.949% | 3.06% | 5.303% | 1.387% | 1.993% | 2.697% | 1.372% | 1.329% | [7 empty] | 0.6576% | 0.9302% | 0.8986% | 1.474%
|  | % Recovery |  |  | 104.1% | 101.8% | 85.51% | 105.2% | 113.4% | 112.5% | 100.3% | 109.4% | 104.8% | 105.3% | 105.9% | 104.1% | 112.5% | 104.7% | 108.8% | 106.4% | 103.7% | 103.6% | 104.7% | 98.15% | 97.48% | 93.53% | 105.7% | 96.53% | 105% | 106.3% | 105.6% | 105% | 108.5% | 108.9% | 110.6% | 84.86% | 70.85% | 76.43% | 73.71% | 75.98% | 70.25% | 74.42% | 101.9% | 102.4% | 94.09% | 103%
| 146 | LFB |  | 06/09/2023 04:15:06 | 9655.3023 | 9422.9929 | 7.968 | 9.4554 | 11.2017 | 10083.7246 | 9111.2353 | 10266.8773 | 9.9671 | 10.052 | 10.5737 | 10.432 | 11.2055 | 9.8508 | 10.3153 | 10.4705 | 10.2236 | 10.2428 | 10.4287 | 9.8336 | 9.127 | 9.4018 | 10.7556 | 9.6608 | 10.8031 | 10.6016 | 10.9115 | 10.3717 | 10.5244 | 10.869 | 11.2639 | [7 empty] | 10.1629 | 10.2121 | 9.5371 | 10.3114
|  | RSD |  |  | 1.269% | 2.343% | 1.815% | 4.982% | 1.072% | 2.292% | 2.476% | 3.998% | 1.207% | 3.174% | 1.434% | 0.9814% | 3.594% | 1.976% | 0.5543% | 2.492% | 0.8441% | 0.7075% | 0.3766% | 1.311% | 6.331% | 1.806% | 1.575% | 1.321% | 6.32% | 1.548% | 1.687% | 3.361% | 9.507% | 0.8262% | 1.807% | [7 empty] | 2.098% | 0.4071% | 1.829% | 1.052%
|  | % Recovery |  |  | 96.55% | 94.23% | 79.68% | 94.55% | 112% | 100.8% | 91.11% | 102.7% | 99.67% | 100.5% | 105.7% | 104.3% | 112.1% | 98.51% | 103.2% | 104.7% | 102.2% | 102.4% | 104.3% | 98.34% | 91.27% | 94.02% | 107.6% | 96.61% | 108% | 106% | 109.1% | 103.7% | 105.2% | 108.7% | 112.6% | 82.89% | 67.17% | 76.28% | 71.85% | 73.86% | 66.55% | 71.39% | 101.6% | 102.1% | 95.37% | 103.1%
| 147 | LFB |  | 06/09/2023 04:21:22 | 9656.6532 | 9273.0567 | 8.353 | 9.8519 | 11.0636 | 10003.9804 | 8980.1949 | 10216.0904 | 10.1707 | 9.9523 | 10.5156 | 10.364 | 11.1921 | 9.9731 | 10.4615 | 10.4314 | 10.0178 | 10.2409 | 10.5259 | 10.0852 | 9.556 | 9.1411 | 10.6072 | 10.1785 | 10.7257 | 10.5576 | 10.7993 | 10.7729 | 10.8282 | 10.8881 | 11.1021 | [7 empty] | 10.101 | 10.267 | 9.6012 | 10.2737
|  | RSD |  |  | 1.556% | 2.512% | 2.297% | 11.82% | 1.804% | 2.885% | 2.136% | 3.916% | 1.577% | 2.587% | 1.014% | 2.177% | 1.768% | 4.919% | 2.86% | 2.969% | 2.783% | 0.9122% | 0.9398% | 2.412% | 1.867% | 2.937% | 1.726% | 1.244% | 4.284% | 1.91% | 5.003% | 4.226% | 1.87% | 1.18% | 1.37% | [7 empty] | 1.071% | 1.201% | 1.171% | 1.935%
|  | % Recovery |  |  | 96.57% | 92.73% | 83.53% | 98.52% | 110.6% | 100% | 89.8% | 102.2% | 101.7% | 99.52% | 105.2% | 103.6% | 111.9% | 99.73% | 104.6% | 104.3% | 100.2% | 102.4% | 105.3% | 100.9% | 95.56% | 91.41% | 106.1% | 101.8% | 107.3% | 105.6% | 108% | 107.7% | 108.3% | 108.9% | 111% | 82.78% | 69.07% | 77% | 72.89% | 74.25% | 69.35% | 74.53% | 101% | 102.7% | 96.01% | 102.7%
| 178 | LFB |  | 06/09/2023 07:24:08 | 10254.4035 | 9776.8176 | 7.8853 | 9.9854 | 11.6517 | 10403.5171 | 9463.4849 | 10132.481 | 10.6011 | 10.5571 | 10.8327 | 10.9165 | 11.4834 | 10.0498 | 10.6216 | 10.4949 | 10.2395 | 10.4629 | 10.6167 | 9.6694 | 9.2012 | 9.2974 | 10.7981 | 9.6129 | 10.6855 | 10.7273 | 10.9476 | 10.3157 | 10.2766 | 10.9803 | 11.2306 | [7 empty] | 10.0459 | 10.144 | 8.9354 | 10.0646
|  | RSD |  |  | 1.994% | 4.674% | 1.721% | 1.702% | 1.205% | 5.903% | 1.74% | 3.07% | 3.286% | 1.969% | 0.202% | 1.012% | 0.9918% | 3.176% | 1.946% | 2.801% | 4.394% | 1.548% | 1.591% | 2.928% | 1.102% | 3.37% | 2.385% | 2.264% | 3.129% | 0.4101% | 2.009% | 3.492% | 2.506% | 2.312% | 0.9785% | [7 empty] | 1.284% | 1.563% | 2.897% | 1.129%
|  | % Recovery |  |  | 102.5% | 97.77% | 78.85% | 99.85% | 116.5% | 104% | 94.63% | 101.3% | 106% | 105.6% | 108.3% | 109.2% | 114.8% | 100.5% | 106.2% | 104.9% | 102.4% | 104.6% | 106.2% | 96.69% | 92.01% | 92.97% | 108% | 96.13% | 106.9% | 107.3% | 109.5% | 103.2% | 102.8% | 109.8% | 112.3% | 84.03% | 65.53% | 73.58% | 69.33% | 70.98% | 64.61% | 70.77% | 100.5% | 101.4% | 89.35% | 100.6%
| 179 | LFB |  | 06/09/2023 07:30:24 | 10017.5401 | 9875.7031 | 8.3478 | 9.8424 | 11.7042 | 10429.6148 | 9441.887 | 10498.9096 | 10.5842 | 10.3864 | 11.0048 | 10.8546 | 11.5103 | 10.3028 | 10.6797 | 10.5309 | 10.309 | 10.4596 | 10.6408 | 10.151 | 9.3122 | 9.3152 | 10.8641 | 10.416 | 11.3736 | 10.5987 | 10.7255 | 10.4707 | 11.0097 | 11.1163 | 11.3147 | [7 empty] | 9.9849 | 10.0809 | 9.3547 | 9.9893
|  | RSD |  |  | 4.066% | 1.718% | 0.3604% | 5.98% | 0.621% | 8.474% | 2.85% | 2.742% | 4.5% | 5.463% | 0.5772% | 0.1967% | 2.567% | 5.021% | 2.151% | 3.469% | 3.617% | 0.5835% | 0.4692% | 2.158% | 5.118% | 1.841% | 2.513% | 4.607% | 6.218% | 1.491% | 2.875% | 2.797% | 2.018% | 1.238% | 3.546% | [7 empty] | 2.159% | 0.2405% | 0.1738% | 0.828%
|  | % Recovery |  |  | 100.2% | 98.76% | 83.48% | 98.42% | 117% | 104.3% | 94.42% | 105% | 105.8% | 103.9% | 110% | 108.5% | 115.1% | 103% | 106.8% | 105.3% | 103.1% | 104.6% | 106.4% | 101.5% | 93.12% | 93.15% | 108.6% | 104.2% | 113.7% | 106% | 107.3% | 104.7% | 110.1% | 111.2% | 113.1% | 80.51% | 65.68% | 73.13% | 68.88% | 71.45% | 62.84% | 71.16% | 99.85% | 100.8% | 93.55% | 99.89%
| 211 | LFB |  | 06/09/2023 10:41:25 | 9832.9683 | 9512.4877 | 7.8929 | 10.3227 | 12.0467 | 10952.736 | 9277.6341 | 10700.7818 | 10.3999 | 10.2808 | 10.8458 | 10.6165 | 11.4097 | 9.9622 | 10.3816 | 10.4985 | 10.1137 | 10.2807 | 10.6458 | 9.8879 | 9.0591 | 8.9786 | 10.8905 | 9.7347 | 10.7941 | 10.8602 | 10.8534 | 10.7314 | 11.0785 | 11.0507 | 11.4922 | [7 empty] | 10.05 | 10.2693 | 9.0812 | 10.0593
|  | RSD |  |  | 0.3722% | 0.7095% | 0.4092% | 2.302% | 1.635% | 2.961% | 2.288% | 2.799% | 2.072% | 0.7586% | 1.42% | 3.045% | 1.008% | 0.2743% | 0.8897% | 2.966% | 1.508% | 1.621% | 1.887% | 0.6751% | 0.5153% | 1.212% | 1.094% | 1.592% | 0.8952% | 2.448% | 3.068% | 3.443% | 3.982% | 1.435% | 1.842% | [7 empty] | 1.486% | 1.009% | 0.4028% | 1.691%
|  | % Recovery |  |  | 98.33% | 95.12% | 78.93% | 103.2% | 120.5% | 109.5% | 92.78% | 107% | 104% | 102.8% | 108.5% | 106.2% | 114.1% | 99.62% | 103.8% | 105% | 101.1% | 102.8% | 106.5% | 98.88% | 90.59% | 89.79% | 108.9% | 97.35% | 107.9% | 108.6% | 108.5% | 107.3% | 110.8% | 110.5% | 114.9% | 75.43% | 58.71% | 68.28% | 62.56% | 64.34% | 58.13% | 66.21% | 100.5% | 102.7% | 90.81% | 100.6%
| 212 | LFB |  | 06/09/2023 10:47:41 | 9938.389 | 9253.4913 | 8.1182 | 9.945 | 12.5979 | 10929.352 | 9350.0958 | 10403.04 | 10.348 | 10.2445 | 10.8123 | 10.6193 | 11.4597 | 9.8958 | 10.5201 | 10.4879 | 10.1432 | 10.3584 | 10.6537 | 10.2057 | 9.4876 | 9.351 | 10.8188 | 10.1654 | 10.7933 | 10.7603 | 11.044 | 10.8664 | 10.9967 | 11.1718 | 11.446 | [7 empty] | 10.1133 | 10.2828 | 9.1207 | 10.1033
|  | RSD |  |  | 1.363% | 0.9576% | 0.4975% | 10.58% | 3.418% | 4.138% | 1.906% | 4.119% | 3.203% | 0.586% | 0.8713% | 0.8031% | 2.468% | 1.838% | 0.9471% | 0.957% | 1.795% | 1.94% | 0.6845% | 1.999% | 6.535% | 4.748% | 1.459% | 1.278% | 1.138% | 2.432% | 2.577% | 0.8326% | 5.092% | 0.8677% | 1.182% | [7 empty] | 1.021% | 1.875% | 1.393% | 1.396%
|  | % Recovery |  |  | 99.38% | 92.53% | 81.18% | 99.45% | 126% | 109.3% | 93.5% | 104% | 103.5% | 102.4% | 108.1% | 106.2% | 114.6% | 98.96% | 105.2% | 104.9% | 101.4% | 103.6% | 106.5% | 102.1% | 94.88% | 93.51% | 108.2% | 101.7% | 107.9% | 107.6% | 110.4% | 108.7% | 110% | 111.7% | 114.5% | 75.3% | 59.84% | 69.47% | 63.35% | 65.19% | 56.98% | 64.4% | 101.1% | 102.8% | 91.21% | 101%
| 235 | LFB |  | 06/09/2023 13:49:03 | 9526.0582 | 9724.8269 | 8.4123 | 8.6757 | 10.5802 | 9700.8945 | 9346.9596 | 9322.9533 | 9.6882 | 9.7567 | 9.794 | 9.6263 | 9.7864 | 9.3537 | 9.5731 | 9.715 | 9.8361 | 9.7629 | 9.8083 | 9.8132 | 10.5522 | 9.1845 | 10.0564 | 9.7069 | 10.5715 | 9.8014 | 9.8874 | 9.8269 | 9.9559 | 9.7918 | 9.6215 | [7 empty] | 9.6619 | 9.7692 | 8.9503 | 9.0958
|  | RSD |  |  | 1.971% | 2.693% | 2.087% | 2.464% | 2.013% | 1.746% | 2.139% | 3.888% | 2.24% | 2.933% | 0.566% | 2.028% | 1.221% | 3.181% | 0.8417% | 2.321% | 0.9363% | 0.3318% | 1.594% | 2.129% | 4.804% | 2.38% | 1.219% | 2.705% | 2.723% | 0.6988% | 1.754% | 3.182% | 5.108% | 0.2897% | 1.027% | [7 empty] | 0.623% | 0.2285% | 0.9035% | 1.009%
|  | % Recovery |  |  | 95.26% | 97.25% | 84.12% | 86.76% | 105.8% | 97.01% | 93.47% | 93.23% | 96.88% | 97.57% | 97.94% | 96.26% | 97.86% | 93.54% | 95.73% | 97.15% | 98.36% | 97.63% | 98.08% | 98.13% | 105.5% | 91.84% | 100.6% | 97.07% | 105.7% | 98.01% | 98.87% | 98.27% | 99.56% | 97.92% | 96.22% | 94.3% | 89.18% | 92.98% | 89.97% | 90.47% | 92.42% | 92.17% | 96.62% | 97.69% | 89.5% | 90.96%
| 268 | LFB |  | 06/09/2023 17:03:50 | 9262.4846 | 9471.6401 | 8.1244 | 9.0627 | 10.738 | 9258.761 | 9226.0767 | 9558.0387 | 9.5624 | 9.4227 | 9.5924 | 9.4899 | 9.6537 | 9.1613 | 9.3395 | 9.3215 | 9.4256 | 9.4542 | 9.5007 | 9.5848 | 10.4672 | 9.226 | 9.914 | 9.6342 | 10.2524 | 9.8394 | 9.8034 | 9.6658 | 10.2339 | 9.6754 | 9.7665 | [7 empty] | 9.7837 | 9.8395 | 9.4397 | 9.6013
|  | RSD |  |  | 3.367% | 2.065% | 0.8945% | 3.709% | 1.459% | 2.052% | 4.04% | 6.098% | 0.2166% | 0.6783% | 1.158% | 1.358% | 0.7748% | 0.8552% | 3.031% | 1.757% | 0.8996% | 1.666% | 2.012% | 2.059% | 8.005% | 1.009% | 0.9985% | 4.012% | 4.183% | 1.46% | 0.7016% | 3.069% | 3.291% | 0.2506% | 1.327% | [7 empty] | 0.4803% | 1.302% | 1.585% | 1.491%
|  | % Recovery |  |  | 92.62% | 94.72% | 81.24% | 90.63% | 107.4% | 92.59% | 92.26% | 95.58% | 95.62% | 94.23% | 95.92% | 94.9% | 96.54% | 91.61% | 93.39% | 93.21% | 94.26% | 94.54% | 95.01% | 95.85% | 104.7% | 92.26% | 99.14% | 96.34% | 102.5% | 98.39% | 98.03% | 96.66% | 102.3% | 96.75% | 97.66% | 90.98% | 81.99% | 86.51% | 83.61% | 85.07% | 81.96% | 83.14% | 97.84% | 98.4% | 94.4% | 96.01%
QC STD 5
|  | Sample Id | R | Acquisition Time | Na KED 23/23 | Mg KED 24/24 | Al 27/27 | Al KED 27/27 | P Oshift 31/47 | S (34) -1 34/34 | K KED 39/39 | Ca KED 43/43 | V KED 51/51 | Cr (52) 52/52 | Cr (53) NH3 53/53 | Fe (54) NH3 54/54 | Mn Oshift 55/71 | Fe (56) KED 56/56 | Co KED 59/59 | Ni KED 60/60 | Cu (63) KED 63/63 | Cu (65) NH3 65/65 | Zn NH3 66/66 | As 75/75 | As KED 75/75 | As RH Oshift 75/91 | As Oshift 75/91 | Se (77) 77/77 | Se (77) Oshift 77/93 | Se (78) Oshift 78/94 | Se (80) Oshift 80/96 | Se (82) 82/82 | Se (82) Oshift 82/98 | Mo (95) NH3 95/95 | Mo (95) OShift 95/127 | Rh (IS1) 103/103 (IS) | Rh (IS2) 103/103 (IS) | Rh (IS3) 103/103 (IS) | Rh (IS4) 103/103 (IS) | Rh (IS5) 103/103 (IS) | Rh (IS6) 103/103 (IS) | Rh OShift (IS1) 103/119 (IS) | Cd (111) NH3 111/111 | Cd (114) NH3 114/114 | Pb 208/208 | Pb NH3 208/208
| 22 | LLOQ 0.01 |  | 06/08/2023 15:33:39 | -1.6627 | -30.1172 | -0.23747 | -0.2305 | -1.0393 | 358.4147 | 18.1369 | -16 | 0.01364 | 0.012039 | 0.010911 | -0.37113 | 0.0079147 | -0.38952 | 0.01076 | 0.01227 | 0.00791 | 0.011354 | -0.038124 | 0.033303 | 0.0045282 | 0.1668 | 0.0078158 | 0.013396 | 0.004558 | 0.015157 | 0.0084529 | -0.0978 | 0.017665 | 0.014279 | 0.012314 | [7 empty] | 0.010372 | 0.010054 | 0.0099926 | 0.0090555
|  | RSD |  |  | 308.6% | 12.31% | 5.451% | 41.13% | 0.9484% | 38.48% | 22.98% | 58.61% | 43.26% | 0.352% | 29.41% | 3.33% | 28.94% | 8.545% | 18.02% | 18.77% | 63.54% | 4.325% | 15.09% | 58.68% | 196.4% | 1.104% | 24.52% | 886.9% | 374.5% | 26.74% | 33.53% | 119.8% | 53.41% | 11.34% | 10.62% | [7 empty] | 13.96% | 1.238% | 6.908% | 1.477%
|  | % Recovery |  |  | -16.63% | -301.2% |  | -2305% | -10393% | 3584% | 181.4% | -160% | 136.4% | 120.4% | 109.1% | -3711% | 79.15% | -3895% | 107.6% | 122.7% | 79.1% | 113.5% | -381.2% | 333% | 45.28% | 1668% | 78.16% | 134% | 45.58% | 151.6% | 84.53% | -978% | 176.6% | 142.8% | 123.1% | 99.38% | 92.53% | 94.36% | 94% | 97.31% | 90.98% | 93.77% | 103.7% | 100.5% | 99.93% | 90.56%
| 236 | LLOQ 0.01 |  | 06/09/2023 13:54:51 | 34.7293 | 24.036 | -0.18318 | -0.13095 | 0.33499 | -45.9485 | 49.555 | 76.3561 | 0.021211 | 0.0062971 | 0.0080325 | -0.36542 | 0.037102 | -0.41536 | 0.0093336 | 0.011693 | 0.0076755 | 0.010278 | 0.025539 | 0.061947 | 0.016909 | 0.27198 | 0.011849 | 0.11901 | -0.037815 | 0.014684 | 0.0096409 | 0.0045345 | -0.028951 | 0.043077 | 0.023431 | [7 empty] | 0.012204 | 0.0080342 | 0.0083319 | 0.0096463
|  | RSD |  |  | 0.2893% | 14.89% | 3.99% | 220.9% | 20.29% | 216.7% | 16.16% | 9.341% | 17.14% | 87.27% | 18.08% | 3.164% | 4.878% | 8.803% | 11.9% | 145.9% | 46.07% | 10.75% | 107.5% | 31.88% | 69.33% | 0.666% | 17.31% | 139.9% | 39.87% | 100.7% | 57.56% | 1442% | 0.8341% | 6.331% | 5.867% | [7 empty] | 7.97% | 15.3% | 2.014% | 4.699%
|  | % Recovery |  |  | 347.3% | 240.4% |  | -1310% | 3350% | -459.5% | 495.6% | 763.6% | 212.1% | 62.97% | 80.32% | -3654% | 371% | -4154% | 93.34% | 116.9% | 76.75% | 102.8% | 255.4% | 619.5% | 169.1% | 2720% | 118.5% | 1190% | -378.1% | 146.8% | 96.41% | 45.35% | -289.5% | 430.8% | 234.3% | 97.94% | 95.36% | 103.2% | 97.37% | 96.99% | 96.06% | 95.85% | 122% | 80.34% | 83.32% | 96.46%
QC STD 6
|  | Sample Id | R | Acquisition Time | Na KED 23/23 | Mg KED 24/24 | Al 27/27 | Al KED 27/27 | P Oshift 31/47 | S (34) -1 34/34 | K KED 39/39 | Ca KED 43/43 | V KED 51/51 | Cr (52) 52/52 | Cr (53) NH3 53/53 | Fe (54) NH3 54/54 | Mn Oshift 55/71 | Fe (56) KED 56/56 | Co KED 59/59 | Ni KED 60/60 | Cu (63) KED 63/63 | Cu (65) NH3 65/65 | Zn NH3 66/66 | As 75/75 | As KED 75/75 | As RH Oshift 75/91 | As Oshift 75/91 | Se (77) 77/77 | Se (77) Oshift 77/93 | Se (78) Oshift 78/94 | Se (80) Oshift 80/96 | Se (82) 82/82 | Se (82) Oshift 82/98 | Mo (95) NH3 95/95 | Mo (95) OShift 95/127 | Rh (IS1) 103/103 (IS) | Rh (IS2) 103/103 (IS) | Rh (IS3) 103/103 (IS) | Rh (IS4) 103/103 (IS) | Rh (IS5) 103/103 (IS) | Rh (IS6) 103/103 (IS) | Rh OShift (IS1) 103/119 (IS) | Cd (111) NH3 111/111 | Cd (114) NH3 114/114 | Pb 208/208 | Pb NH3 208/208
| 23 | CCV 10 ppb MAJORS |  | 06/08/2023 15:39:28 | -12.8383 | -41.1886 | -0.73032 | -0.80788 | -2.3731 | -1892.7639 | -3.8651 | -28.9428 | 0.0032821 | 0.0021771 | 0.0008629 | -0.53724 | -0.025739 | -1.1955 | 0.0040634 | -0.00086363 | -0.0091461 | -0.0092637 | -0.1815 | 0.53901 | -0.0057419 | 0.15706 | -0.0030757 | -1.1802 | -0.01521 | 0.0015557 | 0.00049202 | -0.049839 | 0.0041782 | -0.011053 | -0.013216 | [7 empty] | 0.00036437 | -0.00088094 | -0.0026241 | -0.0023583
|  | RSD |  |  | 8.007% | 6.53% | 0.05544% | 2.902% | 0.2924% | 1.745% | 85.87% | 10.78% | 24.56% | 84.48% | 75.38% | 0.743% | 0.5204% | 1.197% | 58.66% | 345.7% | 19.23% | 3.912% | 0.5899% | 0.559% | 1.756e-06% | 0.4532% | 22.48% | 2.887% | 37.89% | 198% | 324.1% | 50.84% | 116.1% | 2.355% | 0.6336% | [7 empty] | 2.529% | 15.37% | 1.52% | 4.686%
|  | % Recovery |  |  | -0.1284% | -0.4119% | -7.303% | -8.079% | -23.73% | -18.93% | -0.03865% | -0.2894% | 0.03282% | 0.02177% | 0.008629% | -5.372% | -0.2574% | -11.96% | 0.04063% | -0.008636% | -0.09146% | -0.09264% | -1.815% | 5.39% | -0.05742% | 1.571% | -0.03076% | -11.8% | -0.1521% | 0.01556% | 0.00492% | -0.4984% | 0.04178% | -0.1105% | -0.1322% | 107.1% | 103.8% | 108.8% | 112.7% | 111.4% | 98.11% | 101.1% | 0.003644% | -0.008809% | -0.02624% | -0.02358%
| 24 | CCV 10 ppb MAJORS |  | 06/08/2023 15:45:42 | 10570.1578 | 10848.3085 | 8.4876 | 9.8854 | 7.878 | 10722.0026 | 10934.0449 | 10683.3747 | 10.7684 | 10.625 | 9.9121 | 9.8704 | 10.3925 | 10.5423 | 10.752 | 10.8388 | 10.5791 | 9.6384 | 9.7375 | 9.9639 | 9.7083 | 9.3648 | 9.9952 | 9.2899 | 9.8888 | 9.6532 | 9.6943 | 10.0276 | 9.5757 | 10.4552 | 10.571 | [7 empty] | 9.7097 | 9.8231 | 9.3278 | 9.859
|  | RSD |  |  | 3.924% | 3.826% | 1.675% | 7.245% | 0.858% | 5.07% | 4.18% | 7.165% | 5.051% | 3.04% | 0.3671% | 2.073% | 1.539% | 5.295% | 2.08% | 3.351% | 3.396% | 0.5702% | 1.153% | 1.958% | 5.212% | 1.719% | 0.582% | 2.74% | 2.749% | 3.096% | 2.426% | 2.537% | 5.963% | 1.111% | 0.6585% | [7 empty] | 2.287% | 2.446% | 1.225% | 0.7778%
|  | % Recovery |  |  | 105.7% | 108.5% | 84.88% | 98.85% | 78.78% | 107.2% | 109.3% | 106.8% | 107.7% | 106.2% | 99.12% | 98.7% | 103.9% | 105.4% | 107.5% | 108.4% | 105.8% | 96.38% | 97.38% | 99.64% | 97.08% | 93.65% | 99.95% | 92.9% | 98.89% | 96.53% | 96.94% | 100.3% | 95.76% | 104.6% | 105.7% | 88.57% | 83.78% | 83.11% | 85.07% | 86.08% | 82.86% | 83.09% | 97.1% | 98.23% | 93.28% | 98.59%
| 38 | CCV 10 ppb MAJORS |  | 06/08/2023 17:07:25 | 10449.6951 | 10827.3927 | 8.8754 | 9.385 | 8.2383 | 10773.6911 | 10424.1421 | 10695.1868 | 10.6484 | 10.3125 | 10.4227 | 10.1615 | 10.7974 | 10.6119 | 10.7221 | 10.7614 | 10.186 | 10.1361 | 10.2327 | 10.0147 | 9.7206 | 9.9877 | 10.6143 | 9.7898 | 10.0967 | 9.9586 | 10.3682 | 10.7447 | 10.154 | 10.8832 | 10.8953 | [7 empty] | 10.0934 | 10.2379 | 9.6027 | 10.3107
|  | RSD |  |  | 1.398% | 2.301% | 2.136% | 4.211% | 2.162% | 5.318% | 3.828% | 4.548% | 1.319% | 2.001% | 0.6263% | 2.353% | 2.792% | 1.879% | 1.781% | 2.393% | 0.7663% | 1.571% | 0.7093% | 2.017% | 6.087% | 1.938% | 1.772% | 5.454% | 6.628% | 2.415% | 2.156% | 2.226% | 2.962% | 1.783% | 1.95% | [7 empty] | 0.8862% | 1.062% | 2.642% | 0.7133%
|  | % Recovery |  |  | 104.5% | 108.3% | 88.75% | 93.85% | 82.38% | 107.7% | 104.2% | 107% | 106.5% | 103.1% | 104.2% | 101.6% | 108% | 106.1% | 107.2% | 107.6% | 101.9% | 101.4% | 102.3% | 100.1% | 97.21% | 99.88% | 106.1% | 97.9% | 101% | 99.59% | 103.7% | 107.4% | 101.5% | 108.8% | 109% | 81.72% | 75.72% | 80.3% | 77.33% | 80.42% | 75.71% | 74.68% | 100.9% | 102.4% | 96.03% | 103.1%
| 39 | CCV 10 ppb MAJORS |  | 06/08/2023 17:13:42 | 10344.4684 | 10497.8461 | 8.8338 | 9.7847 | 8.2239 | 10880.8404 | 10208.6794 | 10646.9174 | 10.4707 | 10.391 | 10.2279 | 10.1228 | 10.7894 | 10.4285 | 10.7659 | 10.7974 | 10.1789 | 9.9797 | 10.0794 | 9.9816 | 9.4301 | 9.5176 | 10.4027 | 9.9573 | 9.9292 | 10.162 | 10.2351 | 10.7463 | 10.1786 | 10.762 | 10.8587 | [7 empty] | 9.9789 | 10.0871 | 9.5894 | 10.074
|  | RSD |  |  | 2.163% | 3.52% | 2.496% | 1.908% | 3.549% | 3.049% | 3.207% | 3.507% | 2.41% | 1.684% | 1.536% | 2.545% | 1.108% | 1.507% | 1.972% | 0.9588% | 1.03% | 0.3916% | 0.6404% | 2.389% | 2.378% | 2.892% | 2.14% | 3.465% | 3.7% | 1.002% | 4.183% | 3.177% | 0.733% | 0.5248% | 1.059% | [7 empty] | 1.114% | 0.7689% | 1.817% | 2.85%
|  | % Recovery |  |  | 103.4% | 105% | 88.34% | 97.85% | 82.24% | 108.8% | 102.1% | 106.5% | 104.7% | 103.9% | 102.3% | 101.2% | 107.9% | 104.3% | 107.7% | 108% | 101.8% | 99.8% | 100.8% | 99.82% | 94.3% | 95.18% | 104% | 99.57% | 99.29% | 101.6% | 102.4% | 107.5% | 101.8% | 107.6% | 108.6% | 83.52% | 77.58% | 81.94% | 78.91% | 81.81% | 76.17% | 78.78% | 99.79% | 100.9% | 95.89% | 100.7%
| 54 | CCV 10 ppb MAJORS |  | 06/08/2023 18:42:59 | 10458.6717 | 10458.0202 | 8.9889 | 9.0553 | 8.7825 | 10989.2516 | 10201.0749 | 10682.2509 | 10.4222 | 10.3563 | 10.7701 | 10.755 | 11.2439 | 10.4457 | 10.6393 | 10.8685 | 10.4075 | 10.5617 | 10.5997 | 10.1685 | 9.2452 | 9.8126 | 10.6672 | 9.9977 | 10.7093 | 10.3419 | 10.5523 | 10.5936 | 10.8071 | 11.173 | 11.1479 | [7 empty] | 10.2947 | 10.3651 | 9.4158 | 10.1256
|  | RSD |  |  | 3.112% | 4.906% | 1.872% | 10.72% | 1.448% | 9.433% | 3.579% | 4.774% | 4.357% | 3.016% | 1.021% | 1.484% | 0.7007% | 5.581% | 3.58% | 4.051% | 3.982% | 1.078% | 0.2076% | 1.703% | 7.327% | 3.373% | 0.8733% | 3.124% | 4.14% | 0.3015% | 1.247% | 2.745% | 1.721% | 1.023% | 1.688% | [7 empty] | 0.8978% | 0.5974% | 0.7201% | 1.521%
|  | % Recovery |  |  | 104.6% | 104.6% | 89.89% | 90.55% | 87.82% | 109.9% | 102% | 106.8% | 104.2% | 103.6% | 107.7% | 107.5% | 112.4% | 104.5% | 106.4% | 108.7% | 104.1% | 105.6% | 106% | 101.7% | 92.45% | 98.13% | 106.7% | 99.98% | 107.1% | 103.4% | 105.5% | 105.9% | 108.1% | 111.7% | 111.5% | 83.63% | 76.25% | 84.05% | 77.71% | 80.9% | 73.25% | 77% | 102.9% | 103.7% | 94.16% | 101.3%
| 55 | CCV 10 ppb MAJORS |  | 06/08/2023 18:49:21 | 11226.1341 | 11171.3778 | 8.0549 | 10.1827 | 8.6149 | 11890.9129 | 11159.7109 | 11518.3926 | 11.3565 | 11.4204 | 10.7401 | 10.5869 | 11.286 | 11.3798 | 11.5004 | 11.9248 | 11.1092 | 10.4923 | 10.4743 | 9.3896 | 10.3233 | 9.9616 | 10.7008 | 8.8304 | 10.6567 | 10.204 | 10.5193 | 9.8819 | 10.6587 | 11.0819 | 11.1202 | [7 empty] | 10.2097 | 10.2815 | 8.5632 | 10.1849
|  | RSD |  |  | 5.82% | 4.912% | 18.26% | 13.05% | 1.272% | 6.42% | 4.533% | 2.793% | 3.298% | 5.645% | 3.362% | 1.295% | 1.617% | 6.732% | 6.421% | 5.501% | 5.538% | 1.493% | 0.9544% | 15.33% | 7.289% | 1.999% | 1.809% | 16.68% | 3.088% | 1.778% | 1.037% | 16.24% | 1.469% | 0.7366% | 1.877% | [7 empty] | 0.6119% | 0.7003% | 14.85% | 2.218%
|  | % Recovery |  |  | 112.3% | 111.7% | 80.55% | 101.8% | 86.15% | 118.9% | 111.6% | 115.2% | 113.6% | 114.2% | 107.4% | 105.9% | 112.9% | 113.8% | 115% | 119.2% | 111.1% | 104.9% | 104.7% | 93.9% | 103.2% | 99.62% | 107% | 88.3% | 106.6% | 102% | 105.2% | 98.82% | 106.6% | 110.8% | 111.2% | 91.58% | 76.22% | 78.94% | 78.44% | 80.23% | 73.93% | 76% | 102.1% | 102.8% | 85.63% | 101.8%
| 69 | CCV 10 ppb MAJORS |  | 06/08/2023 20:42:27 | 10394.4355 | 10123.1069 | 8.9606 | 10.4524 | 8.5693 | 10636.0102 | 9949.1138 | 10421.0072 | 10.3064 | 10.1256 | 10.7798 | 10.8039 | 11.3846 | 10.0072 | 10.3051 | 10.4742 | 10.0267 | 10.5788 | 10.747 | 10.0081 | 8.8278 | 9.7108 | 10.7809 | 9.6899 | 10.5624 | 10.2631 | 10.5892 | 10.53 | 10.8058 | 11.2381 | 11.3551 | [7 empty] | 10.4028 | 10.5105 | 9.7307 | 10.2385
|  | RSD |  |  | 9.652% | 10.23% | 1.133% | 20.94% | 1.913% | 10.51% | 8.22% | 8.157% | 8.658% | 11.09% | 0.6992% | 2.692% | 1.552% | 12.74% | 12.59% | 9.952% | 10.69% | 0.6192% | 1.034% | 3.569% | 9.276% | 1.619% | 2.043% | 1.661% | 3.644% | 3.145% | 1.369% | 2.275% | 3.362% | 0.9512% | 1.661% | [7 empty] | 1.47% | 1.216% | 2.054% | 1.159%
|  | % Recovery |  |  | 103.9% | 101.2% | 89.61% | 104.5% | 85.69% | 106.4% | 99.49% | 104.2% | 103.1% | 101.3% | 107.8% | 108% | 113.8% | 100.1% | 103.1% | 104.7% | 100.3% | 105.8% | 107.5% | 100.1% | 88.28% | 97.11% | 107.8% | 96.9% | 105.6% | 102.6% | 105.9% | 105.3% | 108.1% | 112.4% | 113.6% | 82.58% | 72.99% | 83.29% | 74.95% | 79.58% | 71.87% | 75.24% | 104% | 105.1% | 97.31% | 102.4%
| 70 | CCV 10 ppb MAJORS |  | 06/08/2023 20:48:44 | 10712.17 | 10386.3941 | 8.9557 | 9.6579 | 8.7643 | 11081.0986 | 10424.5086 | 10876.6374 | 10.7028 | 10.5454 | 10.7862 | 10.679 | 11.3673 | 10.6017 | 11.0031 | 10.9919 | 10.4889 | 10.4813 | 10.6064 | 10.2466 | 9.8342 | 9.9471 | 10.8315 | 10.1333 | 10.4448 | 10.5841 | 10.8201 | 10.7146 | 10.7312 | 11.1392 | 11.438 | [7 empty] | 10.2411 | 10.4511 | 9.6987 | 10.0626
|  | RSD |  |  | 2.531% | 1.324% | 2.353% | 5.286% | 0.4586% | 1.417% | 2.841% | 3.097% | 3.616% | 3.01% | 1.803% | 2.595% | 0.6656% | 5.096% | 1.286% | 1.73% | 2.156% | 1.508% | 1.473% | 2.335% | 6.449% | 0.4683% | 0.5822% | 3.065% | 4.766% | 1.313% | 2.078% | 5.488% | 1.611% | 1.623% | 0.745% | [7 empty] | 1.333% | 0.7408% | 2.958% | 10.63%
|  | % Recovery |  |  | 107.1% | 103.9% | 89.56% | 96.58% | 87.64% | 110.8% | 104.2% | 108.8% | 107% | 105.5% | 107.9% | 106.8% | 113.7% | 106% | 110% | 109.9% | 104.9% | 104.8% | 106.1% | 102.5% | 98.34% | 99.47% | 108.3% | 101.3% | 104.4% | 105.8% | 108.2% | 107.1% | 107.3% | 111.4% | 114.4% | 82.38% | 73.71% | 81.22% | 75.78% | 80.19% | 71.8% | 74.5% | 102.4% | 104.5% | 96.99% | 100.6%
| 85 | CCV 10 ppb MAJORS |  | 06/08/2023 22:16:17 | 10649.4991 | 10628.4209 | 8.8167 | 9.9724 | 9.1292 | 10967.8857 | 10420.3227 | 11135.6009 | 10.6916 | 10.8152 | 11.012 | 10.9275 | 11.4545 | 10.5824 | 11.0306 | 11.03 | 10.424 | 10.6278 | 10.662 | 10.0995 | 9.4107 | 9.7515 | 10.7147 | 10.0855 | 10.2857 | 10.1452 | 10.6641 | 10.4133 | 10.75 | 11.2823 | 11.3646 | [7 empty] | 10.2776 | 10.3963 | 9.593 | 10.2827
|  | RSD |  |  | 6.305% | 2.131% | 3.576% | 5.114% | 1.357% | 3.614% | 4.269% | 5.643% | 2.25% | 2.919% | 0.8148% | 1.892% | 2.444% | 2.07% | 3.013% | 4.064% | 4.566% | 1.909% | 1.324% | 2.204% | 9.699% | 1.49% | 1.926% | 0.9116% | 1.368% | 1.105% | 2.113% | 3.774% | 3.278% | 1.151% | 2.101% | [7 empty] | 0.8158% | 0.4726% | 0.7265% | 0.8947%
|  | % Recovery |  |  | 106.5% | 106.3% | 88.17% | 99.72% | 91.29% | 109.7% | 104.2% | 111.4% | 106.9% | 108.2% | 110.1% | 109.3% | 114.5% | 105.8% | 110.3% | 110.3% | 104.2% | 106.3% | 106.6% | 101% | 94.11% | 97.52% | 107.1% | 100.9% | 102.9% | 101.5% | 106.6% | 104.1% | 107.5% | 112.8% | 113.6% | 83.51% | 74.06% | 81.09% | 75.51% | 79.25% | 72.62% | 75.56% | 102.8% | 104% | 95.93% | 102.8%
| 86 | CCV 10 ppb MAJORS |  | 06/08/2023 22:22:34 | 10400.6221 | 10452.9318 | 8.8292 | 9.3935 | 8.7749 | 10752.0465 | 10239.9338 | 10814.7052 | 10.4035 | 10.4614 | 11.0016 | 10.8312 | 11.4635 | 10.3305 | 10.8005 | 10.5844 | 10.2011 | 10.6095 | 10.6924 | 10.2414 | 9.5224 | 9.7114 | 10.8677 | 10.2623 | 10.7228 | 10.7197 | 10.6732 | 10.8037 | 10.4807 | 11.253 | 11.5059 | [7 empty] | 10.2741 | 10.3278 | 9.6288 | 10.2904
|  | RSD |  |  | 2.333% | 2.473% | 1.703% | 9.724% | 2.115% | 3.677% | 3.352% | 4.043% | 1.159% | 3.066% | 0.1845% | 2.35% | 0.6768% | 4.325% | 3.149% | 4.923% | 4.067% | 1.079% | 1.5% | 1.003% | 4.056% | 3.258% | 1.421% | 2.682% | 5.019% | 2.494% | 1.45% | 2.956% | 4.033% | 1.525% | 1.45% | [7 empty] | 1.205% | 0.8559% | 1.152% | 1.287%
|  | % Recovery |  |  | 104% | 104.5% | 88.29% | 93.93% | 87.75% | 107.5% | 102.4% | 108.1% | 104% | 104.6% | 110% | 108.3% | 114.6% | 103.3% | 108% | 105.8% | 102% | 106.1% | 106.9% | 102.4% | 95.22% | 97.11% | 108.7% | 102.6% | 107.2% | 107.2% | 106.7% | 108% | 104.8% | 112.5% | 115.1% | 83.26% | 73.61% | 82.67% | 76.59% | 79.12% | 72.31% | 76.52% | 102.7% | 103.3% | 96.29% | 102.9%
| 100 | CCV 10 ppb MAJORS |  | 06/08/2023 23:44:46 | 10522.2205 | 10253.3821 | 8.8086 | 9.7325 | 8.9314 | 10993.611 | 9852.9594 | 10755.8497 | 10.6614 | 10.8242 | 10.9445 | 10.8515 | 11.4439 | 10.5267 | 11.0466 | 11.0528 | 10.5473 | 10.6169 | 10.7119 | 10.0794 | 9.3788 | 9.6111 | 10.6378 | 10.2046 | 10.3166 | 10.3514 | 10.6686 | 10.8022 | 10.4152 | 11.2476 | 11.2625 | [7 empty] | 10.2242 | 10.3082 | 9.5753 | 10.3607
|  | RSD |  |  | 3.561% | 2.323% | 0.4597% | 0.5806% | 0.9748% | 0.7811% | 3.265% | 1.227% | 1.906% | 3.073% | 1.074% | 3.368% | 1.353% | 0.9957% | 1.185% | 2.515% | 1.951% | 1.408% | 1.422% | 1.694% | 8.074% | 2.646% | 1.566% | 3.528% | 3.817% | 0.9127% | 3.742% | 3.149% | 2.02% | 1.444% | 1.015% | [7 empty] | 2.369% | 1.789% | 0.8116% | 0.666%
|  | % Recovery |  |  | 105.2% | 102.5% | 88.09% | 97.33% | 89.31% | 109.9% | 98.53% | 107.6% | 106.6% | 108.2% | 109.4% | 108.5% | 114.4% | 105.3% | 110.5% | 110.5% | 105.5% | 106.2% | 107.1% | 100.8% | 93.79% | 96.11% | 106.4% | 102% | 103.2% | 103.5% | 106.7% | 108% | 104.2% | 112.5% | 112.6% | 84.9% | 75.42% | 81.01% | 76.37% | 79.11% | 72.69% | 77.56% | 102.2% | 103.1% | 95.75% | 103.6%
| 101 | CCV 10 ppb MAJORS |  | 06/08/2023 23:51:04 | 10195.537 | 10325.5385 | 8.7585 | 9.3852 | 8.878 | 10646.2834 | 10001.4775 | 10830.1385 | 10.6174 | 10.3862 | 11.0038 | 10.9492 | 11.5407 | 10.3497 | 10.9763 | 10.8296 | 10.279 | 10.5901 | 10.6735 | 10.2518 | 9.8277 | 9.6832 | 10.6885 | 10.3024 | 10.4745 | 10.3285 | 10.5242 | 10.828 | 10.9306 | 11.205 | 11.3154 | [7 empty] | 10.1536 | 10.3769 | 9.6745 | 10.1861
|  | RSD |  |  | 2.099% | 4.97% | 1.829% | 6.916% | 2.236% | 8.17% | 6.59% | 7.024% | 4.519% | 5.354% | 1.677% | 2.201% | 0.9493% | 5.057% | 5.132% | 6.707% | 5.054% | 1.648% | 0.8047% | 1.871% | 7.225% | 0.8138% | 2.029% | 1.196% | 4.677% | 2.953% | 1.326% | 1.363% | 4.715% | 0.7306% | 0.4558% | [7 empty] | 0.88% | 0.6275% | 1.501% | 1.009%
|  | % Recovery |  |  | 102% | 103.3% | 87.58% | 93.85% | 88.78% | 106.5% | 100% | 108.3% | 106.2% | 103.9% | 110% | 109.5% | 115.4% | 103.5% | 109.8% | 108.3% | 102.8% | 105.9% | 106.7% | 102.5% | 98.28% | 96.83% | 106.9% | 103% | 104.7% | 103.3% | 105.2% | 108.3% | 109.3% | 112.1% | 113.2% | 84.93% | 74.76% | 82.61% | 76.88% | 80.24% | 74.11% | 76.62% | 101.5% | 103.8% | 96.75% | 101.9%
| 116 | CCV 10 ppb MAJORS |  | 06/09/2023 01:18:35 | 10506.043 | 10123.676 | 8.4551 | 9.7246 | 8.9667 | 10850.9075 | 10202.4914 | 10915.2735 | 10.9189 | 10.5881 | 10.8122 | 10.7201 | 11.4696 | 10.6245 | 11.0236 | 10.9675 | 10.5965 | 10.5167 | 10.5185 | 10.0309 | 10.0877 | 9.4417 | 10.7086 | 10.1537 | 10.663 | 10.3829 | 10.4867 | 10.428 | 10.6185 | 11.2173 | 11.3052 | [7 empty] | 10.255 | 10.2996 | 9.7402 | 10.5455
|  | RSD |  |  | 0.9866% | 4.566% | 1.7% | 16.49% | 4.33% | 3.829% | 3.383% | 3.527% | 2.789% | 1.129% | 0.3295% | 0.7158% | 1.652% | 1.91% | 1.302% | 1.61% | 1.899% | 1.525% | 1.966% | 1.88% | 3.906% | 1.141% | 0.3744% | 2.105% | 2.899% | 2.222% | 2.557% | 2.685% | 4.619% | 1.626% | 0.8331% | [7 empty] | 2.466% | 2.695% | 0.9918% | 0.4528%
|  | % Recovery |  |  | 105.1% | 101.2% | 84.55% | 97.25% | 89.67% | 108.5% | 102% | 109.2% | 109.2% | 105.9% | 108.1% | 107.2% | 114.7% | 106.2% | 110.2% | 109.7% | 106% | 105.2% | 105.2% | 100.3% | 100.9% | 94.42% | 107.1% | 101.5% | 106.6% | 103.8% | 104.9% | 104.3% | 106.2% | 112.2% | 113.1% | 84.69% | 72.31% | 77.51% | 74.78% | 77.45% | 72.3% | 76.2% | 102.6% | 103% | 97.4% | 105.5%
| 117 | CCV 10 ppb MAJORS |  | 06/09/2023 01:24:53 | 10686.6783 | 10343.3676 | 8.558 | 9.103 | 9.3316 | 11101.8856 | 10156.0171 | 10668.8401 | 10.7468 | 10.7033 | 10.9293 | 11.0091 | 11.5584 | 10.3537 | 11.069 | 11.1126 | 10.583 | 10.6431 | 10.6241 | 10.157 | 9.1867 | 9.286 | 10.754 | 10.1504 | 10.8812 | 10.5997 | 10.8801 | 10.7601 | 10.9323 | 11.2961 | 11.5315 | [7 empty] | 10.3244 | 10.4178 | 9.784 | 10.4368
|  | RSD |  |  | 1.734% | 3.184% | 1.989% | 0.218% | 3.262% | 3.701% | 4.617% | 2.429% | 1.451% | 0.57% | 0.7688% | 1.304% | 0.554% | 0.184% | 0.4318% | 2.314% | 2.807% | 1.181% | 0.31% | 2.289% | 2.125% | 0.889% | 0.3817% | 1.317% | 2.564% | 1.257% | 3.44% | 2.717% | 3.773% | 0.2432% | 0.9816% | [7 empty] | 1.279% | 0.5765% | 0.5871% | 1.856%
|  | % Recovery |  |  | 106.9% | 103.4% | 85.58% | 91.03% | 93.32% | 111% | 101.6% | 106.7% | 107.5% | 107% | 109.3% | 110.1% | 115.6% | 103.5% | 110.7% | 111.1% | 105.8% | 106.4% | 106.2% | 101.6% | 91.87% | 92.86% | 107.5% | 101.5% | 108.8% | 106% | 108.8% | 107.6% | 109.3% | 113% | 115.3% | 84.3% | 71.72% | 78.96% | 75.01% | 77.43% | 70.24% | 77.18% | 103.2% | 104.2% | 97.84% | 104.4%
| 131 | CCV 10 ppb MAJORS |  | 06/09/2023 02:46:38 | 10356.5737 | 9983.1131 | 8.0889 | 9.4782 | 8.916 | 10990.5441 | 9940.5629 | 10699.272 | 10.8473 | 10.5968 | 10.6712 | 10.5622 | 11.5428 | 10.2954 | 10.9612 | 11.0111 | 10.5235 | 10.343 | 10.5023 | 9.9541 | 9.7851 | 9.3474 | 10.796 | 10.2314 | 10.7548 | 10.4601 | 10.8769 | 10.3907 | 10.8284 | 11.1769 | 11.5011 | [7 empty] | 10.1741 | 10.3961 | 9.7876 | 10.6141
|  | RSD |  |  | 2.24% | 3.91% | 0.4469% | 9.831% | 1.266% | 1.989% | 4.328% | 9.195% | 3.94% | 2.605% | 2.505% | 1.788% | 0.6731% | 1.792% | 4.482% | 3.456% | 2.423% | 1.088% | 1.004% | 1.593% | 7.148% | 3.508% | 0.7177% | 0.3774% | 1.179% | 1.739% | 1.423% | 3.432% | 2.419% | 2.421% | 0.6006% | [7 empty] | 1.27% | 2.01% | 1.785% | 1.372%
|  | % Recovery |  |  | 103.6% | 99.83% | 80.89% | 94.78% | 89.16% | 109.9% | 99.41% | 107% | 108.5% | 106% | 106.7% | 105.6% | 115.4% | 103% | 109.6% | 110.1% | 105.2% | 103.4% | 105% | 99.54% | 97.85% | 93.47% | 108% | 102.3% | 107.5% | 104.6% | 108.8% | 103.9% | 108.3% | 111.8% | 115% | 84.44% | 69.31% | 77.46% | 73.94% | 75.36% | 69.45% | 74.44% | 101.7% | 104% | 97.88% | 106.1%
| 132 | CCV 10 ppb MAJORS |  | 06/09/2023 02:52:55 | 10271.5016 | 9936.3669 | 8.4039 | 9.3585 | 9.4053 | 10912.3798 | 9802.8815 | 11043.2201 | 10.6373 | 10.476 | 10.8694 | 10.6794 | 11.2887 | 10.4824 | 10.9061 | 10.7761 | 10.5068 | 10.5846 | 10.5224 | 10.1115 | 10.0138 | 9.4949 | 10.6146 | 10.4333 | 10.5058 | 10.1291 | 10.3115 | 10.9188 | 10.3008 | 11.1412 | 11.3613 | [7 empty] | 10.2073 | 10.2703 | 9.8385 | 10.4226
|  | RSD |  |  | 0.792% | 3.826% | 8.315% | 3.488% | 2.104% | 4.433% | 1.268% | 3.389% | 1.482% | 1.028% | 1.719% | 0.2639% | 1.202% | 0.2959% | 0.985% | 1.616% | 1.573% | 0.2985% | 0.9905% | 8.438% | 3.851% | 1.265% | 0.5623% | 9.027% | 3.159% | 3.873% | 0.7131% | 8.988% | 4.598% | 1.193% | 0.7664% | [7 empty] | 0.8013% | 0.7989% | 8.418% | 1.884%
|  | % Recovery |  |  | 102.7% | 99.36% | 84.04% | 93.59% | 94.05% | 109.1% | 98.03% | 110.4% | 106.4% | 104.8% | 108.7% | 106.8% | 112.9% | 104.8% | 109.1% | 107.8% | 105.1% | 105.8% | 105.2% | 101.1% | 100.1% | 94.95% | 106.1% | 104.3% | 105.1% | 101.3% | 103.1% | 109.2% | 103% | 111.4% | 113.6% | 83.93% | 71.58% | 77.4% | 74.06% | 76.78% | 66.89% | 74.34% | 102.1% | 102.7% | 98.39% | 104.2%
| 148 | CCV 10 ppb MAJORS |  | 06/09/2023 04:27:10 | 9994.5089 | 9616.8679 | 7.8722 | 9.2965 | 8.7872 | 10492.0982 | 9483.7382 | 10467.4739 | 10.5612 | 10.6635 | 11.1277 | 11.1114 | 11.464 | 10.1943 | 10.712 | 10.8583 | 10.2157 | 10.7439 | 10.69 | 9.8211 | 9.5118 | 9.2813 | 10.6774 | 10.0933 | 10.4667 | 10.6094 | 10.6429 | 10.0384 | 10.721 | 11.5107 | 11.4835 | [7 empty] | 10.3355 | 10.4697 | 9.6241 | 10.5172
|  | RSD |  |  | 1.111% | 3.886% | 1.038% | 9.674% | 1.587% | 2.797% | 0.709% | 1.281% | 0.4145% | 1.635% | 3.369% | 4.628% | 1.641% | 2.885% | 0.7588% | 1.798% | 0.7796% | 4.131% | 3.053% | 2.749% | 2.633% | 2.461% | 2.171% | 4.001% | 2.794% | 1.634% | 2.507% | 5.808% | 3.108% | 3.616% | 1.239% | [7 empty] | 3.368% | 3.529% | 2.303% | 2.04%
|  | % Recovery |  |  | 99.95% | 96.17% | 78.72% | 92.96% | 87.87% | 104.9% | 94.84% | 104.7% | 105.6% | 106.6% | 111.3% | 111.1% | 114.6% | 101.9% | 107.1% | 108.6% | 102.2% | 107.4% | 106.9% | 98.21% | 95.12% | 92.81% | 106.8% | 100.9% | 104.7% | 106.1% | 106.4% | 100.4% | 107.2% | 115.1% | 114.8% | 85.76% | 69.23% | 77.78% | 72.55% | 74.68% | 68.82% | 74.03% | 103.4% | 104.7% | 96.24% | 105.2%
| 149 | CCV 10 ppb MAJORS |  | 06/09/2023 04:33:27 | 10179.8295 | 9996.6224 | 8.1827 | 9.6363 | 8.8462 | 10714.2111 | 9654.1533 | 10923.3615 | 10.9967 | 10.6108 | 10.1528 | 9.9857 | 11.4208 | 10.6172 | 10.9366 | 10.8637 | 10.4136 | 9.9154 | 9.8559 | 10.0347 | 9.398 | 9.3666 | 10.6098 | 10.4625 | 10.2342 | 10.3869 | 10.5024 | 10.3776 | 10.6661 | 10.6211 | 11.3745 | [7 empty] | 9.5247 | 9.7712 | 9.7723 | 10.3593
|  | RSD |  |  | 2.395% | 3.242% | 0.7682% | 10.73% | 2.817% | 6.89% | 2.875% | 1.185% | 2.726% | 2.176% | 11.78% | 11.07% | 0.9599% | 3.023% | 3.331% | 3.46% | 1.378% | 11.86% | 11.88% | 0.7606% | 9.008% | 2.177% | 1.059% | 1.853% | 2.614% | 0.6121% | 1.758% | 1.214% | 3.392% | 9.996% | 2.512% | [7 empty] | 11.17% | 11.59% | 1.179% | 2.21%
|  | % Recovery |  |  | 101.8% | 99.97% | 81.83% | 96.36% | 88.46% | 107.1% | 96.54% | 109.2% | 110% | 106.1% | 101.5% | 99.86% | 114.2% | 106.2% | 109.4% | 108.6% | 104.1% | 99.15% | 98.56% | 100.3% | 93.98% | 93.67% | 106.1% | 104.6% | 102.3% | 103.9% | 105% | 103.8% | 106.7% | 106.2% | 113.7% | 84.56% | 70.17% | 77.06% | 78.69% | 77.03% | 68.88% | 73.87% | 95.25% | 97.71% | 97.72% | 103.6%
| 163 | CCV 10 ppb MAJORS |  | 06/09/2023 05:55:13 | 10077.9948 | 9722.7396 | 8.0811 | 9.0427 | 8.8603 | 10644.8836 | 9633.7547 | 10401.7439 | 10.7421 | 10.5474 | 10.7787 | 10.5641 | 11.435 | 10.3795 | 10.753 | 10.4794 | 10.2912 | 10.486 | 10.3246 | 10.0781 | 9.8555 | 9.3768 | 10.6654 | 10.4149 | 9.7581 | 10.3503 | 10.4824 | 10.4773 | 10.3623 | 11.0465 | 11.3729 | [7 empty] | 10.0787 | 10.2221 | 9.812 | 13.6664
|  | RSD |  |  | 1.205% | 3.032% | 1.471% | 14.68% | 0.9839% | 1.378% | 2.402% | 1.374% | 1.625% | 0.8472% | 1.039% | 1.458% | 0.7819% | 2.638% | 1.411% | 1.708% | 1.422% | 0.5708% | 0.5821% | 1.435% | 1.329% | 3.961% | 0.4614% | 1.963% | 3.18% | 1.652% | 0.776% | 2.065% | 6.388% | 0.9623% | 0.5113% | [7 empty] | 1.654% | 1.263% | 1.312% | 23.36%
|  | % Recovery |  |  | 100.8% | 97.23% | 80.81% | 90.43% | 88.6% | 106.4% | 96.34% | 104% | 107.4% | 105.5% | 107.8% | 105.6% | 114.3% | 103.8% | 107.5% | 104.8% | 102.9% | 104.9% | 103.2% | 100.8% | 98.55% | 93.77% | 106.7% | 104.1% | 97.58% | 103.5% | 104.8% | 104.8% | 103.6% | 110.5% | 113.7% | 81.81% | 68.72% | 76.71% | 72.71% | 60.91% | 68.12% | 72.7% | 100.8% | 102.2% | 98.12% | 136.7%
| 164 | CCV 10 ppb MAJORS |  | 06/09/2023 06:01:30 | 10127.2038 | 9660.7643 | 8.3293 | 9.0625 | 8.9396 | 10355.8633 | 9514.7639 | 10478.8626 | 10.4444 | 10.3642 | 10.8827 | 10.6084 | 11.6028 | 10.0025 | 10.6399 | 10.6506 | 10.14 | 10.5524 | 10.5135 | 10.1763 | 9.8881 | 9.3709 | 10.7637 | 10.3398 | 10.5144 | 10.6341 | 10.6843 | 10.4701 | 10.8723 | 11.1997 | 11.5358 | [7 empty] | 10.1199 | 10.2084 | 9.744 | 10.3248
|  | RSD |  |  | 4.962% | 6.161% | 0.569% | 14.89% | 1.428% | 7.478% | 6.147% | 4.902% | 6.247% | 4.351% | 0.1003% | 0.7372% | 0.7321% | 5.53% | 4.061% | 6.237% | 5.87% | 1.386% | 1.012% | 1.673% | 2.445% | 2.186% | 0.4337% | 3.275% | 2.525% | 2.61% | 2.19% | 2.928% | 0.8907% | 0.5577% | 1.048% | [7 empty] | 0.5124% | 0.1178% | 1.098% | 1.932%
|  | % Recovery |  |  | 101.3% | 96.61% | 83.29% | 90.63% | 89.4% | 103.6% | 95.15% | 104.8% | 104.4% | 103.6% | 108.8% | 106.1% | 116% | 100% | 106.4% | 106.5% | 101.4% | 105.5% | 105.1% | 101.8% | 98.88% | 93.71% | 107.6% | 103.4% | 105.1% | 106.3% | 106.8% | 104.7% | 108.7% | 112% | 115.4% | 81.89% | 68.26% | 76.6% | 72.22% | 74.51% | 67.44% | 72.86% | 101.2% | 102.1% | 97.44% | 103.2%
| 180 | CCV 10 ppb MAJORS |  | 06/09/2023 07:36:12 | 10782.0523 | 10337.5341 | 7.9763 | 10.0551 | 9.3865 | 11819.8229 | 10201.5805 | 11595.8589 | 11.5809 | 11.3385 | 11.0865 | 10.9656 | 11.8503 | 11.1059 | 11.603 | 11.524 | 11.058 | 10.6293 | 10.6038 | 9.8505 | 10.1951 | 9.3939 | 10.8976 | 10.0099 | 10.7302 | 10.6745 | 10.6631 | 10.3899 | 11.0971 | 11.2673 | 11.6711 | [7 empty] | 10.2221 | 10.3652 | 9.41 | 10.2653
|  | RSD |  |  | 4.478% | 4.13% | 0.8609% | 5.552% | 0.2169% | 6.816% | 4.667% | 6.945% | 6.073% | 3.249% | 1.391% | 2.069% | 1.191% | 5.693% | 4.701% | 5.062% | 4.557% | 1.038% | 1.771% | 2.275% | 8.046% | 2.05% | 0.4987% | 0.9237% | 4.923% | 1.252% | 1.099% | 4.086% | 6.472% | 1.093% | 0.2538% | [7 empty] | 0.9885% | 1.076% | 1.259% | 1.852%
|  | % Recovery |  |  | 107.8% | 103.4% | 79.76% | 100.6% | 93.86% | 118.2% | 102% | 116% | 115.8% | 113.4% | 110.9% | 109.7% | 118.5% | 111.1% | 116% | 115.2% | 110.6% | 106.3% | 106% | 98.5% | 102% | 93.94% | 109% | 100.1% | 107.3% | 106.7% | 106.6% | 103.9% | 111% | 112.7% | 116.7% | 82.83% | 66.02% | 71.09% | 70.1% | 71.7% | 64.48% | 71.18% | 102.2% | 103.7% | 94.1% | 102.7%
| 181 | CCV 10 ppb MAJORS |  | 06/09/2023 07:42:30 | 10829.0372 | 10386.1016 | 8.1288 | 10.1892 | 9.1513 | 11204.1152 | 10091.0181 | 10646.2779 | 11.0517 | 10.9682 | 11.1525 | 11.1076 | 11.7863 | 10.9493 | 11.3019 | 11.151 | 10.829 | 10.7566 | 10.6965 | 10.0111 | 9.5612 | 9.4705 | 10.8046 | 10.4323 | 10.2157 | 10.3569 | 10.676 | 10.3559 | 10.4126 | 11.4757 | 11.5754 | [7 empty] | 10.309 | 10.3199 | 9.5899 | 11.0177
|  | RSD |  |  | 5.292% | 4.558% | 0.3193% | 3.284% | 1.502% | 5.498% | 6.339% | 4.168% | 5.459% | 5.05% | 0.909% | 0.7093% | 1.981% | 7.816% | 8.37% | 3.609% | 6.844% | 1.137% | 0.9539% | 1.504% | 9.195% | 1.206% | 0.9484% | 2.495% | 5.745% | 1.847% | 1.837% | 3.558% | 2.106% | 0.9962% | 2.408% | [7 empty] | 1.246% | 1.013% | 0.6264% | 4.33%
|  | % Recovery |  |  | 108.3% | 103.9% | 81.29% | 101.9% | 91.51% | 112% | 100.9% | 106.5% | 110.5% | 109.7% | 111.5% | 111.1% | 117.9% | 109.5% | 113% | 111.5% | 108.3% | 107.6% | 107% | 100.1% | 95.61% | 94.7% | 108% | 104.3% | 102.2% | 103.6% | 106.8% | 103.6% | 104.1% | 114.8% | 115.8% | 81.6% | 66.23% | 72.4% | 69.62% | 67.66% | 65.16% | 70.19% | 103.1% | 103.2% | 95.9% | 110.2%
| 195 | CCV 10 ppb MAJORS |  | 06/09/2023 09:05:39 | 10992.7877 | 10545.2027 | 7.9606 | 11.007 | 9.3178 | 11988.486 | 10257.7245 | 11514.3664 | 11.4202 | 11.2655 | 11.0759 | 11.0225 | 11.7624 | 11.1772 | 11.4556 | 11.429 | 11.0224 | 10.6617 | 10.6764 | 10.1097 | 10.0917 | 9.1025 | 10.8117 | 10.5986 | 11.059 | 10.5726 | 10.6964 | 10.5923 | 10.839 | 11.4291 | 11.6291 | [7 empty] | 10.2806 | 10.4932 | 9.6686 | 10.4674
|  | RSD |  |  | 7.743% | 7.759% | 1.75% | 6.668% | 1.424% | 3.139% | 5.494% | 7.472% | 6.522% | 4.361% | 0.3646% | 1.787% | 2.535% | 7.456% | 5.523% | 9.39% | 8.203% | 0.835% | 1.026% | 1.952% | 7.912% | 2.52% | 2.846% | 0.8718% | 2.828% | 4.348% | 3.244% | 3.961% | 2.009% | 0.318% | 2.221% | [7 empty] | 1.541% | 0.9358% | 2.823% | 1.808%
|  | % Recovery |  |  | 109.9% | 105.5% | 79.61% | 110.1% | 93.18% | 119.9% | 102.6% | 115.1% | 114.2% | 112.7% | 110.8% | 110.2% | 117.6% | 111.8% | 114.6% | 114.3% | 110.2% | 106.6% | 106.8% | 101.1% | 100.9% | 91.02% | 108.1% | 106% | 110.6% | 105.7% | 107% | 105.9% | 108.4% | 114.3% | 116.3% | 76.5% | 60.91% | 66.04% | 65.35% | 66.32% | 59.03% | 67.24% | 102.8% | 104.9% | 96.69% | 104.7%
| 196 | CCV 10 ppb MAJORS |  | 06/09/2023 09:11:56 | 10389.8526 | 9967.1841 | 7.8848 | 9.9334 | 9.1764 | 11508.9672 | 9789.0791 | 10913.4873 | 10.8624 | 10.6877 | 11.1521 | 11.1438 | 11.8195 | 10.6551 | 10.862 | 10.8946 | 10.5392 | 10.7941 | 10.7343 | 10.1943 | 9.6135 | 9.5865 | 10.9978 | 10.3312 | 10.5074 | 10.5744 | 10.8123 | 10.8018 | 10.7438 | 11.5463 | 11.7391 | [7 empty] | 10.354 | 10.5002 | 9.5143 | 10.3809
|  | RSD |  |  | 2.254% | 3.334% | 1.767% | 7.041% | 1.681% | 2.113% | 3.32% | 3.161% | 0.2662% | 0.5944% | 1.361% | 0.8754% | 1.332% | 0.7173% | 0.4369% | 1.272% | 0.4594% | 1.424% | 0.8357% | 1.029% | 6.151% | 5.06% | 0.5289% | 1.175% | 5.416% | 3.064% | 2.069% | 4.155% | 8.438% | 0.5508% | 0.3467% | [7 empty] | 0.7146% | 1.249% | 1.135% | 0.4573%
|  | % Recovery |  |  | 103.9% | 99.67% | 78.85% | 99.33% | 91.76% | 115.1% | 97.89% | 109.1% | 108.6% | 106.9% | 111.5% | 111.4% | 118.2% | 106.6% | 108.6% | 108.9% | 105.4% | 107.9% | 107.3% | 101.9% | 96.13% | 95.87% | 110% | 103.3% | 105.1% | 105.7% | 108.1% | 108% | 107.4% | 115.5% | 117.4% | 77.86% | 61.65% | 70.2% | 65.34% | 67.53% | 60.04% | 65.78% | 103.5% | 105% | 95.14% | 103.8%
| 213 | CCV 10 ppb MAJORS |  | 06/09/2023 10:53:35 | 10573.7087 | 10097.63 | 7.9675 | 9.7495 | 9.4746 | 11710.2334 | 9917.0711 | 11235.9814 | 11.097 | 10.8697 | 11.2383 | 11.0506 | 12.0634 | 10.9045 | 11.0704 | 10.927 | 10.5331 | 10.6238 | 10.6343 | 10.1742 | 9.6743 | 9.2865 | 11.1364 | 10.4692 | 11.1464 | 10.938 | 11.154 | 10.8519 | 10.7872 | 11.5057 | 11.9297 | [7 empty] | 10.2002 | 10.353 | 9.3466 | 10.2881
|  | RSD |  |  | 3.116% | 1.419% | 1.892% | 11.33% | 3.979% | 1.882% | 2.378% | 2.882% | 3.499% | 1.75% | 0.9934% | 1.641% | 2.841% | 5.233% | 2.677% | 6.456% | 1.979% | 1.541% | 0.972% | 1.991% | 3.832% | 1.762% | 2.168% | 3.43% | 1.663% | 2.875% | 1.486% | 6.646% | 3.833% | 0.6272% | 1.671% | [7 empty] | 1.844% | 1.14% | 0.7541% | 1.824%
|  | % Recovery |  |  | 105.7% | 101% | 79.67% | 97.5% | 94.75% | 117.1% | 99.17% | 112.4% | 111% | 108.7% | 112.4% | 110.5% | 120.6% | 109% | 110.7% | 109.3% | 105.3% | 106.2% | 106.3% | 101.7% | 96.74% | 92.86% | 111.4% | 104.7% | 111.5% | 109.4% | 111.5% | 108.5% | 107.9% | 115.1% | 119.3% | 78.2% | 59.96% | 69.85% | 65.14% | 66.94% | 59.47% | 66.8% | 102% | 103.5% | 93.47% | 102.9%
| 237 | CCV 10 ppb MAJORS |  | 06/09/2023 14:00:40 | 10350.2821 | 10395.9406 | 8.2917 | 8.2791 | 8.8623 | 10539.3843 | 10188.7254 | 9940.4101 | 10.6015 | 10.5051 | 9.6156 | 9.4123 | 10.2109 | 10.2352 | 10.4859 | 10.5562 | 10.4632 | 9.3715 | 9.2352 | 10.1257 | 11.4212 | 9.7152 | 10.2558 | 9.7888 | 10.1078 | 9.6431 | 9.8179 | 9.8217 | 9.8025 | 9.5089 | 10.0476 | [7 empty] | 9.3278 | 9.4768 | 9.4024 | 9.5241
|  | RSD |  |  | 2.001% | 5.187% | 1.483% | 11.52% | 1.486% | 5.221% | 1.009% | 5.159% | 1.938% | 3.537% | 4.363% | 5.38% | 1.713% | 3.904% | 2.35% | 2.645% | 2.893% | 4.032% | 4.528% | 0.8566% | 5.814% | 1.457% | 1.392% | 2.487% | 8.134% | 3.491% | 1.345% | 3.213% | 6.148% | 3.165% | 0.9913% | [7 empty] | 5.345% | 4.442% | 2.354% | 1.365%
|  | % Recovery |  |  | 103.5% | 104% | 82.92% | 82.79% | 88.62% | 105.4% | 101.9% | 99.4% | 106% | 105.1% | 96.16% | 94.12% | 102.1% | 102.4% | 104.9% | 105.6% | 104.6% | 93.71% | 92.35% | 101.3% | 114.2% | 97.15% | 102.6% | 97.89% | 101.1% | 96.43% | 98.18% | 98.22% | 98.03% | 95.09% | 100.5% | 93.59% | 88.01% | 88.92% | 93.93% | 88.94% | 90.2% | 87.55% | 93.28% | 94.77% | 94.02% | 95.24%
| 238 | CCV 10 ppb MAJORS |  | 06/09/2023 14:06:57 | 9896.7305 | 9948.2934 | 8.4312 | 8.5052 | 8.6271 | 9698.5241 | 9916.7378 | 9574.7797 | 10.2203 | 10.0061 | 10.3564 | 10.2811 | 10.3055 | 9.8619 | 10.1985 | 10.1822 | 10.0987 | 10.0408 | 9.7984 | 10.1363 | 10.6039 | 9.5365 | 10.2556 | 10.169 | 10.0833 | 10.0715 | 9.7977 | 9.7178 | 10.01 | 10.2348 | 10.1442 | [7 empty] | 9.9873 | 10.0255 | 9.4366 | 9.4457
|  | RSD |  |  | 1.903% | 4.626% | 0.4587% | 9.105% | 2.421% | 3.949% | 3.686% | 1.546% | 1.112% | 0.4334% | 0.3089% | 0.6088% | 0.7884% | 3.248% | 2.253% | 0.4193% | 1.106% | 0.2633% | 1.198% | 0.3964% | 4.138% | 1.663% | 0.941% | 1.412% | 7.296% | 0.3244% | 1.709% | 1.227% | 2.412% | 0.3062% | 0.9473% | [7 empty] | 1.181% | 0.304% | 0.8786% | 0.793%
|  | % Recovery |  |  | 98.97% | 99.48% | 84.31% | 85.05% | 86.27% | 96.99% | 99.17% | 95.75% | 102.2% | 100.1% | 103.6% | 102.8% | 103.1% | 98.62% | 102% | 101.8% | 101% | 100.4% | 97.98% | 101.4% | 106% | 95.36% | 102.6% | 101.7% | 100.8% | 100.7% | 97.98% | 97.18% | 100.1% | 102.3% | 101.4% | 93.76% | 88.24% | 92.31% | 89.01% | 90.21% | 90.74% | 89.46% | 99.87% | 100.3% | 94.37% | 94.46%
| 253 | CCV 10 ppb MAJORS |  | 06/09/2023 15:35:51 | 10079.2403 | 10296.5081 | 8.5051 | 9.6642 | 8.4516 | 10013.7953 | 9964.8395 | 9910.8312 | 10.4844 | 10.0899 | 10.2684 | 10.1233 | 11.6735 | 9.6451 | 9.9282 | 10.0368 | 10.0718 | 9.9945 | 9.8268 | 10.1057 | 10.6171 | 10.7958 | 11.7603 | 9.9018 | 11.5121 | 11.2319 | 11.3624 | 9.8211 | 11.813 | 10.2358 | 11.5954 | [7 empty] | 10.1339 | 10.274 | 9.7834 | 9.9058
|  | RSD |  |  | 1.135% | 0.9669% | 1.743% | 9.969% | 0.671% | 4.626% | 3.642% | 3.541% | 1.743% | 1.433% | 0.7051% | 4.496% | 11.34% | 0.5592% | 1.372% | 0.6122% | 1.445% | 2.079% | 0.7372% | 0.9167% | 5.1% | 11.44% | 10.66% | 5.454% | 7.3% | 10.92% | 8.991% | 2.097% | 8.8% | 1.454% | 10.58% | [7 empty] | 0.8174% | 1.27% | 0.7153% | 0.748%
|  | % Recovery |  |  | 100.8% | 103% | 85.05% | 96.64% | 84.52% | 100.1% | 99.65% | 99.11% | 104.8% | 100.9% | 102.7% | 101.2% | 116.7% | 96.45% | 99.28% | 100.4% | 100.7% | 99.94% | 98.27% | 101.1% | 106.2% | 108% | 117.6% | 99.02% | 115.1% | 112.3% | 113.6% | 98.21% | 118.1% | 102.4% | 116% | 91.38% | 76.63% | 87.42% | 84.92% | 86.81% | 86.51% | 78.54% | 101.3% | 102.7% | 97.83% | 99.06%
| 254 | CCV 10 ppb MAJORS |  | 06/09/2023 15:42:08 | 10060.8796 | 10229.1112 | 8.2553 | 9.5989 | 5.9915 | 10299.688 | 9854.0791 | 10178.4009 | 10.3675 | 10.232 | 10.1164 | 10.1097 | 10.3263 | 9.9755 | 10.0747 | 10.3532 | 10.1758 | 9.8394 | 9.6371 | 9.7508 | 10.6018 | 9.6584 | 10.23 | 9.8167 | 10.3005 | 9.8925 | 9.9642 | 9.3935 | 9.948 | 10.0776 | 10.1505 | [7 empty] | 9.9543 | 10.0537 | 9.5638 | 9.8555
|  | RSD |  |  | 1.376% | 0.9306% | 14.27% | 8.328% | 24.56% | 3.595% | 2.187% | 3.918% | 2.521% | 0.413% | 0.4046% | 0.4177% | 0.4863% | 2.569% | 0.5052% | 0.6538% | 0.7631% | 1.024% | 0.4255% | 10.65% | 3.475% | 1.348% | 0.4422% | 16.48% | 3.246% | 2.911% | 1.412% | 11.8% | 5.809% | 0.9523% | 0.4916% | [7 empty] | 1.853% | 0.256% | 12.55% | 2.723%
|  | % Recovery |  |  | 100.6% | 102.3% | 82.55% | 95.99% | 59.92% | 103% | 98.54% | 101.8% | 103.7% | 102.3% | 101.2% | 101.1% | 103.3% | 99.76% | 100.7% | 103.5% | 101.8% | 98.39% | 96.37% | 97.51% | 106% | 96.58% | 102.3% | 98.17% | 103% | 98.93% | 99.64% | 93.94% | 99.48% | 100.8% | 101.5% | 94.68% | 86.93% | 88.77% | 88.2% | 87.97% | 113.8% | 86.77% | 99.54% | 100.5% | 95.64% | 98.56%
| 269 | CCV 10 ppb MAJORS |  | 06/09/2023 17:09:38 | 9677.6664 | 9970.5728 | 8.16 | 8.6489 | 8.2377 | 10400.4483 | 9814.0069 | 9921.448 | 10.2958 | 9.8427 | 10.7959 | 10.7359 | 9.9488 | 9.6644 | 9.8859 | 10.0682 | 9.9351 | 10.4581 | 10.7515 | 9.9187 | 10.7892 | 9.1103 | 9.7958 | 10.2309 | 10.1506 | 9.8266 | 9.5771 | 9.5753 | 9.9265 | 10.6589 | 9.8607 | [7 empty] | 10.6465 | 10.7446 | 9.7305 | 9.8364
|  | RSD |  |  | 4.702% | 5.555% | 3.816% | 10.86% | 1.49% | 4.587% | 4.219% | 2.496% | 3.35% | 3.834% | 10.19% | 10.87% | 5.882% | 4.07% | 5.515% | 4.889% | 4.176% | 9.835% | 9.884% | 2.896% | 3.332% | 3.435% | 5.355% | 7.599% | 4.198% | 7.167% | 5.377% | 3.631% | 11.7% | 10.24% | 4.988% | [7 empty] | 8.805% | 9.728% | 4.397% | 4.964%
|  | % Recovery |  |  | 96.78% | 99.71% | 81.6% | 86.49% | 82.38% | 104% | 98.14% | 99.21% | 103% | 98.43% | 108% | 107.4% | 99.49% | 96.64% | 98.86% | 100.7% | 99.35% | 104.6% | 107.5% | 99.19% | 107.9% | 91.1% | 97.96% | 102.3% | 101.5% | 98.27% | 95.77% | 95.75% | 99.26% | 106.6% | 98.61% | 91.13% | 84.78% | 87.87% | 80.8% | 85.92% | 84.74% | 85.97% | 106.5% | 107.4% | 97.3% | 98.36%
| 270 | CCV 10 ppb MAJORS |  | 06/09/2023 17:15:55 | 9642.6703 | 9692.0215 | 7.9438 | 8.8318 | 8.3094 | 9318.3186 | 9668.4219 | 10123.8807 | 9.798 | 9.5417 | 10.0094 | 9.9519 | 10.1888 | 9.5323 | 9.6938 | 9.8957 | 9.7993 | 9.7489 | 10.0579 | 9.7141 | 10.4829 | 9.2205 | 10.0749 | 9.5894 | 10.3581 | 9.6943 | 9.7155 | 9.4027 | 10.0792 | 9.9916 | 10.0249 | [7 empty] | 9.8276 | 9.9611 | 9.4716 | 9.6015
|  | RSD |  |  | 1.062% | 5.24% | 0.8341% | 3.531% | 1.401% | 3.368% | 3.37% | 2.221% | 1.938% | 0.8451% | 0.5243% | 0.7871% | 1.079% | 2.003% | 0.3734% | 0.8577% | 1.03% | 1.069% | 0.6643% | 1.334% | 1.874% | 1.058% | 1.303% | 1.51% | 3.037% | 2.027% | 2.174% | 3.022% | 1.007% | 0.5006% | 0.45% | [7 empty] | 0.965% | 1.004% | 1.073% | 1.296%
|  | % Recovery |  |  | 96.43% | 96.92% | 79.44% | 88.32% | 83.09% | 93.18% | 96.68% | 101.2% | 97.98% | 95.42% | 100.1% | 99.52% | 101.9% | 95.32% | 96.94% | 98.96% | 97.99% | 97.49% | 100.6% | 97.14% | 104.8% | 92.21% | 100.7% | 95.89% | 103.6% | 96.94% | 97.16% | 94.03% | 100.8% | 99.92% | 100.2% | 91.84% | 84.24% | 90.44% | 85.32% | 86.31% | 86.15% | 86.85% | 98.28% | 99.61% | 94.72% | 96.02%
| 286 | CCV 10 ppb MAJORS |  | 06/09/2023 18:49:43 | 9357.2943 | 9934.5988 | 8.0297 | 7.555 | 8.1385 | 9656.7771 | 9593.9796 | 9754.7959 | 9.9263 | 9.6521 | 9.8255 | 9.8189 | 10.8173 | 9.2958 | 9.6017 | 9.4213 | 9.6505 | 9.5815 | 9.7832 | 9.743 | 10.1818 | 10.2667 | 10.9586 | 9.6974 | 10.5995 | 10.6169 | 10.7044 | 9.5458 | 10.7835 | 9.7857 | 10.8307 | [7 empty] | 9.7434 | 9.9244 | 9.7802 | 9.9968
|  | RSD |  |  | 2.741% | 5.015% | 14.26% | 5.126% | 1.099% | 1.669% | 1.255% | 1.909% | 0.9032% | 1.416% | 1.565% | 1.035% | 7.055% | 2.228% | 2.345% | 1.238% | 1.062% | 0.8639% | 1.541% | 11.02% | 4.703% | 8.674% | 7.185% | 11.69% | 5.974% | 9.169% | 4.742% | 10.36% | 7.422% | 0.9807% | 6.831% | [7 empty] | 1.072% | 0.6601% | 12.15% | 1.204%
|  | % Recovery |  |  | 93.57% | 99.35% | 80.3% | 75.55% | 81.39% | 96.57% | 95.94% | 97.55% | 99.26% | 96.52% | 98.26% | 98.19% | 108.2% | 92.96% | 96.02% | 94.21% | 96.5% | 95.81% | 97.83% | 97.43% | 101.8% | 102.7% | 109.6% | 96.97% | 106% | 106.2% | 107% | 95.46% | 107.8% | 97.86% | 108.3% | 88.8% | 74.18% | 84.58% | 81.52% | 82.34% | 83.03% | 74.64% | 97.43% | 99.24% | 97.8% | 99.97%
| 287 | CCV 10 ppb MAJORS |  | 06/09/2023 18:56:01 | 9359.2825 | 9758.1549 | 8.2043 | 8.6707 | 8.0438 | 9660.2451 | 9328.1657 | 9747.3858 | 9.7615 | 9.5959 | 10.1773 | 10.1279 | 10.1694 | 9.3479 | 9.6939 | 9.4181 | 9.5992 | 9.934 | 10.1681 | 10.0508 | 9.7274 | 9.4649 | 10.1579 | 9.8794 | 9.8433 | 9.6168 | 9.9831 | 9.6275 | 9.9185 | 10.1765 | 10.0221 | [7 empty] | 9.986 | 10.1948 | 9.7992 | 9.0283
|  | RSD |  |  | 2.859% | 2.127% | 1.786% | 12.44% | 1.209% | 2.23% | 2.564% | 1.483% | 0.1942% | 0.5899% | 4.962% | 3.548% | 2.187% | 1.563% | 1.088% | 1.485% | 1.025% | 3.64% | 3.864% | 1.278% | 3.4% | 0.3443% | 1.051% | 0.5177% | 3.997% | 0.6025% | 2.794% | 4.309% | 1.816% | 3.564% | 0.8846% | [7 empty] | 4.109% | 3.562% | 0.4475% | 11.43%
|  | % Recovery |  |  | 93.59% | 97.58% | 82.04% | 86.71% | 80.44% | 96.6% | 93.28% | 97.47% | 97.61% | 95.96% | 101.8% | 101.3% | 101.7% | 93.48% | 96.94% | 94.18% | 95.99% | 99.34% | 101.7% | 100.5% | 97.27% | 94.65% | 101.6% | 98.79% | 98.43% | 96.17% | 99.83% | 96.27% | 99.19% | 101.8% | 100.2% | 88.17% | 80.45% | 87.35% | 80.9% | 89.66% | 85.46% | 81.4% | 99.86% | 101.9% | 97.99% | 90.28%
QC STD 7
|  | Sample Id | R | Acquisition Time | Na KED 23/23 | Mg KED 24/24 | Al 27/27 | Al KED 27/27 | P Oshift 31/47 | S (34) -1 34/34 | K KED 39/39 | Ca KED 43/43 | V KED 51/51 | Cr (52) 52/52 | Cr (53) NH3 53/53 | Fe (54) NH3 54/54 | Mn Oshift 55/71 | Fe (56) KED 56/56 | Co KED 59/59 | Ni KED 60/60 | Cu (63) KED 63/63 | Cu (65) NH3 65/65 | Zn NH3 66/66 | As 75/75 | As KED 75/75 | As RH Oshift 75/91 | As Oshift 75/91 | Se (77) 77/77 | Se (77) Oshift 77/93 | Se (78) Oshift 78/94 | Se (80) Oshift 80/96 | Se (82) 82/82 | Se (82) Oshift 82/98 | Mo (95) NH3 95/95 | Mo (95) OShift 95/127 | Rh (IS1) 103/103 (IS) | Rh (IS2) 103/103 (IS) | Rh (IS3) 103/103 (IS) | Rh (IS4) 103/103 (IS) | Rh (IS5) 103/103 (IS) | Rh (IS6) 103/103 (IS) | Rh OShift (IS1) 103/119 (IS) | Cd (111) NH3 111/111 | Cd (114) NH3 114/114 | Pb 208/208 | Pb NH3 208/208
| 25 | LLOQ 0.1 |  | 06/08/2023 15:51:31 | 99.835 | 75.8092 | 0.065718 | -0.051843 | -0.56644 | 279.1031 | 103.7612 | 103.8505 | 0.096542 | 0.11222 | 0.10337 | -0.098896 | 0.089628 | -0.1612 | 0.097228 | 0.11966 | 0.11388 | 0.10932 | 0.10192 | 0.15234 | 0.05636 | 0.24575 | 0.095374 | 0.14192 | 0.036329 | 0.11013 | 0.10514 | 0.1039 | 0.12918 | 0.09714 | 0.096442 | [7 empty] | 0.094455 | 0.09834 | 0.094652 | 0.099408
|  | RSD |  |  | 10% | 18.83% | 7.743% | 257.7% | 11.72% | 53.16% | 25.89% | 4.92% | 6.66% | 2.601% | 4.678% | 19.24% | 2.346% | 65.44% | 6.931% | 19.68% | 9.698% | 1.077% | 1.126% | 25.91% | 28.79% | 2.059% | 5.295% | 27.78% | 31.83% | 23.22% | 24.96% | 23.65% | 26.52% | 4.191% | 3.158% | [7 empty] | 5.915% | 4.322% | 2.776% | 1.306%
|  | % Recovery |  |  | 99.83% | 75.81% | 65.72% | -51.84% | -566.4% | 279.1% | 103.8% | 103.9% | 96.54% | 112.2% | 103.4% | -98.9% | 89.63% | -161.2% | 97.23% | 119.7% | 113.9% | 109.3% | 101.9% | 152.3% | 56.36% | 245.7% | 95.37% | 141.9% | 36.33% | 110.1% | 105.1% | 103.9% | 129.2% | 97.14% | 96.44% | 94.78% | 90.18% | 96.15% | 92.76% | 94.03% | 88.56% | 91.39% | 94.46% | 98.34% | 94.65% | 99.41%
| 239 | LLOQ 0.1 |  | 06/09/2023 14:12:45 | 133.0755 | 130.3552 | -0.42096 | -0.35051 | -0.37528 | 202.177 | 139.043 | 145.5491 | 0.10403 | 0.10336 | 0.10204 | -0.17154 | 0.10846 | -0.051706 | 0.11566 | 0.10179 | 0.11814 | 0.10638 | 0.45672 | 0.2031 | 0.15196 | 0.35565 | 0.10433 | 0.2092 | 0.094324 | 0.091263 | 0.1066 | 0.10935 | 0.072271 | 0.11944 | 0.10119 | [7 empty] | 0.09904 | 0.097686 | 0.095693 | 0.095319
|  | RSD |  |  | 10.65% | 2.291% | 1.396% | 7.909% | 9.729% | 25.1% | 9.767% | 22.46% | 5.976% | 7.188% | 2.736% | 2.304% | 3.603% | 177.7% | 6.558% | 21.27% | 4.665% | 4.967% | 44.59% | 17.17% | 58.26% | 0.5793% | 2.907% | 32.04% | 74.45% | 10.41% | 22.28% | 129.9% | 46.92% | 1.277% | 3.202% | [7 empty] | 5.171% | 1.98% | 1.465% | 1.867%
|  | % Recovery |  |  | 133.1% | 130.4% | -421% | -350.5% | -375.3% | 202.2% | 139% | 145.5% | 104% | 103.4% | 102% | -171.5% | 108.5% | -51.71% | 115.7% | 101.8% | 118.1% | 106.4% | 456.7% | 203.1% | 152% | 355.6% | 104.3% | 209.2% | 94.32% | 91.26% | 106.6% | 109.3% | 72.27% | 119.4% | 101.2% | 97.83% | 95.77% | 97.96% | 95.98% | 97.76% | 97.33% | 96.95% | 99.04% | 97.69% | 95.69% | 95.32%
QC Duplicates (% Difference)
|  | Sample Id | R | Acquisition Time | Na KED 23/23 | Mg KED 24/24 | Al 27/27 | Al KED 27/27 | P Oshift 31/47 | S (34) -1 34/34 | K KED 39/39 | Ca KED 43/43 | V KED 51/51 | Cr (52) 52/52 | Cr (53) NH3 53/53 | Fe (54) NH3 54/54 | Mn Oshift 55/71 | Fe (56) KED 56/56 | Co KED 59/59 | Ni KED 60/60 | Cu (63) KED 63/63 | Cu (65) NH3 65/65 | Zn NH3 66/66 | As 75/75 | As KED 75/75 | As RH Oshift 75/91 | As Oshift 75/91 | Se (77) 77/77 | Se (77) Oshift 77/93 | Se (78) Oshift 78/94 | Se (80) Oshift 80/96 | Se (82) 82/82 | Se (82) Oshift 82/98 | Mo (95) NH3 95/95 | Mo (95) OShift 95/127 | Rh (IS1) 103/103 (IS) | Rh (IS2) 103/103 (IS) | Rh (IS3) 103/103 (IS) | Rh (IS4) 103/103 (IS) | Rh (IS5) 103/103 (IS) | Rh (IS6) 103/103 (IS) | Rh OShift (IS1) 103/119 (IS) | Cd (111) NH3 111/111 | Cd (114) NH3 114/114 | Pb 208/208 | Pb NH3 208/208
| 31 | 8_2_21_UF_6_SP_LD, Duplicate of Row Index 29 |  | 06/08/2023 16:26:19 | 0.3377% | 1.378% | 15.39% | 7.913% | 0.1995% | 2.661% | 2.259% | 2.404% | 0.3165% | 1.566% | 0.3986% | 8.963% | 2.73% | 9.457% | 2.877% | 0.3249% | 1.299% | 0.2675% | 1.519% | 1.202% | 2.277% | 1.538% | 0.01891% | 0.7673% | 6.233% | 1.596% | 1.72% | 5.555% | 0.08432% | 1.269% | 1.626% | [7 empty] | 0.9455% | 1.073% | 0.6216% | 1.278%
| 47 | 8_2_21_F_8_DUP_LD, Duplicate of Row Index 46 |  | 06/08/2023 18:01:24 | 20.9% | 15.23% | 56.82% | 48.91% | 43.08% | 23.1% | 17.67% | 20.29% | 14.39% | 27.73% | 34.64% | 77.59% | 1.634% | 58.25% | 21.11% | 4.402% | 17.57% | 5.147% | 6.366% | 1.244% | 19.16% | 2.202% | 1.765% | 1.664% | 12.36% | 7.632% | 13.01% | 1.281% | 12.19% | 5.48% | 2.61% | [7 empty] | 22.39% | 16.35% | 8.051% | 4.137%
| 48 | 8_2_21_F_8_DUP_LD, Duplicate of Row Index 46 |  | 06/08/2023 18:07:40 | 16.95% | 16.64% | 59.1% | 53.31% | 23.72% | 19.37% | 15.09% | 19.92% | 13.87% | 35.11% | 35.27% | 79.94% | 2.878% | 63.02% | 27.93% | 7.196% | 15.39% | 1.595% | 9.857% | 0.3919% | 21.85% | 2.729% | 0.4818% | 17% | 0.7134% | 4.823% | 15.07% | 16.22% | 57.6% | 4.286% | 2.265% | [7 empty] | 20.64% | 6.103% | 5.609% | 6.851%
| 94 | 8_16_21_F_1_DUP_LD, Duplicate of Row Index 93 |  | 06/08/2023 23:08:58 | 6.015% | 8.856% | 56.07% | 27.12% | 12.84% | 4.843% | 5.239% | 5.69% | 8.419% | 2.647% | 8.62% | 2.447% | 0.2098% | 5.442% | 13.97% | 3.812% | 3.316% | 2.769% | 7.623% | 3.833% | 4.899% | 1.332% | 1.386% | 2.979% | 24.66% | 30.82% | 21.43% | 15.42% | 24.04% | 1.557% | 1.824% | [7 empty] | 4.872% | 13.69% | 1.535% | 4.161%
| 95 | 8_16_21_F_1_DUP_LD, Duplicate of Row Index 93 |  | 06/08/2023 23:15:14 | 1.283% | 2.088% | 56.27% | 39.53% | 14.47% | 2.476% | 0.4002% | 1.1% | 3.569% | 13.17% | 9.174% | 2.848% | 0.2407% | 6.615% | 6.768% | 17.34% | 2.328% | 2.081% | 5.85% | 1.343% | 1.782% | 4.919% | 0.3446% | 23.99% | 26.29% | 4.617% | 21.44% | 27.7% | 44.52% | 1.169% | 0.8583% | [7 empty] | 28.26% | 14.19% | 11.75% | 4.134%
| 141 | MT 32 D LD, Duplicate of Row Index 140 |  | 06/09/2023 03:45:07 | 0.3733% | 0% | 77.72% | 102.6% | 17.5% | 1.028% | 159.1% | 33.67% | 36.51% | 35.95% | 24.85% | 12.39% | 34.3% | 19.97% | 13.95% | 6.437% | 27.52% | 26.1% | 25.67% | 228.5% | 224.1% | 0.5815% | 68.49% | 32.36% | 30.71% | 236.9% | 206.8% | 7.41% | 0% | 6.886% | 3.295% | [7 empty] | 272.5% | 570.2% | 3.793% | 0.3897%
| 142 | MT 32 D LD, Duplicate of Row Index 140 |  | 06/09/2023 03:51:24 | 3.974% | 2.507% | 84.65% | 78.63% | 6.067% | 33.77% | 159.9% | 33.63% | 16.76% | 19.52% | 21.99% | 13.9% | 35.54% | 22.5% | 0% | 21.85% | 38.56% | 27.96% | 25.84% | 21.67% | 224.1% | 0.418% | 55.92% | 3.857% | 0.7125% | 237.8% | 197.2% | 318.9% | 243.8% | 4.168% | 2.154% | [7 empty] | 109.3% | 14.19% | 2.283% | 4.235%
| 183 | 8_30_21_W_6_LD, Duplicate of Row Index 182 |  | 06/09/2023 07:54:05 | 4.988% | 7.161% | 28.65% | 24.56% | 41.53% | 3.67% | 3.605% | 4.553% | 4.531% | 31.81% | 26.96% | 1.353% | 9.571% | 4.027% | 8.657% | 8.835% | 5.435% | 1.43% | 0.439% | 1.807% | 7.354% | 9.796% | 8.469% | 20.1% | 186.1% | 9.923% | 27.37% | 158.2% | 69.42% | 1.61% | 9.823% | [7 empty] | 54.51% | 40.2% | 5.993% | 5.348%
| 184 | 8_30_21_W_6_LD, Duplicate of Row Index 182 |  | 06/09/2023 08:00:20 | 0.9767% | 1.676% | 25.06% | 29.7% | 51.96% | 6.32% | 4.972% | 6.873% | 2.429% | 19.48% | 28.72% | 0.8138% | 0.5965% | 6.68% | 1.115% | 6.184% | 3.884% | 0.7469% | 1.683% | 0.5424% | 14.34% | 0.2789% | 0.4542% | 6.126% | 85.94% | 15.69% | 26.43% | 128.2% | 35.17% | 0.02727% | 0.4844% | [7 empty] | 27.04% | 2.595% | 9.143% | 14.25%
| 247 | 8_30_21_UF_9_LD, Duplicate of Row Index 246 |  | 06/09/2023 15:00:03 | 4.545% | 0.1405% | 123.2% | 139.2% | 20.14% | 0.3233% | 1.392% | 0.63% | 0.9295% | 17.71% | 3.143% | 89.49% | 2.901% | 22.73% | 11.57% | 15.14% | 6.847% | 4.435% | 69.36% | 1.023% | 22.57% | 0.54% | 2.428% | 2.21% | 50.74% | 3.358% | 25.21% | 24.8% | 3.298% | 4.88% | 2.78% | [7 empty] | 5.731% | 6.393% | 9.387% | 5.977%
| 248 | 8_30_21_UF_9_LD, Duplicate of Row Index 246 |  | 06/09/2023 15:06:19 | 5.373% | 0.3409% | 123.8% | 135% | 11.86% | 1.578% | 0.3418% | 2.759% | 2.295% | 1.177% | 8.044% | 46.48% | 1.872% | 27.68% | 50.75% | 1.887% | 4.862% | 1.408% | 110.1% | 2.483% | 11.78% | 1.572% | 0.5461% | 28.22% | 96.39% | 11.99% | 6.618% | 34.3% | 118.5% | 1.545% | 1.302% | [7 empty] | 9.931% | 18.48% | 6.184% | 3.249%
| 281 | 9_13_21_F_11_LD, Duplicate of Row Index 280 |  | 06/09/2023 18:19:44 | 0.9177% | 1.896% | 11.91% | 19.18% | 36.24% | 1.261% | 1.892% | 1.444% | 7.303% | 6.629% | 8.915% | 7.432% | 1.765% | 1.821% | 20.96% | 1.947% | 1.632% | 2.716% | 6.336% | 2.702% | 4.083% | 0.3675% | 0.8683% | 7.789% | 293.5% | 6.139% | 7.697% | 28.43% | 65.66% | 4.152% | 0.8553% | [7 empty] | 9.025% | 11.25% | 1.61% | 1.326%
| 282 | 9_13_21_F_11_LD, Duplicate of Row Index 280 |  | 06/09/2023 18:26:00 | 1.946% | 0.9471% | 14.45% | 14.68% | 38.23% | 6.038% | 0.5109% | 2.432% | 7.846% | 0.1857% | 16.11% | 9.94% | 1.013% | 5.562% | 13.92% | 8.55% | 1.393% | 5.27% | 8.34% | 0.1266% | 1.065% | 0.6928% | 1.041% | 0.8991% | 86.52% | 8.745% | 7.278% | 4.573% | 53.7% | 6.695% | 0.02148% | [7 empty] | 3.684% | 17.17% | 3.581% | 2.836%
Samples, Blanks & Calibration Standards
|  | Sample Id | R | Acquisition Time | Na KED 23/23 | Mg KED 24/24 | Al 27/27 | Al KED 27/27 | P Oshift 31/47 | S (34) -1 34/34 | K KED 39/39 | Ca KED 43/43 | V KED 51/51 | Cr (52) 52/52 | Cr (53) NH3 53/53 | Fe (54) NH3 54/54 | Mn Oshift 55/71 | Fe (56) KED 56/56 | Co KED 59/59 | Ni KED 60/60 | Cu (63) KED 63/63 | Cu (65) NH3 65/65 | Zn NH3 66/66 | As 75/75 | As KED 75/75 | As RH Oshift 75/91 | As Oshift 75/91 | Se (77) 77/77 | Se (77) Oshift 77/93 | Se (78) Oshift 78/94 | Se (80) Oshift 80/96 | Se (82) 82/82 | Se (82) Oshift 82/98 | Mo (95) NH3 95/95 | Mo (95) OShift 95/127 | Rh (IS1) 103/103 (IS) | Rh (IS2) 103/103 (IS) | Rh (IS3) 103/103 (IS) | Rh (IS4) 103/103 (IS) | Rh (IS5) 103/103 (IS) | Rh (IS6) 103/103 (IS) | Rh OShift (IS1) 103/119 (IS) | Cd (111) NH3 111/111 | Cd (114) NH3 114/114 | Pb 208/208 | Pb NH3 208/208
| 166 | std 50 |  | 06/09/2023 06:13:06 | 51651.8561 | 48467.7539 | 55.9609 | 50.8192 | 60.0335 | 53601.7771 | 50084.3154 | 56040.8476 | 56.0659 | 54.2931 | 58.1352 | 54.7715 | 63.5806 | 53.763 | 54.8729 | 52.9413 | 50.4784 | 52.6223 | 54.5789 | 53.1201 | 48.1276 | 52.5513 | 60.13 | 51.8461 | 59.4216 | 59.8176 | 60.1396 | 55.8621 | 60.35 | 59.4817 | 63.8111 | [7 empty] | 51.0029 | 51.2617 | 46.2789 | 49.1607
|  | IS % Recovery | [33 empty] | 75.09% | 58.15% | 69.78% | 63.71% | 65.7% | 60.31% | 60.94%
|  | Sample RSD |  |  | 0.7463% | 0.3294% | 2.691% | 2.032% | 1.259% | 2.406% | 1.939% | 1.504% | 1.572% | 2.351% | 0.9776% | 1.227% | 3.2% | 0.9811% | 2.075% | 1.414% | 2.094% | 0.2198% | 1.525% | 1.457% | 3.145% | 3.163% | 3.343% | 0.7246% | 4.44% | 2.618% | 2.156% | 1.403% | 1.497% | 0.6338% | 2.722% | [7 empty] | 0.4949% | 1.412% | 2.071% | 1.59%
|  | Sample Concentration SD |  |  | 385.5019 | 159.6628 | 1.506 | 1.0328 | 0.75575 | 1289.6392 | 971.0027 | 842.6086 | 0.88145 | 1.2766 | 0.56833 | 0.6718 | 2.0345 | 0.52746 | 1.1384 | 0.74865 | 1.0572 | 0.11568 | 0.8323 | 0.77417 | 1.5135 | 1.6623 | 2.0099 | 0.37569 | 2.6382 | 1.566 | 1.2969 | 0.78369 | 0.90368 | 0.37701 | 1.737 | [7 empty] | 0.25243 | 0.72369 | 0.95834 | 0.78145
| 168 | std 100 |  | 06/09/2023 06:25:11 | 112871.2702 | 104438.8781 | 119.8729 | 108.1934 | 122.6976 | 120006.7696 | 112264.9622 | 122218.2656 | 123.8248 | 118.1572 | 122.1673 | 111.7135 | 125.6833 | 113.8752 | 117.0495 | 111.5168 | 106.2146 | 103.9908 | 108.2579 | 109.2092 | 102.5304 | 99.6412 | 113.5297 | 104.3797 | 117.7598 | 114.707 | 116.6555 | 115.5131 | 116.6196 | 123.3314 | 125.0769 | [7 empty] | 100.9788 | 101.576 | 88.5866 | 95.0166
|  | IS % Recovery | [33 empty] | 67.03% | 55.16% | 61.17% | 56.92% | 59.11% | 54.45% | 57.49%
|  | Sample RSD |  |  | 2.768% | 2.826% | 0.8988% | 2.491% | 0.9708% | 3.456% | 1.161% | 2.354% | 3.269% | 3.302% | 0.8056% | 1.222% | 0.7578% | 2.6% | 2.587% | 2.338% | 1.254% | 1.729% | 1.25% | 0.9519% | 4.169% | 1.415% | 0.6251% | 1.051% | 2.572% | 1.7% | 1.15% | 1.106% | 1.769% | 0.2886% | 0.5067% | [7 empty] | 1.259% | 0.8655% | 1.199% | 0.291%
|  | Sample Concentration SD |  |  | 3124.1874 | 2951.8517 | 1.0775 | 2.6949 | 1.1911 | 4146.8462 | 1303.4086 | 2877.0605 | 4.0484 | 3.9015 | 0.98421 | 1.3654 | 0.95245 | 2.961 | 3.0284 | 2.6072 | 1.3314 | 1.7985 | 1.3537 | 1.0395 | 4.2747 | 1.4099 | 0.70968 | 1.0969 | 3.0283 | 1.9497 | 1.3414 | 1.2778 | 2.0625 | 0.3559 | 0.63378 | [7 empty] | 1.2716 | 0.87918 | 1.0626 | 0.27654
| 169 | std 100 |  | 06/09/2023 06:31:28 | 114581.4942 | 107924.3964 | 119.3748 | 109.6745 | 123.6732 | 124004.7376 | 113937.6549 | 126412.622 | 128.7455 | 122.1822 | 122.1227 | 113.51 | 124.9597 | 118.1089 | 119.4415 | 112.6024 | 109.9544 | 104.9428 | 107.6217 | 107.7039 | 105.1992 | 97.7652 | 113.9056 | 102.8396 | 116.2066 | 114.668 | 116.4887 | 114.8819 | 116.5623 | 123.9658 | 124.9704 | [7 empty] | 100.0453 | 100.768 | 88.0225 | 107.195
|  | IS % Recovery | [33 empty] | 68.64% | 55.86% | 60.32% | 57.32% | 59.31% | 54.65% | 59.54%
|  | Sample RSD |  |  | 1.599% | 2.293% | 0.7628% | 2.212% | 0.4017% | 3.062% | 1.315% | 3.065% | 3.353% | 3.193% | 1.015% | 0.3053% | 0.1406% | 2.001% | 2.443% | 2.066% | 2.468% | 0.3275% | 0.4105% | 1.041% | 3.07% | 1.333% | 0.4331% | 1.422% | 0.899% | 1.1% | 0.728% | 0.8383% | 1.541% | 1.058% | 1.215% | [7 empty] | 0.4738% | 0.5588% | 1.111% | 9.818%
|  | Sample Concentration SD |  |  | 1832.1689 | 2475.1653 | 0.91063 | 2.4259 | 0.49685 | 3796.7299 | 1498.4319 | 3874.3512 | 4.3173 | 3.901 | 1.2394 | 0.34658 | 0.17563 | 2.3638 | 2.9174 | 2.3267 | 2.714 | 0.34369 | 0.44176 | 1.1217 | 3.2291 | 1.3035 | 0.49333 | 1.4625 | 1.0447 | 1.2614 | 0.84804 | 0.96302 | 1.7958 | 1.312 | 1.5188 | [7 empty] | 0.47405 | 0.56308 | 0.97789 | 10.5249
| 187 | 8_30_21_W_6_DUP_LSP |  | 06/09/2023 08:17:43 | 24410.2748 | 22811.342 | 29.2755 | 36.2223 | 26.8324 | 28535.2011 | 21352.6907 | 32565.2048 | 23.717 | 22.9373 | 22.4266 | 33.6525 | 27.1012 | 34.6269 | 23.114 | 23.4389 | 22.9904 | 21.925 | 21.8919 | 21.955 | 21.0684 | 19.1227 | 22.7403 | 20.8789 | 22.4305 | 21.4571 | 21.5996 | 21.79 | 22.1116 | 24.9539 | 25.2518 | [7 empty] | 20.5686 | 20.6769 | 19.5534 | 20.5707
|  | IS % Recovery | [33 empty] | 76.7% | 59.98% | 62.48% | 63.08% | 64.6% | 58.64% | 65.69%
|  | Sample RSD |  |  | 3.001% | 3.341% | 27.15% | 8.439% | 0.6343% | 6.243% | 5.854% | 4.977% | 4.712% | 6.375% | 1.388% | 0.7777% | 0.7105% | 4.254% | 5.191% | 4.995% | 4.63% | 0.7684% | 1.008% | 26.43% | 8.012% | 0.8958% | 0.3974% | 30.91% | 1.618% | 1.383% | 2.027% | 24.76% | 2.694% | 0.4598% | 1.386% | [7 empty] | 0.8698% | 0.6223% | 26.52% | 1.845%
|  | Sample Concentration SD |  |  | 732.5529 | 762.054 | 7.9494 | 3.0569 | 0.17021 | 1781.3519 | 1249.9078 | 1620.7138 | 1.1176 | 1.4621 | 0.31128 | 0.26172 | 0.19254 | 1.4729 | 1.1999 | 1.1707 | 1.0645 | 0.16847 | 0.22076 | 5.8032 | 1.688 | 0.17131 | 0.090372 | 6.4533 | 0.36301 | 0.29668 | 0.43787 | 5.3942 | 0.59579 | 0.11474 | 0.35004 | [7 empty] | 0.17891 | 0.12866 | 5.1861 | 0.37949
| 188 | 8_30_21_W_6_DUP_LSP |  | 06/09/2023 08:24:01 | 23297.3304 | 22080.5706 | 29.1137 | 34.8303 | 27.1284 | 27411.4871 | 20719.8489 | 31220.0385 | 23.0276 | 22.4004 | 22.4249 | 33.4705 | 27.5932 | 33.3037 | 22.5586 | 22.6137 | 22.2724 | 22.1001 | 22.1271 | 21.7824 | 21.1742 | 19.9588 | 23.1001 | 20.637 | 21.8095 | 21.8018 | 22.021 | 21.6341 | 22.2679 | 24.9207 | 25.629 | [7 empty] | 20.3124 | 20.5536 | 18.8463 | 20.3084
|  | IS % Recovery | [33 empty] | 76.04% | 60.61% | 66.26% | 64.13% | 65.73% | 59.9% | 64.6%
|  | Sample RSD |  |  | 3.492% | 3.523% | 3.049% | 6.238% | 0.9339% | 6.973% | 2.351% | 3.663% | 4.047% | 3.548% | 1.192% | 1.517% | 0.5419% | 3.566% | 4.574% | 3.429% | 4.801% | 0.8837% | 0.6958% | 0.859% | 4.563% | 2.481% | 1.286% | 0.503% | 0.7655% | 2.404% | 1.345% | 1.831% | 2.647% | 1.32% | 0.4473% | [7 empty] | 1.804% | 0.9137% | 1.416% | 1.536%
|  | Sample Concentration SD |  |  | 813.4913 | 777.8766 | 0.88764 | 2.1726 | 0.25335 | 1911.3495 | 487.1928 | 1143.4632 | 0.93198 | 0.79468 | 0.26732 | 0.5079 | 0.14952 | 1.1877 | 1.0319 | 0.77553 | 1.0693 | 0.19529 | 0.15396 | 0.1871 | 0.96616 | 0.49524 | 0.29706 | 0.1038 | 0.16694 | 0.52401 | 0.29619 | 0.39614 | 0.58946 | 0.32898 | 0.11464 | [7 empty] | 0.36647 | 0.1878 | 0.26688 | 0.3119
| 189 | 8_30_21_W_8 |  | 06/09/2023 08:29:49 | 1906.6865 | 1437.7037 | 8.7656 | 11.4516 | 5.2063 | 4353.9732 | 391.7948 | 7581.3015 | 0.17632 | 0.047681 | 0.039718 | 7.9169 | 4.0568 | 8.4035 | 0.011605 | 0.066095 | 1.0286 | 0.98613 | 0.35984 | 1.0637 | 0.95114 | 1.1036 | 1.1419 | 0.33572 | -0.011255 | 0.023672 | 0.023889 | 0.094173 | 0.025466 | 1.6545 | 1.6208 | [7 empty] | 0.0046769 | 0.0037236 | 0.061714 | 0.065929
|  | IS % Recovery | [33 empty] | 81.34% | 65.75% | 70.13% | 67.02% | 71.55% | 52.74% | 72.69%
|  | Sample RSD |  |  | 13.86% | 6.71% | 0.8871% | 12.22% | 8.516% | 13.8% | 12.48% | 12.29% | 2.526% | 9.115% | 11.13% | 2.942% | 0.7847% | 11.26% | 27.98% | 13.98% | 10.53% | 2.496% | 4.542% | 4.261% | 27% | 3.703% | 0.4845% | 8.316% | 163.2% | 20.94% | 25.97% | 32.11% | 52.82% | 2.549% | 1.609% | [7 empty] | 40.09% | 25.43% | 3.151% | 3.287%
|  | Sample Concentration SD |  |  | 264.2432 | 96.4709 | 0.077757 | 1.3999 | 0.44339 | 600.719 | 48.9103 | 932.009 | 0.0044539 | 0.0043463 | 0.0044208 | 0.2329 | 0.031833 | 0.94614 | 0.0032467 | 0.0092429 | 0.1083 | 0.024618 | 0.016345 | 0.045322 | 0.25677 | 0.040869 | 0.0055325 | 0.027918 | 0.018364 | 0.0049578 | 0.0062052 | 0.030243 | 0.01345 | 0.042171 | 0.026082 | [7 empty] | 0.0018749 | 0.00094683 | 0.0019445 | 0.0021669
| 200 | 8_30_21_F_1 |  | 06/09/2023 09:35:08 | 7546.9993 | 5893.4038 | 4.1234 | 5.4336 | 8.2326 | 23277.9562 | 1918.9465 | 39001.853 | 0.36695 | 0.13047 | 0.13971 | 3.8402 | 12.858 | 4.2528 | 0.035297 | 0.4696 | 0.91798 | 0.92522 | 5.6603 | 4.7334 | 4.1202 | 4.211 | 5.0723 | 0.58023 | 0.043061 | 0.04526 | 0.051136 | 0.40318 | 0.028399 | 16.5593 | 17.217 | [7 empty] | 0.0076911 | 0.006764 | 0.094584 | 0.1036
|  | IS % Recovery | [33 empty] | 75.23% | 59.18% | 65.74% | 63.28% | 65.05% | 59.99% | 67.7%
|  | Sample RSD |  |  | 1.818% | 4.395% | 1.748% | 12.91% | 2.608% | 3.488% | 5.899% | 3.35% | 11.41% | 10.3% | 3.948% | 1.156% | 1.699% | 7.664% | 17.86% | 4.757% | 6.039% | 1.576% | 0.517% | 0.3426% | 4.485% | 3.034% | 0.5212% | 17.29% | 84.8% | 26.32% | 31.89% | 13.73% | 52.06% | 2.078% | 2.078% | [7 empty] | 16.56% | 10.51% | 4.856% | 1.446%
|  | Sample Concentration SD |  |  | 137.1797 | 259.0168 | 0.072072 | 0.70135 | 0.21467 | 811.9052 | 113.2056 | 1306.4793 | 0.041863 | 0.01344 | 0.0055165 | 0.044388 | 0.21851 | 0.32594 | 0.0063032 | 0.02234 | 0.055438 | 0.014582 | 0.029266 | 0.016219 | 0.18477 | 0.12778 | 0.026437 | 0.10032 | 0.036516 | 0.011913 | 0.016307 | 0.055364 | 0.014786 | 0.34403 | 0.35779 | [7 empty] | 0.0012733 | 0.00071122 | 0.0045934 | 0.0014985
| 202 | 8_30_21_F_1_DUP |  | 06/09/2023 09:47:13 | 7752.4812 | 5762.7149 | 4.9108 | 5.3855 | 7.619 | 23318.8905 | 1795.3641 | 37666.5611 | 0.38171 | 0.13172 | 0.12302 | 3.8841 | 12.5475 | 4.2022 | 0.03405 | 0.47364 | 0.91571 | 0.94654 | 5.6642 | 5.1486 | 4.0809 | 4.4115 | 5.0008 | 0.72635 | 0.035867 | 0.056912 | 0.04277 | 0.4554 | 0.051839 | 16.6007 | 17.0399 | [7 empty] | 0.0075863 | 0.006655 | 0.10496 | 0.10493
|  | IS % Recovery | [33 empty] | 69.54% | 60.69% | 67.16% | 63.29% | 66.22% | 59.43% | 65.19%
|  | Sample RSD |  |  | 6.548% | 5.553% | 17.84% | 13.89% | 2.308% | 5.251% | 5.929% | 4.311% | 5.597% | 2.617% | 1.93% | 0.6876% | 0.87% | 8.976% | 6.471% | 10.12% | 4.245% | 0.6914% | 0.883% | 14.17% | 10.45% | 1.903% | 1.672% | 54.13% | 74.29% | 39.4% | 9.34% | 25.91% | 70.1% | 0.1152% | 1.045% | [7 empty] | 25.82% | 21.32% | 13.01% | 1.424%
|  | Sample Concentration SD |  |  | 507.6378 | 320.0173 | 0.87625 | 0.74791 | 0.17586 | 1224.5761 | 106.4466 | 1623.7946 | 0.021363 | 0.0034469 | 0.0023742 | 0.026705 | 0.10916 | 0.37719 | 0.0022033 | 0.047923 | 0.038871 | 0.006544 | 0.050012 | 0.7296 | 0.42651 | 0.083942 | 0.083627 | 0.39319 | 0.026647 | 0.022421 | 0.0039948 | 0.11799 | 0.036339 | 0.01913 | 0.17812 | [7 empty] | 0.0019584 | 0.001419 | 0.013651 | 0.0014939
| 203 | 8_30_21_F_1_DUP |  | 06/09/2023 09:53:30 | 7650.2474 | 5679.3873 | 4.371 | 5.7819 | 7.4682 | 22870.8844 | 1842.7719 | 36766.0404 | 0.37471 | 0.12309 | 0.12712 | 3.8458 | 12.7895 | 3.9592 | 0.03163 | 0.45206 | 0.87886 | 0.94584 | 5.6941 | 4.6671 | 4.0617 | 4.4278 | 5.057 | 0.49434 | 0.018815 | 0.042613 | 0.04766 | 0.39072 | 0.067201 | 16.5282 | 17.0674 | [7 empty] | 0.0080843 | 0.0059629 | 0.095829 | 0.10651
|  | IS % Recovery | [33 empty] | 74.41% | 60.17% | 67.53% | 62.93% | 65.38% | 59.07% | 65.11%
|  | Sample RSD |  |  | 5.558% | 4.213% | 7.668% | 2.941% | 1.862% | 2.888% | 6.903% | 5.911% | 4.218% | 13.3% | 9.249% | 0.8739% | 1.148% | 2.167% | 29.69% | 7.366% | 2.644% | 0.5441% | 0.3573% | 6.312% | 4.046% | 1.089% | 1.206% | 24.15% | 52.57% | 13.15% | 51.78% | 7.264% | 34.23% | 1.222% | 0.8524% | [7 empty] | 28.48% | 10.96% | 8.257% | 5.012%
|  | Sample Concentration SD |  |  | 425.1767 | 239.2478 | 0.33517 | 0.17007 | 0.13903 | 660.4859 | 127.2011 | 2173.113 | 0.015806 | 0.016367 | 0.011757 | 0.033607 | 0.14683 | 0.085808 | 0.0093916 | 0.033299 | 0.023235 | 0.0051464 | 0.020343 | 0.2946 | 0.16434 | 0.04822 | 0.060975 | 0.11938 | 0.0098906 | 0.0056056 | 0.024679 | 0.028382 | 0.023 | 0.20193 | 0.14548 | [7 empty] | 0.0023024 | 0.00065362 | 0.0079129 | 0.0053389
| 204 | 8_30_21_F_6 |  | 06/09/2023 09:59:18 | 8437.8915 | 6658.332 | 2.5772 | 2.7705 | 0.89808 | 19039.1386 | 1830.7598 | 36241.4473 | 0.72013 | 0.049118 | 0.063106 | 6.3566 | 12.4064 | 6.0683 | 0.028888 | 0.25293 | 2.7501 | 2.9508 | 3.3973 | 4.889 | 4.3966 | 4.629 | 5.2435 | 0.51356 | 0.053797 | 0.10099 | 0.083519 | 0.36134 | 0.13043 | 9.4547 | 9.5542 | [7 empty] | 0.02032 | 0.016284 | 0.12611 | 0.16512
|  | IS % Recovery | [33 empty] | 73.67% | 60.41% | 68.62% | 62.3% | 55.02% | 59.39% | 64.83%
|  | Sample RSD |  |  | 2.939% | 7.033% | 1.349% | 4.032% | 4.89% | 3.684% | 4.951% | 4.135% | 5.602% | 12.3% | 5.338% | 1.149% | 1.894% | 3.482% | 17.24% | 8.813% | 4.528% | 1.61% | 1.743% | 0.6164% | 6.319% | 5.037% | 1.5% | 10.73% | 83.09% | 21.13% | 21% | 9.295% | 23.68% | 1.589% | 1.746% | [7 empty] | 10.85% | 10.89% | 1.269% | 6.986%
|  | Sample Concentration SD |  |  | 247.9776 | 468.2738 | 0.034762 | 0.11172 | 0.043917 | 701.3455 | 90.6477 | 1498.7273 | 0.040342 | 0.0060409 | 0.0033685 | 0.073017 | 0.23496 | 0.21127 | 0.0049814 | 0.022291 | 0.12451 | 0.047499 | 0.059202 | 0.030136 | 0.27783 | 0.23316 | 0.078639 | 0.055127 | 0.044702 | 0.021335 | 0.017535 | 0.033587 | 0.030886 | 0.15026 | 0.16683 | [7 empty] | 0.0022055 | 0.0017735 | 0.0016002 | 0.011536
| 206 | 8_30_21_F_8 |  | 06/09/2023 10:11:26 | 9010.4641 | 6602.8872 | 1.5135 | 1.8563 | 2.452 | 18134.7389 | 1923.611 | 36632.3517 | 0.90541 | 0.068773 | 0.062577 | 6.8386 | 15.1166 | 6.8628 | 0.026937 | 0.2484 | 3.1665 | 3.3355 | 1.6868 | 5.28 | 4.4256 | 5.3196 | 6.1959 | 0.5456 | 0.041655 | 0.099006 | 0.090011 | 0.42065 | 0.090699 | 8.5001 | 9.4844 | [7 empty] | 0.017546 | 0.013855 | 0.10288 | 0.11192
|  | IS % Recovery | [33 empty] | 73.74% | 55.67% | 68.63% | 61.82% | 64.27% | 59.12% | 61%
|  | Sample RSD |  |  | 3.255% | 3.768% | 3.273% | 1.158% | 2.254% | 2.003% | 5.937% | 2.208% | 10.5% | 27.42% | 4.147% | 1.184% | 8.02% | 6.344% | 23.99% | 19.65% | 4.337% | 0.9351% | 3.194% | 0.5447% | 10.08% | 7.39% | 8.555% | 35.17% | 72.19% | 34.21% | 15.73% | 5.41% | 60.17% | 1.083% | 8.542% | [7 empty] | 13.17% | 3.299% | 3.91% | 1.675%
|  | Sample Concentration SD |  |  | 293.3108 | 248.7658 | 0.049534 | 0.021491 | 0.055258 | 363.2922 | 114.1958 | 808.9085 | 0.095053 | 0.018858 | 0.0025948 | 0.080953 | 1.2124 | 0.43536 | 0.0064621 | 0.048821 | 0.13732 | 0.03119 | 0.053871 | 0.028759 | 0.44612 | 0.39309 | 0.53006 | 0.19186 | 0.03007 | 0.033868 | 0.014163 | 0.022758 | 0.054578 | 0.092056 | 0.81019 | [7 empty] | 0.0023101 | 0.00045707 | 0.0040228 | 0.0018751
| 207 | 8_30_21_F_8 |  | 06/09/2023 10:17:43 | 9078.2803 | 6883.7996 | 1.5155 | 1.2477 | 2.3067 | 18594.5226 | 1935.3788 | 38345.1762 | 0.90178 | 0.067133 | 0.063359 | 6.8157 | 13.4181 | 7.1124 | 0.028799 | 0.22869 | 3.2746 | 3.2554 | 1.6816 | 5.3109 | 4.9143 | 4.8249 | 5.4584 | 0.59737 | 0.087629 | 0.09677 | 0.092591 | 0.45284 | 0.085654 | 8.4679 | 8.398 | [7 empty] | 0.015406 | 0.014741 | 0.10661 | 0.11109
|  | IS % Recovery | [33 empty] | 73.75% | 60.95% | 66.15% | 62.27% | 65.57% | 58.39% | 65.08%
|  | Sample RSD |  |  | 2.693% | 4.847% | 2.374% | 28.26% | 3.888% | 3.363% | 2.204% | 1.808% | 2.666% | 5.608% | 1.908% | 1.509% | 5.159% | 2.629% | 5.586% | 13.8% | 0.2135% | 1.566% | 0.8485% | 3.101% | 6.826% | 3.873% | 5.116% | 7.968% | 17.03% | 20.58% | 25.93% | 14.28% | 4.72% | 0.502% | 5.52% | [7 empty] | 21.67% | 8.397% | 2.136% | 2.39%
|  | Sample Concentration SD |  |  | 244.4623 | 333.6669 | 0.035974 | 0.35265 | 0.089694 | 625.3482 | 42.6483 | 693.1464 | 0.024042 | 0.0037647 | 0.0012089 | 0.10285 | 0.69227 | 0.187 | 0.0016086 | 0.031567 | 0.0069919 | 0.050974 | 0.014269 | 0.16466 | 0.33545 | 0.18686 | 0.27924 | 0.0476 | 0.014928 | 0.019916 | 0.024009 | 0.064672 | 0.0040433 | 0.042509 | 0.46355 | [7 empty] | 0.0033381 | 0.0012378 | 0.0022773 | 0.0026553

## Calibration Curves 14
Calibration Table - Na KED 23/23
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 0.66667 | 173.2051 | 0 |  | 0 | 0
Cal. Std.4 | 0.067659 | 9.7064 | 500 | 498.4736 | -1.5264 | -0.30528
Cal. Std.5 | 0.13488 | 5.6906 | 1000 | 1007.0998 | 7.0998 | 0.70998
Cal. Std.6 | 0.65964 | 4.0162 | 5000 | 4977.807 | -22.193 | -0.44386
Cal. Std.7 | 1.3162 | 0.15691 | 10000 | 9946.064 | -53.936 | -0.53936
Cal. Std.8 | 3.3594 | 1.7658 | 25000 | 25406.3918 | 406.3918 | 1.6256
Cal. Std.9 | 6.5352 | 3.2525 | 50000 | 49436.9049 | -563.0951 | -1.1262
Cal. Std.10 | 13.228 | 1.0757 | 100000 | 100079.1431 | 79.1431 | 0.079143
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99995 | 0.00013216 | 0.0017826 | 6.0638e-07 | 0.00071377 | 0.3659 | 1.9013
Calibration Table - Mg KED 24/24
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 0 |  | 0 |  | 0 | 0
Cal. Std.4 | 0.029335 | 2.0699 | 500 | 485.3308 | -14.6692 | -2.9338
Cal. Std.5 | 0.06095 | 9.4037 | 1000 | 1054.4398 | 54.4398 | 5.444
Cal. Std.6 | 0.28232 | 3.062 | 5000 | 5039.4171 | 39.4171 | 0.78834
Cal. Std.7 | 0.57525 | 1.2033 | 10000 | 10312.481 | 312.481 | 3.1248
Cal. Std.8 | 1.4062 | 0.60685 | 25000 | 25270.3596 | 270.3596 | 1.0814
Cal. Std.9 | 2.7392 | 5.3702 | 50000 | 49266.5198 | -733.4802 | -1.467
Cal. Std.10 | 5.2221 | 0.81279 | 100000 | 93962.2249 | -6037.7751 | -6.0378
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99907 | 5.5551e-05 | 0.0023743 | 1.0715e-06 | 0.0012613 | 0 | 0
Calibration Table - Al 27/27
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 3125.6753 | 1.8553 | 0 |  | 0 | 0
Cal. Std.4 | 0.54334 | 0.84256 | 0.5 | 0.47868 | -0.021319 | -4.2637
Cal. Std.5 | 0.82088 | 2.0436 | 1 | 1.0892 | 0.089218 | 8.9218
Cal. Std.6 | 2.6337 | 3.8349 | 5 | 5.0771 | 0.077086 | 1.5417
Cal. Std.7 | 4.6405 | 1.0911 | 10 | 9.4917 | -0.50827 | -5.0827
Cal. Std.8 | 10.8929 | 1.4504 | 25 | 23.246 | -1.754 | -7.0161
Cal. Std.9 | 23.6723 | 1.0326 | 50 | 51.3582 | 1.3582 | 2.7165
Cal. Std.10 | 47.2306 | 0.56987 | 100 | 103.1824 | 3.1824 | 3.1824
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99798 | 0.45458 | 0.32574 | 0.012946 | 0.015239 | 0.038176 | 0.0014106
Calibration Table - Al KED 27/27
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 0.66667 | 173.2051 | 0 |  | 0 | 0
Cal. Std.4 | 0.0042787 | 32.6156 | 0.5 | 0.47216 | -0.027842 | -5.5683
Cal. Std.5 | 0.0063906 | 21.4345 | 1 | 1.1036 | 0.10364 | 10.3639
Cal. Std.6 | 0.020356 | 4.2899 | 5 | 5.2792 | 0.27917 | 5.5835
Cal. Std.7 | 0.035472 | 5.7512 | 10 | 9.7988 | -0.20117 | -2.0117
Cal. Std.8 | 0.08523 | 2.7095 | 25 | 24.6767 | -0.32328 | -1.2931
Cal. Std.9 | 0.16654 | 5.3351 | 50 | 48.987 | -1.013 | -2.026
Cal. Std.10 | 0.32026 | 1.0156 | 100 | 94.9518 | -5.0482 | -5.0482
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99785 | 0.0033445 | 0.0026995 | 9.8214e-05 | 0.00011561 | 0.014168 | 0.073621
Calibration Table - P Oshift 31/47
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 1866.122 | 4.2227 | 0 |  | 0 | 0
Cal. Std.1 | 0.015805 | 2.1833 | 0.01 | 0.0018259 | -0.0081741 | -81.7407
Cal. Std.2 | 0.026814 | 3.2204 | 0.05 | 0.20978 | 0.15978 | 319.5697
Cal. Std.3 | 0.031078 | 5.624 | 0.1 | 0.29033 | 0.19033 | 190.327
Cal. Std.4 | 0.037774 | 1.9744 | 0.5 | 0.41682 | -0.083184 | -16.6368
Cal. Std.5 | 0.03598 | 0.61515 | 1 | 0.38292 | -0.61708 | -61.7082
Cal. Std.6 | 0.099592 | 2.2624 | 5 | 1.5846 | -3.4154 | -68.3088
Cal. Std.7 | 0.17184 | 0.79518 | 10 | 2.9493 | -7.0507 | -70.5075
Cal. Std.8 | 0.41718 | 1.485 | 25 | 7.5839 | -17.4161 | -69.6645
Cal. Std.9 | 0.8018 | 0.6558 | 50 | 14.8493 | -35.1507 | -70.3013
Cal. Std.10 | 1.5494 | 1.0296 | 100 | 28.971 | -71.029 | -71.029
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.5689 | 0.052938 | 0.015709 | 0.027056 | 0.00083478 | 0.15293 | 0.017561
Calibration Table - S (34) -1 34/34
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 636.0142 | 5.5902 | 0 |  | 0 | 0
Cal. Std.5 | 0.019425 | 13.8701 | 1000 | 1007.023 | 7.023 | 0.7023
Cal. Std.6 | 0.099651 | 7.3881 | 5000 | 4877.3394 | -122.6606 | -2.4532
Cal. Std.7 | 0.19793 | 1.8831 | 10000 | 9618.3635 | -381.6365 | -3.8164
Cal. Std.8 | 0.53608 | 2.6343 | 25000 | 25931.7876 | 931.7876 | 3.7272
Cal. Std.9 | 1.0376 | 4.4481 | 50000 | 50124.5599 | 124.5599 | 0.24912
Cal. Std.10 | 2.1044 | 1.8011 | 100000 | 101591.0065 | 1591.0065 | 1.591
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99955 | 2.0729e-05 | -0.0014491 | 3.0935e-07 | 0.00073876 | 2194.029 | 423.1612
Calibration Table - K KED 39/39
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 17.3333 | 40.5218 | 0 |  | 0 | 0
Cal. Std.4 | 0.032215 | 13.3903 | 500 | 508.0443 | 8.0443 | 1.6089
Cal. Std.5 | 0.06226 | 6.2734 | 1000 | 976.9374 | -23.0626 | -2.3063
Cal. Std.6 | 0.30558 | 3.3251 | 5000 | 4774.2096 | -225.7904 | -4.5158
Cal. Std.7 | 0.63169 | 0.85023 | 10000 | 9863.6386 | -136.3614 | -1.3636
Cal. Std.8 | 1.6256 | 0.65231 | 25000 | 25375.2676 | 375.2676 | 1.5011
Cal. Std.9 | 3.2588 | 3.9849 | 50000 | 50863.5671 | 863.5671 | 1.7271
Cal. Std.10 | 6.6219 | 1.881 | 100000 | 103348.6199 | 3348.6199 | 3.3486
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99951 | 6.4076e-05 | -0.00033835 | 8.9744e-07 | 0.0010564 | 19.2815 | 23.0107
Calibration Table - Ca KED 43/43
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 0.66667 | 173.2051 | 0 |  | 0 | 0
Cal. Std.4 | 0.013254 | 22.0623 | 500 | 524.8403 | 24.8403 | 4.9681
Cal. Std.5 | 0.022617 | 6.2068 | 1000 | 915.9157 | -84.0843 | -8.4084
Cal. Std.6 | 0.11072 | 3.9347 | 5000 | 4595.9485 | -404.0515 | -8.081
Cal. Std.7 | 0.23574 | 0.80551 | 10000 | 9817.8288 | -182.1712 | -1.8217
Cal. Std.8 | 0.61981 | 1.4546 | 25000 | 25859.6813 | 859.6813 | 3.4387
Cal. Std.9 | 1.2386 | 1.7585 | 50000 | 51704.8596 | 1704.8596 | 3.4097
Cal. Std.10 | 2.5503 | 1.8948 | 100000 | 106494.6693 | 6494.6693 | 6.4947
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99759 | 2.3942e-05 | 0.00068842 | 7.443e-07 | 0.00087612 | 1.9728 | 10.2508
Calibration Table - V KED 51/51
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 0.66667 | 173.2051 | 0 |  | 0 | 0
Cal. Std.1 | 0.00062795 | 36.2739 | 0.01 | 0.010117 | 0.00011728 | 1.1728
Cal. Std.2 | 0.0042491 | 16.6661 | 0.05 | 0.049247 | -0.00075294 | -1.5059
Cal. Std.3 | 0.008006 | 16.1039 | 0.1 | 0.089844 | -0.010156 | -10.1562
Cal. Std.4 | 0.048565 | 7.3371 | 0.5 | 0.52812 | 0.028121 | 5.6242
Cal. Std.5 | 0.096201 | 5.9336 | 1 | 1.0429 | 0.042871 | 4.2871
Cal. Std.6 | 0.433 | 2.9794 | 5 | 4.6822 | -0.31776 | -6.3552
Cal. Std.7 | 0.9177 | 0.58851 | 10 | 9.9199 | -0.080061 | -0.80061
Cal. Std.8 | 2.3818 | 2.1479 | 25 | 25.7403 | 0.74031 | 2.9613
Cal. Std.9 | 4.8476 | 1.9014 | 50 | 52.3862 | 2.3862 | 4.7724
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99843 | 0.092542 | -0.00030832 | 0.0019597 | 5.7361e-05 | 0.00051205 | 0.0026607
Calibration Table - Cr (52) 52/52
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 6.6667 | 45.8258 | 0 |  | 0 | 0
Cal. Std.1 | 0.0013032 | 15.4483 | 0.01 | 0.0097977 | -0.00020229 | -2.0229
Cal. Std.2 | 0.0074039 | 9.9329 | 0.05 | 0.05682 | 0.0068199 | 13.6398
Cal. Std.3 | 0.012041 | 9.2029 | 0.1 | 0.092559 | -0.0074409 | -7.4409
Cal. Std.4 | 0.06634 | 5.3402 | 0.5 | 0.51108 | 0.011079 | 2.2158
Cal. Std.5 | 0.13056 | 3.9473 | 1 | 1.0061 | 0.0060522 | 0.60522
Cal. Std.6 | 0.62285 | 1.6439 | 5 | 4.8005 | -0.19948 | -3.9895
Cal. Std.7 | 1.2538 | 1.3552 | 10 | 9.6636 | -0.33639 | -3.3639
Cal. Std.8 | 3.2344 | 0.53567 | 25 | 24.9296 | -0.070439 | -0.28176
Cal. Std.9 | 6.5284 | 3.9604 | 50 | 50.3191 | 0.31906 | 0.63811
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.9981 | 0.12974 | 3.2052e-05 | 0.0030274 | 8.8611e-05 | 0.0036614 | 0.0049333
Calibration Table - Cr (53) NH3 53/53
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 24.6667 | 51.4934 | 0 |  | 0 | 0
Cal. Std.1 | 0.00035141 | 20.0046 | 0.01 | 0.0099211 | -7.8903e-05 | -0.78903
Cal. Std.2 | 0.0019311 | 15.8724 | 0.05 | 0.052285 | 0.0022853 | 4.5706
Cal. Std.3 | 0.0036888 | 2.2031 | 0.1 | 0.09942 | -0.00058028 | -0.58028
Cal. Std.4 | 0.018279 | 2.879 | 0.5 | 0.49069 | -0.0093134 | -1.8627
Cal. Std.5 | 0.036357 | 2.4527 | 1 | 0.97549 | -0.024508 | -2.4508
Cal. Std.6 | 0.18079 | 0.19789 | 5 | 4.8487 | -0.15133 | -3.0267
Cal. Std.7 | 0.37046 | 1.682 | 10 | 9.9353 | -0.064651 | -0.64651
Cal. Std.8 | 0.95305 | 1.0854 | 25 | 25.5587 | 0.55869 | 2.2348
Cal. Std.9 | 1.912 | 1.5885 | 50 | 51.2753 | 1.2753 | 2.5506
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99964 | 0.037289 | -1.8543e-05 | 0.00037744 | 1.1048e-05 | 0.0023955 | 0.0037094
Calibration Table - Fe (54) NH3 54/54
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 819.0943 | 2.0297 | 0 |  | 0 | 0
Cal. Std.3 | 0.0092849 | 11.8985 | 0.1 | 0.09179 | -0.0082097 | -8.2097
Cal. Std.4 | 0.02406 | 3.268 | 0.5 | 0.632 | 0.132 | 26.3991
Cal. Std.5 | 0.043361 | 2.569 | 1 | 1.3377 | 0.33767 | 33.7675
Cal. Std.6 | 0.1248 | 0.61263 | 5 | 4.3151 | -0.68486 | -13.6973
Cal. Std.7 | 0.24985 | 1.9053 | 10 | 8.8872 | -1.1128 | -11.1275
Cal. Std.8 | 0.63196 | 1.0794 | 25 | 22.8583 | -2.1417 | -8.5666
Cal. Std.9 | 1.2539 | 1.2187 | 50 | 45.5966 | -4.4034 | -8.8067
Cal. Std.10 | 2.4749 | 2.0851 | 100 | 90.2413 | -9.7587 | -9.7587
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.98101 | 0.027351 | 0.0067744 | 0.0022077 | 0.00060923 | 0.10837 | 0.0062739
Calibration Table - Mn Oshift 55/71
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 13.3333 | 22.9129 | 0 |  | 0 | 0
Cal. Std.1 | 0.0023566 | 5.8706 | 0.01 | 0.010906 | 0.00090608 | 9.0608
Cal. Std.2 | 0.0034571 | 5.8201 | 0.05 | 0.028375 | -0.021625 | -43.2508
Cal. Std.3 | 0.0076336 | 14.1456 | 0.1 | 0.094666 | -0.0053335 | -5.3335
Cal. Std.4 | 0.033858 | 2.919 | 0.5 | 0.51092 | 0.010917 | 2.1833
Cal. Std.5 | 0.068312 | 1.0529 | 1 | 1.0578 | 0.057795 | 5.7795
Cal. Std.6 | 0.33281 | 1.0485 | 5 | 5.2561 | 0.25607 | 5.1214
Cal. Std.7 | 0.66096 | 1.7171 | 10 | 10.4646 | 0.46464 | 4.6464
Cal. Std.8 | 1.734 | 1.2902 | 25 | 27.4965 | 2.4965 | 9.9859
Cal. Std.9 | 3.5237 | 3.8073 | 50 | 55.9035 | 5.9035 | 11.8071
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.98458 | 0.063001 | 0.0016695 | 0.0042305 | 0.00012383 | 0.00083164 | 0.00060419
Calibration Table - Fe (56) KED 56/56
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 1537.4161 | 3.0493 | 0 |  | 0 | 0
Cal. Std.3 | 0.03176 | 20.5582 | 0.1 | 0.087268 | -0.012732 | -12.7321
Cal. Std.4 | 0.11641 | 10.7052 | 0.5 | 0.73154 | 0.23154 | 46.3075
Cal. Std.5 | 0.20512 | 2.3589 | 1 | 1.4067 | 0.40674 | 40.6737
Cal. Std.6 | 0.56583 | 2.1682 | 5 | 4.1521 | -0.84788 | -16.9577
Cal. Std.7 | 1.1209 | 1.3729 | 10 | 8.3763 | -1.6237 | -16.2365
Cal. Std.8 | 2.8653 | 0.85558 | 25 | 21.6535 | -3.3465 | -13.3859
Cal. Std.9 | 5.6697 | 3.9498 | 50 | 42.9976 | -7.0024 | -14.0047
Cal. Std.10 | 11.3638 | 1.0224 | 100 | 86.3357 | -13.6643 | -13.6643
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.962 | 0.13139 | 0.020294 | 0.015225 | 0.0042013 | 0.83653 | 0.089134
Calibration Table - Co KED 59/59
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 3.3333 | 91.6515 | 0 |  | 0 | 0
Cal. Std.1 | 0.0020768 | 26.5287 | 0.01 | 0.010156 | 0.00015627 | 1.5627
Cal. Std.2 | 0.011407 | 8.3232 | 0.05 | 0.045229 | -0.0047708 | -9.5416
Cal. Std.3 | 0.026673 | 11.3532 | 0.1 | 0.10261 | 0.0026108 | 2.6108
Cal. Std.4 | 0.13826 | 5.4947 | 0.5 | 0.52205 | 0.022047 | 4.4093
Cal. Std.5 | 0.2658 | 1.6993 | 1 | 1.0015 | 0.0014547 | 0.14547
Cal. Std.6 | 1.2815 | 2.1325 | 5 | 4.8195 | -0.18047 | -3.6093
Cal. Std.7 | 2.6884 | 1.9613 | 10 | 10.1078 | 0.10778 | 1.0778
Cal. Std.8 | 6.7843 | 2.498 | 25 | 25.5041 | 0.50413 | 2.0165
Cal. Std.9 | 13.4778 | 3.9004 | 50 | 50.6641 | 0.66413 | 1.3283
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99904 | 0.26603 | -0.0006251 | 0.0043979 | 0.00012873 | 0.00089789 | 0.0024562
Calibration Table - Ni KED 60/60
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 10.6667 | 60.2728 | 0 |  | 0 | 0
Cal. Std.1 | 0.0006839 | 76.8965 | 0.01 | 0.010002 | 2.026e-06 | 0.02026
Cal. Std.2 | 0.0037178 | 37.2687 | 0.05 | 0.051672 | 0.001672 | 3.344
Cal. Std.3 | 0.0066861 | 6.1109 | 0.1 | 0.092441 | -0.0075594 | -7.5594
Cal. Std.4 | 0.036867 | 7.0764 | 0.5 | 0.50696 | 0.006962 | 1.3924
Cal. Std.5 | 0.075699 | 2.0934 | 1 | 1.0403 | 0.040305 | 4.0305
Cal. Std.6 | 0.36201 | 3.0924 | 5 | 4.9727 | -0.027286 | -0.54573
Cal. Std.7 | 0.72622 | 0.9786 | 10 | 9.975 | -0.025031 | -0.25031
Cal. Std.8 | 1.8337 | 2.5396 | 25 | 25.1854 | 0.18542 | 0.74169
Cal. Std.9 | 3.5977 | 4.6379 | 50 | 49.4133 | -0.58674 | -1.1735
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99939 | 0.072808 | -4.433e-05 | 0.00095932 | 2.8079e-05 | 0.010525 | 0.019461
Calibration Table - Cu (63) KED 63/63
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 12 | 16.6667 | 0 |  | 0 | 0
Cal. Std.1 | 0.0033747 | 28.1595 | 0.01 | 0.0099265 | -7.3465e-05 | -0.73465
Cal. Std.2 | 0.010775 | 2.8643 | 0.05 | 0.051378 | 0.0013781 | 2.7562
Cal. Std.3 | 0.019441 | 3.2334 | 0.1 | 0.099923 | -7.6962e-05 | -0.076962
Cal. Std.4 | 0.09753 | 8.0475 | 0.5 | 0.53735 | 0.037351 | 7.4703
Cal. Std.5 | 0.18989 | 1.8049 | 1 | 1.0547 | 0.054708 | 5.4708
Cal. Std.6 | 0.86246 | 1.8156 | 5 | 4.8222 | -0.17778 | -3.5556
Cal. Std.7 | 1.7161 | 1.4989 | 10 | 9.604 | -0.39602 | -3.9602
Cal. Std.8 | 4.3625 | 1.3546 | 25 | 24.4282 | -0.57182 | -2.2873
Cal. Std.9 | 8.4739 | 3.0116 | 50 | 47.4587 | -2.5413 | -5.0826
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99895 | 0.17852 | 0.0016026 | 0.0030902 | 9.0451e-05 | 0.0048107 | 0.0025632
Calibration Table - Cu (65) NH3 65/65
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 74.0002 | 28.0874 | 0 |  | 0 | 0
Cal. Std.1 | 0.0014182 | 9.6613 | 0.01 | 0.0098239 | -0.00017606 | -1.7606
Cal. Std.2 | 0.0046606 | 5.398 | 0.05 | 0.053643 | 0.0036426 | 7.2851
Cal. Std.3 | 0.0082308 | 7.1962 | 0.1 | 0.10189 | 0.0018925 | 1.8925
Cal. Std.4 | 0.039555 | 2.1735 | 0.5 | 0.52521 | 0.025213 | 5.0427
Cal. Std.5 | 0.076643 | 1.5393 | 1 | 1.0264 | 0.02643 | 2.643
Cal. Std.6 | 0.35814 | 1.0384 | 5 | 4.8307 | -0.16933 | -3.3866
Cal. Std.7 | 0.71214 | 1.0821 | 10 | 9.6147 | -0.38528 | -3.8528
Cal. Std.8 | 1.7822 | 0.55445 | 25 | 24.0762 | -0.92381 | -3.6952
Cal. Std.9 | 3.5463 | 2.0684 | 50 | 47.916 | -2.084 | -4.168
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99898 | 0.073996 | 0.00069125 | 0.0012651 | 3.7029e-05 | 0.0036213 | 0.0030974
Calibration Table - Zn NH3 66/66
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 358.6712 | 9.0317 | 0 |  | 0 | 0
Cal. Std.1 | 0.0042953 | 2.4513 | 0.01 | 0.011821 | 0.001821 | 18.2103
Cal. Std.2 | 0.0035461 | 4.9587 | 0.05 | -0.010984 | -0.060984 | -121.9687
Cal. Std.3 | 0.009071 | 2.1057 | 0.1 | 0.1572 | 0.057205 | 57.2049
Cal. Std.4 | 0.023211 | 2.3319 | 0.5 | 0.58767 | 0.087672 | 17.5343
Cal. Std.5 | 0.039931 | 0.82743 | 1 | 1.0967 | 0.096662 | 9.6662
Cal. Std.6 | 0.17535 | 0.51537 | 5 | 5.219 | 0.21904 | 4.3808
Cal. Std.7 | 0.3439 | 1.8823 | 10 | 10.3502 | 0.35019 | 3.5019
Cal. Std.8 | 0.8723 | 0.99115 | 25 | 26.436 | 1.436 | 5.7438
Cal. Std.9 | 1.7294 | 2.1304 | 50 | 52.5284 | 2.5284 | 5.0567
Cal. Std.10 | 3.3108 | 0.89714 | 100 | 100.6699 | 0.66987 | 0.66987
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.90226 | 0.032849 | 0.003907 | 0.0055503 | 0.00017124 | 0.039506 | 0.0105
Calibration Table - As 75/75
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | -2829.9787 | 10.311 | 0 |  | 0 | 0
Cal. Std.1 | 0.00058488 | 170.9038 | 0.01 | 0.0098238 | -0.00017623 | -1.7623
Cal. Std.2 | 0.0024827 | 19.3554 | 0.05 | 0.061122 | 0.011122 | 22.2441
Cal. Std.3 | 0.002988 | 22.6936 | 0.1 | 0.07478 | -0.02522 | -25.2197
Cal. Std.4 | 0.017124 | 7.9402 | 0.5 | 0.45688 | -0.043119 | -8.6237
Cal. Std.5 | 0.037036 | 6.2023 | 1 | 0.99509 | -0.0049141 | -0.49141
Cal. Std.6 | 0.19493 | 4.1963 | 5 | 5.2629 | 0.2629 | 5.258
Cal. Std.7 | 0.36865 | 0.41807 | 10 | 9.9586 | -0.041425 | -0.41425
Cal. Std.8 | 0.96704 | 0.33957 | 25 | 26.1327 | 1.1327 | 4.5306
Cal. Std.9 | 1.9329 | 1.5382 | 50 | 52.2393 | 2.2393 | 4.4786
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99138 | 0.036997 | 0.00022143 | 0.0018485 | 5.4107e-05 | -0.42454 | 0.12281
Calibration Table - As KED 75/75
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 0.66667 | 173.2051 | 0 |  | 0 | 0
Cal. Std.1 | 9.7081e-05 | 148.2525 | 0.01 | 0.0093443 | -0.00065571 | -6.5571
Cal. Std.2 | 0.00059586 | 39.8171 | 0.05 | 0.061334 | 0.011334 | 22.6675
Cal. Std.3 | 0.001174 | 28.8855 | 0.1 | 0.12159 | 0.021594 | 21.5939
Cal. Std.4 | 0.0046034 | 9.1133 | 0.5 | 0.47906 | -0.020945 | -4.189
Cal. Std.5 | 0.0094293 | 18.6919 | 1 | 0.98208 | -0.017918 | -1.7918
Cal. Std.6 | 0.041825 | 4.2926 | 5 | 4.3589 | -0.64115 | -12.823
Cal. Std.7 | 0.092094 | 3.9902 | 10 | 9.5986 | -0.40145 | -4.0145
Cal. Std.8 | 0.22258 | 1.8671 | 25 | 23.1992 | -1.8008 | -7.2031
Cal. Std.9 | 0.44284 | 2.855 | 50 | 46.1585 | -3.8415 | -7.683
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99099 | 0.0095938 | 7.4339e-06 | 0.00049019 | 1.4348e-05 | 0.0049392 | 0.025665
Calibration Table - As RH Oshift 75/91
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 27.3334 | 27.702 | 0 |  | 0 | 0
Cal. Std.1 | 0.020421 | 22.5679 | 0.01 | 0.16912 | 0.15912 | 1591.2355
Cal. Std.2 | 0.085586 | 3.5702 | 0.05 | 0.20366 | 0.15366 | 307.3285
Cal. Std.3 | 0.18132 | 13.0577 | 0.1 | 0.25441 | 0.15441 | 154.4083
Cal. Std.4 | 0.86037 | 2.639 | 0.5 | 0.61434 | 0.11434 | 22.8679
Cal. Std.5 | 1.7672 | 5.1834 | 1 | 1.095 | 0.094999 | 9.4999
Cal. Std.6 | 8.8731 | 0.38783 | 5 | 4.8615 | -0.13847 | -2.7695
Cal. Std.7 | 17.2916 | 1.2738 | 10 | 9.3238 | -0.67622 | -6.7622
Cal. Std.8 | 46.8386 | 5.416 | 25 | 24.9852 | -0.014784 | -0.059135
Cal. Std.9 | 94.3201 | 4.0329 | 50 | 50.1529 | 0.15295 | 0.30589
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99987 | 1.8866 | -0.29865 | 0.011424 | 0.21714 | 0.0028507 | 0.0023005
Calibration Table - As Oshift 75/91
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 27.3334 | 27.702 | 0 |  | 0 | 0
Cal. Std.1 | 0.00040199 | 20.6038 | 0.01 | 0.010071 | 7.086e-05 | 0.7086
Cal. Std.2 | 0.0016891 | 5.7083 | 0.05 | 0.047607 | -0.0023926 | -4.7852
Cal. Std.3 | 0.0035802 | 11.2975 | 0.1 | 0.10276 | 0.002761 | 2.761
Cal. Std.4 | 0.017107 | 4.8972 | 0.5 | 0.49726 | -0.0027434 | -0.54868
Cal. Std.5 | 0.033966 | 2.6265 | 1 | 0.98893 | -0.011073 | -1.1073
Cal. Std.6 | 0.16741 | 0.057062 | 5 | 4.8808 | -0.11917 | -2.3834
Cal. Std.7 | 0.33206 | 1.9313 | 10 | 9.6827 | -0.31729 | -3.1729
Cal. Std.8 | 0.8836 | 0.36 | 25 | 25.7677 | 0.76769 | 3.0708
Cal. Std.9 | 1.808 | 3.3751 | 50 | 52.7285 | 2.7285 | 5.4571
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.9994 | 0.034289 | 5.6671e-05 | 0.00044849 | 1.3127e-05 | 0.0031337 | 0.0027428
Calibration Table - Se (77) 77/77
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 1025.3702 | 6.9306 | 0 |  | 0 | 0
Cal. Std.2 | 0.00027647 | 96.6098 | 0.05 | 0.044258 | -0.0057421 | -11.4842
Cal. Std.3 | 0.00058516 | 48.0334 | 0.1 | 0.12234 | 0.022338 | 22.3379
Cal. Std.4 | 0.0021941 | 12.3176 | 0.5 | 0.52929 | 0.029292 | 5.8585
Cal. Std.5 | 0.0038487 | 16.1032 | 1 | 0.94781 | -0.052195 | -5.2195
Cal. Std.6 | 0.020449 | 5.3896 | 5 | 5.1466 | 0.14658 | 2.9317
Cal. Std.7 | 0.037662 | 1.4593 | 10 | 9.5002 | -0.49976 | -4.9976
Cal. Std.8 | 0.094454 | 1.4519 | 25 | 23.8652 | -1.1348 | -4.5394
Cal. Std.9 | 0.18812 | 2.3681 | 50 | 47.5563 | -2.4437 | -4.8874
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99342 | 0.0039536 | 0.00010149 | 0.00018609 | 2.342e-05 | 1.4405 | 0.33098
Calibration Table - Se (77) Oshift 77/93
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 4.6667 | 24.7436 | 0 |  | 0 | 0
Cal. Std.2 | 3.2578e-05 | 36.6045 | 0.05 | 0.046236 | -0.0037643 | -7.5287
Cal. Std.3 | 8.5637e-05 | 73.9659 | 0.1 | 0.11725 | 0.017253 | 17.2531
Cal. Std.4 | 0.00033729 | 17.8242 | 0.5 | 0.45408 | -0.04592 | -9.184
Cal. Std.5 | 0.00072061 | 8.9271 | 1 | 0.96714 | -0.032856 | -3.2856
Cal. Std.6 | 0.0036954 | 2.2273 | 5 | 4.9489 | -0.051129 | -1.0226
Cal. Std.7 | 0.0072456 | 1.8499 | 10 | 9.7007 | -0.2993 | -2.993
Cal. Std.8 | 0.01928 | 2.3217 | 25 | 25.8087 | 0.8087 | 3.2348
Cal. Std.9 | 0.038671 | 5.0306 | 50 | 51.763 | 1.763 | 3.526
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99586 | 0.00074712 | -1.9656e-06 | 2.7825e-05 | 3.502e-06 | 0.024545 | 0.019275
Calibration Table - Se (78) Oshift 78/94
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 1.3333 | 86.6025 | 0 |  | 0 | 0
Cal. Std.2 | 0.00011248 | 20.5263 | 0.05 | 0.04831 | -0.0016901 | -3.3802
Cal. Std.3 | 0.00024871 | 32.5112 | 0.1 | 0.10607 | 0.0060659 | 6.0659
Cal. Std.4 | 0.0012323 | 9.1073 | 0.5 | 0.52311 | 0.023107 | 4.6214
Cal. Std.5 | 0.0023321 | 6.4277 | 1 | 0.98939 | -0.010606 | -1.0606
Cal. Std.6 | 0.011471 | 0.83625 | 5 | 4.8641 | -0.13585 | -2.717
Cal. Std.7 | 0.021641 | 5.5987 | 10 | 9.1762 | -0.82382 | -8.2382
Cal. Std.8 | 0.059989 | 1.1825 | 25 | 25.435 | 0.43504 | 1.7402
Cal. Std.9 | 0.12143 | 4.3432 | 50 | 51.4843 | 1.4843 | 2.9686
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99864 | 0.0023586 | -1.4598e-06 | 5.0239e-05 | 6.3229e-06 | 0.0022 | 0.0057162
Calibration Table - Se (80) Oshift 80/96
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 2.6667 | 114.5644 | 0 |  | 0 | 0
Cal. Std.2 | 0.00025487 | 31.8466 | 0.05 | 0.050293 | 0.0002929 | 0.5858
Cal. Std.3 | 0.00049354 | 19.191 | 0.1 | 0.098736 | -0.0012639 | -1.2639
Cal. Std.4 | 0.0025285 | 7.7738 | 0.5 | 0.51177 | 0.01177 | 2.354
Cal. Std.5 | 0.0047754 | 7.8561 | 1 | 0.96784 | -0.032164 | -3.2164
Cal. Std.6 | 0.024221 | 2.1242 | 5 | 4.9147 | -0.085297 | -1.7059
Cal. Std.7 | 0.047121 | 4.1422 | 10 | 9.5627 | -0.43729 | -4.3729
Cal. Std.8 | 0.12674 | 0.82044 | 25 | 25.7233 | 0.72325 | 2.893
Cal. Std.9 | 0.25799 | 4.1511 | 50 | 52.3632 | 2.3632 | 4.7263
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99939 | 0.0049268 | 7.0888e-06 | 7.0179e-05 | 8.8325e-06 | 0.0021387 | 0.0074192
Calibration Table - Se (82) 82/82
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 37.4705 | 24.7288 | 0 |  | 0 | 0
Cal. Std.2 | 0.0003067 | 39.8882 | 0.05 | 0.064389 | 0.014389 | 28.7787
Cal. Std.3 | 0.00022757 | 77.0352 | 0.1 | 0.041378 | -0.058622 | -58.6222
Cal. Std.4 | 0.0018241 | 8.3996 | 0.5 | 0.50568 | 0.0056802 | 1.136
Cal. Std.5 | 0.0037315 | 8.2727 | 1 | 1.0604 | 0.060392 | 6.0392
Cal. Std.6 | 0.019051 | 7.2303 | 5 | 5.5156 | 0.51565 | 10.313
Cal. Std.7 | 0.034112 | 3.014 | 10 | 9.8956 | -0.10437 | -1.0437
Cal. Std.8 | 0.09081 | 1.5593 | 25 | 26.3849 | 1.3849 | 5.5395
Cal. Std.9 | 0.18552 | 2.9073 | 50 | 53.9298 | 3.9298 | 7.8596
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.96329 | 0.0034385 | 8.5294e-05 | 0.00039124 | 4.924e-05 | 0.060553 | 0.045543
Calibration Table - Se (82) Oshift 82/98
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 0 |  | 0 |  | 0 | 0
Cal. Std.2 | 4.2864e-05 | 10.4112 | 0.05 | 0.046867 | -0.003133 | -6.2661
Cal. Std.3 | 0.00010147 | 7.6349 | 0.1 | 0.11289 | 0.012887 | 12.8872
Cal. Std.4 | 0.00046051 | 5.7893 | 0.5 | 0.51732 | 0.017317 | 3.4634
Cal. Std.5 | 0.00080189 | 8.9508 | 1 | 0.90186 | -0.098137 | -9.8137
Cal. Std.6 | 0.0044141 | 6.1317 | 5 | 4.9708 | -0.029192 | -0.58385
Cal. Std.7 | 0.0082357 | 3.0423 | 10 | 9.2756 | -0.72442 | -7.2442
Cal. Std.8 | 0.022492 | 0.78884 | 25 | 25.3348 | 0.33479 | 1.3392
Cal. Std.9 | 0.047149 | 3.5409 | 50 | 53.109 | 3.109 | 6.2181
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99651 | 0.00088775 | 1.2575e-06 | 3.0366e-05 | 3.8218e-06 | 0 | 0
Calibration Table - Mo (95) NH3 95/95
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 6 | 33.3333 | 0 |  | 0 | 0
Cal. Std.1 | 0.0014894 | 4.0855 | 0.01 | 0.010528 | 0.00052771 | 5.2771
Cal. Std.2 | 0.0032758 | 1.2829 | 0.05 | 0.038235 | -0.011765 | -23.5307
Cal. Std.3 | 0.0068791 | 4.3074 | 0.1 | 0.094119 | -0.0058811 | -5.8811
Cal. Std.4 | 0.032769 | 2.4468 | 0.5 | 0.49566 | -0.0043447 | -0.86893
Cal. Std.5 | 0.066797 | 1.7821 | 1 | 1.0234 | 0.023405 | 2.3405
Cal. Std.6 | 0.32728 | 1.1004 | 5 | 5.0634 | 0.06336 | 1.2672
Cal. Std.7 | 0.66859 | 2.3554 | 10 | 10.3569 | 0.35687 | 3.5687
Cal. Std.8 | 1.7316 | 1.5672 | 25 | 26.843 | 1.843 | 7.372
Cal. Std.9 | 3.5617 | 1.5568 | 50 | 55.2276 | 5.2276 | 10.4553
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99457 | 0.064477 | 0.00081057 | 0.0025505 | 7.4654e-05 | 0.0003367 | 0.00033556
Calibration Table - Mo (95) OShift 95/127
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 11.3333 | 36.7353 | 0 |  | 0 | 0
Cal. Std.1 | 0.0015006 | 4.2637 | 0.01 | 0.010484 | 0.00048401 | 4.8401
Cal. Std.2 | 0.0033439 | 4.4091 | 0.05 | 0.039375 | -0.010625 | -21.2505
Cal. Std.3 | 0.0068072 | 8.5681 | 0.1 | 0.093657 | -0.0063434 | -6.3434
Cal. Std.4 | 0.032927 | 4.8223 | 0.5 | 0.50304 | 0.0030414 | 0.60827
Cal. Std.5 | 0.066159 | 1.2696 | 1 | 1.0239 | 0.023894 | 2.3894
Cal. Std.6 | 0.32498 | 1.0535 | 5 | 5.0805 | 0.080468 | 1.6094
Cal. Std.7 | 0.6391 | 1.4036 | 10 | 10.0038 | 0.0038266 | 0.038266
Cal. Std.8 | 1.7049 | 0.63434 | 25 | 26.7079 | 1.7079 | 6.8315
Cal. Std.9 | 3.5507 | 2.6572 | 50 | 55.6385 | 5.6385 | 11.2769
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99526 | 0.063802 | 0.00083166 | 0.0023565 | 6.8975e-05 | 0.00069482 | 0.00074075
Calibration Table - Cd (111) NH3 111/111
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 1.3333 | 86.6025 | 0 |  | 0 | 0
Cal. Std.1 | 0.00032219 | 13.8733 | 0.01 | 0.0098595 | -0.00014052 | -1.4052
Cal. Std.2 | 0.0017438 | 10.4265 | 0.05 | 0.052823 | 0.0028231 | 5.6462
Cal. Std.3 | 0.0034071 | 2.2732 | 0.1 | 0.10309 | 0.0030938 | 3.0938
Cal. Std.4 | 0.016462 | 0.73644 | 0.5 | 0.49764 | -0.0023554 | -0.47109
Cal. Std.5 | 0.032535 | 1.4032 | 1 | 0.98343 | -0.01657 | -1.657
Cal. Std.6 | 0.1608 | 0.23223 | 5 | 4.8599 | -0.1401 | -2.8021
Cal. Std.7 | 0.32471 | 2.3207 | 10 | 9.8137 | -0.18629 | -1.8629
Cal. Std.8 | 0.83243 | 0.46667 | 25 | 25.1583 | 0.1583 | 0.63319
Cal. Std.9 | 1.6349 | 1.1205 | 50 | 49.4126 | -0.58745 | -1.1749
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99959 | 0.033088 | -4.0346e-06 | 0.00035686 | 1.0445e-05 | 0.00014594 | 0.00037918
Calibration Table - Cd (114) NH3 114/114
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 8 | 25 | 0 |  | 0 | 0
Cal. Std.1 | 0.00083445 | 6.6896 | 0.01 | 0.0099355 | -6.4458e-05 | -0.64458
Cal. Std.2 | 0.0040819 | 3.6776 | 0.05 | 0.051858 | 0.0018585 | 3.717
Cal. Std.3 | 0.0077321 | 2.5747 | 0.1 | 0.098982 | -0.0010184 | -1.0184
Cal. Std.4 | 0.038877 | 0.20648 | 0.5 | 0.50105 | 0.0010483 | 0.20966
Cal. Std.5 | 0.077707 | 1.626 | 1 | 1.0023 | 0.002333 | 0.2333
Cal. Std.6 | 0.38178 | 1.502 | 5 | 4.9278 | -0.072195 | -1.4439
Cal. Std.7 | 0.76959 | 2.548 | 10 | 9.9343 | -0.065715 | -0.65715
Cal. Std.8 | 1.9485 | 1.3201 | 25 | 25.1537 | 0.15371 | 0.61485
Cal. Std.9 | 3.834 | 1.662 | 50 | 49.4946 | -0.50538 | -1.0108
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99987 | 0.077461 | 6.4835e-05 | 0.00047562 | 1.3921e-05 | 0.00037369 | 0.0002792
Calibration Table - Pb 208/208
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 35.3334 | 26.1442 | 0 |  | 0 | 0
Cal. Std.1 | 0.0020678 | 7.2115 | 0.01 | 0.0099215 | -7.8462e-05 | -0.78462
Cal. Std.2 | 0.0091888 | 1.2188 | 0.05 | 0.050434 | 0.00043419 | 0.86839
Cal. Std.3 | 0.018836 | 5.5314 | 0.1 | 0.10532 | 0.0053196 | 5.3196
Cal. Std.4 | 0.090908 | 0.59319 | 0.5 | 0.51535 | 0.015352 | 3.0704
Cal. Std.5 | 0.17807 | 0.91962 | 1 | 1.0112 | 0.011222 | 1.1222
Cal. Std.6 | 0.93466 | 3.571 | 5 | 5.3156 | 0.31563 | 6.3126
Cal. Std.7 | 1.7108 | 0.4661 | 10 | 9.7315 | -0.26853 | -2.6853
Cal. Std.8 | 4.1861 | 0.80748 | 25 | 23.8138 | -1.1862 | -4.7449
Cal. Std.9 | 8.0438 | 1.0683 | 50 | 45.7608 | -4.2392 | -8.4784
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99876 | 0.17577 | 0.0003239 | 0.0033139 | 9.6998e-05 | 0.0011151 | 0.0008541
Calibration Table - Pb NH3 208/208
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 69.3334 | 13.6322 | 0 |  | 0 | 0
Cal. Std.1 | 0.0035424 | 3.9883 | 0.01 | 0.0099589 | -4.1138e-05 | -0.41138
Cal. Std.2 | 0.015696 | 2.0123 | 0.05 | 0.0496 | -0.00040031 | -0.80062
Cal. Std.3 | 0.032396 | 2.95 | 0.1 | 0.10407 | 0.0040669 | 4.0669
Cal. Std.4 | 0.16135 | 0.91935 | 0.5 | 0.52464 | 0.024639 | 4.9278
Cal. Std.5 | 0.32903 | 2.3749 | 1 | 1.0715 | 0.07154 | 7.154
Cal. Std.6 | 1.5291 | 1.611 | 5 | 4.9856 | -0.014398 | -0.28797
Cal. Std.7 | 3.0813 | 0.97034 | 10 | 10.0482 | 0.048184 | 0.48184
Cal. Std.8 | 6.8361 | 1.9216 | 25 | 22.2944 | -2.7056 | -10.8222
Cal. Std.9 | 14.6703 | 0.71298 | 50 | 47.8458 | -2.1542 | -4.3083
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99844 | 0.30661 | 0.00048895 | 0.0064848 | 0.00018981 | 0.0010272 | 0.00043261

## Calibration Curves 227
Calibration Table - Na KED 23/23
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 0.66667 | 173.2051 | 0 |  | 0 | 0
Cal. Std.4 | 0.057662 | 2.6106 | 500 | 495.3934 | -4.6066 | -0.92132
Cal. Std.5 | 0.12386 | 3.1256 | 1000 | 1033.9805 | 33.9805 | 3.398
Cal. Std.6 | 0.57098 | 2.4965 | 5000 | 4671.4489 | -328.5511 | -6.571
Cal. Std.7 | 1.1984 | 3.1012 | 10000 | 9775.9075 | -224.0925 | -2.2409
Cal. Std.8 | 2.9794 | 4.0203 | 25000 | 24264.5626 | -735.4374 | -2.9417
Cal. Std.9 | 6.1906 | 1.9246 | 50000 | 50389.3228 | 389.3228 | 0.77865
Cal. Std.10 | 13.3334 | 9.9177 | 100000 | 108498.323 | 8498.323 | 8.4983
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99851 | 0.00012292 | -0.0032318 | 3.0058e-06 | 0.0035381 | 0.4337 | 2.2536
Calibration Table - Mg KED 24/24
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 0.66667 | 173.2051 | 0 |  | 0 | 0
Cal. Std.4 | 0.023293 | 9.0126 | 500 | 497.1769 | -2.8231 | -0.56462
Cal. Std.5 | 0.048444 | 2.8206 | 1000 | 1019.6441 | 19.6441 | 1.9644
Cal. Std.6 | 0.23104 | 1.471 | 5000 | 4812.6647 | -187.3353 | -3.7467
Cal. Std.7 | 0.47789 | 2.5177 | 10000 | 9940.3967 | -59.6033 | -0.59603
Cal. Std.8 | 1.1834 | 3.075 | 25000 | 24596.0447 | -403.9553 | -1.6158
Cal. Std.9 | 2.389 | 2.7671 | 50000 | 49640.771 | -359.229 | -0.71846
Cal. Std.10 | 5.0674 | 9.2945 | 100000 | 105277.2281 | 5277.2281 | 5.2772
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99948 | 4.814e-05 | -0.0006407 | 6.9516e-07 | 0.00081827 | 1.1078 | 5.7561
Calibration Table - Al 27/27
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 12476.8066 | 8.6575 | 0 |  | 0 | 0
Cal. Std.4 | 0.39798 | 6.5659 | 0.5 | 0.50416 | 0.004158 | 0.8316
Cal. Std.5 | 0.60264 | 5.231 | 1 | 1.0309 | 0.030873 | 3.0873
Cal. Std.6 | 1.8251 | 1.8063 | 5 | 4.177 | -0.82301 | -16.4602
Cal. Std.7 | 3.3977 | 1.8172 | 10 | 8.2242 | -1.7758 | -17.7584
Cal. Std.8 | 9.5805 | 7.9106 | 25 | 24.1363 | -0.86375 | -3.455
Cal. Std.9 | 21.924 | 0.96919 | 50 | 55.9032 | 5.9032 | 11.8065
Cal. Std.10 | 47.5866 | 15.5342 | 100 | 121.9483 | 21.9483 | 21.9483
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.98731 | 0.38856 | 0.20208 | 0.027952 | 0.032902 | 0.18254 | 0.048092
Calibration Table - Al KED 27/27
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 8.6667 | 26.6469 | 0 |  | 0 | 0
Cal. Std.4 | 0.0034904 | 27.9474 | 0.5 | 0.47438 | -0.025617 | -5.1234
Cal. Std.5 | 0.0055044 | 40.8791 | 1 | 1.0999 | 0.099922 | 9.9922
Cal. Std.6 | 0.01869 | 7.3149 | 5 | 5.1956 | 0.19555 | 3.911
Cal. Std.7 | 0.033295 | 9.8593 | 10 | 9.7316 | -0.26836 | -2.6836
Cal. Std.8 | 0.077814 | 10.767 | 25 | 23.5596 | -1.4404 | -5.7618
Cal. Std.9 | 0.15883 | 3.7514 | 50 | 48.7241 | -1.2759 | -2.5519
Cal. Std.10 | 0.33106 | 7.3902 | 100 | 102.2174 | 2.2174 | 2.2174
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99797 | 0.0032196 | 0.0019631 | 9.1811e-05 | 0.00010807 | 0.21647 | 0.17519
Calibration Table - P Oshift 31/47
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 547.3438 | 4.3595 | 0 |  | 0 | 0
Cal. Std.3 | 0.016815 | 2.2529 | 0.1 | 0.078811 | -0.021189 | -21.1893
Cal. Std.4 | 0.036425 | 2.2355 | 0.5 | 1.0945 | 0.59451 | 118.9027
Cal. Std.5 | 0.030595 | 1.8605 | 1 | 0.79253 | -0.20747 | -20.7474
Cal. Std.6 | 0.10213 | 0.81082 | 5 | 4.4976 | -0.5024 | -10.048
Cal. Std.7 | 0.1689 | 2.0847 | 10 | 7.9561 | -2.0439 | -20.4387
Cal. Std.8 | 0.41772 | 0.58178 | 25 | 20.8438 | -4.1562 | -16.6247
Cal. Std.9 | 0.83759 | 2.4206 | 50 | 42.5906 | -7.4094 | -14.8189
Cal. Std.10 | 1.6557 | 0.66143 | 100 | 84.9644 | -15.0356 | -15.0356
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.89153 | 0.019307 | 0.015294 | 0.0040046 | 0.0011051 | 0.1647 | 0.026817
Calibration Table - S (34) -1 34/34
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 706.0176 | 11.5388 | 0 |  | 0 | 0
Cal. Std.5 | 0.016997 | 0.80356 | 1000 | 1020.9912 | 20.9912 | 2.0991
Cal. Std.6 | 0.090226 | 6.0783 | 5000 | 4530.9264 | -469.0736 | -9.3815
Cal. Std.7 | 0.19714 | 9.4796 | 10000 | 9655.611 | -344.389 | -3.4439
Cal. Std.8 | 0.51753 | 5.7665 | 25000 | 25012.154 | 12.154 | 0.048616
Cal. Std.9 | 1.0522 | 1.3455 | 50000 | 50641.9098 | 641.9098 | 1.2838
Cal. Std.10 | 2.278 | 8.8238 | 100000 | 109393.8023 | 9393.8023 | 9.3938
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99771 | 2.0863e-05 | -0.0043037 | 7.0793e-07 | 0.0016906 | 2720.7428 | 997.153
Calibration Table - K KED 39/39
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 10 | 40 | 0 |  | 0 | 0
Cal. Std.4 | 0.028033 | 6.9895 | 500 | 519.1502 | 19.1502 | 3.83
Cal. Std.5 | 0.052074 | 6.086 | 1000 | 938.557 | -61.443 | -6.1443
Cal. Std.6 | 0.26405 | 3.1069 | 5000 | 4636.5605 | -363.4395 | -7.2688
Cal. Std.7 | 0.56426 | 4.1281 | 10000 | 9873.7988 | -126.2012 | -1.262
Cal. Std.8 | 1.4032 | 5.5811 | 25000 | 24509.5362 | -490.4638 | -1.9619
Cal. Std.9 | 2.9061 | 3.8215 | 50000 | 50728.439 | 728.439 | 1.4569
Cal. Std.10 | 6.381 | 10.1413 | 100000 | 111350.031 | 11350.031 | 11.35
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99747 | 5.7322e-05 | -0.0017252 | 1.828e-06 | 0.0021518 | 13.9927 | 16.5686
Calibration Table - Ca KED 43/43
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 0 |  | 0 |  | 0 | 0
Cal. Std.4 | 0.0109 | 6.2022 | 500 | 508.0748 | 8.0748 | 1.615
Cal. Std.5 | 0.022452 | 13.4255 | 1000 | 976.7367 | -23.2633 | -2.3263
Cal. Std.6 | 0.11633 | 2.4354 | 5000 | 4785.2026 | -214.7974 | -4.2959
Cal. Std.7 | 0.24425 | 2.6867 | 10000 | 9974.9268 | -25.0732 | -0.25073
Cal. Std.8 | 0.59907 | 5.7864 | 25000 | 24369.9104 | -630.0896 | -2.5204
Cal. Std.9 | 1.2354 | 1.8765 | 50000 | 50184.4092 | 184.4092 | 0.36882
Cal. Std.10 | 2.6459 | 7.9534 | 100000 | 107409.5804 | 7409.5804 | 7.4096
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99908 | 2.4649e-05 | -0.0016238 | 4.7418e-07 | 0.00055815 | 0 | 0
Calibration Table - V KED 51/51
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 0 |  | 0 |  | 0 | 0
Cal. Std.1 | 0.00074076 | 87.2195 | 0.01 | 0.0099299 | -7.0073e-05 | -0.70073
Cal. Std.2 | 0.0047 | 8.5898 | 0.05 | 0.055829 | 0.0058286 | 11.6572
Cal. Std.3 | 0.0071925 | 14.6783 | 0.1 | 0.084724 | -0.015276 | -15.2762
Cal. Std.4 | 0.040851 | 10.3629 | 0.5 | 0.47493 | -0.025074 | -5.0148
Cal. Std.5 | 0.086064 | 5.8893 | 1 | 0.99907 | -0.000932 | -0.0932
Cal. Std.6 | 0.41803 | 0.86602 | 5 | 4.8475 | -0.15255 | -3.051
Cal. Std.7 | 0.87383 | 3.7723 | 10 | 10.1315 | 0.13149 | 1.3149
Cal. Std.8 | 2.2305 | 2.0151 | 25 | 25.8586 | 0.85863 | 3.4345
Cal. Std.9 | 4.6463 | 1.9434 | 50 | 53.8646 | 3.8646 | 7.7292
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99675 | 0.08626 | -0.0001158 | 0.0026355 | 7.7142e-05 | 0 | 0
Calibration Table - Cr (52) 52/52
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 7.3333 | 15.7459 | 0 |  | 0 | 0
Cal. Std.1 | 0.0016825 | 6.3438 | 0.01 | 0.010011 | 1.1299e-05 | 0.11299
Cal. Std.2 | 0.0061929 | 17.66 | 0.05 | 0.047303 | -0.0026971 | -5.3942
Cal. Std.3 | 0.013857 | 5.3543 | 0.1 | 0.11067 | 0.010666 | 10.6661
Cal. Std.4 | 0.058043 | 6.4807 | 0.5 | 0.476 | -0.024 | -4.8
Cal. Std.5 | 0.12185 | 5.5763 | 1 | 1.0035 | 0.0035186 | 0.35186
Cal. Std.6 | 0.57854 | 1.6686 | 5 | 4.7794 | -0.22057 | -4.4114
Cal. Std.7 | 1.2063 | 4.2102 | 10 | 9.9694 | -0.0306 | -0.306
Cal. Std.8 | 3.0396 | 3.8016 | 25 | 25.1273 | 0.1273 | 0.50919
Cal. Std.9 | 6.2458 | 1.4567 | 50 | 51.6358 | 1.6358 | 3.2716
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99866 | 0.12095 | 0.00047162 | 0.002371 | 6.9398e-05 | 0.0048745 | 0.0023606
Calibration Table - Cr (53) NH3 53/53
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 21.3333 | 21.6506 | 0 |  | 0 | 0
Cal. Std.1 | 0.00044429 | 26.208 | 0.01 | 0.0099801 | -1.9899e-05 | -0.19899
Cal. Std.2 | 0.0019112 | 8.9791 | 0.05 | 0.050739 | 0.00073883 | 1.4777
Cal. Std.3 | 0.0036953 | 3.7632 | 0.1 | 0.10031 | 0.00031021 | 0.31021
Cal. Std.4 | 0.01715 | 8.8564 | 0.5 | 0.47416 | -0.02584 | -5.1679
Cal. Std.5 | 0.035264 | 2.2751 | 1 | 0.97747 | -0.022528 | -2.2528
Cal. Std.6 | 0.17649 | 0.75528 | 5 | 4.9016 | -0.098447 | -1.9689
Cal. Std.7 | 0.36068 | 1.4715 | 10 | 10.0193 | 0.019318 | 0.19318
Cal. Std.8 | 0.92443 | 0.59744 | 25 | 25.6835 | 0.68351 | 2.734
Cal. Std.9 | 1.8873 | 2.3036 | 50 | 52.4368 | 2.4368 | 4.8735
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99953 | 0.03599 | 8.5104e-05 | 0.00041874 | 1.2257e-05 | 0.0025791 | 0.0017432
Calibration Table - Fe (54) NH3 54/54
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 648.0744 | 1.481 | 0 |  | 0 | 0
Cal. Std.3 | 0.0075477 | 8.4437 | 0.1 | 0.090388 | -0.0096123 | -9.6123
Cal. Std.4 | 0.023637 | 9.6065 | 0.5 | 0.70103 | 0.20103 | 40.2061
Cal. Std.5 | 0.036679 | 1.2646 | 1 | 1.196 | 0.19597 | 19.5966
Cal. Std.6 | 0.12155 | 3.3327 | 5 | 4.4172 | -0.5828 | -11.656
Cal. Std.7 | 0.24379 | 1.3567 | 10 | 9.0564 | -0.94356 | -9.4356
Cal. Std.8 | 0.62351 | 0.21235 | 25 | 23.4674 | -1.5326 | -6.1304
Cal. Std.9 | 1.1428 | 2.2149 | 50 | 43.1758 | -6.8242 | -13.6483
Cal. Std.10 | 2.3945 | 0.36772 | 100 | 90.6799 | -9.3201 | -9.3201
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.97976 | 0.026349 | 0.0051661 | 0.0021975 | 0.00060642 | 0.10688 | 0.0098275
Calibration Table - Mn Oshift 55/71
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 15.3333 | 7.5307 | 0 |  | 0 | 0
Cal. Std.1 | 0.0024008 | 7.4251 | 0.01 | 0.011118 | 0.0011178 | 11.1782
Cal. Std.2 | 0.0033623 | 7.2298 | 0.05 | 0.02682 | -0.02318 | -46.3598
Cal. Std.3 | 0.0065846 | 1.1023 | 0.1 | 0.079448 | -0.020552 | -20.5522
Cal. Std.4 | 0.033014 | 2.1567 | 0.5 | 0.5111 | 0.011099 | 2.2198
Cal. Std.5 | 0.067084 | 1.0342 | 1 | 1.0675 | 0.067535 | 6.7535
Cal. Std.6 | 0.33556 | 1.9863 | 5 | 5.4524 | 0.4524 | 9.0479
Cal. Std.7 | 0.68194 | 1.0635 | 10 | 11.1095 | 1.1095 | 11.0951
Cal. Std.8 | 1.7243 | 0.9862 | 25 | 28.133 | 3.133 | 12.5319
Cal. Std.9 | 3.4944 | 0.90854 | 50 | 57.0428 | 7.0428 | 14.0857
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.97811 | 0.061229 | 0.0017201 | 0.0049237 | 0.00014412 | 0.0012659 | 0.00027195
Calibration Table - Fe (56) KED 56/56
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 1244.7209 | 1.6176 | 0 |  | 0 | 0
Cal. Std.3 | 0.031696 | 14.0136 | 0.1 | 0.089586 | -0.010414 | -10.4144
Cal. Std.4 | 0.10746 | 4.1475 | 0.5 | 0.73094 | 0.23094 | 46.1882
Cal. Std.5 | 0.1587 | 7.7782 | 1 | 1.1647 | 0.16473 | 16.4726
Cal. Std.6 | 0.53342 | 3.772 | 5 | 4.337 | -0.66303 | -13.2607
Cal. Std.7 | 1.0419 | 4.7129 | 10 | 8.6418 | -1.3582 | -13.5817
Cal. Std.8 | 2.6267 | 4.5172 | 25 | 22.0579 | -2.9421 | -11.7683
Cal. Std.9 | 5.4121 | 2.5385 | 50 | 45.6379 | -4.3621 | -8.7243
Cal. Std.10 | 11.2535 | 9.7588 | 100 | 95.0887 | -4.9113 | -4.9113
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.97617 | 0.11812 | 0.021114 | 0.01072 | 0.0029582 | 0.84676 | 0.044085
Calibration Table - Co KED 59/59
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 2 | 100 | 0 |  | 0 | 0
Cal. Std.1 | 0.0023213 | 21.8481 | 0.01 | 0.0098963 | -0.00010371 | -1.0371
Cal. Std.2 | 0.013123 | 4.8199 | 0.05 | 0.052132 | 0.002132 | 4.2641
Cal. Std.3 | 0.025994 | 5.3701 | 0.1 | 0.10246 | 0.0024596 | 2.4596
Cal. Std.4 | 0.12441 | 7.6981 | 0.5 | 0.48728 | -0.01272 | -2.5441
Cal. Std.5 | 0.2554 | 2.8202 | 1 | 0.99944 | -0.0005616 | -0.05616
Cal. Std.6 | 1.2118 | 1.0841 | 5 | 4.7388 | -0.26116 | -5.2233
Cal. Std.7 | 2.5486 | 2.2247 | 10 | 9.9661 | -0.033891 | -0.33891
Cal. Std.8 | 6.4126 | 4.9587 | 25 | 25.0744 | 0.074368 | 0.29747
Cal. Std.9 | 13.0659 | 1.949 | 50 | 51.0892 | 1.0892 | 2.1784
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99956 | 0.25575 | -0.00020965 | 0.0028611 | 8.3745e-05 | 0.00062835 | 0.0018761
Calibration Table - Ni KED 60/60
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 11.3333 | 36.7353 | 0 |  | 0 | 0
Cal. Std.1 | 0.00056481 | 111.194 | 0.01 | 0.0096857 | -0.00031426 | -3.1426
Cal. Std.2 | 0.0039474 | 39.878 | 0.05 | 0.058461 | 0.008461 | 16.922
Cal. Std.3 | 0.0067195 | 3.4122 | 0.1 | 0.098433 | -0.0015668 | -1.5668
Cal. Std.4 | 0.034311 | 4.605 | 0.5 | 0.49629 | -0.0037145 | -0.7429
Cal. Std.5 | 0.065135 | 8.3183 | 1 | 0.94075 | -0.059246 | -5.9246
Cal. Std.6 | 0.32624 | 1.3341 | 5 | 4.7057 | -0.29431 | -5.8862
Cal. Std.7 | 0.70171 | 4.0687 | 10 | 10.1198 | 0.11983 | 1.1983
Cal. Std.8 | 1.7059 | 3.8553 | 25 | 24.5999 | -0.40008 | -1.6003
Cal. Std.9 | 3.4932 | 2.4302 | 50 | 50.3715 | 0.37153 | 0.74306
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99745 | 0.069351 | -0.00010691 | 0.0018741 | 5.4856e-05 | 0.013125 | 0.014334
Calibration Table - Cu (63) KED 63/63
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 6.6667 | 17.3205 | 0 |  | 0 | 0
Cal. Std.1 | 0.0035297 | 25.6201 | 0.01 | 0.0099481 | -5.1923e-05 | -0.51923
Cal. Std.2 | 0.01034 | 3.014 | 0.05 | 0.051314 | 0.0013138 | 2.6277
Cal. Std.3 | 0.018219 | 5.6547 | 0.1 | 0.099169 | -0.00083056 | -0.83056
Cal. Std.4 | 0.087301 | 0.18698 | 0.5 | 0.51878 | 0.018779 | 3.7557
Cal. Std.5 | 0.1678 | 4.8349 | 1 | 1.0078 | 0.0077634 | 0.77634
Cal. Std.6 | 0.80859 | 0.85318 | 5 | 4.8999 | -0.10008 | -2.0016
Cal. Std.7 | 1.6282 | 3.5426 | 10 | 9.8785 | -0.12149 | -1.2149
Cal. Std.8 | 4.0369 | 4.5685 | 25 | 24.5089 | -0.49112 | -1.9645
Cal. Std.9 | 8.1819 | 2.0536 | 50 | 49.6855 | -0.31447 | -0.62894
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99978 | 0.16463 | 0.0018919 | 0.0013085 | 3.8299e-05 | 0.0032524 | 0.0016464
Calibration Table - Cu (65) NH3 65/65
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 87.3336 | 18.3683 | 0 |  | 0 | 0
Cal. Std.1 | 0.0012329 | 4.1153 | 0.01 | 0.0099485 | -5.1465e-05 | -0.51465
Cal. Std.2 | 0.0040891 | 2.6122 | 0.05 | 0.050989 | 0.00098944 | 1.9789
Cal. Std.3 | 0.0074875 | 4.1721 | 0.1 | 0.099823 | -0.00017716 | -0.17716
Cal. Std.4 | 0.037592 | 10.143 | 0.5 | 0.5324 | 0.0324 | 6.48
Cal. Std.5 | 0.071227 | 1.4093 | 1 | 1.0157 | 0.015718 | 1.5718
Cal. Std.6 | 0.33827 | 0.94002 | 5 | 4.8529 | -0.1471 | -2.9419
Cal. Std.7 | 0.68506 | 0.40214 | 10 | 9.8361 | -0.16386 | -1.6386
Cal. Std.8 | 1.7179 | 0.52142 | 25 | 24.677 | -0.32295 | -1.2918
Cal. Std.9 | 3.3595 | 2.115 | 50 | 48.2667 | -1.7333 | -3.4665
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.9995 | 0.069593 | 0.00054056 | 0.00083536 | 2.4451e-05 | 0.0054611 | 0.0032406
Calibration Table - Zn NH3 66/66
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 2218.8526 | 34.4216 | 0 |  | 0 | 0
Cal. Std.3 | -0.00011766 | 180.0347 | 0.1 | 0.098964 | -0.001036 | -1.036
Cal. Std.4 | 0.013858 | 13.8373 | 0.5 | 0.53189 | 0.031887 | 6.3773
Cal. Std.5 | 0.028445 | 2.009 | 1 | 0.98372 | -0.016281 | -1.6281
Cal. Std.6 | 0.15308 | 1.1148 | 5 | 4.8444 | -0.15562 | -3.1124
Cal. Std.7 | 0.3132 | 2.0483 | 10 | 9.8043 | -0.19574 | -1.9574
Cal. Std.8 | 0.81542 | 0.58195 | 25 | 25.361 | 0.36101 | 1.444
Cal. Std.9 | 1.6029 | 2.2035 | 50 | 49.7547 | -0.24533 | -0.49065
Cal. Std.10 | 3.238 | 0.44826 | 100 | 100.4032 | 0.40319 | 0.40319
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99952 | 0.032283 | -0.0033125 | 0.00040814 | 0.00011263 | 0.29755 | 0.29655
Calibration Table - As 75/75
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | -3006.808 | 4.4982 | 0 |  | 0 | 0
Cal. Std.1 | -9.8126e-05 | 1988.4909 | 0.01 | 0.0088698 | -0.0011302 | -11.3017
Cal. Std.2 | 0.0024345 | 56.915 | 0.05 | 0.076359 | 0.026359 | 52.7173
Cal. Std.3 | 0.0037661 | 21.96 | 0.1 | 0.11184 | 0.011841 | 11.8411
Cal. Std.4 | 0.015791 | 13.1077 | 0.5 | 0.43226 | -0.06774 | -13.548
Cal. Std.5 | 0.032642 | 2.033 | 1 | 0.88131 | -0.11869 | -11.8686
Cal. Std.6 | 0.16388 | 1.8959 | 5 | 4.3785 | -0.62147 | -12.4294
Cal. Std.7 | 0.33259 | 1.3635 | 10 | 8.8741 | -1.1259 | -11.2594
Cal. Std.8 | 0.97422 | 6.2141 | 25 | 25.9717 | 0.97171 | 3.8869
Cal. Std.9 | 1.7251 | 0.43309 | 50 | 45.9809 | -4.0191 | -8.0381
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.9754 | 0.037527 | -0.00043099 | 0.0032057 | 9.3831e-05 | -0.45554 | 0.070342
Calibration Table - As KED 75/75
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 0 |  | 0 |  | 0 | 0
Cal. Std.1 | 0 |  | 0.01 | 0.010141 | 0.00014133 | 1.4133
Cal. Std.2 | 0.00032843 | 50.2367 | 0.05 | 0.052645 | 0.0026452 | 5.2903
Cal. Std.3 | 0.00048939 | 32.3247 | 0.1 | 0.073475 | -0.026525 | -26.5249
Cal. Std.4 | 0.0040803 | 9.9415 | 0.5 | 0.53819 | 0.038185 | 7.6371
Cal. Std.5 | 0.0079092 | 13.5038 | 1 | 1.0337 | 0.033701 | 3.3701
Cal. Std.6 | 0.036897 | 13.4632 | 5 | 4.7851 | -0.21491 | -4.2983
Cal. Std.7 | 0.076492 | 2.0125 | 10 | 9.9092 | -0.090759 | -0.90759
Cal. Std.8 | 0.20646 | 5.134 | 25 | 26.7291 | 1.7291 | 6.9162
Cal. Std.9 | 0.41373 | 3.2874 | 50 | 53.5519 | 3.5519 | 7.1038
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99376 | 0.0077272 | -7.8364e-05 | 0.00032786 | 9.5965e-06 | 0 | 0
Calibration Table - As RH Oshift 75/91
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 3.3333 | 34.641 | 0 |  | 0 | 0
Cal. Std.1 | 0.014565 | 29.5136 | 0.01 | 0.27042 | 0.26042 | 2604.1508
Cal. Std.2 | 0.073199 | 14.7912 | 0.05 | 0.30683 | 0.25683 | 513.6623
Cal. Std.3 | 0.13293 | 5.7828 | 0.1 | 0.34393 | 0.24393 | 243.9266
Cal. Std.4 | 0.68358 | 2.0108 | 0.5 | 0.68592 | 0.18592 | 37.1845
Cal. Std.5 | 1.4445 | 3.3853 | 1 | 1.1585 | 0.15852 | 15.852
Cal. Std.6 | 7.3489 | 2.8888 | 5 | 4.8256 | -0.17439 | -3.4877
Cal. Std.7 | 14.7666 | 1.0373 | 10 | 9.4325 | -0.5675 | -5.675
Cal. Std.8 | 38.2578 | 1.9662 | 25 | 24.0223 | -0.97768 | -3.9107
Cal. Std.9 | 81.0733 | 2.9333 | 50 | 50.614 | 0.61395 | 1.2279
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99958 | 1.6101 | -0.42083 | 0.017599 | 0.33449 | 0.00048289 | 0.00050545
Calibration Table - As Oshift 75/91
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 3.3333 | 34.641 | 0 |  | 0 | 0
Cal. Std.1 | 0.0003153 | 31.7597 | 0.01 | 0.010054 | 5.4131e-05 | 0.54131
Cal. Std.2 | 0.0016145 | 16.5741 | 0.05 | 0.050625 | 0.00062486 | 1.2497
Cal. Std.3 | 0.0029767 | 4.7144 | 0.1 | 0.093163 | -0.0068375 | -6.8375
Cal. Std.4 | 0.015222 | 3.7894 | 0.5 | 0.47554 | -0.024462 | -4.8925
Cal. Std.5 | 0.03162 | 2.0962 | 1 | 0.98759 | -0.012408 | -1.2408
Cal. Std.6 | 0.15913 | 2.0724 | 5 | 4.9692 | -0.030797 | -0.61594
Cal. Std.7 | 0.32216 | 2.1698 | 10 | 10.0603 | 0.060322 | 0.60322
Cal. Std.8 | 0.84642 | 1.327 | 25 | 26.4312 | 1.4312 | 5.7249
Cal. Std.9 | 1.6887 | 1.0176 | 50 | 52.7338 | 2.7338 | 5.4675
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99906 | 0.032024 | -6.6755e-06 | 0.00052504 | 1.5368e-05 | 0.0005258 | 0.00054191
Calibration Table - Se (77) 77/77
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 1070.7069 | 6.6947 | 0 |  | 0 | 0
Cal. Std.2 | 0.00028323 | 53.3124 | 0.05 | 0.033696 | -0.016304 | -32.6075
Cal. Std.3 | 0.0007653 | 17.5001 | 0.1 | 0.16985 | 0.069851 | 69.8505
Cal. Std.4 | 0.0017174 | 26.1088 | 0.5 | 0.43876 | -0.061238 | -12.2476
Cal. Std.5 | 0.002975 | 8.6213 | 1 | 0.79395 | -0.20605 | -20.6049
Cal. Std.6 | 0.017013 | 3.45 | 5 | 4.7587 | -0.24126 | -4.8252
Cal. Std.7 | 0.033296 | 3.6494 | 10 | 9.3575 | -0.64247 | -6.4247
Cal. Std.8 | 0.098232 | 8.4962 | 25 | 27.6977 | 2.6977 | 10.7909
Cal. Std.9 | 0.17024 | 3.1447 | 50 | 48.0343 | -1.9657 | -3.9315
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.94668 | 0.0035407 | 0.00016392 | 0.00049192 | 6.1912e-05 | 1.7195 | 0.37844
Calibration Table - Se (77) Oshift 77/93
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 0 |  | 0 |  | 0 | 0
Cal. Std.2 | 2.7406e-05 | 79.0564 | 0.05 | 0.038717 | -0.011283 | -22.5651
Cal. Std.3 | 0.00010741 | 20.7658 | 0.1 | 0.1467 | 0.046704 | 46.7039
Cal. Std.4 | 0.00036051 | 5.7845 | 0.5 | 0.48833 | -0.011671 | -2.3342
Cal. Std.5 | 0.00067309 | 19.8657 | 1 | 0.91025 | -0.089748 | -8.9748
Cal. Std.6 | 0.0033801 | 6.281 | 5 | 4.5641 | -0.43585 | -8.7171
Cal. Std.7 | 0.0072002 | 0.85954 | 10 | 9.7204 | -0.27959 | -2.7959
Cal. Std.8 | 0.018107 | 6.168 | 25 | 24.4428 | -0.5572 | -2.2288
Cal. Std.9 | 0.03738 | 0.59533 | 50 | 50.456 | 0.45597 | 0.91194
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.97616 | 0.00074086 | -1.278e-06 | 6.7254e-05 | 8.4644e-06 | 0 | 0
Calibration Table - Se (78) Oshift 78/94
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 1.3333 | 86.6025 | 0 |  | 0 | 0
Cal. Std.2 | 0.00010902 | 38.5223 | 0.05 | 0.050493 | 0.00049332 | 0.98663
Cal. Std.3 | 0.00021547 | 41.9255 | 0.1 | 0.097739 | -0.002261 | -2.261
Cal. Std.4 | 0.0011469 | 1.8443 | 0.5 | 0.51115 | 0.011148 | 2.2297
Cal. Std.5 | 0.0022317 | 2.6631 | 1 | 0.99264 | -0.0073618 | -0.73618
Cal. Std.6 | 0.010902 | 4.3708 | 5 | 4.8409 | -0.15914 | -3.1829
Cal. Std.7 | 0.021582 | 0.10685 | 10 | 9.581 | -0.41899 | -4.1899
Cal. Std.8 | 0.058184 | 0.71828 | 25 | 25.827 | 0.82699 | 3.308
Cal. Std.9 | 0.11698 | 1.2369 | 50 | 51.9228 | 1.9228 | 3.8456
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99944 | 0.002253 | -4.742e-06 | 3.0805e-05 | 3.877e-06 | 0.0029892 | 0.0077666
Calibration Table - Se (80) Oshift 80/96
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 2.6667 | 43.3013 | 0 |  | 0 | 0
Cal. Std.2 | 0.00021796 | 27.3054 | 0.05 | 0.050559 | 0.00055921 | 1.1184
Cal. Std.3 | 0.00044332 | 5.8589 | 0.1 | 0.097601 | -0.0023987 | -2.3987
Cal. Std.4 | 0.0024104 | 4.4093 | 0.5 | 0.50821 | 0.0082147 | 1.6429
Cal. Std.5 | 0.0047054 | 8.0035 | 1 | 0.98729 | -0.012712 | -1.2712
Cal. Std.6 | 0.023655 | 5.4389 | 5 | 4.9428 | -0.057228 | -1.1446
Cal. Std.7 | 0.046392 | 2.7928 | 10 | 9.689 | -0.31104 | -3.1104
Cal. Std.8 | 0.1222 | 1.3386 | 25 | 25.513 | 0.51305 | 2.0522
Cal. Std.9 | 0.24696 | 2.0249 | 50 | 51.5557 | 1.5557 | 3.1113
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99969 | 0.0047906 | -2.4244e-05 | 4.8872e-05 | 6.1509e-06 | 0.0028178 | 0.0036988
Calibration Table - Se (82) 82/82
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 59.5411 | 40.3621 | 0 |  | 0 | 0
Cal. Std.2 | 6.2855e-06 | 1029.952 | 0.05 | 0.047628 | -0.0023719 | -4.7439
Cal. Std.3 | 0.00021708 | 53.7168 | 0.1 | 0.11245 | 0.012454 | 12.4541
Cal. Std.4 | 0.0012961 | 33.3112 | 0.5 | 0.4443 | -0.055705 | -11.1409
Cal. Std.5 | 0.002864 | 6.6686 | 1 | 0.92644 | -0.073555 | -7.3555
Cal. Std.6 | 0.015869 | 5.1007 | 5 | 4.9258 | -0.074187 | -1.4837
Cal. Std.7 | 0.031346 | 1.2992 | 10 | 9.6856 | -0.31442 | -3.1442
Cal. Std.8 | 0.092412 | 6.4707 | 25 | 28.4648 | 3.4648 | 13.8592
Cal. Std.9 | 0.16497 | 2.3153 | 50 | 50.7775 | 0.77753 | 1.5551
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99518 | 0.0032517 | -0.00014859 | 0.00013076 | 1.6458e-05 | 0.10426 | 0.12837
Calibration Table - Se (82) Oshift 82/98
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 0 |  | 0 |  | 0 | 0
Cal. Std.2 | 8.1967e-05 | 45.6483 | 0.05 | 0.056507 | 0.0065069 | 13.0139
Cal. Std.3 | 9.6819e-05 | 11.2702 | 0.1 | 0.074294 | -0.025706 | -25.7061
Cal. Std.4 | 0.00043963 | 21.5619 | 0.5 | 0.48483 | -0.015168 | -3.0336
Cal. Std.5 | 0.00089169 | 8.4217 | 1 | 1.0262 | 0.026202 | 2.6202
Cal. Std.6 | 0.0041653 | 2.5335 | 5 | 4.9466 | -0.053389 | -1.0678
Cal. Std.7 | 0.0083849 | 4.3767 | 10 | 9.9999 | -0.00014907 | -0.0014907
Cal. Std.8 | 0.022589 | 0.18808 | 25 | 27.0101 | 2.0101 | 8.0405
Cal. Std.9 | 0.044347 | 1.5583 | 50 | 53.0673 | 3.0673 | 6.1345
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99191 | 0.00083503 | 3.4782e-05 | 4.3634e-05 | 5.4916e-06 | 0 | 0
Calibration Table - Mo (95) NH3 95/95
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 8.6667 | 26.6469 | 0 |  | 0 | 0
Cal. Std.1 | 0.0013853 | 8.0294 | 0.01 | 0.01049 | 0.00048985 | 4.8985
Cal. Std.2 | 0.0032307 | 10.3344 | 0.05 | 0.040352 | -0.0096478 | -19.2957
Cal. Std.3 | 0.006251 | 4.8754 | 0.1 | 0.089227 | -0.010773 | -10.7729
Cal. Std.4 | 0.031729 | 11.1695 | 0.5 | 0.50151 | 0.0015125 | 0.3025
Cal. Std.5 | 0.063711 | 0.96032 | 1 | 1.0191 | 0.019051 | 1.9051
Cal. Std.6 | 0.31399 | 0.92366 | 5 | 5.0691 | 0.069133 | 1.3827
Cal. Std.7 | 0.65048 | 0.9702 | 10 | 10.5141 | 0.51409 | 5.1409
Cal. Std.8 | 1.6682 | 0.61992 | 25 | 26.9833 | 1.9833 | 7.9334
Cal. Std.9 | 3.3534 | 1.8911 | 50 | 54.2528 | 4.2528 | 8.5056
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99538 | 0.061797 | 0.00073705 | 0.0022528 | 6.5938e-05 | 0.00060995 | 0.00049682
Calibration Table - Mo (95) OShift 95/127
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 4 | 86.6025 | 0 |  | 0 | 0
Cal. Std.1 | 0.001657 | 8.6478 | 0.01 | 0.010688 | 0.00068785 | 6.8785
Cal. Std.2 | 0.0031844 | 0.71997 | 0.05 | 0.035084 | -0.014916 | -29.8311
Cal. Std.3 | 0.0066462 | 0.68224 | 0.1 | 0.090377 | -0.009623 | -9.623
Cal. Std.4 | 0.032046 | 3.3242 | 0.5 | 0.49607 | -0.0039283 | -0.78566
Cal. Std.5 | 0.066596 | 0.62671 | 1 | 1.0479 | 0.047916 | 4.7916
Cal. Std.6 | 0.32373 | 1.9098 | 5 | 5.1549 | 0.1549 | 3.098
Cal. Std.7 | 0.66389 | 1.5561 | 10 | 10.5881 | 0.58813 | 5.8813
Cal. Std.8 | 1.7093 | 0.85967 | 25 | 27.2854 | 2.2854 | 9.1415
Cal. Std.9 | 3.4585 | 0.67197 | 50 | 55.2244 | 5.2244 | 10.4488
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99129 | 0.062608 | 0.0009878 | 0.0031446 | 9.2041e-05 | 0.00032271 | 0.00083847
Calibration Table - Cd (111) NH3 111/111
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 0 |  | 0 |  | 0 | 0
Cal. Std.1 | 0.0002725 | 3.5838 | 0.01 | 0.0099797 | -2.0338e-05 | -0.20338
Cal. Std.2 | 0.0014691 | 4.1129 | 0.05 | 0.049769 | -0.00023062 | -0.46123
Cal. Std.3 | 0.0030648 | 4.9843 | 0.1 | 0.10283 | 0.0028271 | 2.8271
Cal. Std.4 | 0.015007 | 9.0386 | 0.5 | 0.49994 | -6.4337e-05 | -0.012867
Cal. Std.5 | 0.03054 | 2.4707 | 1 | 1.0164 | 0.016404 | 1.6404
Cal. Std.6 | 0.14911 | 1.5949 | 5 | 4.9591 | -0.040924 | -0.81848
Cal. Std.7 | 0.29633 | 1.4302 | 10 | 9.8544 | -0.14563 | -1.4563
Cal. Std.8 | 0.75799 | 1.2968 | 25 | 25.2051 | 0.20509 | 0.82037
Cal. Std.9 | 1.4685 | 2.1783 | 50 | 48.8322 | -1.1678 | -2.3357
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99986 | 0.030074 | -2.7629e-05 | 0.00018738 | 5.4846e-06 | 0 | 0
Calibration Table - Cd (114) NH3 114/114
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 7.3333 | 56.7727 | 0 |  | 0 | 0
Cal. Std.1 | 0.00079445 | 8.6516 | 0.01 | 0.0099566 | -4.3353e-05 | -0.43353
Cal. Std.2 | 0.0037003 | 5.4159 | 0.05 | 0.051083 | 0.0010827 | 2.1653
Cal. Std.3 | 0.007171 | 0.79298 | 0.1 | 0.1002 | 0.0002036 | 0.2036
Cal. Std.4 | 0.035227 | 8.0223 | 0.5 | 0.49727 | -0.0027271 | -0.54542
Cal. Std.5 | 0.070097 | 0.79966 | 1 | 0.99079 | -0.0092079 | -0.92079
Cal. Std.6 | 0.35315 | 0.87608 | 5 | 4.9968 | -0.0032214 | -0.064428
Cal. Std.7 | 0.70789 | 0.82135 | 10 | 10.0174 | 0.017397 | 0.17397
Cal. Std.8 | 1.792 | 0.26037 | 25 | 25.3602 | 0.36018 | 1.4407
Cal. Std.9 | 3.4616 | 2.2944 | 50 | 48.9903 | -1.0097 | -2.0195
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99992 | 0.070656 | 9.0947e-05 | 0.00034581 | 1.0122e-05 | 0.00044941 | 0.00075674
Calibration Table - Pb 208/208
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 20.6667 | 53.2991 | 0 |  | 0 | 0
Cal. Std.1 | 0.0019799 | 8.8311 | 0.01 | 0.0099525 | -4.7545e-05 | -0.47545
Cal. Std.2 | 0.0078521 | 2.0903 | 0.05 | 0.049616 | -0.0003843 | -0.76861
Cal. Std.3 | 0.016127 | 5.0763 | 0.1 | 0.1055 | 0.0055044 | 5.5044
Cal. Std.4 | 0.077348 | 5.3034 | 0.5 | 0.51902 | 0.01902 | 3.8039
Cal. Std.5 | 0.14975 | 1.0133 | 1 | 1.008 | 0.0080363 | 0.80363
Cal. Std.6 | 0.72992 | 1.4258 | 5 | 4.9268 | -0.073248 | -1.465
Cal. Std.7 | 1.4414 | 0.64729 | 10 | 9.7324 | -0.26757 | -2.6757
Cal. Std.8 | 3.9164 | 5.1446 | 25 | 26.4491 | 1.4491 | 5.7963
Cal. Std.9 | 6.6241 | 0.46251 | 50 | 44.7382 | -5.2618 | -10.5236
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99864 | 0.14805 | 0.00050638 | 0.0029267 | 8.5665e-05 | 0.00079348 | 0.0012676
Calibration Table - Pb NH3 208/208
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 56.6667 | 25.0398 | 0 |  | 0 | 0
Cal. Std.1 | 0.0033125 | 3.8708 | 0.01 | 0.0099419 | -5.8116e-05 | -0.58116
Cal. Std.2 | 0.015534 | 1.8456 | 0.05 | 0.051294 | 0.0012936 | 2.5871
Cal. Std.3 | 0.028983 | 1.8284 | 0.1 | 0.096802 | -0.0031978 | -3.1978
Cal. Std.4 | 0.17701 | 16.3617 | 0.5 | 0.59767 | 0.09767 | 19.534
Cal. Std.5 | 0.29787 | 1.9552 | 1 | 1.0066 | 0.0066233 | 0.66233
Cal. Std.6 | 1.433 | 1.6442 | 5 | 4.8475 | -0.15252 | -3.0504
Cal. Std.7 | 2.8455 | 0.83733 | 10 | 9.6269 | -0.37314 | -3.7314
Cal. Std.8 | 6.8343 | 2.2932 | 25 | 23.1234 | -1.8766 | -7.5065
Cal. Std.9 | 14.0806 | 8.1769 | 50 | 47.6419 | -2.3581 | -4.7162
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99659 | 0.29554 | 0.00037423 | 0.0092503 | 0.00027076 | 0.0010189 | 0.00076116