## Raw Intensities
| Sample Id | R | Acquisition Time | QC Status | Dataset File | Method File | Na KED 23/23 Helium KED_LowLow (cps) | Mg KED 24/24 Helium KED_LowLow (cps) | Al KED 27/27 Helium KED_LowLow (cps) | P Oshift 31/47 Oxygen DRC (cps) | S (34) -1 34/34 Helium KED_LowLow (cps) | K KED 39/39 Helium KED_LowLow (cps) | Ca KED 43/43 Helium KED_LowLow (cps) | V KED 51/51 Helium KED_LowLow (cps) | Cr (52) 52/52 Helium KED_LowLow (cps) | Cr (53) NH3 53/53 Ammonia DRC (cps) | Fe (54) NH3 54/54 Ammonia DRC (cps) | Mn Oshift 55/71 Oxygen DRC (cps) | Fe (56) KED 56/56 Helium KED_LowLow (cps) | Co KED 59/59 Helium KED_LowLow (cps) | Ni KED 60/60 Helium KED_LowLow (cps) | Cu (63) KED 63/63 Helium KED_LowLow (cps) | Cu (65) NH3 65/65 Ammonia DRC (cps) | Zn NH3 66/66 Ammonia DRC (cps) | As 75/75 (cps) | As KED 75/75 Helium KED_LowLow (cps) | As Oshift 75/91 Oxygen DRC (cps) | Se (77) 77/77 (cps) | Se (77) Oshift 77/93 Oxygen DRC (cps) | Se (78) Oshift 78/94 Oxygen DRC (cps) | Se (80) Oshift 80/96 Oxygen DRC (cps) | Se (82) 82/82 (cps) | Se (82) Oshift 82/98 Oxygen DRC (cps) | Mo (95) NH3 95/95 Ammonia DRC (cps) | Mo (95) OShift 95/127 Oxygen DRC (cps) | Rh (IS1) 103/103 (IS) (cps) | Rh (IS2) 103/103 (IS) Oxygen DRC (cps) | Rh (IS4) 103/103 (IS) Ammonia DRC (cps) | Rh (IS5) 103/103 (IS) Ammonia DRC (cps) | Rh (IS6) 103/103 (IS) Oxygen DRC (cps) | Rh (IS7) 103/103 (IS) Helium KED_LowLow (cps) | Cd (111) NH3 111/111 Ammonia DRC (cps) | Cd (114) NH3 114/114 Ammonia DRC (cps) | Pb 208/208 (cps) | Pb NH3 208/208 Ammonia DRC (cps)
1 | sac blank |  | 06/16/2023 11:58:56 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\sac blank .650 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 5.3333 | 0 | 68.6668 | 1100.0424 | 2904.9621 | 46.0001 | 0.66667 | 0.66667 | 36 | 34 | 646.1479 | 20.6667 | 10006.837 | 2 | 26 | 27.3334 | 135.334 | 2227.5091 | -4100.674 | 1.3333 | 16 | 1387.4008 | 0 | 2.6667 | 6.6667 | 60.7165 | 0 | 274.6693 | 20 | 271255.1708 | 364338.5337 | 410502.873 | 325134.5776 | 318004.0787 | 34778.2898 | 1.3333 | 22.6667 | 34.6667 | 62.6667
2 | sac blank cal |  | 06/16/2023 12:04:57 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\sac blank cal.651 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4.6667 | 0.66667 | 57.3335 | 1104.7094 | 2758.9331 | 50.6668 | 0 | 3.3333 | 28.6667 | 21.3334 | 653.5937 | 25.3334 | 9852.0654 | 4.6667 | 28 | 22 | 118.6672 | 1820.7828 | -3859.13 | 2 | 14 | 1368.7323 | 1.3333 | 4.6667 | 3.3333 | 50.6298 | 2 | 259.3357 | 7.3333 | 273640.0501 | 363100.0755 | 410914.6013 | 323638.2545 | 316582.3718 | 33547.3484 | 0.66667 | 36.6667 | 49.3334 | 111.3335
3 | sac blank cal |  | 06/16/2023 12:10:57 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\sac blank cal.652 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4.6667 | 0 | 30.6667 | 1023.37 | 2780.9375 | 55.3334 | 0 | 4 | 22 | 18.6667 | 653.4244 | 18 | 9682.6143 | 6 | 31.3334 | 21.3334 | 126.0006 | 3588.4537 | -3957.3866 | 0 | 10.6667 | 1390.0677 | 0 | 0.66667 | 5.3333 | 69.8335 | 0 | 252.0022 | 5.3333 | 269150.2831 | 360078.3844 | 404036.9377 | 319318.0668 | 309692.239 | 33259.3435 | 2.6667 | 30 | 58 | 88.0001
4 | CAL BLANK |  | 06/16/2023 12:16:57 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CAL BLANK.653 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4.6667 | 0 | 50.6668 | 950.0317 | 2914.298 | 44.6667 | 0.66667 | 6 | 27.3334 | 20 | 663.2662 | 25.3334 | 9780.6812 | 5.3333 | 26.6667 | 16 | 127.3339 | 1445.4065 | -3944.1929 | 0.66667 | 14 | 1375.3996 | 0 | 0.66667 | 6 | 63.8984 | 1.3333 | 231.3352 | 4 | 269907.9681 | 354160.9715 | 399133.1355 | 321765.3432 | 314307.4835 | 32774.2249 | 2.6667 | 29.3334 | 51.3334 | 95.3335
5 | LOW 0.01 |  | 06/16/2023 12:22:58 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LOW 0.01.654 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 37.3334 | 19.3333 | 158.0009 | 5823.8549 | 2862.2868 | 76.6669 | 21.3334 | 28.6667 | 101.3337 | 144.6674 | 3296.2355 | 474.6746 | 11708.8136 | 111.3338 | 44.0001 | 129.3339 | 603.3461 | 2600.2374 | -3749.4175 | 5.3333 | 112.6671 | 1418.7373 | 0.66667 | 8.6667 | 10.6667 | 69.8876 | 3.3333 | 511.3425 | 255.3356 | 270980.1273 | 354851.3171 | 400108.6326 | 321453.915 | 308628.908 | 33364.9273 | 77.3335 | 251.3356 | 578.0047 | 1151.3512
6 | LOW 0.05 |  | 06/16/2023 12:28:58 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LOW 0.05.655 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 209.3349 | 89.3336 | 136.6673 | 6237.3675 | 2933.6345 | 180.0012 | 58.6668 | 146.6674 | 205.3348 | 490.6751 | 1737.3449 | 780.6881 | 10305.0488 | 337.3373 | 108.0004 | 244.0021 | 1178.0486 | 2032.1446 | -3621.0163 | 12 | 413.3393 | 1466.742 | 8 | 27.3334 | 69.3335 | 94.5654 | 14.6667 | 1050.7053 | 735.3523 | 268088.4116 | 348270.1943 | 395087.317 | 318408.0119 | 299948.2338 | 32527.6586 | 408.0059 | 1074.0404 | 1566.7008 | 3367.4881
7 | LOW 0.1 |  | 06/16/2023 12:34:59 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LOW 0.1.656 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 372.0049 | 172.6677 | 51.3334 | 2926.2997 | 2784.9384 | 279.3361 | 137.334 | 243.3354 | 348.671 | 892.6946 | 2425.0478 | 1520.7476 | 10716.6857 | 614.0133 | 189.3346 | 415.3394 | 2034.8116 | 2344.1923 | -3607.6547 | 26.6667 | 739.3525 | 1511.4133 | 12 | 45.3334 | 114.0005 | 107.2052 | 19.3333 | 1738.7725 | 1412.7366 | 265582.7134 | 347157.0686 | 400478.0263 | 321727.2865 | 302713.9605 | 32481.5538 | 812.0231 | 1964.1354 | 2958.1227 | 6454.5731
8 | LOW 0.5 |  | 06/16/2023 12:40:59 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LOW 0.5.657 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1905.4609 | 895.3615 | 262.0024 | 7618.0314 | 3193.6904 | 1178.7155 | 590.0123 | 1225.3859 | 1646.7618 | 4420.0176 | 6751.0466 | 7157.7932 | 13073.3133 | 3048.3253 | 833.3577 | 1984.8047 | 9774.0104 | 6478.1363 | -972.643 | 132.6673 | 3545.7734 | 1692.1002 | 84.0002 | 238.6687 | 535.3434 | 327.2452 | 99.3337 | 8341.1012 | 7068.4152 | 260297.4376 | 339285.2492 | 385459.7385 | 314549.4697 | 295286.5087 | 32322.5262 | 4033.9031 | 9580.545 | 14546.2277 | 31903.9654
9 | STD 1 |  | 06/16/2023 12:46:58 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 1.658 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3827.8463 | 1710.7691 | 294.003 | 5903.8865 | 3846.5179 | 2386.1994 | 1095.3755 | 2431.5405 | 3257.7047 | 8420.4809 | 10339.4742 | 14093.6171 | 15142.6884 | 5909.2219 | 1579.4207 | 3639.1305 | 17636.2153 | 9441.7862 | 1271.3786 | 232.0019 | 6893.6634 | 2012.1419 | 148.6675 | 487.3417 | 1052.0387 | 533.9296 | 175.3344 | 15880.8228 | 14054.244 | 256927.841 | 338303.9383 | 382744.0335 | 310102.4133 | 296139.2927 | 32209.6051 | 7692.0706 | 18452.5781 | 27425.0089 | 62309.4235
10 | STD 5 |  | 06/16/2023 12:52:58 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 5.659 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 18468.5972 | 8144.9886 | 948.6982 | 18379.1498 | 7475.9561 | 10991.5611 | 5537.0731 | 11834.2334 | 15447.681 | 39541.9835 | 29269.6188 | 68156.2047 | 25556.181 | 28289.3201 | 7309.2032 | 16908.6731 | 80925.9021 | 39376.1958 | 23808.5334 | 1124.711 | 33116.3412 | 4225.9585 | 747.3529 | 2366.196 | 4846.1553 | 2654.0049 | 900.0284 | 76033.8083 | 66289.4481 | 253318.9374 | 322069.7942 | 368756.3281 | 298152.7386 | 287394.9296 | 30266.7021 | 37461.7243 | 89060.7553 | 129775.485 | 293187.7839
11 | STD 10 |  | 06/16/2023 12:58:58 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 10.660 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 36354.2013 | 16140.4489 | 1633.4268 | 28195.1323 | 11899.6223 | 21753.22 | 11044.9347 | 23652.2308 | 30294.0888 | 77718.166 | 55110.6175 | 136802.5803 | 39727.8371 | 55563.1854 | 14144.3348 | 32596.4812 | 155860.317 | 76070.6676 | 51456.631 | 2148.8283 | 64866.9384 | 7505.305 | 1415.4035 | 4570.0644 | 9653.2612 | 5074.1805 | 1728.7714 | 149949.5634 | 132254.0487 | 251310.7966 | 316131.6182 | 361552.0644 | 294804.2408 | 281081.1205 | 29857.8363 | 72362.1457 | 172538.3481 | 240456.8359 | 561793.9752
12 | STD 25 |  | 06/16/2023 13:04:58 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 25.661 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 92457.5698 | 39275.9179 | 3862.5222 | 67540.7423 | 25411.2514 | 54841.7409 | 26737.6651 | 58716.4196 | 74560.7545 | 192475.3035 | 132237.0919 | 334186.4236 | 82493.5034 | 138356.141 | 34806.3513 | 79760.7161 | 375704.3454 | 180574.7734 | 140920.8353 | 5488.3876 | 163950.7622 | 16371.3759 | 3687.8094 | 11488.6181 | 24718.701 | 13158.2053 | 4437.3557 | 369976.4138 | 329101.1521 | 250484.9366 | 298535.2123 | 336580.588 | 276834.7038 | 292829.6889 | 28656.0456 | 174907.6082 | 416050.292 | 533396.2897 | 1312711.6311
13 | STD 50 |  | 06/16/2023 13:10:58 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 50.662 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 177624.2761 | 73197.73 | 7227.1621 | 123855.954 | 45565.9019 | 107924.8603 | 52431.3919 | 114919.7532 | 144433.3121 | 365089.75 | 242055.5503 | 625051.394 | 155361.6357 | 262760.6042 | 64773.8585 | 147222.7296 | 684039.3706 | 330052.6661 | 273194.0211 | 10100.9049 | 308711.4467 | 29696.8358 | 6937.0181 | 21852.7026 | 46216.6504 | 24870.7528 | 8347.7743 | 698783.3691 | 624893.4231 | 237649.7514 | 280575.4282 | 316455.4454 | 261750.9622 | 253532.0934 | 27131.7462 | 318930.804 | 753160.4395 | 793996.3566 | 2321006.6929
14 | STD 100 |  | 06/16/2023 13:16:57 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 100.663 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 345077.232 | 139833.7653 | 13771.3016 | 229938.5101 | 82806.6332 | 212571.1363 | 101062.9537 | 226808.3452 | 280403.9198 | 689788.2108 | 430944.1601 | 1172055.9329 | 270250.7532 | 499456.1773 | 122640.8556 | 272544.5907 | 1215676.5993 | 578935.912 | 523683.4614 | 19051.364 | 567154.0318 | 53996.5315 | 12827.7585 | 40869.044 | 86436.0385 | 47847.9721 | 15730.6564 | 1295778.4247 | 1181199.0135 | 210906.948 | 255858.9654 | 277990.8731 | 230754.8096 | 230674.1552 | 24521.027 | 560029.9298 | 1314367.2737 | 1156979.2929 | 4496557.2426
15 | Calibration Curves |  | 06/16/2023 13:16:57
16 | LBLANK |  | 06/16/2023 13:22:58 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.664 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 27.3334 | 7.3333 | 21.3333 | 1730.7715 | 3063.6624 | 78.0002 | 4 | 5.3333 | 52.0001 | 62.0001 | 559.0423 | 37.3334 | 10218.3203 | 19.3333 | 79.3336 | 35.3334 | 160.0009 | 5230.2951 | -3737.8997 | 0 | 23.3334 | 1395.4015 | 2 | 12 | 6 | 66.4893 | 0 | 3459.0854 | 163.3343 | 264900.2734 | 344637.1589 | 383179.046 | 314055.9082 | 301031.8543 | 31338.3419 | 6 | 40.6667 | 37.3334 | 90.0001
17 | LBLANK |  | 06/16/2023 13:41:50 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.665 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10 | 1.3333 | 46.6668 | 741.3526 | 2342.1965 | 68.6668 | 0.66667 | 6 | 42.6667 | 49.3334 | 556.8861 | 18.6667 | 9491.1642 | 6 | 37.3334 | 18.6667 | 122.0005 | 1410.0772 | -3758.8296 | 0 | 19.3333 | 1292.7252 | 1.3333 | 8 | 4.6667 | 67.248 | 0 | 1020.7031 | 84.6669 | 254503.0387 | 322216.0874 | 364269.8879 | 305105.7248 | 279979.7305 | 32704.8747 | 1.3333 | 25.3334 | 18.6667 | 51.3334
18 | ICV PE |  | 06/16/2023 13:47:50 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\ICV PE.666 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 36947.06 | 16230.5483 | 1626.7593 | 50369.9785 | 11509.9686 | 21749.2123 | 10734.6985 | 22778.1505 | 29610.6559 | 77278.4744 | 55019.8525 | 132295.7637 | 39101.4485 | 53502.6694 | 13632.5034 | 32058.5984 | 152458.0792 | 70854.6257 | 49903.6395 | 1938.1316 | 63390.3394 | 5330.328 | 1385.4006 | 4462.6969 | 9519.171 | 4839.881 | 1652.7623 | 138700.1196 | 121796.3519 | 245191.604 | 300675.269 | 345722.3099 | 284714.6029 | 265426.4691 | 28577.8896 | 67625.6889 | 161797.1324 | 225287.7191 | 520975.7883
19 | CCV 1ppb TRACE |  | 06/16/2023 13:53:50 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.667 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9.3333 | 1.3333 | 8.6667 | 314.0035 | 385.3386 | 30.6667 | 2 | 2.6667 | 24 | 29.3334 | 432.9562 | 8 | 4530.0545 | 8.6667 | 21.3333 | 16 | 102.0004 | 707.351 | 84.4113 | 0.66667 | 6.6667 | 262.6691 | 0 | 5.3333 | 3.3333 | 18.1663 | 0 | 796.6889 | 4.6667 | 275520.5968 | 357683.5542 | 404966.6031 | 332299.3953 | 314707.8476 | 36459.1362 | 2 | 16 | 15.3333 | 28
20 | CCV 1ppb TRACE |  | 06/16/2023 14:00:07 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.668 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3871.8582 | 1764.7758 | 364.6713 | 3375.0656 | 3844.5175 | 2492.8843 | 1085.3746 | 2511.5541 | 3281.0435 | 8507.8661 | 8317.1101 | 14118.3068 | 14188.3764 | 5883.2112 | 1548.7506 | 3660.4691 | 17325.166 | 8431.8214 | 1476.0578 | 212.6683 | 6631.5391 | 2049.4805 | 137.334 | 463.3409 | 1040.0379 | 535.4786 | 175.3344 | 15944.8935 | 13360.9125 | 249367.3099 | 319820.7585 | 366417.1397 | 299374.3473 | 284383.4349 | 30032.8719 | 7306.5354 | 17358.5414 | 25701.0181 | 57111.9197
21 | LFB |  | 06/16/2023 14:06:07 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.669 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 36072.1644 | 15755.3536 | 1616.0915 | 31506.7061 | 11501.2962 | 21747.2132 | 10401.1187 | 22624.5784 | 29451.0214 | 74861.0116 | 51497.402 | 127229.3949 | 38168.9432 | 53788.4656 | 13603.8128 | 31829.4425 | 148263.436 | 72310.5911 | 49974.4883 | 2203.5039 | 61986.1911 | 7118.4402 | 1393.4013 | 4378.0046 | 9174.2785 | 5047.2691 | 1713.4361 | 138998.6283 | 122634.1419 | 240822.2342 | 295243.5241 | 340267.0087 | 278862.1119 | 260573.1265 | 26507.9964 | 67178.3009 | 158042.3154 | 218214.3708 | 505261.1607
22 | LLOQ 0.01 |  | 06/16/2023 14:12:05 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LLOQ 0.01.670 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 58.6668 | 19.3333 | 215.3352 | 6129.9818 | 3018.9859 | 82.0002 | 12 | 30 | 81.3336 | 159.3342 | 3346.3578 | 429.3398 | 11441.2466 | 98.0003 | 54.6668 | 123.3339 | 610.013 | 2210.8378 | -3659.5427 | 3.3333 | 86.0003 | 1387.4008 | 2 | 11.3333 | 12 | 76.6842 | 4 | 953.3652 | 274.0026 | 251360.2066 | 316677.6067 | 367645.3383 | 296732.1147 | 276145.5118 | 30093.6642 | 72.0002 | 212.0016 | 478.0033 | 1070.6824
23 | CCV 10 ppb MAJORS |  | 06/16/2023 14:18:05 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.671 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 36298.7293 | 15651.2373 | 1570.0863 | 25418.5951 | 11582.6951 | 21595.644 | 10249.0095 | 22507.7171 | 29008.0896 | 76653.7797 | 52666.6729 | 129678.5925 | 37911.5749 | 54022.6275 | 13970.1614 | 32061.9411 | 150974.2554 | 73002.0546 | 49833.7552 | 2134.1597 | 62066.5376 | 7138.4504 | 1335.3959 | 4205.9527 | 9149.5965 | 4878.9237 | 1632.76 | 142788.1001 | 125084.5744 | 238342.5087 | 287577.2647 | 332470.9726 | 276625.3969 | 258561.7143 | 27893.2048 | 66868.8021 | 159375.4311 | 220830.0208 | 514189.9914
24 | LLOQ 0.1 |  | 06/16/2023 14:24:05 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LLOQ 0.1.672 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 398.0056 | 165.3343 | 48.6668 | 1508.7468 | 3027.6546 | 290.0029 | 102.0004 | 248.6688 | 316.6702 | 865.3596 | 2253.6225 | 1483.4105 | 10860.7945 | 568.0114 | 176.6678 | 417.3394 | 1980.804 | 2314.8546 | -3201.3829 | 23.3334 | 664.6821 | 1453.4074 | 13.3333 | 50.0001 | 102.0004 | 89.7634 | 26.6667 | 2240.1757 | 1336.0625 | 251609.074 | 315533.9998 | 362521.6662 | 297856.501 | 271411.4699 | 29644.0604 | 721.3516 | 1699.4346 | 2512.0872 | 5650.4385
25 | 9_13_21_F_11_DUP |  | 06/16/2023 14:30:04 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_F_11_DUP.673 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 23465.2569 | 12855.7856 | 1392.0679 | 9238.3201 | 15118.6629 | 3983.8889 | 34013.1108 | 1652.0956 | 367.3381 | 856.6924 | 24234.7973 | 60597.5962 | 23696.9725 | 124.0005 | 342.6708 | 6781.6095 | 32251.0341 | 25093.3564 | 17429.94 | 812.6898 | 24515.6848 | 1365.3987 | 12.6667 | 47.3334 | 91.3336 | 244.5466 | 13.3333 | 31387.7863 | 27113.0388 | 222221.8074 | 278349.2177 | 318240.2778 | 264312.8247 | 250563.0459 | 27121.0586 | 80.6669 | 217.335 | 2854.7769 | 6247.2034
26 | LBLANK |  | 06/16/2023 14:47:08 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.675 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 7.3333 | 2 | 27.3334 | 548.6772 | 2630.2422 | 54.0001 | 2 | 4.6667 | 48.6668 | 42.6667 | 508.849 | 14.6667 | 8976.1526 | 8.6667 | 23.3334 | 16 | 110.0004 | 748.0199 | -3533.6284 | 2 | 16 | 1213.385 | 0.66667 | 2.6667 | 4.6667 | 48.6839 | 0.66667 | 446.6737 | 34 | 232895.7495 | 295279.9247 | 337579.0948 | 277993.1585 | 251865.0368 | 27638.0454 | 0.66667 | 22.6667 | 21.3333 | 46
27 | 0.5 STD |  | 06/16/2023 14:53:09 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\0.5 STD.676 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1808.1146 | 808.0229 | 219.335 | 6119.9781 | 3097.0024 | 1165.3809 | 534.01 | 1106.0428 | 1497.4118 | 3911.8691 | 6180.5957 | 6469.466 | 11696.7867 | 2636.9101 | 712.0178 | 1762.7755 | 8777.3648 | 5843.1951 | -827.9804 | 99.3337 | 3135.011 | 1484.0772 | 68.6668 | 228.0018 | 468.6744 | 264.1041 | 93.3336 | 7410.5886 | 6200.0118 | 230199.9828 | 290615.7396 | 333024.3242 | 273633.7973 | 252893.5011 | 27255.314 | 3387.0682 | 8158.9966 | 12077.9824 | 27116.7994
28 | 0.1 STD |  | 06/16/2023 14:59:12 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\0.1 STD.677 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 369.3381 | 156.0009 | 35.3334 | 2612.2392 | 2807.6093 | 262.0024 | 116.0005 | 200.0014 | 323.337 | 786.0217 | 2112.2344 | 1312.727 | 9785.3532 | 554.0107 | 164.6676 | 356.6712 | 1782.1112 | 1831.4508 | -2780.8681 | 28.6667 | 643.3478 | 1220.0522 | 18 | 46.0001 | 98.667 | 88.5329 | 10.6667 | 1729.438 | 1250.7216 | 232108.9808 | 290921.0448 | 337273.0849 | 275960.6031 | 251580.6985 | 27957.9979 | 720.0181 | 1693.4337 | 2468.7496 | 5259.7159
29 | CCV 10 ppb MAJORS |  | 06/16/2023 15:05:12 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.678 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 33715.0722 | 14700.8951 | 1354.7312 | 23464.5899 | 10895.4871 | 20478.6678 | 9871.41 | 21243.1171 | 27118.3815 | 71727.6243 | 49614.5523 | 118680.9492 | 35635.7305 | 49969.2712 | 12901.8257 | 29762.9841 | 141866.9482 | 68146.8315 | 47602.6412 | 2017.4758 | 57287.9822 | 6596.8573 | 1284.7245 | 3960.5491 | 8521.8745 | 4700.1808 | 1514.747 | 133908.6976 | 116119.3569 | 226803.0185 | 265194.2076 | 312687.2907 | 258218.9534 | 241226.3203 | 26232.0635 | 63210.885 | 150898.8389 | 210596.6962 | 485072.6851
30 | 9_13_21_F_11_DUP |  | 06/16/2023 15:11:13 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_F_11_DUP.679 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 22671.3095 | 12880.4716 | 1420.7375 | 8628.6072 | 14741.6027 | 3853.1863 | 32591.1348 | 1523.4146 | 360.0045 | 819.3569 | 23443.4799 | 57970.0531 | 23049.2463 | 119.3338 | 342.0041 | 6528.8252 | 30942.1439 | 23901.979 | 16479.8966 | 744.0194 | 23366.4281 | 1293.392 | 12 | 37.3334 | 74.0002 | 213.8414 | 8.6667 | 30017.5127 | 26189.9853 | 218252.5864 | 268588.916 | 306604.6407 | 253331.8242 | 241789.1264 | 26547.3105 | 65.3335 | 187.3346 | 2702.7668 | 5981.818
31 | 9_13_21_F_11_DUP_LSP |  | 06/16/2023 15:17:13 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_F_11_DUP_LSP.680 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 52943.257 | 24977.1497 | 2531.5577 | 32271.746 | 21092.894 | 23313.0118 | 38590.0523 | 21736.5246 | 26075.1091 | 68182.3261 | 67955.7258 | 165785.7687 | 46914.2441 | 47516.241 | 12173.1848 | 33171.1357 | 154422.883 | 81496.4761 | 63347.9481 | 2588.2351 | 75098.8782 | 6249.3668 | 1175.3818 | 3763.1623 | 7764.1095 | 4631.3256 | 1392.7346 | 151584.1376 | 134553.3611 | 215057.5576 | 259936.8998 | 295256.0017 | 244398.65 | 230444.1962 | 25108.7185 | 58887.139 | 139815.5558 | 199971.2084 | 452526.2681
32 | 9_13_21_UF_1 |  | 06/16/2023 15:23:13 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_UF_1.681 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 28419.5918 | 8671.9648 | 294.6697 | 20484.0121 | 19547.3703 | 4162.6065 | 32875.1193 | 654.015 | 231.3352 | 426.6731 | 4886.1652 | 151557.0409 | 12677.6251 | 163.3343 | 646.6813 | 1745.44 | 8888.0979 | 9297.0273 | 15875.5242 | 708.6843 | 22162.5127 | 1355.3978 | 12.6667 | 14.6667 | 43.3334 | 210.5134 | 11.3333 | 199195.1787 | 177925.9155 | 217805.1482 | 267313.8922 | 306000.5528 | 256214.671 | 241074.402 | 25727.1456 | 44.6667 | 122.0005 | 1144.6847 | 2492.7515
33 | 9_13_21_UF_6 |  | 06/16/2023 15:29:13 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_UF_6.682 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 30291.4153 | 10120.2524 | 147.3341 | 8156.9977 | 14920.4568 | 3768.4971 | 33081.6013 | 1333.3956 | 264.6691 | 447.3404 | 3550.2885 | 138193.2083 | 12229.2322 | 118.6672 | 346.0042 | 5244.296 | 24770.7916 | 15318.8762 | 16240.8317 | 750.6864 | 23052.5863 | 1384.7338 | 13.3333 | 37.3334 | 74.0002 | 258.5145 | 14 | 75573.3825 | 67848.744 | 217645.5528 | 266102.2501 | 299295.5547 | 255242.904 | 241211.4325 | 25851.3699 | 92.0003 | 240.002 | 1516.0321 | 3332.1499
34 | 9_13_21_UF_6_DUP |  | 06/16/2023 15:35:13 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_UF_6_DUP.683 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 29528.5068 | 9870.076 | 142.6674 | 7877.5048 | 14725.588 | 3623.1261 | 31953.6975 | 1332.0621 | 315.3368 | 649.3481 | 4264.632 | 133892.5981 | 12015.0508 | 112.0004 | 348.0043 | 5169.6019 | 24726.0493 | 12793.7262 | 16423.3764 | 744.0194 | 22901.6761 | 1336.0625 | 8.6667 | 25.3334 | 70.0002 | 245.1484 | 12 | 74765.8075 | 66274.7215 | 211424.5194 | 260531.6015 | 299978.8753 | 245936.9483 | 232790.1873 | 25386.5418 | 98.667 | 215.335 | 1554.034 | 3241.4771
35 | 9_13_21_UF_8 |  | 06/16/2023 15:41:13 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_UF_8.684 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 30233.294 | 9535.1829 | 136.0006 | 8942.7986 | 13290.8492 | 3786.5018 | 30725.0076 | 1683.4326 | 234.0019 | 419.3395 | 3134.7299 | 118838.2878 | 11097.6425 | 92.0003 | 336.004 | 4508.7116 | 21375.3177 | 7865.4982 | 18222.7606 | 745.3528 | 24676.6301 | 1246.7211 | 8.6667 | 30.6667 | 68.6668 | 201.0663 | 11.3333 | 62448.2253 | 55388.5019 | 210079.0281 | 259007.5555 | 299086.0065 | 247261.6142 | 232864.9443 | 24993.1778 | 65.3335 | 146.0008 | 1275.3559 | 2708.0989
36 | 9_13_21_UF_9 |  | 06/16/2023 15:47:13 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_UF_9.685 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 27946.6422 | 9800.0274 | 172.6677 | 8402.4708 | 12522.8221 | 3599.7869 | 28155.7188 | 1472.0759 | 298.6698 | 551.3441 | 3861.1219 | 43018.0097 | 11724.812 | 83.3336 | 274.6693 | 4198.6171 | 19562.7188 | 6310.0601 | 14702.7169 | 680.0162 | 21162.335 | 1363.3986 | 10 | 27.3334 | 60.6668 | 214.5729 | 14.6667 | 47216.568 | 42211.6222 | 212480.5758 | 258893.0249 | 297856.341 | 244576.0713 | 231758.2798 | 25092.6904 | 39.3334 | 94.0003 | 1299.3567 | 2653.4286
37 | LBLANK |  | 06/16/2023 15:53:13 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.686 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.6667 | 0 | 12.6667 | 520.6762 | 2544.8939 | 42.0001 | 1.3333 | 2 | 24 | 36 | 506.9349 | 23.3334 | 8355.1094 | 1.3333 | 22.6667 | 20 | 92.667 | 485.3416 | -3087.3044 | 9.3333 | 256.669 | 1225.3859 | 0 | 3.3333 | 4.6667 | 53.3669 | 0 | 106.6671 | 22.6667 | 226720.5403 | 278322.546 | 325386.1147 | 268735.6278 | 243668.5438 | 26821.1566 | 0.66667 | 8 | 27.3333 | 40
38 | CCV 1ppb TRACE |  | 06/16/2023 15:59:13 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.687 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3699.8124 | 1606.0904 | 301.3365 | 2738.2625 | 3589.785 | 2178.8329 | 1037.3711 | 2239.509 | 2872.956 | 7369.9008 | 7223.8483 | 12141.8245 | 12471.4419 | 5201.6153 | 1403.4025 | 3217.6957 | 15164.7114 | 7447.2749 | 1351.0476 | 210.6682 | 5915.2245 | 1788.7787 | 127.3339 | 420.6729 | 839.358 | 509.6605 | 156.0009 | 13277.5012 | 11807.5463 | 224402.0785 | 278187.3128 | 318651.5832 | 262640.4992 | 245157.8257 | 25306.4862 | 6539.4968 | 15201.4208 | 23588.9466 | 51513.0124
39 | CCV 1ppb TRACE |  | 06/16/2023 16:05:32 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.688 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3490.4265 | 1576.0871 | 318.6702 | 2808.276 | 3577.1145 | 2214.8384 | 1024.0367 | 2260.1789 | 2882.9577 | 7675.3948 | 7674.8918 | 12386.7011 | 12905.1597 | 5224.2885 | 1461.4081 | 3176.3532 | 15367.5947 | 7494.6323 | 1144.9651 | 218.6683 | 5980.585 | 1829.4505 | 121.3339 | 412.0059 | 856.0256 | 482.1587 | 172.6677 | 13623.8265 | 11905.6279 | 224690.3416 | 278094.7005 | 322519.6244 | 265604.4616 | 245624.4224 | 26899.9785 | 6568.177 | 15444.3448 | 23933.2525 | 51619.7504
40 | CCV 10 ppb MAJORS |  | 06/16/2023 16:11:33 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.689 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 32901.1814 | 14745.6072 | 1446.74 | 22292.7143 | 10934.8523 | 20272.3847 | 9675.2754 | 20779.1027 | 26481.1889 | 69845.6723 | 49632.524 | 115214.8006 | 35097.731 | 49026.6515 | 12423.4001 | 28986.044 | 136703.6293 | 66868.8029 | 45338.5175 | 1970.8027 | 55059.2416 | 6543.4984 | 1213.385 | 3887.1957 | 8230.3704 | 4504.3584 | 1467.4087 | 128089.0226 | 111832.7 | 223053.9793 | 261302.1483 | 306860.724 | 252052.8155 | 233908.1594 | 25695.0913 | 61667.4922 | 147709.0483 | 213475.8557 | 479107.1083
41 | CCV 10 ppb MAJORS |  | 06/16/2023 16:17:52 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.690 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 33645.5762 | 15132.0104 | 1416.737 | 22882.9816 | 11031.5938 | 20377.8587 | 9696.6231 | 21022.7941 | 27060.9397 | 71045.5677 | 49122.6719 | 117784.9063 | 35621.6902 | 49555.812 | 12483.4521 | 29495.0855 | 140076.7558 | 68282.1363 | 46780.0192 | 1922.7961 | 56226.4427 | 6541.4975 | 1218.7188 | 3877.8596 | 8144.3216 | 4633.3696 | 1474.0761 | 130528.9779 | 113425.1778 | 224935.1818 | 263575.3844 | 308180.7727 | 254476.1549 | 238497.7411 | 25829.9986 | 61433.1452 | 147252.372 | 212036.0831 | 479424.1948
42 | 9_13_21_UF_10 |  | 06/16/2023 16:23:53 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_UF_10.691 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 26156.5942 | 12279.9426 | 210.0015 | 10646.6328 | 14335.1902 | 3913.2028 | 32594.4756 | 1577.4205 | 241.3354 | 406.6725 | 2513.9929 | 49456.7966 | 11666.0949 | 78.6669 | 265.3358 | 3400.4048 | 16065.0299 | 6989.7103 | 16807.0076 | 756.6867 | 23151.4116 | 1231.3864 | 8 | 28.6667 | 68.6668 | 243.2294 | 11.3333 | 30714.3173 | 27358.1724 | 212735.826 | 260791.7426 | 295509.905 | 249160.768 | 234713.7987 | 25505.418 | 56.0001 | 104.6671 | 1363.3584 | 2989.4524
43 | 9_13_21_UF_11 |  | 06/16/2023 16:29:52 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_UF_11.692 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 25749.186 | 12406.7194 | 152.0008 | 10225.6584 | 13906.7658 | 3869.8574 | 32392.6841 | 1534.0824 | 463.3409 | 987.3675 | 6550.6837 | 53338.7236 | 13272.1661 | 71.3335 | 256.0023 | 4026.5675 | 18919.1872 | 8774.6944 | 16278.5911 | 697.3504 | 22326.1001 | 1236.0538 | 8 | 35.3334 | 68.0002 | 217.164 | 14 | 28652.708 | 25089.3552 | 208468.1962 | 254132.0577 | 294973.6357 | 240084.9302 | 228577.5079 | 25088.678 | 45.3334 | 142.6674 | 1568.7001 | 3289.4804
44 | 9_13_21_LAB_A |  | 06/16/2023 16:35:53 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_LAB_A.693 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 36 | 0 | 133.334 | 1087.3747 | 2649.579 | 44.0001 | 3.3333 | 2 | 83.3336 | 146.0008 | 1584.8661 | 477.3413 | 8978.1548 | 6.6667 | 60.0001 | 1841.452 | 8633.2748 | 2762.2672 | -3107.6649 | 4.6667 | 165.3343 | 1204.0509 | 2 | 2 | 5.3333 | 70.0932 | 0 | 91.3336 | 9.3333 | 226541.1767 | 275132.9634 | 320898.8992 | 265179.6474 | 239875.8733 | 27457.696 | 0.66667 | 16.6667 | 664.6731 | 1464.0296
45 | 9_13_21_BB |  | 06/16/2023 16:41:52 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_BB.694 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10 | 0 | 158.0009 | 1058.7059 | 2500.886 | 48.6668 | 6.6667 | 4.6667 | 72.0002 | 158.0009 | 1494.3733 | 51.3334 | 8745.3429 | 2.6667 | 30.6667 | 408.6725 | 1974.1366 | 1666.0973 | -3310.178 | 3.3333 | 105.3337 | 1169.3814 | 0 | 1.3333 | 5.3333 | 39.9901 | 0.66667 | 76.0002 | 23.3334 | 227480.8662 | 273392.5949 | 320221.2874 | 263843.1728 | 237195.6658 | 26411.0608 | 2 | 24.6667 | 695.34 | 1504.6976
46 | 9_13_21_FB |  | 06/16/2023 16:47:52 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_FB.695 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 587.3454 | 59.3335 | 2010.8082 | 3859.8549 | 2677.5843 | 169.3344 | 218.6683 | 32 | 234.6686 | 605.3462 | 6743.6361 | 3575.1141 | 11762.1745 | 57.3334 | 79.3336 | 640.681 | 3783.8346 | 122285.8304 | -3392.933 | 12.6667 | 309.3367 | 1237.3869 | 1.3333 | 3.3333 | 16 | 54.7598 | 1.3333 | 1421.4041 | 1190.7164 | 219852.1767 | 269126.1513 | 315132.5781 | 259127.0575 | 237578.0783 | 24873.6706 | 59.3335 | 149.3341 | 2194.0655 | 4862.3288
47 | 5_18_23_MFB |  | 06/16/2023 16:53:52 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_MFB.696 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 15892.1698 | 7135.1154 | 762.687 | 21548.2429 | 6508.1493 | 9615.2349 | 4736.1203 | 10094.899 | 13367.5877 | 35402.4843 | 24161.453 | 57289.9841 | 21407.3625 | 24034.8688 | 6274.0442 | 14708.2362 | 70102.2564 | 32559.7316 | 19594.0298 | 963.3658 | 26992.8133 | 3709.1485 | 625.347 | 1879.4571 | 4053.9085 | 2199.7299 | 750.0197 | 62635.0164 | 54191.2622 | 238030.4131 | 262239.7433 | 308394.5686 | 255513.517 | 232536.0075 | 25712.4526 | 30459.1044 | 72755.4659 | 108986.9226 | 247462.1911
48 | 5_18_23_DB |  | 06/16/2023 16:59:52 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_DB.697 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 54.0001 | 18 | 109.3338 | 12203.2108 | 2763.6007 | 70.6668 | 20.6667 | 31.3334 | 82.6669 | 176.6678 | 1767.8712 | 216.6684 | 9415.1015 | 76.0002 | 36.6667 | 271.3359 | 1313.3937 | 1246.7211 | -3570.3485 | 4.6667 | 117.3338 | 1292.7253 | 6 | 6.6667 | 20.6667 | 58.0119 | 3.3333 | 238.002 | 155.3342 | 230308.4754 | 274486.113 | 326597.392 | 265475.7653 | 240996.5053 | 26869.9135 | 94.667 | 223.3351 | 814.6758 | 1815.3783
49 | 5_18_23_T_239 |  | 06/16/2023 17:05:52 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_T_239.698 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 140043.7836 | 16674.3929 | 14257.1119 | 26316.2184 | 2918.2983 | 7111.1034 | 44306.6092 | 3571.7799 | 3387.7358 | 9080.2186 | 180963.3552 | 37085.408 | 108271.4323 | 6210.017 | 1967.4688 | 2059.482 | 10637.2943 | 36378.2615 | 7879.763 | 391.3387 | 12035.7359 | 3513.0987 | 224.6684 | 672.0158 | 1391.4011 | 937.6933 | 255.3356 | 12907.8297 | 10914.835 | 194973.0541 | 232993.3409 | 278031.1657 | 232347.7448 | 205432.5785 | 23385.799 | 19174.8608 | 44901.7947 | 27220.1506 | 56672.0393
50 | 5_18_23_T_239 |  | 06/16/2023 17:11:52 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_T_239.699 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 139649.2559 | 16599.6394 | 14326.5146 | 27567.9094 | 2947.6375 | 7057.7443 | 45347.1975 | 3605.1218 | 3593.1186 | 8872.7554 | 177293.3299 | 36196.4651 | 109514.8578 | 6289.3851 | 2002.1403 | 2185.5005 | 10838.1103 | 35963.2111 | 7973.1715 | 409.3392 | 11597.3723 | 3453.7508 | 204.6681 | 610.013 | 1380.7334 | 903.0186 | 250.0022 | 13093.3318 | 10770.7269 | 204143.7208 | 231693.2066 | 276709.5398 | 231500.0199 | 205744.9947 | 22539.1292 | 18978.5983 | 43995.6444 | 26310.1309 | 55187.6703
51 | 5_18_23_M_224 |  | 06/16/2023 17:17:52 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_M_224.700 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 13394.944 | 910.0291 | 1141.3789 | 39435.6937 | 3024.3201 | 969.3662 | 2882.9576 | 2112.1562 | 132.6673 | 258.669 | 1666.4813 | 60042.5908 | 8908.1109 | 69.3335 | 86.6669 | 236.002 | 1078.0407 | 6228.691 | -2312.5902 | 10 | 343.3375 | 1316.7274 | 7.3333 | 7.3333 | 22.6667 | 385.2767 | 8 | 3257.7048 | 2208.1707 | 219624.0834 | 252525.7153 | 306452.7819 | 254962.3906 | 226682.3356 | 25922.1656 | 16.6667 | 56.6668 | 444.0027 | 864.6768
52 | LBLANK |  | 06/16/2023 17:23:52 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.701 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 7.3333 | 0.66667 | 5.3333 | 2049.4804 | 2424.2066 | 52.0001 | 2 | 4 | 30.6667 | 30.6667 | 463.8152 | 8.6667 | 8230.3707 | 3.3333 | 28 | 22.6667 | 103.3337 | 506.009 | -3269.2611 | 0.66667 | 28 | 1198.717 | 0 | 1.3333 | 3.3333 | 43.3289 | 0 | 400.0056 | 10.6667 | 201553.7404 | 265923.1425 | 330118.1227 | 264892.8544 | 235002.586 | 26214.0318 | 0 | 14.6667 | 21.3333 | 42.6667
53 | CCV 1ppb TRACE |  | 06/16/2023 17:29:53 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.702 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3140.3453 | 1435.4055 | 311.3367 | 2935.6353 | 3182.3545 | 2042.146 | 922.0298 | 2114.8235 | 2745.5973 | 7417.2584 | 7368.8774 | 11657.4219 | 11936.3189 | 4891.504 | 1244.0545 | 3148.347 | 14864.3969 | 7323.8773 | 1345.7252 | 192.0013 | 5494.3899 | 1762.1087 | 147.3341 | 408.0058 | 818.6901 | 492.2994 | 143.3341 | 13418.9664 | 11385.2022 | 223263.1519 | 266509.0913 | 315013.4182 | 257338.7071 | 230993.4473 | 26008.9883 | 6177.3355 | 15051.2614 | 23197.3071 | 50028.328
54 | CCV 1ppb TRACE |  | 06/16/2023 17:36:12 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.703 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3279.7098 | 1497.4118 | 296.0031 | 2768.2683 | 3354.3941 | 2062.1489 | 966.0327 | 2116.8238 | 2736.929 | 7393.2471 | 7157.0427 | 11689.4472 | 12652.9354 | 4982.2021 | 1364.0651 | 3218.3627 | 14839.0384 | 7309.8701 | 1014.2633 | 200.6681 | 5619.1052 | 1829.4509 | 111.3338 | 370.6715 | 842.0249 | 506.7997 | 145.3341 | 13370.9238 | 11497.2929 | 224805.8413 | 267475.5532 | 315418.1545 | 262800.0217 | 234862.9989 | 26553.3296 | 6256.0365 | 15105.9835 | 23867.1404 | 51126.4486
55 | LFB |  | 06/16/2023 17:42:12 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.704 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 30080.9725 | 13290.18 | 1412.7365 | 26983.4649 | 10396.4488 | 18427.8796 | 9162.9377 | 19252.2988 | 24838.911 | 65574.8245 | 46454.3936 | 106721.149 | 32979.3577 | 45783.9179 | 11684.1102 | 27609.3254 | 130741.5463 | 62065.1985 | 43746.7732 | 1884.1243 | 52114.8867 | 6518.1544 | 1225.3859 | 3626.4606 | 7836.1488 | 4462.1315 | 1367.3988 | 121974.5723 | 104473.2913 | 220002.5068 | 251315.3867 | 302674.9263 | 248180.5763 | 224792.9774 | 25123.407 | 58733.1627 | 139417.7011 | 205601.9968 | 458998.3734
56 | CCV 10 ppb MAJORS |  | 06/16/2023 17:48:11 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.705 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 31921.6252 | 14302.4894 | 1352.064 | 21974.8908 | 10683.3269 | 19859.7962 | 9564.5347 | 20315.1024 | 26096.4844 | 70287.8295 | 49313.219 | 112534.844 | 34840.4337 | 48789.8451 | 12221.893 | 28807.6859 | 137472.9591 | 66998.7464 | 44600.1696 | 1902.1266 | 53237.6817 | 6611.5301 | 1178.7153 | 3680.4742 | 8012.2463 | 4379.781 | 1371.3992 | 128922.4921 | 109834.6122 | 223989.8079 | 253676.5238 | 304676.7752 | 249380.5831 | 227631.292 | 25757.8687 | 60882.7988 | 145396.161 | 212136.8486 | 474949.5576
57 | CCV 10 ppb MAJORS |  | 06/16/2023 17:54:30 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.706 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 32340.5706 | 14386.5744 | 1420.7374 | 22458.9803 | 11036.93 | 19689.5613 | 9689.2854 | 20820.4976 | 26329.5782 | 71384.5862 | 49291.1896 | 114513.8593 | 35193.3017 | 49505.646 | 12639.59 | 29585.9413 | 139035.9942 | 67019.5105 | 46741.4308 | 1902.7934 | 53791.0828 | 6740.9241 | 1164.7142 | 3816.5099 | 8016.9156 | 4657.4021 | 1487.4108 | 130226.2395 | 111809.8636 | 225456.5945 | 259229.737 | 306434.3765 | 250498.6629 | 231380.4208 | 25567.5293 | 60722.1238 | 145878.3846 | 214521.5894 | 473337.4678
58 | 5_18_23_MFB |  | 06/16/2023 18:00:31 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_MFB.707 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 16726.4536 | 7389.912 | 828.024 | 21657.0711 | 6881.6573 | 9988.8251 | 5010.879 | 10591.9252 | 13692.5605 | 36851.4895 | 26299.9559 | 58423.8954 | 22653.2806 | 25701.1002 | 6573.5126 | 15418.9849 | 72827.2211 | 33638.907 | 20680.0205 | 1016.0364 | 27527.8364 | 3851.8526 | 584.6786 | 1892.7921 | 3965.884 | 2337.0174 | 730.0187 | 65505.2153 | 55110.0962 | 227168.7348 | 263210.706 | 299416.7669 | 259611.1886 | 234635.4219 | 26291.5097 | 31322.3131 | 74920.0417 | 111535.4125 | 242411.6111
59 | 10_1_21_W_1 |  | 06/16/2023 18:06:31 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_W_1.708 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 5923.2289 | 1948.133 | 837.358 | 9506.4966 | 6876.9885 | 942.6978 | 7712.7482 | 112.6671 | 78.0002 | 158.0009 | 37407.0552 | 65481.7325 | 29837.8014 | 45.3334 | 123.3339 | 1402.7355 | 6506.1484 | 2878.2901 | -35.4688 | 132.6673 | 3775.1655 | 1275.3904 | 3.3333 | 8.6667 | 8 | 88.5978 | 3.3333 | 45707.6727 | 39348.7838 | 222341.6447 | 262292.9322 | 306837.6297 | 256352.9768 | 231197.2607 | 26107.8416 | 13.3333 | 46.0001 | 748.6745 | 1548.0328
60 | 10_1_21_W_6 |  | 06/16/2023 18:12:31 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_W_6.709 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 5957.9094 | 2278.1816 | 1358.7313 | 6301.3895 | 5807.1804 | 843.3583 | 7501.9699 | 300.0032 | 86.6669 | 163.3343 | 46305.9469 | 36339.5003 | 34642.6261 | 42.6667 | 116.6671 | 2503.5529 | 11679.4404 | 6017.2679 | 256.5394 | 129.3339 | 4259.9686 | 1222.0524 | 4.6667 | 5.3333 | 20.6667 | 78.6464 | 2.6667 | 20881.2511 | 18076.0956 | 220208.9965 | 255835.2487 | 310648.0455 | 230935.2044 | 227870.6403 | 25965.5768 | 33.3334 | 69.3335 | 1720.7066 | 3737.5226
61 | 10_1_21_W_8 |  | 06/16/2023 18:18:32 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_W_8.710 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 5850.5318 | 2076.8177 | 582.6786 | 6615.5318 | 5270.3062 | 818.6901 | 6987.0416 | 321.337 | 99.3337 | 175.3344 | 25588.5548 | 26706.2729 | 22368.1659 | 20.6667 | 76.6669 | 2344.1924 | 11015.5792 | 4956.86 | 518.1759 | 132.6673 | 4588.7369 | 1239.3875 | 3.3333 | 10 | 14.6667 | 79.9472 | 2 | 18316.4015 | 15516.4229 | 217964.5017 | 252784.2394 | 297880.2667 | 251690.5409 | 224484.0621 | 25045.2699 | 15.3333 | 44.6667 | 824.009 | 1837.3808
62 | 10_1_21_W_9 |  | 06/16/2023 18:24:32 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_W_9.711 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 5017.5481 | 1932.1307 | 834.6911 | 6117.3095 | 4783.4676 | 758.6868 | 6211.3503 | 274.6693 | 162.0009 | 348.6709 | 27109.1019 | 19839.7678 | 23396.4833 | 24 | 62.6668 | 1653.4292 | 7718.7515 | 3014.3183 | -273.2957 | 130.6673 | 3581.1156 | 1256.0552 | 1.3333 | 5.3333 | 14 | 76.6355 | 3.3333 | 12841.1044 | 10853.4552 | 220411.3985 | 255191.7939 | 304229.5321 | 251776.7868 | 228135.3368 | 25629.6391 | 24 | 53.3334 | 804.009 | 1796.7107
63 | 10_1_21_W_10 |  | 06/16/2023 18:30:32 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_W_10.712 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4374.6699 | 2368.1963 | 4123.9291 | 9713.3014 | 4954.8592 | 752.6868 | 6646.2124 | 294.6697 | 119.3338 | 256.0023 | 69896.0165 | 41051.5655 | 46442.7046 | 51.3334 | 84.6669 | 1794.7797 | 8620.6004 | 5985.9216 | 180.8942 | 120.0005 | 4033.2359 | 1304.7263 | 1.3333 | 11.3333 | 12 | 91.3348 | 2 | 7579.3437 | 6500.8127 | 221263.8085 | 250960.0586 | 301984.7588 | 252930.7804 | 223834.7418 | 25336.4487 | 22 | 79.3336 | 2450.0812 | 5180.3634
64 | 10_1_21_W_11 |  | 06/16/2023 18:36:31 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_W_11.713 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4162.6064 | 2270.1804 | 2905.6291 | 9147.5945 | 4890.8373 | 746.0195 | 6388.0953 | 337.3373 | 168.6677 | 312.0034 | 81595.7213 | 42687.6844 | 52082.1056 | 53.3334 | 89.3336 | 2256.1785 | 10999.5668 | 7069.7501 | 368.3581 | 136.6673 | 4259.3017 | 1296.0588 | 2.6667 | 9.3333 | 12 | 53.8982 | 4.6667 | 7170.4663 | 6073.9578 | 221139.1867 | 254335.5104 | 302494.4431 | 252005.3675 | 226552.985 | 25152.7947 | 26 | 85.3336 | 4888.3296 | 10348.1444
65 | 10_1_21_W_11_LD |  | 06/16/2023 18:42:31 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_W_11_LD.714 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.3333 | 0 | 2.6667 | 455.3406 | 374.0049 | 18.6667 | 0.66667 | 0.66667 | 14.6667 | 24 | 374.2797 | 10.6667 | 3563.1118 | 1.3333 | 14 | 13.3333 | 69.3335 | 499.3421 | 189.8128 | 1.3333 | 6.6667 | 278.6694 | 0 | 0 | 2 | 30.2912 | 0 | 153.3342 | 4 | 259144.9361 | 320374.5675 | 365527.382 | 309786.5814 | 275561.4134 | 33300.1863 | 0.66667 | 4.6667 | 8 | 25.3333
66 | 10_1_21_W_11_LD |  | 06/16/2023 18:48:48 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_W_11_LD.715 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4217.956 | 2340.1918 | 3061.6615 | 7487.2956 | 4851.4905 | 762.0203 | 6591.5241 | 322.0037 | 170.6677 | 384.6718 | 80946.5196 | 43528.216 | 52636.1339 | 64.6668 | 83.3336 | 2215.5052 | 10747.3765 | 7028.3955 | 472.7362 | 168.001 | 4475.3679 | 1246.7211 | 3.3333 | 8 | 9.3333 | 65.3832 | 1.3333 | 7260.5128 | 6204.6806 | 213700.7956 | 248734.0528 | 294694.6534 | 243946.5561 | 218436.1539 | 25387.2046 | 32.6667 | 90.667 | 4859.6529 | 10238.7791
67 | 10_1_21_W_11_DUP |  | 06/16/2023 18:54:48 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_W_11_DUP.716 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4245.2978 | 2275.5148 | 3086.3335 | 7325.2125 | 4718.1125 | 693.3502 | 6300.0559 | 303.3366 | 112.6671 | 216.6683 | 83396.0373 | 44408.2567 | 53427.0616 | 46.6667 | 74.6669 | 2246.1769 | 10810.0902 | 7065.7483 | 381.0018 | 138.0007 | 4180.6117 | 1180.7155 | 3.3333 | 11.3333 | 20 | 102.8414 | 2.6667 | 6952.3587 | 5929.8976 | 215824.816 | 243311.1972 | 294246.7851 | 245492.7531 | 217355.1481 | 24546.4058 | 26.6667 | 102.667 | 3050.1273 | 6269.8779
68 | LBLANK |  | 06/16/2023 19:00:48 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.717 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.3333 | 0.66667 | 8 | 1088.0415 | 2430.2067 | 34.6667 | 0 | 4 | 32 | 34.6667 | 404.8588 | 12.6667 | 7986.8987 | 0.66667 | 24.6667 | 33.3334 | 95.3337 | 447.3403 | -3109.9699 | 0 | 8.6667 | 1190.0497 | 0 | 0.66667 | 0.66667 | 57.3886 | 2.6667 | 78.6669 | 14 | 220811.6288 | 252707.7376 | 303732.6637 | 248627.6351 | 221990.164 | 24474.952 | 0.66667 | 14 | 18.6667 | 35.3334
69 | CCV 1ppb TRACE |  | 06/16/2023 19:06:48 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.718 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3030.3222 | 1354.731 | 314.6701 | 2616.9065 | 3215.6954 | 1863.4549 | 891.3612 | 2027.4773 | 2636.2433 | 7183.8081 | 7065.7179 | 10797.4125 | 11427.237 | 4741.4541 | 1224.0524 | 2884.958 | 14110.9684 | 6996.3796 | 1011.9497 | 198.0014 | 5170.9358 | 1664.7637 | 119.3338 | 360.6712 | 774.6877 | 455.0722 | 154.6675 | 12662.2772 | 10794.0766 | 213475.1851 | 251145.0825 | 297457.6162 | 245559.1745 | 217306.6575 | 24708.0163 | 5972.582 | 14277.1313 | 23208.6924 | 48776.55
70 | CCV 1ppb TRACE |  | 06/16/2023 19:13:07 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.719 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3118.3405 | 1474.0761 | 299.3365 | 2584.9007 | 3333.0556 | 1929.4638 | 927.3635 | 2003.4738 | 2701.589 | 7156.4604 | 7209.9577 | 11012.2434 | 11620.7252 | 4734.1178 | 1228.7196 | 3038.9899 | 14170.3584 | 7025.7273 | 1269.6917 | 178.0011 | 5259.6349 | 1688.7665 | 123.3339 | 355.3378 | 773.3544 | 488.9497 | 148.0008 | 12780.3845 | 10718.0195 | 216313.3968 | 251321.4315 | 299332.4744 | 247438.8538 | 222111.5086 | 25124.0735 | 5953.2405 | 14703.5632 | 23677.6591 | 49988.3901
71 | CCV 10 ppb MAJORS |  | 06/16/2023 19:19:08 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.720 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 29628.0257 | 13262.8208 | 1307.3932 | 20098.7969 | 10193.6368 | 18098.1246 | 8981.4896 | 18820.3894 | 24044.2194 | 65589.566 | 45413.4101 | 103598.2861 | 32045.2396 | 44386.8557 | 11273.1134 | 26756.3708 | 127304.0702 | 61913.8867 | 43375.868 | 1781.4445 | 49302.2628 | 6402.1028 | 1082.7077 | 3433.0793 | 7371.9024 | 4438.0443 | 1270.7232 | 119373.3487 | 101477.1493 | 211084.8035 | 234238.106 | 285234.9467 | 236019.0841 | 210393.4368 | 23924.6837 | 56274.6224 | 135653.0797 | 210160.8325 | 458453.6631
72 | CCV 10 ppb MAJORS |  | 06/16/2023 19:25:27 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.721 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 30535.9419 | 13327.5485 | 1336.0626 | 20757.0697 | 10458.494 | 18622.1301 | 9281.0144 | 19661.5215 | 25474.6945 | 68428.1727 | 47078.0867 | 108691.9867 | 33692.352 | 46494.8746 | 11743.4916 | 27767.6272 | 130965.5997 | 63734.5243 | 44351.1479 | 1852.7868 | 51018.9412 | 6515.4855 | 1160.7138 | 3604.4547 | 7632.0383 | 4599.8194 | 1419.4039 | 123067.843 | 104752.0022 | 216719.5152 | 240521.9679 | 291033.3941 | 241381.0105 | 215514.8586 | 24525.704 | 58181.5772 | 138119.1533 | 208989.658 | 460605.4908
73 | 10_1_21_W_11_DUP_LSP |  | 06/16/2023 19:31:27 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_W_11_DUP_LSP.722 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 63950.1817 | 27839.1001 | 6324.7333 | 49397.9374 | 19953.9259 | 37320.0217 | 24199.8137 | 38929.643 | 49272.8297 | 130111.9745 | 171583.0137 | 255578.6978 | 103277.9789 | 90235.4317 | 22753.4399 | 54500.4324 | 254738.5524 | 123133.7556 | 97267.6307 | 3614.4573 | 104876.2305 | 11357.8467 | 2257.5119 | 6891.6625 | 14900.4344 | 9018.4742 | 2721.5926 | 244866.1699 | 211303.2497 | 214672.9548 | 235448.4978 | 282882.0065 | 233593.0776 | 208420.0414 | 23512.6675 | 109982.4094 | 262853.014 | 371395.5631 | 858119.506
74 | 10_1_21_F_1 |  | 06/16/2023 19:37:27 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_F_1.723 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 26819.8196 | 8354.4426 | 525.343 | 14742.2716 | 21829.3327 | 4220.6237 | 34206.9081 | 442.0068 | 222.0017 | 489.3417 | 50350.317 | 186750.7577 | 37991.7859 | 134.0006 | 516.0093 | 2292.1842 | 11261.7717 | 86301.2273 | 10223.207 | 484.0082 | 14796.3254 | 1328.7285 | 7.3333 | 15.3333 | 29.3334 | 242.6114 | 8 | 200880.512 | 178568.4632 | 205124.6338 | 238381.1234 | 279514.2026 | 232699.2542 | 214816.4371 | 23883.9498 | 58.6668 | 128.0006 | 1306.0234 | 2760.7711
75 | 10_1_21_F_1_DUP |  | 06/16/2023 19:43:27 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_F_1_DUP.724 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 28160.3977 | 8758.0171 | 491.3418 | 13754.6192 | 22199.2371 | 4372.6697 | 35079.6881 | 421.3396 | 192.0013 | 479.3414 | 36555.7565 | 191041.5872 | 30002.1445 | 161.3343 | 575.3449 | 2135.493 | 10939.5209 | 83161.3526 | 10566.1522 | 490.6751 | 15282.1698 | 1336.7293 | 5.3333 | 21.3333 | 38.6667 | 217.782 | 10.6667 | 206328.7425 | 182798.1338 | 204207.2745 | 236946.9859 | 275806.337 | 229754.8895 | 214827.3163 | 23660.2446 | 54.6668 | 119.3338 | 1281.3558 | 2652.7611
76 | 10_1_21_F_6 |  | 06/16/2023 19:49:26 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_F_6.725 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 28228.5316 | 9924.7801 | 495.3419 | 6638.2095 | 17072.8635 | 3839.1825 | 35212.6801 | 1153.3799 | 206.0015 | 486.0083 | 32803.4779 | 116360.6457 | 27958.6669 | 99.3337 | 348.0043 | 5694.4688 | 26914.6641 | 31495.3476 | 12860.727 | 555.3442 | 17957.2794 | 1277.3907 | 8 | 32 | 52.6668 | 219.1316 | 13.3333 | 94494.857 | 84093.5111 | 205881.6289 | 239090.1854 | 277634.1111 | 230949.2563 | 213765.6342 | 23623.5168 | 82.0002 | 214.0016 | 2129.3946 | 4523.6131
77 | 10_1_21_F_6_DUP |  | 06/16/2023 19:55:26 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_F_6_DUP.726 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 27600.6405 | 9882.0845 | 514.0093 | 6469.4646 | 17078.8713 | 3737.1555 | 34029.8167 | 1180.0489 | 228.0018 | 476.6746 | 32723.4016 | 114836.3918 | 27996.0732 | 100.667 | 304.6699 | 5645.7828 | 27024.2042 | 35294.2118 | 12932.9633 | 566.0112 | 17783.0617 | 1294.0587 | 9.3333 | 23.3334 | 63.3335 | 259.0785 | 10 | 92560.9009 | 82996.4069 | 204082.9322 | 235478.2253 | 275983.7174 | 228008.5396 | 213804.7565 | 23679.6103 | 76.0002 | 164.6676 | 2266.0704 | 4756.3128
78 | 10_1_21_F_8 |  | 06/16/2023 20:01:27 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_F_8.727 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 29440.9746 | 10046.8668 | 484.0082 | 6717.5801 | 15804.0735 | 3837.8491 | 33904.8601 | 1495.4116 | 184.0012 | 390.672 | 37737.3241 | 109960.2521 | 30272.0403 | 90.0003 | 266.0025 | 6206.6815 | 29809.0724 | 33971.0118 | 15329.2395 | 664.0155 | 20478.6717 | 1372.0659 | 9.3333 | 30.6667 | 70.6668 | 233.8148 | 10.6667 | 83941.941 | 74992.3298 | 204764.4579 | 236861.4969 | 277169.6527 | 232381.5749 | 214158.7051 | 23370.4403 | 48.6668 | 160.6676 | 4446.9374 | 9379.8707
79 | 10_1_21_F_9 |  | 06/16/2023 20:07:27 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_F_9.728 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 25138.766 | 9531.8456 | 510.6758 | 8059.607 | 13878.0717 | 3443.0816 | 30422.3653 | 1218.052 | 316.0035 | 796.6889 | 27144.3513 | 54405.4046 | 24412.843 | 85.3336 | 270.6693 | 5845.8628 | 27400.2525 | 41480.1373 | 11338.6887 | 515.3427 | 15920.8696 | 1413.4033 | 9.3333 | 34.6667 | 60.0001 | 231.9066 | 16.6667 | 58458.0319 | 51897.4321 | 207947.347 | 238050.3564 | 279222.4364 | 233246.0702 | 215219.2342 | 23671.5984 | 58.6668 | 164.0009 | 11939.9317 | 24589.6149
80 | 10_1_21_F_9_DUP |  | 06/16/2023 20:13:27 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_F_9_DUP.729 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 24276.6112 | 9334.382 | 612.6798 | 8073.6149 | 13525.0672 | 3369.0647 | 29611.3274 | 1254.7218 | 269.3359 | 712.0178 | 25852.2112 | 54040.0184 | 23487.961 | 149.3341 | 301.3365 | 5846.5303 | 27000.8284 | 41397.2295 | 10836.2667 | 517.3427 | 15812.0793 | 1352.7307 | 5.3333 | 20 | 62.0001 | 212.5512 | 12 | 58342.9051 | 51170.8232 | 204967.39 | 237165.7943 | 279319.6481 | 231094.1699 | 213783.2324 | 23489.9637 | 56.6668 | 158.6675 | 2071.3913 | 4353.5921
81 | 10_1_21_F_10 |  | 06/16/2023 20:19:27 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_F_10.730 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 21917.4689 | 11336.4974 | 1027.3703 | 9847.3936 | 13800.6647 | 3504.4299 | 31806.0347 | 1347.397 | 172.6677 | 393.3388 | 24853.861 | 58109.2862 | 23302.9965 | 92.0003 | 227.3351 | 4374.6707 | 20495.3606 | 44715.206 | 14365.1209 | 584.6787 | 18257.6596 | 1500.7455 | 7.3333 | 21.3334 | 76.0002 | 239.2347 | 8.6667 | 35320.2769 | 31358.3804 | 216898.2326 | 234249.6623 | 273951.1177 | 229752.1803 | 213063.8334 | 23677.6079 | 42.6667 | 125.3339 | 403596.3132 | 912439.7857
82 | 10_1_21_F_11 |  | 06/16/2023 20:25:27 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_F_11.731 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 21249.795 | 11227.7458 | 854.0255 | 7826.8107 | 13922.115 | 3400.4047 | 31287.563 | 1367.3988 | 255.3356 | 638.0143 | 30469.4384 | 50857.702 | 26098.4874 | 94.0003 | 238.6687 | 4137.9327 | 20274.3769 | 42267.7711 | 13651.8243 | 581.3452 | 19048.0244 | 1336.7293 | 9.3333 | 29.3334 | 66.0002 | 188.5237 | 8.6667 | 33607.489 | 29475.711 | 203191.7289 | 234705.0101 | 272726.1868 | 227456.5354 | 210235.6963 | 23301.6594 | 46.0001 | 110.6671 | 2330.0755 | 4805.6561
83 | LBLANK |  | 06/16/2023 20:31:27 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.732 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 5.3333 | 0.66667 | 5.3333 | 384.0052 | 2482.8825 | 40.6667 | 0.66667 | 0.66667 | 32.6667 | 34 | 462.2102 | 13.3333 | 7954.2145 | 0 | 22.6667 | 16 | 104.6671 | 418.6728 | -3519.5924 | 0 | 7.3333 | 1187.3828 | 0 | 2 | 0 | 47.4372 | 0.66667 | 54.6668 | 18.6667 | 217429.1389 | 255193.3839 | 301075.6912 | 244986.8911 | 222540.6439 | 24425.535 | 0 | 10 | 22.6667 | 34.6667
84 | CCV 1ppb TRACE |  | 06/16/2023 20:37:27 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.733 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3151.0149 | 1460.0747 | 298.0031 | 2592.9022 | 3178.3537 | 1979.4706 | 862.0261 | 2090.8197 | 2662.2481 | 7183.1408 | 7110.8918 | 11041.5989 | 11604.7127 | 4748.1223 | 1195.3834 | 2972.976 | 13947.4724 | 6932.3498 | 1109.2068 | 192.0013 | 5268.3048 | 1696.1008 | 114.0005 | 356.6711 | 762.0203 | 471.1376 | 128.0006 | 12160.5069 | 10599.9311 | 213637.239 | 251752.975 | 293376.5585 | 243083.4035 | 221706.6148 | 24801.5115 | 5763.8292 | 14177.0321 | 22907.171 | 47671.1158
85 | CCV 1ppb TRACE |  | 06/16/2023 20:43:46 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.734 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3145.0129 | 1486.0773 | 288.6696 | 2522.2227 | 3293.0463 | 1966.1353 | 969.3663 | 2042.8129 | 2652.2463 | 7131.1131 | 7189.4368 | 11225.0751 | 11775.5184 | 4737.4521 | 1293.3919 | 3010.3176 | 13949.4749 | 6933.6823 | 1052.2783 | 176.0011 | 5284.3107 | 1746.7736 | 128.0006 | 346.6709 | 750.0197 | 472.2661 | 138.6673 | 12361.3461 | 10812.0907 | 218104.9314 | 250040.5897 | 295752.8513 | 244913.0413 | 218465.805 | 24658.5982 | 5914.5575 | 14088.9442 | 22766.4501 | 48247.2114
86 | LFB |  | 06/16/2023 20:49:46 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.735 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 28923.9189 | 12859.1191 | 1304.7264 | 26067.0943 | 9983.4883 | 17596.1694 | 8683.305 | 18156.8649 | 23617.5073 | 62719.3789 | 43788.7618 | 100441.884 | 30942.1431 | 42947.1346 | 10931.5142 | 25728.4816 | 122094.8686 | 58239.8081 | 42173.0007 | 1782.1112 | 48052.0182 | 6214.0186 | 1078.0407 | 3397.7375 | 7297.8647 | 4253.2307 | 1284.7245 | 113329.1111 | 97383.4684 | 205684.7127 | 234097.2258 | 282494.6079 | 231977.0383 | 208758.9221 | 23774.4335 | 54261.5258 | 130002.8568 | 199091.8657 | 429599.4265
87 | CCV 10 ppb MAJORS |  | 06/16/2023 20:55:46 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.736 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 30382.2807 | 13485.0298 | 1368.7323 | 21257.8046 | 10231.6634 | 18419.8688 | 9146.2607 | 19417.8547 | 24790.8258 | 66453.5632 | 45603.0382 | 105811.756 | 32128.7559 | 44957.9655 | 11383.8675 | 26624.7879 | 127035.6767 | 62181.7028 | 43677.1368 | 1826.7835 | 49516.3368 | 6320.7326 | 1101.3758 | 3485.7587 | 7437.9359 | 4544.4578 | 1323.3947 | 117868.306 | 101906.1877 | 214245.3488 | 234992.3377 | 280397.7682 | 234565.6582 | 211960.8813 | 23856.5699 | 55893.8009 | 133466.5692 | 206109.5704 | 449133.5821
88 | 10_1_21_UF_1 |  | 06/16/2023 21:01:46 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_UF_1.737 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 29150.3782 | 8230.3707 | 173.3344 | 12175.8537 | 21064.1848 | 4195.9503 | 34173.4963 | 438.0068 | 237.3353 | 427.3398 | 4281.2054 | 179396.8824 | 11832.2338 | 130.6673 | 460.6741 | 1556.0849 | 7750.7693 | 70908.2138 | 9457.2654 | 468.6744 | 13983.509 | 1339.3961 | 6 | 16 | 36 | 223.2235 | 8.6667 | 195951.4114 | 171842.7459 | 206878.5613 | 232866.6779 | 273077.9848 | 226776.564 | 209068.9974 | 23713.6671 | 60.0001 | 112.0004 | 693.3401 | 1520.0314
89 | 10_1_21_UF_6 |  | 06/16/2023 21:07:46 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_UF_6.738 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 31087.7895 | 9977.4832 | 188.0012 | 8411.8092 | 17002.1143 | 3845.1845 | 34994.8159 | 1142.7124 | 261.3357 | 526.0097 | 3861.5241 | 113821.6521 | 11475.9419 | 95.3337 | 308.67 | 3722.4853 | 18664.1846 | 26187.3238 | 12443.0051 | 578.0117 | 17256.4183 | 1366.0655 | 6.6667 | 26 | 66.6668 | 231.1913 | 16 | 93075.564 | 81076.103 | 212243.7496 | 231230.4399 | 271488.7743 | 228232.0462 | 208406.5412 | 22879.6433 | 65.3335 | 156.6675 | 1215.3536 | 2628.7603
90 | 10_1_21_UF_8 |  | 06/16/2023 21:13:46 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_UF_8.739 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 30419.6954 | 9614.5678 | 189.3346 | 11336.4963 | 15276.8348 | 3753.16 | 33121.0209 | 1483.4104 | 419.3395 | 999.3683 | 6287.7799 | 101409.3327 | 12851.7796 | 107.3337 | 264.0024 | 3358.3952 | 16500.1903 | 26059.0812 | 14328.6807 | 670.0157 | 19141.4822 | 1354.0642 | 8 | 26 | 61.3335 | 237.2347 | 11.3333 | 79987.9762 | 69173.7423 | 202752.9135 | 227831.0851 | 271251.6785 | 224406.3274 | 208813.039 | 22999.8355 | 60.0001 | 114.0005 | 1161.352 | 2464.7492
91 | 10_1_21_UF_8_LD |  | 06/16/2023 21:19:45 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_UF_8_LD.740 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 30917.4204 | 9791.3547 | 297.3364 | 17361.21 | 14887.0873 | 3736.4887 | 32382.6657 | 1458.0746 | 416.6728 | 921.363 | 6273.7149 | 100543.9189 | 12476.7794 | 84.6669 | 280.6694 | 3359.0621 | 16258.5824 | 25720.4672 | 14449.1769 | 628.6805 | 19144.8206 | 1439.4059 | 6 | 30.6667 | 60.0001 | 231.1589 | 13.3333 | 79790.2136 | 69916.0049 | 202584.9798 | 229231.1851 | 270205.4583 | 219655.0453 | 204412.7913 | 23067.9435 | 55.3334 | 118.6672 | 1120.0172 | 2444.7497
92 | 10_1_21_UF_8_LD |  | 06/16/2023 21:26:03 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_UF_8_LD.741 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 30545.9579 | 9482.4807 | 297.3365 | 18124.1627 | 15057.9327 | 3831.1805 | 32481.5529 | 1514.747 | 444.0069 | 941.3643 | 6158.6279 | 100333.8049 | 12324.6475 | 86.0003 | 290.0029 | 3305.7159 | 16312.6437 | 25600.9187 | 14151.3479 | 610.0131 | 18889.8166 | 1349.3971 | 5.3333 | 29.3334 | 62.0001 | 242.5627 | 8.6667 | 79241.8371 | 69514.7217 | 201739.1843 | 226874.062 | 269808.8308 | 225407.0857 | 204114.8025 | 23126.0368 | 60.6668 | 141.334 | 1138.018 | 2420.7464
93 | 10_1_21_UF_8_DUP |  | 06/16/2023 21:32:02 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_UF_8_DUP.742 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 29935.3433 | 9430.4455 | 180.6678 | 12473.4437 | 14784.9803 | 3699.8127 | 31996.461 | 1417.4038 | 676.6827 | 1596.0892 | 9303.5047 | 99196.5492 | 13990.1813 | 98.0003 | 274.0026 | 3531.7699 | 16514.8773 | 27930.6133 | 14407.0541 | 582.6786 | 18811.0443 | 1282.0578 | 10 | 28 | 70.6668 | 253.2189 | 12.6667 | 78844.3236 | 68700.131 | 198818.6442 | 224087.8791 | 264249.7311 | 221903.6464 | 201488.4129 | 22394.206 | 49.3334 | 152.0008 | 1274.0229 | 2594.0907
94 | 10_1_21_UF_8_DUP_LSP |  | 06/16/2023 21:38:02 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_UF_8_DUP_LSP.743 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 55450.0893 | 20471.9923 | 1414.7369 | 33370.9321 | 20674.2833 | 20739.0433 | 36745.21 | 19596.7654 | 23318.3498 | 61774.6248 | 51610.4237 | 186758.1575 | 36567.4084 | 41868.6036 | 10476.5081 | 26981.4583 | 127059.8801 | 76805.924 | 56459.3977 | 2220.1727 | 62424.0965 | 6120.6444 | 1021.3699 | 3311.7175 | 6910.3377 | 4411.3712 | 1198.717 | 179570.8951 | 155840.1009 | 201124.5892 | 221769.4493 | 260233.6254 | 215018.7401 | 198589.5924 | 22183.8785 | 50533.2413 | 120588.8283 | 187892.7454 | 403005.31
95 | 10_1_21_UF_9 |  | 06/16/2023 21:44:02 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_UF_9.744 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 27294.7214 | 9177.614 | 138.0007 | 13067.3077 | 13409.624 | 3316.385 | 29938.674 | 1215.385 | 228.0018 | 392.0054 | 4983.7881 | 50303.0848 | 11913.6324 | 84.6669 | 248.6688 | 3512.4322 | 16726.4538 | 34856.4784 | 10794.0126 | 523.3429 | 15113.9935 | 1338.7296 | 7.3333 | 33.3334 | 60.6668 | 207.2719 | 12.6667 | 56770.578 | 49179.8499 | 206863.8133 | 226839.5365 | 269885.5523 | 221371.4075 | 207318.1731 | 23136.7225 | 52.0001 | 156.6675 | 1066.0155 | 2245.4017
96 | 10_1_21_UF_10 |  | 06/16/2023 21:50:03 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_UF_10.745 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 24137.0472 | 10866.1311 | 126.0006 | 11019.5822 | 13354.2397 | 3357.7281 | 30886.0198 | 1326.7284 | 704.6841 | 1729.4381 | 9422.1475 | 52822.1478 | 13978.836 | 76.0002 | 233.3353 | 2645.5786 | 12930.516 | 29503.7686 | 12666.657 | 550.0106 | 17417.278 | 1366.0654 | 9.3333 | 31.3334 | 60.6668 | 221.928 | 12.6667 | 34547.0599 | 30079.6349 | 218643.3073 | 225333.413 | 269321.7942 | 224555.9367 | 203461.9909 | 22688.002 | 35.3334 | 95.3337 | 1495.3644 | 2984.7868
97 | 10_1_21_UF_11 |  | 06/16/2023 21:56:03 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_UF_11.746 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 23008.5148 | 10984.2214 | 120.6672 | 9659.9331 | 13464.3424 | 3357.7281 | 30938.137 | 1312.727 | 198.6681 | 398.6723 | 9207.5411 | 45082.3567 | 13926.7855 | 76.6669 | 207.3348 | 2451.5438 | 12125.8111 | 30589.3827 | 12729.7609 | 580.0118 | 17895.8692 | 1367.3988 | 11.3333 | 30.6667 | 69.3335 | 184.4209 | 12 | 32299.8082 | 27639.3812 | 215727.6239 | 226880.1134 | 271269.3551 | 226255.1187 | 204109.8743 | 22993.8248 | 38.0001 | 98.667 | 1223.3542 | 2604.7597
98 | LBLANK |  | 06/16/2023 22:02:04 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.747 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 5.3333 | 2.6667 | 18 | 468.6744 | 2628.9091 | 36.6667 | 2.6667 | 1.3333 | 32 | 33.3334 | 442.7176 | 10 | 7632.7052 | 1.3333 | 30 | 13.3333 | 96.667 | 412.6726 | -3236.2215 | 1.3333 | 112.6671 | 1180.7156 | 0.66667 | 3.3333 | 3.3333 | 35.3992 | 0 | 44.6667 | 20 | 215915.3063 | 242124.6624 | 292825.4174 | 241198.5049 | 214592.4169 | 23877.9398 | 2 | 12.6667 | 12.6667 | 33.3334
99 | CCV 1ppb TRACE |  | 06/16/2023 22:08:04 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.748 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2995.6476 | 1393.4014 | 272.6693 | 2374.8642 | 3423.7436 | 1857.4541 | 910.6957 | 1989.4721 | 2546.8939 | 6927.6799 | 6876.2666 | 10573.9125 | 10854.1229 | 4600.7413 | 1216.7186 | 2880.2904 | 13709.2415 | 6602.1921 | 1144.4856 | 180.6678 | 5032.8864 | 1746.1067 | 124.0006 | 352.671 | 735.3523 | 445.1257 | 134.6673 | 11730.8148 | 10387.7756 | 212561.1809 | 241807.3418 | 283916.2958 | 236973.4264 | 209015.4544 | 23843.2157 | 5696.469 | 13397.6128 | 23481.6107 | 47498.8649
100 | CCV 1ppb TRACE |  | 06/16/2023 22:14:23 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.749 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3022.3198 | 1390.7344 | 264.6691 | 2292.1839 | 3267.7073 | 1856.1206 | 930.0303 | 2001.4737 | 2548.8941 | 6894.997 | 6770.1519 | 10602.6009 | 11125.6642 | 4781.4669 | 1238.0537 | 2849.6176 | 13669.2043 | 6593.5213 | 933.9083 | 181.3345 | 5068.2325 | 1776.7772 | 114.0005 | 338.004 | 738.0191 | 453.0559 | 148.6674 | 11887.6111 | 10323.7289 | 212830.4044 | 239561.9603 | 287831.474 | 255690.0581 | 207811.3303 | 23647.5602 | 5726.4813 | 13822.6843 | 23172.7232 | 46370.6587
101 | CCV 10 ppb MAJORS |  | 06/16/2023 22:20:23 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.750 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 28143.6982 | 12470.7763 | 1239.3875 | 19271.6574 | 9798.6929 | 17435.9693 | 8499.1944 | 18210.9385 | 23627.5262 | 63534.9724 | 43947.8694 | 99356.3303 | 30368.9171 | 42925.0657 | 10707.3445 | 25484.0452 | 121566.3859 | 59358.402 | 42005.6093 | 1662.0968 | 46815.2572 | 6161.3284 | 1054.7056 | 3224.364 | 6770.2711 | 4250.7553 | 1253.3884 | 113779.9787 | 96157.2207 | 207076.0329 | 218468.5647 | 270299.8949 | 220860.8839 | 197825.6871 | 22507.7187 | 53962.3943 | 128995.7801 | 202625.4788 | 432480.4756
102 | 10_1_21_BB |  | 06/16/2023 22:26:24 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_BB.751 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4.6667 | 1.3333 | 101.3337 | 962.6992 | 2664.2486 | 36 | 4 | 2 | 78.6669 | 201.3348 | 3306.8487 | 64.6668 | 9395.7555 | 7.3333 | 24 | 262.6691 | 1278.0573 | 1602.0899 | -3337.7573 | 3.3333 | 55.3334 | 1151.3797 | 0.66667 | 2.6667 | 2 | 49.3127 | 0.66667 | 80.6669 | 25.3334 | 218195.1558 | 240728.0062 | 294017.1978 | 239892.9251 | 209727.6287 | 24340.054 | 2.6667 | 18.6667 | 682.6729 | 1439.3616
103 | 10_1_21_FB |  | 06/16/2023 22:32:24 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_FB.752 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 594.6791 | 1.3333 | 665.3488 | 2808.276 | 2432.8741 | 45.3334 | 12 | 6 | 157.3342 | 400.0056 | 4906.6567 | 432.0065 | 9932.7854 | 11.3333 | 80.6669 | 647.348 | 3162.3505 | 26043.7279 | -3525.1759 | 2.6667 | 54.0001 | 1207.3846 | 0.66667 | 2 | 4.6667 | 45.4426 | 0 | 124.6672 | 71.3335 | 213267.6927 | 235447.7732 | 286657.5471 | 236200.0046 | 206355.0413 | 23634.2024 | 2 | 19.3333 | 650.6723 | 1342.025
104 | 10_1_21_LAB_A |  | 06/16/2023 22:38:24 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_LAB_A.753 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 8.6667 | 0 | 104.0004 | 820.0236 | 2431.5404 | 36.6667 | 6 | 1.3333 | 67.3335 | 136.0007 | 2398.1619 | 94.667 | 8173.6714 | 4 | 27.3334 | 265.3358 | 1216.7185 | 2506.2198 | -3070.5841 | 2 | 31.3334 | 1162.714 | 0.66667 | 4 | 2.6667 | 53.3777 | 0.66667 | 44.0001 | 19.3333 | 212866.2801 | 233946.7053 | 289338.1168 | 236296.2581 | 203284.539 | 23661.5802 | 0 | 12.6667 | 624.0056 | 1231.354
105 | 10_1_21_LAB_B |  | 06/16/2023 22:44:24 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_LAB_B.754 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9.3333 | 0.66667 | 88.0003 | 952.6986 | 2334.1907 | 28 | 2.6667 | 1.3333 | 148.0008 | 337.3373 | 2587.4107 | 120.6672 | 8430.4876 | 4 | 29.3334 | 288.6696 | 1380.7334 | 2284.8494 | -3129.5097 | 0.66667 | 28.6667 | 1077.374 | 0 | 0.66667 | 3.3333 | 37.3884 | 0.66667 | 35.3334 | 10.6667 | 214285.4084 | 232737.1938 | 288530.6415 | 237085.8194 | 202243.0071 | 23642.887 | 0.66667 | 16 | 1108.0168 | 2186.0655
106 | 5_18_23_DB |  | 06/16/2023 22:50:24 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_DB.755 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 112970.9811 | 13450.9973 | 11900.9562 | 19668.1967 | 2518.8891 | 5731.1495 | 37472.4169 | 3112.3391 | 2898.2946 | 7960.2176 | 156589.618 | 29857.8399 | 88092.1398 | 4992.2057 | 1602.7566 | 1759.442 | 9423.1071 | 30652.1871 | 6516.9366 | 325.3371 | 9623.2404 | 3161.6833 | 183.3345 | 516.0093 | 1118.0438 | 766.5257 | 221.335 | 10836.1084 | 8840.0678 | 189543.9762 | 193157.7526 | 244004.9463 | 206489.0341 | 171780.7009 | 19993.9824 | 16287.2808 | 38034.5745 | 24631.677 | 50549.8289
107 | 5_18_23_DB |  | 06/16/2023 22:56:42 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_DB.756 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 114304.1585 | 13841.3699 | 11693.4515 | 21229.0964 | 2411.5369 | 5797.1762 | 38439.6499 | 3002.9825 | 2853.6183 | 7828.8114 | 156724.2006 | 29961.3894 | 88423.4831 | 5165.6009 | 1561.4187 | 1732.7718 | 9583.2144 | 31055.7224 | 6607.6245 | 323.337 | 9510.4986 | 3201.0253 | 174.0011 | 534.01 | 1130.7114 | 846.5361 | 210.0015 | 11172.3671 | 8992.1632 | 184970.4632 | 194397.7609 | 248972.8791 | 205198.5688 | 171766.4695 | 20424.5949 | 16453.4703 | 38543.9273 | 25074.6296 | 51024.7441
108 | 5_18_23_T_239 |  | 06/16/2023 23:02:42 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_T_239.757 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 113920.4743 | 13640.5104 | 11925.643 | 21703.8097 | 2435.5412 | 5882.5442 | 38509.1697 | 3019.6525 | 2822.9456 | 8054.2701 | 157427.0674 | 30352.8785 | 89064.1228 | 5116.9165 | 1582.0876 | 1723.4373 | 9332.3818 | 31327.6481 | 6475.0388 | 325.337 | 9648.5907 | 3267.0402 | 169.3343 | 544.0104 | 1138.0454 | 823.9755 | 191.3346 | 11513.9716 | 8875.4231 | 186275.9332 | 194429.5228 | 246050.9063 | 208871.9738 | 170610.0704 | 20140.8548 | 16309.3045 | 38398.8763 | 24962.5075 | 51508.2523
109 | 5_18_23_T_239 |  | 06/16/2023 23:09:00 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_T_239.758 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 112313.1218 | 13412.9624 | 12001.0388 | 20982.0643 | 2540.2259 | 5851.8663 | 38410.9089 | 2994.9807 | 2840.2824 | 7782.1194 | 155476.2298 | 29680.8025 | 88263.829 | 5028.8854 | 1674.0981 | 1698.7677 | 9232.3165 | 30651.5154 | 6536.8473 | 321.337 | 9356.3965 | 3165.6842 | 155.3342 | 530.6766 | 1100.7092 | 815.7921 | 200.6681 | 11517.3098 | 8754.682 | 183970.3873 | 199381.8258 | 247156.3664 | 203305.3163 | 171066.1834 | 19937.904 | 16277.9357 | 38295.9286 | 24753.7221 | 51149.7101
110 | 5_18_23_T_239 |  | 06/16/2023 23:15:00 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_T_239.759 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 11064.9507 | 772.021 | 1006.7022 | 32727.4573 | 2826.2798 | 818.6901 | 2638.2437 | 1772.7767 | 105.3337 | 260.6691 | 1541.1573 | 50535.2255 | 7520.6458 | 69.3335 | 70.0002 | 166.6677 | 1050.7054 | 4930.1842 | -2280.1909 | 7.3333 | 264.6692 | 1238.0537 | 1.3333 | 6 | 16.6667 | 384.0944 | 4.6667 | 2827.6132 | 1804.114 | 203896.8299 | 215545.2695 | 270531.5858 | 237180.181 | 190534.3413 | 21644.3991 | 22 | 37.3334 | 397.3355 | 853.3433
111 | 5_18_23_M_224 |  | 06/16/2023 23:20:59 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_M_224.760 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 12763.0321 | 5665.7911 | 635.3475 | 19297.0258 | 5775.8342 | 8064.9428 | 4152.6035 | 8673.299 | 11122.9956 | 31014.9656 | 21984.221 | 47094.1658 | 18266.3396 | 20395.2185 | 5165.6005 | 12487.4571 | 61091.6873 | 29217.1836 | 17738.3748 | 806.6894 | 22008.2742 | 3622.4596 | 508.6758 | 1453.4073 | 3235.6998 | 2018.1369 | 600.6793 | 55551.807 | 45392.6737 | 187241.5425 | 216581.442 | 273923.2546 | 225922.5181 | 191519.9975 | 22631.9132 | 26658.8558 | 63796.1323 | 110989.6398 | 227964.2829
112 | LBLANK |  | 06/16/2023 23:27:00 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.761 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4 | 0.66667 | 8 | 2612.9056 | 2378.8647 | 32.6667 | 1.3333 | 4.6667 | 34.6667 | 23.3334 | 431.5639 | 16 | 7560.0003 | 8.6667 | 16.6667 | 11.3333 | 85.3336 | 422.6729 | -3342.9402 | 0 | 18 | 1126.0444 | 1.3333 | 2 | 0.66667 | 51.4751 | 1.3333 | 321.337 | 21.3333 | 215594.6652 | 230541.4221 | 288532.7914 | 236399.3492 | 198474.5954 | 23353.7413 | 0.66667 | 17.3333 | 19.3333 | 35.3334
113 | LBLANK |  | 06/16/2023 23:33:18 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.762 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 6 | 0.66667 | 6.6667 | 736.6857 | 2422.8722 | 32.6667 | 1.3333 | 2.6667 | 34 | 29.3334 | 456.597 | 14.6667 | 7514.6429 | 2.6667 | 16.6667 | 8.6667 | 76.6669 | 468.0078 | -3112.1083 | 0.66667 | 11.3333 | 1120.044 | 0 | 1.3333 | 0.66667 | 32.7596 | 0 | 244.6688 | 12.6667 | 213159.4411 | 228269.3137 | 304774.8757 | 236837.7964 | 197036.3344 | 23218.8548 | 1.3333 | 7.3333 | 30.6667 | 48.6667
114 | CCV 1ppb TRACE |  | 06/16/2023 23:39:19 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.763 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2752.2655 | 1258.7221 | 254.669 | 2336.8578 | 3016.3185 | 1720.7705 | 856.6924 | 1864.7884 | 2378.8648 | 6702.24 | 6523.6734 | 9782.0148 | 10662.6447 | 4228.6259 | 1082.7078 | 2658.9142 | 13112.0155 | 6336.7385 | 1079.7791 | 173.3344 | 4568.7321 | 1779.4442 | 92.667 | 318.0035 | 695.3503 | 451.0234 | 121.3338 | 11930.3132 | 9525.175 | 211419.7332 | 226471.1679 | 287143.4317 | 236891.3317 | 195948.0669 | 23145.4025 | 5625.7742 | 13543.752 | 24096.6153 | 49553.8271
115 | CCV 1ppb TRACE |  | 06/16/2023 23:45:38 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.764 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2764.9342 | 1179.3821 | 264.6691 | 2218.1722 | 3145.0131 | 1696.1008 | 903.3621 | 1815.4487 | 2411.5369 | 6604.8599 | 6666.0279 | 9657.264 | 10568.5749 | 4253.2999 | 1126.0444 | 2604.9043 | 13203.4325 | 6455.4583 | 1193.7399 | 156.0009 | 4668.0961 | 1690.1 | 112.0005 | 350.671 | 691.3501 | 463.0833 | 118.6672 | 11900.9554 | 9590.5523 | 184000.8965 | 225039.417 | 279711.9124 | 232508.1963 | 195322.2392 | 22886.3201 | 5467.714 | 13334.8876 | 23971.129 | 47646.264
116 | LFB |  | 06/16/2023 23:51:38 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.765 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 24644.5774 | 10922.8417 | 1138.7124 | 22293.3817 | 8973.4841 | 15079.9562 | 7814.1369 | 16237.8913 | 21020.7906 | 57859.6015 | 40174.4426 | 88190.0614 | 26963.4309 | 37774.5439 | 9814.7037 | 23101.9979 | 112514.0239 | 54189.2492 | 39104.3193 | 1555.4181 | 42742.5193 | 5963.2444 | 924.6966 | 2952.3051 | 6498.8114 | 3990.1462 | 1142.0457 | 106439.7371 | 87548.7865 | 206742.4862 | 209474.6286 | 261882.0107 | 217709.1266 | 184799.0158 | 21661.0774 | 51047.0432 | 122097.5741 | 193806.2825 | 419854.6623
117 | LFB |  | 06/16/2023 23:57:55 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.766 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 25125.409 | 11372.5256 | 1166.7143 | 23065.9403 | 9180.2822 | 15644.5622 | 8060.9403 | 16689.0774 | 21745.8721 | 58893.8149 | 41128.9894 | 90123.4078 | 27374.8705 | 39319.3727 | 10029.5211 | 23674.2677 | 115961.434 | 54946.1352 | 39410.3839 | 1560.0852 | 44036.4389 | 5989.2559 | 982.7005 | 3159.6828 | 6656.2173 | 4019.3598 | 1204.7175 | 108097.4955 | 88892.3891 | 203689.2196 | 211835.7965 | 265950.8103 | 220853.4208 | 190234.2682 | 21916.1337 | 51872.0117 | 124334.0677 | 195602.4955 | 422091.7558
118 | CCV 10 ppb MAJORS |  | 06/17/2023 00:03:55 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.767 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 26587.3871 | 12023.7246 | 1204.0509 | 18977.2638 | 9419.7728 | 16539.5696 | 8529.8795 | 17722.3204 | 22798.178 | 62178.356 | 43304.6483 | 94768.6362 | 29690.1543 | 41582.4422 | 10476.5077 | 24909.0362 | 120438.9194 | 58351.597 | 41337.5549 | 1565.4192 | 44900.4537 | 6167.998 | 948.0315 | 3084.3329 | 6728.2513 | 4113.1237 | 1250.7215 | 113204.7796 | 93123.8797 | 199498.7918 | 215007.9239 | 271098.8805 | 222373.471 | 191990.8429 | 22623.2326 | 53375.5281 | 128119.9532 | 206518.9583 | 440384.785
119 | 5_18_23_MFB |  | 06/17/2023 00:09:55 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_MFB.768 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9.3333 | 0 | 58.0001 | 7593.3512 | 2417.5382 | 42.0001 | 7.3333 | 4 | 54.6668 | 92.0003 | 1062.218 | 40.6667 | 8315.7534 | 5.3333 | 21.3334 | 223.3351 | 1082.7078 | 2198.8359 | -3281.4662 | 0.66667 | 14.6667 | 1252.7216 | 0.66667 | 3.3333 | 3.3333 | 36.7326 | 0 | 374.0049 | 17.3333 | 223592.8997 | 236764.5677 | 298383.3793 | 243014.8884 | 208377.4212 | 24080.9459 | 2.6667 | 12 | 435.336 | 882.6772
120 | 10_15_21_W_1 |  | 06/17/2023 00:15:55 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_W_1.769 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 5759.1612 | 1762.1088 | 632.6807 | 8984.825 | 6795.6165 | 928.6969 | 7595.3519 | 88.6669 | 108.6671 | 211.3349 | 47810.3277 | 55878.4042 | 32990.717 | 44.6667 | 108.0004 | 800.6891 | 4150.6029 | 2560.2295 | -418.994 | 108.6671 | 3098.3361 | 1255.3886 | 2.6667 | 6 | 11.3333 | 98.0773 | 1.3333 | 47945.6626 | 39491.8504 | 204898.3655 | 217310.3825 | 273595.1005 | 228692.6433 | 194188.3075 | 22828.227 | 19.3333 | 40.6667 | 794.6752 | 1542.699
121 | 10_15_21_W_6 |  | 06/17/2023 00:21:56 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_W_6.770 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 5126.2533 | 1800.7802 | 1845.4526 | 5777.8353 | 5471.7146 | 791.3553 | 6876.3213 | 228.6685 | 98.667 | 343.3382 | 71268.9933 | 35279.5096 | 44683.1104 | 44.6667 | 80.0002 | 2840.2827 | 13785.3163 | 5692.4676 | -269.3716 | 107.3337 | 3185.0217 | 1160.7139 | 3.3333 | 10 | 14.6667 | 87.4646 | 0.66667 | 26257.4446 | 21357.2862 | 203600.0275 | 214328.0093 | 269876.8198 | 223530.7375 | 192750.2981 | 22452.9663 | 42.0001 | 78.6669 | 2703.4362 | 5515.0807
122 | 10_15_21_W_8 |  | 06/17/2023 00:27:56 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_W_8.771 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 5115.5826 | 1686.7667 | 809.3563 | 5484.3864 | 5088.2394 | 736.6857 | 6619.5335 | 246.0021 | 79.3336 | 138.6673 | 34923.0262 | 23809.1595 | 25321.7574 | 34 | 76.6669 | 1710.7693 | 8597.253 | 3114.3395 | 108.4197 | 112.6671 | 3297.0472 | 1206.7178 | 2.6667 | 6 | 12.6667 | 77.3563 | 4 | 23113.3563 | 18932.5382 | 203553.291 | 213357.473 | 265460.5259 | 224321.5326 | 190566.6289 | 22238.6298 | 13.3333 | 46.0001 | 1291.3557 | 2575.4258
123 | 10_15_21_W_9 |  | 06/17/2023 00:33:56 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_W_9.772 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4402.0115 | 1583.4211 | 1239.3871 | 5299.6498 | 4633.4181 | 636.6809 | 5833.1909 | 208.6682 | 113.3338 | 242.0021 | 38930.8794 | 21154.9926 | 27219.2412 | 26.6667 | 80.0002 | 1514.7471 | 7551.9964 | 3043.6577 | -434.3484 | 93.3336 | 2642.2444 | 1246.0544 | 1.3333 | 6 | 12 | 93.4376 | 3.3333 | 16238.5589 | 13302.1911 | 173400.8493 | 211414.2617 | 269675.6208 | 224335.0365 | 190470.0886 | 22432.9341 | 22 | 46.0001 | 1299.3565 | 2576.0914
124 | 10_15_21_W_9 |  | 06/17/2023 00:40:15 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_W_9.773 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4357.3313 | 1526.7483 | 1122.7109 | 5322.3247 | 4510.0459 | 634.0141 | 5723.1466 | 207.3349 | 108.0004 | 258.669 | 38435.7494 | 20719.015 | 26886.6131 | 28 | 57.3334 | 1558.7517 | 7612.0275 | 3072.9972 | -579.7937 | 80.0002 | 2674.2512 | 1194.7166 | 0.66667 | 4 | 10.6667 | 94.1097 | 3.3333 | 16245.9016 | 13105.3425 | 202042.7188 | 210628.1957 | 265529.2567 | 221470.3836 | 188652.8415 | 21938.8346 | 15.3333 | 42.6667 | 1269.3554 | 2626.096
125 | 10_15_21_W_10 |  | 06/17/2023 00:46:15 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_W_10.774 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3445.749 | 1778.7774 | 3583.116 | 7177.1368 | 4119.2605 | 614.0132 | 5441.7037 | 263.3358 | 115.3338 | 237.3354 | 103520.3103 | 46963.7473 | 60036.5671 | 60.6668 | 69.3335 | 2135.4931 | 10553.231 | 7471.9537 | -219.7194 | 105.3337 | 2959.6403 | 1168.0478 | 2.6667 | 7.3333 | 13.3333 | 67.367 | 0.66667 | 7258.5111 | 5927.2294 | 200447.9352 | 209412.3814 | 264989.7963 | 221720.7232 | 186431.2964 | 21636.3725 | 34 | 94.0003 | 3785.5288 | 7454.0985
126 | 10_15_21_W_10 |  | 06/17/2023 00:52:33 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_W_10.775 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3413.0744 | 1748.107 | 3653.1338 | 7028.3957 | 4126.5963 | 555.3442 | 5553.7463 | 262.0024 | 98.667 | 225.3351 | 104323.0752 | 47704.8526 | 60785.7175 | 60.6668 | 61.3335 | 2127.4918 | 10569.242 | 7626.7025 | -426.3306 | 80.6669 | 3058.9944 | 1144.7125 | 2.6667 | 5.3333 | 14 | 74.1636 | 2 | 7307.8688 | 5893.2157 | 203380.8894 | 210734.9695 | 267997.7885 | 224432.5744 | 189220.9215 | 21733.1868 | 32.6667 | 81.3336 | 3729.5212 | 7608.7891
127 | 10_15_21_W_10_DUP |  | 06/17/2023 00:58:33 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_W_10_DUP.776 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3325.7205 | 1698.101 | 3401.7384 | 7091.7602 | 4140.6007 | 585.3454 | 5439.0355 | 252.6689 | 168.001 | 415.3394 | 103095.1202 | 45991.9174 | 58660.8745 | 56.0001 | 65.3335 | 2115.49 | 10718.0196 | 7040.4011 | -476.8534 | 90.667 | 3025.6538 | 1230.7197 | 0 | 3.3333 | 7.3333 | 78.8249 | 2 | 7249.1725 | 5734.4843 | 204563.5737 | 212257.3243 | 271447.7379 | 220441.8605 | 188169.1834 | 21888.0909 | 37.3334 | 86.0003 | 3511.5043 | 7200.7118
128 | 10_15_21_W_10_DUP |  | 06/17/2023 01:04:52 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_W_10_DUP.777 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3292.3797 | 1700.1012 | 3191.6901 | 6969.7 | 4055.9091 | 579.3451 | 5173.6034 | 248.6688 | 159.3342 | 370.6715 | 101838.7293 | 45122.4881 | 57792.6648 | 54.6668 | 68.0002 | 2086.8191 | 10341.0754 | 6792.9481 | -363.2684 | 98.0003 | 3048.3255 | 1160.7139 | 0 | 6 | 16.6667 | 96.8089 | 1.3333 | 6962.3632 | 5737.8188 | 198667.8981 | 208045.9729 | 265149.3989 | 219736.5918 | 184411.9914 | 21351.2783 | 38.0001 | 78.6669 | 3385.4931 | 6932.6661
129 | LBLANK |  | 06/17/2023 01:10:52 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.778 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2 | 0.66667 | 6.6667 | 997.3682 | 2225.5069 | 35.3334 | 0.66667 | 2.6667 | 22.6667 | 24 | 376.7451 | 9.3333 | 6668.8896 | 0.66667 | 18 | 15.3333 | 60.6668 | 353.3377 | -3203.5162 | 0 | 8.6667 | 1094.0419 | 0 | 1.3333 | 0.66667 | 40.83 | 0.66667 | 100.0004 | 12.6667 | 203686.4868 | 213597.7438 | 274826.4025 | 222238.6765 | 186069.2562 | 21781.9287 | 0 | 11.3333 | 21.3333 | 33.3333
130 | LBLANK |  | 06/17/2023 01:17:12 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.779 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.3333 | 0 | 11.3333 | 509.3424 | 2240.1758 | 37.3334 | 0.66667 | 0.66667 | 26.6667 | 25.3334 | 372.9602 | 10 | 6864.983 | 2 | 20 | 14 | 77.3335 | 388.672 | -3037.9556 | 0 | 6.6667 | 1050.7054 | 0 | 0.66667 | 4 | 50.8193 | 0 | 88.6669 | 12.6667 | 212315.6613 | 213476.5702 | 269695.9722 | 221435.7216 | 184472.1457 | 21629.0293 | 0.66667 | 9.3333 | 17.3333 | 34
131 | CCV 1ppb TRACE |  | 06/17/2023 01:23:12 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.780 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2530.2241 | 1202.0509 | 257.3357 | 2065.4827 | 2852.9519 | 1685.433 | 808.0229 | 1728.7713 | 2206.8372 | 6227.3574 | 6216.8177 | 9057.5379 | 9615.9019 | 3981.8882 | 1016.0361 | 2468.8802 | 12398.712 | 5984.5877 | 865.4396 | 171.3344 | 4283.9758 | 1584.7547 | 84.6669 | 318.0035 | 635.3475 | 422.5658 | 128.0006 | 11129.0009 | 9102.2325 | 197327.5187 | 210885.8297 | 266356.2786 | 218383.2445 | 183216.0742 | 21309.2149 | 5284.9775 | 12950.5352 | 22951.8896 | 45834.1321
132 | CCV 1ppb TRACE |  | 06/17/2023 01:29:31 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.781 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2550.8945 | 1137.3787 | 242.0021 | 2107.4889 | 2943.6369 | 1649.4286 | 822.0237 | 1762.7754 | 2241.5093 | 6341.4077 | 6420.9563 | 9132.9184 | 9760.6668 | 3985.2226 | 1054.039 | 2462.8793 | 12372.0229 | 6187.3397 | 945.1926 | 166.6677 | 4319.9867 | 1614.758 | 93.3336 | 316.6702 | 656.0151 | 421.7104 | 127.3339 | 11244.4237 | 8912.7802 | 197167.3539 | 210078.1429 | 268014.1955 | 215480.9901 | 182430.8213 | 21646.3886 | 5336.3301 | 12790.3903 | 22772.355 | 47300.6153
133 | CCV 10 ppb MAJORS |  | 06/17/2023 01:35:32 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.782 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 24230.5347 | 10794.7445 | 1053.3723 | 16631.6764 | 8717.3259 | 15065.2735 | 7800.1297 | 16237.2238 | 20802.4697 | 57668.8358 | 40641.5302 | 87154.42 | 26542.6362 | 37749.1433 | 9455.1296 | 22440.2825 | 110553.4784 | 54116.3068 | 38509.1005 | 1534.0824 | 40792.1582 | 5626.4414 | 912.0292 | 2842.2831 | 6129.9817 | 3898.5301 | 1078.0407 | 104574.0439 | 86085.932 | 193907.2248 | 196667.8291 | 251590.9151 | 208486.3137 | 175430.596 | 20501.3671 | 50102.3784 | 120044.9366 | 199589.8235 | 422993.1104
134 | CCV 10 ppb MAJORS |  | 06/17/2023 01:41:51 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.783 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 25194.2024 | 11237.0845 | 1174.715 | 17112.2431 | 9027.5182 | 15444.3441 | 8136.3167 | 16924.6864 | 21285.8464 | 60629.7246 | 41984.4779 | 89630.9683 | 27610.6569 | 38995.1497 | 9927.4488 | 22997.8314 | 115482.9445 | 56014.2842 | 39804.8423 | 1546.0838 | 42190.8774 | 5852.5324 | 948.6982 | 2950.305 | 6332.0698 | 4151.0405 | 1140.7123 | 108207.6228 | 88266.5132 | 200768.2984 | 203428.9291 | 259118.1023 | 213831.8813 | 180850.2197 | 21287.182 | 51527.4564 | 123075.2709 | 203284.3743 | 428434.0173
135 | 10_15_21_W_11 |  | 06/17/2023 01:47:51 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_W_11.784 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3291.046 | 1816.1155 | 4969.5314 | 7070.4161 | 4202.6186 | 620.6802 | 5553.0792 | 306.0033 | 139.334 | 312.6701 | 136472.4731 | 52341.0604 | 78962.9646 | 76.0002 | 80.6669 | 2584.2341 | 12977.8927 | 10882.1433 | -378.8969 | 101.3337 | 3155.6824 | 1177.3818 | 1.3333 | 2 | 15.3333 | 60.6732 | 2 | 6422.7777 | 5028.2181 | 220652.5668 | 212295.18 | 266047.1578 | 222890.0609 | 188363.1115 | 22175.198 | 30.6667 | 123.3339 | 4282.9176 | 9030.4504
136 | 10_15_21_W_11 |  | 06/17/2023 01:54:10 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_W_11.785 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3200.3586 | 1764.7757 | 4736.7853 | 7330.5477 | 3999.8933 | 566.6779 | 5455.7083 | 304.6699 | 130.6673 | 296.6698 | 135087.8073 | 50537.906 | 76498.2799 | 73.3335 | 77.3336 | 2468.8803 | 12691.6357 | 10653.3049 | 152.3104 | 123.3339 | 3145.0129 | 1128.0446 | 2.6667 | 7.3333 | 8.6667 | 66.077 | 2 | 6278.7129 | 5038.8885 | 198411.044 | 206185.2847 | 261352.452 | 218750.9363 | 184408.6813 | 21879.409 | 47.3334 | 106.0004 | 4437.5987 | 8785.0573
137 | 10_15_21_F_1 |  | 06/17/2023 02:00:09 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_1.786 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 26230.065 | 7465.2841 | 310.0034 | 12577.5349 | 22516.3985 | 4275.9732 | 34388.6754 | 312.0034 | 172.6677 | 450.6738 | 47722.6703 | 168951.1873 | 31891.559 | 121.3339 | 402.6723 | 1764.7757 | 9385.7489 | 7922.8632 | 9091.5025 | 386.6719 | 11698.1216 | 1333.3956 | 8 | 20 | 36 | 247.2565 | 7.3333 | 224339.1208 | 182122.2173 | 193546.9848 | 196659.7437 | 249778.0868 | 230545.4449 | 174882.702 | 19965.2755 | 41.3334 | 128.6673 | 69185.3693 | 137503.9756
138 | 10_15_21_F_1 |  | 06/17/2023 02:06:28 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_1.787 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 27797.0261 | 7992.235 | 325.3371 | 12933.1869 | 23483.2856 | 4552.059 | 36007.9976 | 344.0041 | 180.0011 | 447.3403 | 48811.5351 | 176770.28 | 33887.4787 | 123.3339 | 370.0048 | 1818.1157 | 9490.4853 | 8145.6551 | 8987.8349 | 399.3389 | 12157.8381 | 1392.7346 | 7.3333 | 21.3334 | 38.0001 | 225.8416 | 10.6667 | 224274.1169 | 189743.1596 | 197274.196 | 204001.5929 | 250903.5038 | 208673.008 | 182555.111 | 20812.4836 | 50.0001 | 133.334 | 68496.6091 | 140385.9801
139 | 10_15_21_F_1_DUP |  | 06/17/2023 02:12:28 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_1_DUP.788 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 27395.5789 | 7742.7646 | 242.0021 | 11010.2414 | 22888.3215 | 4380.672 | 34882.5386 | 349.3376 | 156.0009 | 319.3369 | 46189.6738 | 172622.0649 | 33373.6056 | 124.6672 | 391.3387 | 1735.4389 | 8885.43 | 8063.6089 | 8649.6382 | 396.0055 | 11946.9936 | 1366.0654 | 7.3333 | 17.3333 | 40.6667 | 240.4654 | 10.6667 | 217144.5686 | 183313.3426 | 192971.307 | 207345.9406 | 253887.5612 | 208702.7761 | 185122.4165 | 21524.871 | 54.0001 | 113.3338 | 1185.3525 | 2356.7432
140 | 10_15_21_F_1_DUP |  | 06/17/2023 02:18:46 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_1_DUP.789 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 27770.9682 | 7880.1742 | 245.3355 | 11173.0361 | 22890.9935 | 4422.6846 | 35041.5929 | 302.6699 | 140.0007 | 318.6702 | 46591.2716 | 172882.501 | 33304.7803 | 114.6671 | 381.3384 | 1766.7759 | 8836.7323 | 8022.2519 | 8710.4018 | 364.0047 | 11928.9786 | 1356.0645 | 4.6667 | 26 | 33.3334 | 230.4598 | 7.3333 | 216218.6511 | 184095.9721 | 193906.7242 | 208955.1989 | 255447.0057 | 211483.2433 | 186161.798 | 21309.2176 | 45.3334 | 120.0005 | 1137.3508 | 2236.7356
141 | 10_15_21_F_6 |  | 06/17/2023 02:24:45 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_6.790 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 26167.9462 | 8853.4094 | 1356.7313 | 6150.6575 | 18160.8726 | 3635.796 | 33548.0128 | 956.0322 | 209.3349 | 515.3426 | 72690.2679 | 74082.9304 | 48094.8239 | 92.0003 | 308.67 | 4768.1299 | 23767.7566 | 27257.3134 | 9486.8495 | 426.0064 | 12732.3384 | 1368.7327 | 12.6667 | 30.6667 | 72.6669 | 239.8636 | 6 | 121472.305 | 103578.1423 | 193767.2872 | 208603.3735 | 252372.5941 | 233819.3169 | 187628.8413 | 21143.6375 | 74.6669 | 187.3346 | 1966.0531 | 3782.2
142 | 10_15_21_F_6 |  | 06/17/2023 02:31:03 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_6.791 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 25294.3741 | 8675.3008 | 1316.7273 | 6013.2656 | 17933.2502 | 3669.1378 | 33191.1829 | 901.3618 | 210.6682 | 526.6764 | 71221.599 | 72590.6374 | 46252.7552 | 109.3338 | 333.3372 | 4745.4549 | 22943.7445 | 27441.6703 | 8982.8055 | 398.0056 | 12785.0523 | 1280.0574 | 15.3333 | 23.3334 | 51.3334 | 222.5839 | 13.3333 | 119673.848 | 102489.6677 | 191007.8457 | 204389.0758 | 248821.558 | 206274.5545 | 183409.9169 | 20870.5678 | 76.6669 | 218.0017 | 1964.7193 | 3878.8719
143 | 10_15_21_F_6_LD |  | 06/17/2023 02:37:02 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_6_LD.792 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 25319.0848 | 8483.8527 | 1431.405 | 10025.5182 | 17401.9284 | 3511.7651 | 32639.9133 | 881.3605 | 218.6683 | 554.0108 | 59802.2368 | 72530.995 | 40312.8026 | 122.6672 | 331.3372 | 4735.4515 | 22838.2414 | 26917.3388 | 9126.6248 | 410.6726 | 12562.8559 | 1326.7283 | 8 | 29.3334 | 48.6667 | 231.7986 | 12 | 120228.4859 | 101747.0536 | 191104.4572 | 204396.6525 | 249191.3654 | 211886.5064 | 182950.7 | 21143.6366 | 79.3336 | 209.3349 | 1908.7162 | 3878.2033
144 | 10_15_21_F_6_LD |  | 06/17/2023 02:43:20 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_6_LD.793 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 25238.9434 | 8487.8548 | 1395.4016 | 11098.3124 | 17638.2159 | 3563.111 | 33141.7333 | 910.6957 | 212.6683 | 513.3426 | 59456.1373 | 71967.4995 | 40411.0856 | 132.0006 | 329.3372 | 4734.7847 | 22993.8276 | 26626.1249 | 8852.613 | 418.6729 | 12515.4801 | 1275.3904 | 8 | 26 | 65.3335 | 207.9547 | 10.6667 | 120305.1207 | 101363.6742 | 198174.0775 | 203468.0756 | 243402.743 | 204426.2661 | 181156.1891 | 20796.4606 | 86.0003 | 193.3346 | 1890.715 | 3801.5298
145 | LBLANK |  | 06/17/2023 02:49:20 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.794 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.3333 | 0 | 7.3333 | 1790.1122 | 2246.8434 | 35.3334 | 2 | 1.3333 | 28 | 34.6667 | 367.6643 | 13.3333 | 7238.5006 | 2.6667 | 23.3334 | 21.3334 | 78.6669 | 381.3384 | -3260.0339 | 1.3333 | 10 | 1200.0505 | 0 | 2 | 1.3333 | 58.1689 | 0.66667 | 62.6668 | 31.3334 | 224135.2586 | 220899.4888 | 272027.0046 | 222536.1512 | 192963.1635 | 21982.9049 | 0 | 6 | 16.6667 | 38
146 | CCV 1ppb TRACE |  | 06/17/2023 02:55:21 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.795 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2704.256 | 1258.0554 | 270.0026 | 2618.9068 | 3133.6772 | 1740.1061 | 851.3587 | 1901.4602 | 2382.1988 | 6608.8617 | 6613.6186 | 9566.5358 | 10186.9646 | 4066.5789 | 1157.3802 | 2589.568 | 12865.7916 | 6272.0433 | 731.1743 | 168.001 | 4554.7262 | 1656.7629 | 90.667 | 308.67 | 637.3476 | 429.7701 | 116.6671 | 11329.8256 | 9511.1654 | 199252.2171 | 219038.4525 | 270028.9737 | 226232.675 | 189364.1506 | 21976.8917 | 5511.7303 | 12985.2333 | 22529.0223 | 46818.1752
147 | CCV 1ppb TRACE |  | 06/17/2023 03:01:39 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.796 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2786.9385 | 1250.0548 | 254.0023 | 2242.843 | 3094.3354 | 1663.4302 | 848.6919 | 1799.4467 | 2305.5194 | 6549.5013 | 6403.5535 | 9464.4676 | 10424.4693 | 4285.9765 | 1123.3776 | 2592.9023 | 12775.7108 | 6073.9584 | 1037.3485 | 160.6676 | 4530.7186 | 1636.0937 | 82.6669 | 323.337 | 672.6825 | 390.2462 | 116.0005 | 11201.7239 | 9165.6067 | 209653.5902 | 214659.4635 | 265138.5341 | 219450.7477 | 189853.2564 | 21880.7445 | 5186.9415 | 12516.8153 | 21952.5784 | 45353.9911
148 | LFB |  | 06/17/2023 03:07:39 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.797 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 24219.848 | 10547.2262 | 1122.0442 | 21751.881 | 8888.098 | 14762.9587 | 7775.4495 | 15820.7564 | 20470.6602 | 56368.9977 | 38468.8643 | 85108.7795 | 26557.3304 | 36949.0599 | 9385.0832 | 22237.9627 | 107642.7197 | 51695.366 | 38317.7297 | 1512.0801 | 40945.262 | 5663.7895 | 986.7008 | 2924.2993 | 6236.0282 | 4009.271 | 1126.0444 | 100760.0958 | 82909.898 | 192954.7021 | 197904.1426 | 249662.7403 | 197932.1121 | 177883.3084 | 20665.605 | 48574.4459 | 116118.0238 | 189571.9001 | 397034.7566
149 | LFB |  | 06/17/2023 03:13:57 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.798 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 23956.7382 | 10958.2028 | 1160.0471 | 22210.5864 | 8816.053 | 15062.6034 | 7555.3313 | 16183.164 | 20577.4772 | 56352.267 | 38947.3463 | 85722.4519 | 26896.6347 | 37428.9697 | 9236.3185 | 22328.1038 | 108822.9127 | 51810.4486 | 38079.248 | 1466.742 | 41200.6601 | 5720.4785 | 910.0291 | 2976.9768 | 6300.7225 | 3953.6406 | 1103.376 | 100258.5966 | 84140.4048 | 195712.8745 | 199577.7699 | 250319.4151 | 205886.4019 | 177171.2535 | 20783.7753 | 48640.6767 | 116927.9231 | 187494.3804 | 397163.9171
150 | CCV 10 ppb MAJORS |  | 06/17/2023 03:19:56 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.799 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 25400.5651 | 11416.5602 | 1125.3777 | 18364.4645 | 9040.861 | 15914.8636 | 8185.0124 | 17098.2281 | 21665.7511 | 59673.0431 | 42114.6467 | 90203.8937 | 28365.4697 | 39187.6811 | 9813.3697 | 23412.5038 | 114282.6599 | 55653.5319 | 39365.8013 | 1607.424 | 42097.269 | 5989.2553 | 927.3636 | 2954.3056 | 6351.4121 | 4082.6297 | 1195.3833 | 106103.2535 | 88296.6981 | 198631.3787 | 201385.5535 | 254420.2451 | 209643.7831 | 180958.9035 | 20809.817 | 50489.7323 | 120294.4225 | 197409.7371 | 417080.4654
151 | CCV 10 ppb MAJORS |  | 06/17/2023 03:26:14 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.800 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 25888.7739 | 11591.3709 | 1174.715 | 18120.1557 | 9222.9768 | 16079.7113 | 8269.0604 | 17572.1354 | 22079.0494 | 60687.9683 | 42069.6939 | 91895.3095 | 28469.0065 | 40247.2969 | 10054.8708 | 23731.6955 | 114811.5195 | 56158.1695 | 39578.0653 | 1556.7515 | 42586.0535 | 5891.2151 | 923.3632 | 2938.3023 | 6425.4456 | 3949.5658 | 1163.3807 | 107774.3436 | 89045.3332 | 203321.6472 | 202666.1177 | 254713.4193 | 211442.6284 | 182760.9652 | 21309.8854 | 50659.0032 | 121533.4551 | 197108.9677 | 419892.7218
152 | 10_15_21_F_6_DUP |  | 06/17/2023 03:32:14 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_6_DUP.801 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 23173.4507 | 7818.1389 | 838.0246 | 5843.1949 | 16727.7891 | 3355.0612 | 31006.9548 | 822.6905 | 206.0015 | 540.0103 | 39528.6156 | 65875.537 | 28966.0036 | 124.6672 | 330.0038 | 4574.0661 | 22505.7141 | 26019.6753 | 8679.1313 | 406.0058 | 11814.2167 | 1272.7235 | 9.3333 | 19.3333 | 52.0001 | 215.9603 | 8 | 116142.2291 | 95964.6004 | 186482.1123 | 197330.8742 | 245336.76 | 203163.7983 | 175566.9552 | 20475.9977 | 94.0003 | 204.0015 | 2137.3958 | 4428.2652
153 | 10_15_21_F_6_DUP |  | 06/17/2023 03:38:32 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_6_DUP.802 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 23160.0956 | 7792.1246 | 868.6932 | 5831.19 | 16504.1957 | 3313.0508 | 30773.111 | 810.6898 | 226.6685 | 562.0111 | 38704.9348 | 65611.6596 | 28877.1612 | 130.6673 | 304.6699 | 4523.3828 | 22217.2648 | 26043.0504 | 9048.3577 | 402.0057 | 11743.4926 | 1223.3857 | 4.6667 | 22.6667 | 64.6668 | 225.82 | 9.3333 | 113539.4254 | 95322.3137 | 173656.9988 | 194276.2278 | 246234.6617 | 195043.7801 | 176240.4757 | 20234.3202 | 90.0003 | 216.0016 | 2202.0678 | 4498.9422
154 | 10_15_21_F_6_DUP_LSP |  | 06/17/2023 03:44:32 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_6_DUP_LSP.803 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 44539.3299 | 17488.0319 | 1846.1193 | 25500.7419 | 21689.7884 | 17803.7542 | 34936.0003 | 16755.1536 | 20232.3183 | 55286.7786 | 73731.4596 | 143973.1968 | 46200.5912 | 36423.0419 | 9184.2851 | 24794.8329 | 120944.4941 | 71385.2557 | 48178.0135 | 1762.7756 | 49903.012 | 5621.1061 | 894.028 | 2792.9397 | 5961.911 | 3954.8592 | 998.7017 | 201414.6206 | 169434.8774 | 186851.4267 | 190635.5384 | 237305.4434 | 196237.3754 | 172860.2828 | 19682.2178 | 45374.6155 | 108947.8593 | 182390.2169 | 382538.1728
155 | 10_15_21_F_6_DUP_LSP |  | 06/17/2023 03:50:52 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_6_DUP_LSP.804 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 46175.1767 | 17839.7995 | 1949.4664 | 27172.4869 | 22179.2038 | 18298.3799 | 36318.1086 | 17277.1082 | 21114.2589 | 56190.9789 | 74917.9976 | 148083.5611 | 47385.7934 | 37657.5682 | 9435.1151 | 25572.8735 | 122893.0646 | 72222.1263 | 48393.7153 | 1769.443 | 50204.7313 | 5669.7923 | 928.6969 | 2834.948 | 5909.2221 | 3998.6926 | 1086.7081 | 204016.5173 | 170761.1958 | 188808.8838 | 192455.744 | 237971.7547 | 197266.0206 | 174389.98 | 19127.4725 | 46438.6932 | 110611.9313 | 183236.7684 | 385080.4504
156 | 10_15_21_F_9 |  | 06/17/2023 03:56:52 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_9.805 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 21225.0909 | 7748.7679 | 626.0137 | 6950.357 | 13665.2012 | 3047.6585 | 28230.5393 | 940.031 | 261.3357 | 715.3513 | 31051.2746 | 40297.4235 | 24591.1473 | 78.0002 | 250.6689 | 5407.6904 | 25841.3515 | 21944.1754 | 8191.4914 | 402.0057 | 11484.615 | 1214.0517 | 6 | 24 | 48.0001 | 190.4805 | 8 | 75371.6597 | 63364.2227 | 188708.3298 | 197495.1606 | 242370.1145 | 201409.8741 | 177979.1903 | 20403.2266 | 62.0001 | 131.3339 | 1835.3791 | 3692.8514
157 | 10_15_21_F_9 |  | 06/17/2023 04:03:12 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_9.806 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 21121.6038 | 7652.0494 | 618.6801 | 5961.9105 | 13416.2979 | 2961.6411 | 27835.0947 | 901.3618 | 252.0022 | 699.3505 | 30745.3354 | 40361.6049 | 23772.43 | 65.3335 | 237.3353 | 5278.3089 | 25976.2653 | 21442.0805 | 8302.6712 | 392.0054 | 11394.5426 | 1298.7257 | 6 | 20.6667 | 36.6667 | 206.5782 | 10.6667 | 74785.2483 | 62817.8243 | 187910.4881 | 196260.9633 | 242056.1358 | 200537.7646 | 176582.665 | 20184.9177 | 73.3335 | 136.6673 | 1866.0477 | 3747.5252
158 | 10_15_21_F_10 |  | 06/17/2023 04:09:13 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_10.807 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 16236.5571 | 8395.1327 | 674.0159 | 8425.1504 | 11094.3083 | 2661.5813 | 26027.0214 | 1022.0366 | 300.6699 | 792.6887 | 28018.6977 | 43546.9387 | 22771.47 | 74.0002 | 343.3375 | 3161.6832 | 15723.984 | 17521.4069 | 9277.2789 | 408.6725 | 12550.1788 | 1248.7213 | 12 | 22 | 46.0001 | 180.6209 | 4 | 32815.6476 | 27877.1811 | 186444.4845 | 187557.0127 | 242890.1485 | 203588.5054 | 174248.9405 | 20114.1506 | 39.3334 | 92.0003 | 4594.2897 | 9128.4653
159 | 10_15_21_F_10 |  | 06/17/2023 04:15:32 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_10.808 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 16004.9606 | 8059.6063 | 708.0176 | 8161.6642 | 10952.1977 | 2593.5689 | 25572.8691 | 975.3667 | 292.6697 | 788.0218 | 29004.5293 | 42473.7227 | 22242.6357 | 62.0001 | 333.3372 | 3172.3523 | 16342.6925 | 17949.942 | 9337.6278 | 406.0058 | 12249.2514 | 1192.0498 | 8.6667 | 22 | 52.0001 | 165.9107 | 10 | 33894.8758 | 27318.0973 | 185670.1133 | 193683.5081 | 216102.7213 | 199739.3711 | 175076.9819 | 20168.8991 | 31.3334 | 91.3336 | 4550.2863 | 9153.1574
160 | 10_15_21_F_10_DUP |  | 06/17/2023 04:21:32 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_10_DUP.809 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 15713.3039 | 8084.2877 | 739.3525 | 7660.0533 | 10573.2464 | 2660.9145 | 25894.7827 | 1008.0357 | 232.0019 | 553.3441 | 27520.6571 | 41373.1629 | 21978.2345 | 73.3335 | 165.3343 | 2927.6336 | 14467.323 | 14508.6972 | 7607.4621 | 392.6721 | 10836.1088 | 1320.0611 | 4 | 16.6667 | 34 | 161.9701 | 5.3333 | 31838.1065 | 26714.2888 | 189429.7684 | 192149.0368 | 246129.6734 | 204096.3558 | 173793.4784 | 19722.9391 | 40.6667 | 78.6669 | 7392.0803 | 14267.4569
161 | 10_15_21_F_10_DUP |  | 06/17/2023 04:27:51 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_10_DUP.810 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 15466.3691 | 7998.2386 | 694.0169 | 7555.998 | 10489.1828 | 2609.5717 | 25627.64 | 977.3668 | 237.3354 | 552.0107 | 27499.6594 | 41191.9691 | 21886.0881 | 56.0001 | 172.6677 | 2862.9536 | 14255.7768 | 14403.258 | 7467.7832 | 349.3376 | 10740.7029 | 1221.3856 | 4 | 16.6667 | 34.6667 | 194.0247 | 6.6667 | 32243.6828 | 26626.7915 | 192713.4984 | 192140.8715 | 240841.9472 | 201098.9556 | 171937.2976 | 19877.1526 | 40.0001 | 80.0002 | 7098.0268 | 14453.5318
162 | LBLANK |  | 06/17/2023 04:33:51 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.811 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4 | 0.66667 | 6 | 315.3368 | 2166.8311 | 44.6667 | 3.3333 | 0 | 24.6667 | 31.3334 | 374.0408 | 14 | 6636.8748 | 0.66667 | 29.3334 | 16 | 116.6672 | 336.004 | -3112.6995 | 0 | 6 | 1050.0387 | 0.66667 | 2.6667 | 2.6667 | 57.4589 | 0 | 59.3335 | 16 | 199352.0712 | 204548.7001 | 259672.8039 | 212781.108 | 178507.6969 | 21126.9495 | 0.66667 | 8.6667 | 20.6667 | 40.6667
163 | LBLANK |  | 06/17/2023 04:40:11 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.812 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.3333 | 0.66667 | 4 | 266.6692 | 2136.8265 | 29.3334 | 0.66667 | 2 | 26 | 32.6667 | 373.6751 | 5.3333 | 6444.1201 | 0.66667 | 19.3333 | 16 | 70.6668 | 360.0045 | -3019.0825 | 0.66667 | 11.3333 | 1047.3717 | 0 | 0 | 0 | 44.738 | 0 | 47.3334 | 10 | 199958.9003 | 202728.4426 | 258474.8407 | 210818.1775 | 177057.1467 | 20707.672 | 0 | 6 | 20 | 45.3334
164 | CCV 1ppb TRACE |  | 06/17/2023 04:46:11 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.813 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2452.8774 | 1194.7166 | 234.6686 | 1922.7961 | 2786.2718 | 1575.4205 | 808.0229 | 1700.7679 | 2158.163 | 6094.6332 | 5996.317 | 8963.4782 | 9228.3139 | 3728.4871 | 1066.04 | 2382.1987 | 11986.3605 | 5887.2128 | 945.5771 | 154.0008 | 4124.5954 | 1624.7591 | 86.0003 | 294.003 | 626.6805 | 447.2228 | 109.3338 | 10721.3551 | 8585.9132 | 190370.8516 | 201549.1784 | 256809.8192 | 210705.9054 | 174259.0792 | 20409.9041 | 5092.908 | 12151.1672 | 22273.5096 | 45089.3127
165 | CCV 1ppb TRACE |  | 06/17/2023 04:52:31 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.814 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2469.547 | 1117.3771 | 229.3352 | 1965.4686 | 2925.6329 | 1621.4255 | 785.355 | 1690.1001 | 2173.4989 | 6182.0041 | 6232.4641 | 9064.8752 | 9532.5129 | 3939.8774 | 1000.7018 | 2395.5343 | 12166.5122 | 5919.2262 | 648.4623 | 138.0007 | 4198.6169 | 1673.4315 | 88.0003 | 294.6697 | 631.3473 | 360.2666 | 113.3338 | 10722.023 | 8736.0038 | 185118.4009 | 201034.6777 | 255695.4727 | 210696.4303 | 176147.9965 | 20547.4372 | 5097.5763 | 12341.9958 | 22426.2362 | 45078.3526
166 | CCV 10 ppb MAJORS |  | 06/17/2023 04:58:31 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.815 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 22806.1934 | 10225.6585 | 1019.3698 | 16012.3038 | 8401.8052 | 14498.6881 | 7525.3166 | 15403.6349 | 19895.8481 | 55015.0699 | 38312.1317 | 81098.2078 | 24840.9136 | 35188.6206 | 8924.1209 | 21235.7725 | 104889.0366 | 50925.9524 | 36754.3424 | 1424.071 | 38301.2774 | 5525.0682 | 847.3585 | 2635.5768 | 5732.484 | 3836.1393 | 1017.3697 | 99058.2561 | 80254.1319 | 186159.8284 | 184408.604 | 237596.2725 | 194845.1367 | 164992.6608 | 19545.362 | 47262.7278 | 112784.8663 | 194618.836 | 404357.268
167 | CCV 10 ppb MAJORS |  | 06/17/2023 05:04:50 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.816 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 23475.2725 | 10350.4153 | 1072.7069 | 15914.8605 | 8521.2078 | 14494.0164 | 7507.9727 | 15594.5078 | 20028.6981 | 55745.8929 | 38668.1591 | 83321.6336 | 25715.124 | 35989.9525 | 9168.2759 | 21264.483 | 106355.7703 | 52062.03 | 37715.377 | 1470.7426 | 39241.1565 | 5606.4336 | 905.3621 | 2690.9201 | 5820.519 | 3807.7891 | 1058.0392 | 100868.871 | 81395.8949 | 188959.2321 | 187782.8129 | 241779.6095 | 199005.2106 | 168244.2933 | 19477.2732 | 47933.6161 | 115078.3187 | 195523.3266 | 406133.9501
168 | 10_15_21_F_11 |  | 06/17/2023 05:10:50 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_11.817 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 14140.9966 | 7812.8031 | 662.6821 | 6410.1045 | 10232.9983 | 2460.2118 | 24175.7735 | 943.3645 | 128.0006 | 308.0033 | 25225.5401 | 34984.7861 | 20340.472 | 68.6668 | 170.6677 | 2934.9683 | 14633.4919 | 20291.735 | 9357.5449 | 404.0057 | 11929.6464 | 1201.3839 | 8.6667 | 18.6667 | 50.6668 | 218.0847 | 10.6667 | 27223.9174 | 22092.4031 | 182129.6387 | 184794.8906 | 238429.9551 | 198138.5998 | 167672.2531 | 19754.3164 | 28.6667 | 78.6669 | 1552.0329 | 3074.1292
169 | 10_15_21_F_11 |  | 06/17/2023 05:17:09 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_11.818 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 14112.9678 | 7672.7267 | 714.6846 | 6438.7845 | 10351.0824 | 2382.1987 | 23775.7701 | 983.3672 | 130.0006 | 324.6704 | 24882.3301 | 34425.4301 | 19753.6483 | 50.6668 | 172.001 | 2854.2852 | 14532.0563 | 20269.7035 | 9038.7186 | 417.3394 | 11709.4637 | 1166.0476 | 4.6667 | 25.3334 | 44.6667 | 204.7294 | 6.6667 | 27208.5592 | 21870.729 | 180979.1681 | 183052.0541 | 235529.0909 | 229846.293 | 164905.6958 | 19309.0413 | 36.6667 | 89.3336 | 1515.3644 | 2861.4448
170 | 10_15_21_UF_1 |  | 06/17/2023 05:23:09 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_1.819 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 26288.8338 | 6804.954 | 96.0003 | 8667.9624 | 20342.4751 | 3965.884 | 31745.9045 | 271.3359 | 317.3369 | 736.6857 | 5491.6603 | 149481.3365 | 10431.8079 | 100.0004 | 319.3369 | 1145.3793 | 6282.0483 | 9699.9588 | 7355.94 | 357.3378 | 9872.7438 | 1270.0565 | 2.6667 | 12.6667 | 32 | 241.3104 | 2.6667 | 198805.2722 | 161251.6462 | 181252.6531 | 183301.8357 | 233677.7124 | 193795.1117 | 164006.0642 | 19147.4905 | 42.6667 | 101.3337 | 691.34 | 1389.3606
171 | 10_15_21_UF_1 |  | 06/17/2023 05:29:28 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_1.820 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 27006.845 | 6886.3261 | 109.3338 | 9212.9703 | 20757.7369 | 4105.9233 | 32502.9335 | 287.3362 | 337.3373 | 737.3524 | 5438.5162 | 153870.905 | 10822.7649 | 109.3338 | 361.3379 | 1161.3806 | 6393.4306 | 9692.6209 | 7215.5892 | 340.6707 | 10084.8924 | 1309.3934 | 3.3333 | 14 | 28.6667 | 226.611 | 5.3333 | 200461.4209 | 165214.5616 | 186684.5401 | 185737.0912 | 235209.185 | 194777.5356 | 166703.7364 | 19241.6168 | 41.3334 | 106.0004 | 696.0065 | 1342.6913
172 | 10_15_21_UF_6 |  | 06/17/2023 05:35:27 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_6.821 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 24581.1306 | 7377.9064 | 112.0004 | 6437.4503 | 16077.0451 | 3259.0384 | 30179.8514 | 792.022 | 172.6677 | 388.0053 | 8508.0821 | 61689.5807 | 12384.0352 | 82.6669 | 261.3358 | 2709.5903 | 13874.7368 | 20072.7608 | 7746.1162 | 354.0044 | 10754.0471 | 1268.723 | 8 | 22.6667 | 49.3334 | 205.9709 | 8 | 108063.1954 | 90156.94 | 182740.1365 | 188840.6959 | 236473.7996 | 196665.1468 | 168618.6759 | 19559.3821 | 63.3335 | 186.0012 | 1013.3476 | 1878.0473
173 | 10_15_21_UF_6 |  | 06/17/2023 05:41:46 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_6.822 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 24076.2733 | 7399.9161 | 122.0005 | 6340.0741 | 15821.4249 | 3242.3685 | 29506.4407 | 747.3529 | 178.6678 | 377.3383 | 8437.7899 | 61473.9845 | 12137.154 | 78.0002 | 243.3354 | 2810.2765 | 14006.8638 | 20242.3312 | 7797.1745 | 406.0058 | 10449.8208 | 1237.387 | 6 | 27.3334 | 49.3334 | 224.6272 | 6 | 108041.0702 | 89190.8963 | 178991.4257 | 187702.4185 | 239357.2954 | 195930.4964 | 169390.3892 | 19699.5733 | 86.6669 | 181.3345 | 988.0137 | 1946.7178
174 | 10_15_21_UF_8 |  | 06/17/2023 05:47:46 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_8.823 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 25114.0583 | 7485.9614 | 247.3355 | 6017.2671 | 15321.5452 | 3351.0597 | 29303.3585 | 1007.3689 | 231.3352 | 445.3403 | 5034.7408 | 62572.0684 | 10408.4589 | 78.0002 | 255.3356 | 2505.5533 | 12963.88 | 18050.063 | 10422.8792 | 420.0063 | 12646.2622 | 1230.0531 | 10 | 22 | 45.3334 | 222.0631 | 9.3333 | 98881.7525 | 81464.2816 | 179865.2216 | 184609.844 | 234867.6569 | 193577.3932 | 166714.4945 | 19214.2467 | 45.3334 | 120.0005 | 898.6772 | 1760.7085
175 | 10_15_21_UF_8 |  | 06/17/2023 05:54:04 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_8.824 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 25108.0456 | 7543.9915 | 252.0022 | 6080.6277 | 15202.7525 | 3223.0303 | 29112.9696 | 1051.372 | 238.002 | 469.341 | 4925.1378 | 62176.3497 | 10253.0117 | 80.0002 | 240.6687 | 2508.8871 | 12791.0572 | 18030.7068 | 9945.9144 | 423.3396 | 12518.8162 | 1277.3906 | 7.3333 | 20.6667 | 55.3334 | 171.1306 | 6.6667 | 98316.5118 | 81050.6127 | 180036.1256 | 184530.8636 | 233117.384 | 195151.8135 | 164856.4671 | 18947.2249 | 42.0001 | 110.6671 | 937.3455 | 1777.3768
176 | 10_15_21_UF_8_LD |  | 06/17/2023 06:00:03 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_8_LD.825 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 25142.1052 | 7560.0001 | 402.0057 | 9908.7691 | 15322.2131 | 3267.7073 | 29202.4842 | 1069.3734 | 262.6691 | 531.3432 | 6289.5567 | 60781.7306 | 11086.3002 | 93.3336 | 259.3357 | 2594.2357 | 13543.0843 | 37336.7285 | 9968.9329 | 412.0059 | 12582.8725 | 1251.3882 | 3.3333 | 20 | 43.3334 | 248.0583 | 8.6667 | 97666.0596 | 79873.3886 | 181436.2512 | 191076.7761 | 233074.6351 | 193066.5251 | 165331.2197 | 19024.6607 | 47.3334 | 120.6672 | 1052.0157 | 2046.0569
177 | 10_15_21_UF_8_LD |  | 06/17/2023 06:06:21 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_8_LD.826 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 25025.234 | 7377.905 | 416.6727 | 10602.6 | 15160.7076 | 3349.0592 | 29207.8331 | 1065.3731 | 244.6688 | 542.0103 | 6323.2024 | 61592.4923 | 10920.1729 | 100.0004 | 267.3359 | 2738.9292 | 13594.4667 | 37503.1628 | 9715.6675 | 432.0066 | 12696.3061 | 1214.7189 | 10.6667 | 27.3334 | 50.6668 | 257.4188 | 9.3333 | 97329.1483 | 80514.2544 | 180026.5989 | 184000.7286 | 231985.3448 | 192938.8837 | 163386.3846 | 19237.6116 | 46.0001 | 113.3338 | 996.6804 | 2017.3894
178 | LBLANK |  | 06/17/2023 06:12:21 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.827 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.6667 | 1.3333 | 6 | 1000.7017 | 2176.8329 | 25.3334 | 2 | 2 | 22.6667 | 22 | 369.7879 | 7.3333 | 6526.1573 | 0 | 20.6667 | 18 | 68.0002 | 358.0046 | -3070.6094 | 2.6667 | 64.0001 | 1065.3731 | 0 | 0.66667 | 4 | 54.8842 | 0.66667 | 48.6668 | 28 | 200966.8193 | 198296.7837 | 254416.8174 | 208455.225 | 173353.5086 | 20285.7266 | 2 | 6 | 12.6667 | 35.3334
179 | LBLANK |  | 06/17/2023 06:18:40 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.828 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4 | 0.66667 | 15.3333 | 354.6711 | 2098.1547 | 28.6667 | 1.3333 | 2 | 22.6667 | 28 | 370.3417 | 12.6667 | 6466.1298 | 2.6667 | 16 | 8.6667 | 70.0002 | 351.3377 | -3221.0749 | 0.66667 | 37.3334 | 1098.0422 | 0 | 0.66667 | 5.3333 | 42.1634 | 0 | 32 | 16.6667 | 190866.0026 | 197206.5071 | 250830.2545 | 209042.9933 | 170652.5948 | 20193.5976 | 0.66667 | 6 | 18.6667 | 36.6667
180 | CCV 1ppb TRACE |  | 06/17/2023 06:24:40 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.829 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2381.5321 | 1055.3724 | 246.6688 | 1870.1224 | 2736.9291 | 1526.7483 | 734.0189 | 1692.767 | 2166.831 | 5995.925 | 5954.7011 | 8603.9238 | 9063.5434 | 3696.4786 | 972.0331 | 2252.1776 | 11700.1239 | 5781.8368 | 796.4364 | 152.6675 | 4127.9297 | 1560.752 | 100.667 | 290.003 | 566.0112 | 390.5915 | 125.3339 | 10625.9505 | 8177.674 | 191272.9224 | 193909.8438 | 246706.6842 | 203618.8993 | 168932.3219 | 20154.2069 | 4962.1951 | 11832.8989 | 22036.6418 | 43881.7543
181 | CCV 1ppb TRACE |  | 06/17/2023 06:31:00 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.830 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2421.5386 | 1098.0423 | 234.6686 | 1905.4604 | 2876.2895 | 1508.0796 | 774.021 | 1652.0956 | 2194.1687 | 5949.9057 | 5852.1108 | 8621.2678 | 9184.2853 | 3758.4944 | 983.3672 | 2375.5308 | 11638.7396 | 5720.479 | 583.8055 | 158.6675 | 3958.5485 | 1586.0882 | 77.3335 | 285.3362 | 596.0125 | 424.0611 | 111.3338 | 10387.1087 | 8514.5371 | 194706.8485 | 192286.1883 | 246594.7966 | 204136.1845 | 168369.8052 | 20022.6905 | 4848.8229 | 11670.7669 | 21823.1733 | 44448.5198
182 | LFB |  | 06/17/2023 06:37:00 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.831 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 21492.1567 | 9351.7266 | 1006.0356 | 19409.1786 | 8117.6397 | 13279.5027 | 6904.3352 | 14369.8934 | 18686.8829 | 51546.1654 | 35634.8916 | 75324.7333 | 23493.3059 | 32980.0342 | 8328.427 | 19955.9292 | 99718.8281 | 48227.9561 | 35813.9148 | 1370.7325 | 36737.18 | 5174.9372 | 806.0228 | 2569.5644 | 5473.7152 | 3680.4726 | 982.0338 | 92728.6571 | 75064.709 | 181138.4955 | 177307.5431 | 228563.6534 | 191457.111 | 160398.107 | 18649.5001 | 44887.0763 | 107635.9715 | 182132.8959 | 381529.7533
183 | LFB |  | 06/17/2023 06:43:18 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.832 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 21544.2352 | 9425.7753 | 998.0349 | 19544.694 | 7917.527 | 13410.2929 | 7017.7234 | 14534.0586 | 18614.1198 | 52237.3348 | 35809.9102 | 76342.108 | 23801.8168 | 33531.9772 | 8490.5233 | 20159.5497 | 100584.8646 | 47320.2496 | 35190.9566 | 1375.3997 | 36728.4888 | 5200.2799 | 849.3586 | 2606.238 | 5695.802 | 3697.2078 | 992.0345 | 93572.8028 | 74617.0287 | 187461.283 | 180219.6815 | 236882.5577 | 180644.0781 | 160326.0397 | 18918.5208 | 45402.7018 | 108684.5347 | 183004.2593 | 388923.6422
184 | CCV 10 ppb MAJORS |  | 06/17/2023 06:49:17 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.833 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 22862.2807 | 10140.932 | 1015.3695 | 16174.485 | 8177.6733 | 14200.3878 | 7476.6231 | 15555.7983 | 19620.1314 | 54897.2991 | 39224.7885 | 80850.1536 | 25086.0158 | 35043.5967 | 8810.05 | 21175.6828 | 105220.7757 | 51358.1534 | 36076.2928 | 1404.069 | 38156.894 | 5439.0354 | 834.0244 | 2722.9262 | 5541.7414 | 3758.1172 | 1039.3712 | 98807.1961 | 79893.4482 | 186718.8349 | 182251.8645 | 236351.1477 | 197875.063 | 163652.7113 | 19117.4505 | 47120.2609 | 113505.8157 | 190691.2481 | 402125.5602
185 | CCV 10 ppb MAJORS |  | 06/17/2023 06:55:37 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.834 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 23439.2149 | 10279.034 | 1094.042 | 16033.6608 | 8487.1878 | 14663.5234 | 7472.6207 | 15810.0776 | 20307.0904 | 56136.0892 | 38809.252 | 83306.8858 | 26100.4919 | 36054.1067 | 9086.8896 | 21558.255 | 107111.399 | 51588.9844 | 37357.5614 | 1467.4089 | 38363.4456 | 5434.367 | 885.3608 | 2740.2629 | 5701.1374 | 3842.7983 | 1045.3716 | 101329.4894 | 81088.8253 | 187258.0078 | 183939.9481 | 240665.0947 | 197206.5111 | 164958.973 | 19572.0666 | 47583.7885 | 114400.2361 | 194496.7139 | 409272.4925
186 | 10_15_21_UF_8_DUP |  | 06/17/2023 07:01:37 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_8_DUP.835 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 24213.1704 | 7219.8244 | 100.667 | 5720.4788 | 14843.7078 | 3180.354 | 28404.8804 | 1004.0353 | 268.0025 | 536.0101 | 7646.1819 | 58771.3144 | 11568.6823 | 72.6669 | 236.002 | 2550.8946 | 12713.6553 | 13686.5539 | 9646.754 | 397.3389 | 12280.6117 | 1210.0514 | 6 | 34 | 56.0001 | 218.0198 | 8.6667 | 96267.2996 | 77792.5808 | 175619.5521 | 177620.059 | 226994.0516 | 189489.8634 | 160156.7652 | 18140.8502 | 42.6667 | 110.6671 | 929.3452 | 1896.0488
187 | 10_15_21_UF_8_DUP |  | 06/17/2023 07:07:56 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_8_DUP.836 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 23373.7764 | 6941.6863 | 96.667 | 5913.2245 | 14628.1561 | 3133.6771 | 27939.9655 | 1006.0356 | 253.3356 | 549.3439 | 7666.1559 | 58198.3246 | 11252.4305 | 66.6668 | 233.3352 | 2473.5475 | 12360.6786 | 13635.1736 | 9583.5213 | 412.006 | 11783.5258 | 1181.3823 | 6 | 20.6667 | 54.0001 | 179.3199 | 8.6667 | 94285.4661 | 76581.3981 | 174848.8704 | 176115.6196 | 226146.1251 | 187589.0043 | 157829.0767 | 18461.9239 | 56.6668 | 114.6671 | 956.0122 | 1844.7135
188 | 10_15_21_UF_8_DUP_LSP |  | 06/17/2023 07:13:55 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_8_DUP_LSP.837 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 41734.2087 | 15459.0275 | 1118.0438 | 24788.852 | 18758.3087 | 16070.3688 | 31327.6513 | 15198.7477 | 18413.1934 | 50841.6459 | 41870.0317 | 126445.8417 | 27986.7205 | 32222.3019 | 8136.9838 | 20844.5305 | 104061.6412 | 53632.4879 | 44924.9335 | 1687.433 | 45698.98 | 5130.2545 | 818.0235 | 2478.8818 | 5238.2939 | 3654.0486 | 924.6966 | 176061.6607 | 143112.634 | 175805.8329 | 172510.6873 | 221580.5565 | 182451.1219 | 153653.2783 | 17739.6751 | 42640.2096 | 101585.2442 | 178320.6363 | 364977.619
189 | 10_15_21_UF_8_DUP_LSP |  | 06/17/2023 07:20:13 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_8_DUP_LSP.838 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 43387.1223 | 15537.1115 | 1144.0458 | 25446.6459 | 19123.4579 | 16354.0225 | 32193.5689 | 15784.0497 | 18787.6801 | 52062.6989 | 42100.435 | 128685.6979 | 28651.3707 | 33255.3292 | 8401.1365 | 20945.3437 | 105397.3742 | 54421.464 | 45900.4316 | 1674.7649 | 46203.9342 | 5255.6337 | 828.024 | 2506.22 | 5491.7221 | 3754.692 | 966.0327 | 179616.7978 | 145685.0897 | 174802.7817 | 175803.1371 | 224809.1359 | 184127.0192 | 158617.061 | 18410.5234 | 43593.7473 | 104171.0983 | 179320.909 | 369674.8348
190 | 10_15_21_UF_9 |  | 06/17/2023 07:26:13 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_9.839 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 20968.7114 | 6690.2335 | 1357.3978 | 12163.8431 | 12347.3338 | 2779.6039 | 25624.2958 | 829.3574 | 248.0022 | 595.3457 | 4058.4961 | 33748.484 | 9492.486 | 60.6668 | 190.6679 | 2160.8301 | 11055.611 | 14670.1954 | 7350.3383 | 344.0042 | 9813.3702 | 1252.055 | 5.3333 | 21.3334 | 41.3334 | 193.9707 | 12 | 67735.5447 | 55049.8678 | 176957.9175 | 179070.0154 | 228511.8481 | 189262.1326 | 161040.6892 | 18805.7037 | 50.6668 | 131.3339 | 695.3402 | 1235.3539
191 | 10_15_21_UF_9 |  | 06/17/2023 07:32:31 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_9.840 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 20493.3555 | 6660.886 | 1299.3924 | 10872.1375 | 12313.3046 | 2666.2489 | 25120.0668 | 790.022 | 254.669 | 572.6782 | 3892.9832 | 33526.6344 | 9226.3119 | 58.6668 | 173.3344 | 2085.4856 | 10950.8624 | 14658.1831 | 6965.2252 | 330.6705 | 9810.7017 | 1171.3814 | 3.3333 | 21.3333 | 34.6667 | 169.336 | 9.3333 | 67042.2923 | 54485.71 | 176461.9026 | 176592.1157 | 225989.6366 | 187774.694 | 157785.9396 | 18560.0513 | 46.0001 | 122.0005 | 664.6732 | 1292.6895
192 | 10_15_21_UF_10 |  | 06/17/2023 07:38:31 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_10.841 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 16222.5396 | 7277.1866 | 96.0003 | 8245.7129 | 9962.1394 | 2294.851 | 23439.8821 | 898.695 | 223.3351 | 462.6742 | 3246.2221 | 35087.0382 | 8880.7603 | 51.3334 | 172.001 | 1674.0982 | 8732.6687 | 11215.0674 | 8250.5472 | 391.3387 | 10740.0361 | 1184.0491 | 11.3333 | 20 | 46.0001 | 170.6585 | 11.3333 | 29772.3307 | 24053.568 | 176045.412 | 177743.0057 | 226734.5776 | 195773.579 | 158987.1306 | 18311.7288 | 34 | 66.0002 | 698.0068 | 1410.0272
193 | 10_15_21_UF_10 |  | 06/17/2023 07:44:50 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_10.842 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 16152.4613 | 7064.4131 | 83.3336 | 7746.7676 | 9898.7637 | 2386.1994 | 23251.5767 | 880.6938 | 216.6683 | 468.6744 | 3169.4434 | 34720.8098 | 8695.9793 | 49.3334 | 139.334 | 1672.7646 | 8660.6245 | 11327.156 | 8070.4257 | 387.3386 | 10677.3235 | 1216.7185 | 5.3333 | 25.3334 | 39.3334 | 182.8264 | 4.6667 | 29523.1452 | 23835.8701 | 176471.9983 | 175324.6821 | 225631.2105 | 210859.4073 | 156111.7018 | 18480.6137 | 32.6667 | 68.0002 | 729.3409 | 1334.6913
194 | 10_15_21_UF_11 |  | 06/17/2023 07:50:50 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_11.843 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 15525.0982 | 7274.5184 | 97.3337 | 6685.5646 | 9843.391 | 2242.1761 | 23071.9494 | 862.6927 | 170.001 | 312.6701 | 7107.643 | 30404.9904 | 10599.2642 | 50.6668 | 124.0005 | 1646.7616 | 8618.599 | 15551.1264 | 8634.6153 | 382.0051 | 11109.6521 | 1084.7079 | 3.3333 | 18.6667 | 36 | 176.0893 | 11.3333 | 25190.8588 | 20659.5969 | 173534.3168 | 176096.0343 | 226270.534 | 195551.5092 | 157579.0269 | 18452.5771 | 22.6667 | 84.6669 | 750.0079 | 1428.6948
195 | 10_15_21_UF_11 |  | 06/17/2023 07:57:09 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_11.844 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 15570.4813 | 7099.7643 | 108.6671 | 6758.2651 | 9760.6672 | 2159.4965 | 22362.1564 | 846.0251 | 175.3344 | 342.0042 | 7064.5745 | 30186.5259 | 10407.1228 | 52.0001 | 159.3342 | 1668.0974 | 8691.3103 | 15498.4032 | 8287.3984 | 378.005 | 11022.2507 | 1166.7143 | 5.3333 | 25.3334 | 42.6667 | 176.637 | 8.6667 | 24810.1939 | 20150.8696 | 174433.8234 | 172161.851 | 223393.1077 | 183754.9223 | 153532.6838 | 18143.517 | 28 | 80.0002 | 770.0083 | 1458.6951
196 | LBLANK |  | 06/17/2023 08:03:10 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.845 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.3333 | 0.66667 | 4 | 340.0041 | 2006.141 | 29.3334 | 0 | 0.66667 | 28 | 22 | 370.0325 | 12 | 6095.9672 | 0 | 16.6667 | 18 | 69.3335 | 329.3371 | -2929.9935 | 0.66667 | 53.3334 | 1041.3713 | 0 | 3.3333 | 2.6667 | 42.8841 | 1.3333 | 45.3334 | 13.3333 | 189282.6021 | 185827.5427 | 244047.783 | 198986.36 | 161653.5274 | 19002.6303 | 0 | 6.6667 | 16.6667 | 38.6667
197 | LBLANK |  | 06/17/2023 08:09:28 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.846 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.66667 | 0 | 6 | 234.6686 | 1970.1359 | 35.3334 | 2.6667 | 0.66667 | 24 | 22.6667 | 358.2473 | 12 | 6083.9622 | 2.6667 | 24 | 10.6667 | 70.0002 | 322.6703 | -2907.0245 | 1.3333 | 54.6668 | 1009.369 | 0.66667 | 2 | 3.3333 | 35.4587 | 0.66667 | 28.6667 | 11.3333 | 170330.3624 | 184724.6494 | 245022.0128 | 197140.9919 | 159579.0776 | 19329.0679 | 2.6667 | 8.6667 | 23.3333 | 34.6667
198 | CCV 1ppb TRACE |  | 06/17/2023 08:15:29 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.847 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2294.8511 | 1018.703 | 230.6685 | 1764.776 | 2733.5952 | 1441.4061 | 723.3517 | 1600.7564 | 2041.4793 | 5565.0838 | 5676.1858 | 8000.9077 | 8624.6034 | 3637.7965 | 930.0304 | 2204.8369 | 11068.2898 | 5324.3253 | 601.1323 | 152.0008 | 3829.18 | 1514.747 | 76.6669 | 257.3357 | 565.3445 | 374.4938 | 106.0004 | 9903.4331 | 7881.507 | 179175.3089 | 182495.5995 | 234603.3359 | 195131.6463 | 158877.9284 | 18716.9201 | 4842.1539 | 11345.8378 | 21515.0369 | 43327.0665
199 | CCV 1ppb TRACE |  | 06/17/2023 08:21:48 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.848 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2304.186 | 1026.7036 | 242.6688 | 1744.7733 | 2690.2537 | 1504.746 | 675.3493 | 1591.4221 | 2048.1469 | 5560.4155 | 5558.8764 | 8049.6008 | 8719.3268 | 3563.7779 | 886.0275 | 2178.8329 | 11059.6132 | 5310.9874 | 789.4852 | 132.0006 | 3724.4855 | 1483.4104 | 88.667 | 266.0025 | 582.6786 | 376.4938 | 98.0003 | 9810.7011 | 7798.7951 | 176383.5824 | 179315.0525 | 232796.8527 | 192144.925 | 158625.1506 | 18897.1576 | 4752.1236 | 11258.4357 | 21053.3543 | 42492.1291
200 | CCV 10 ppb MAJORS |  | 06/17/2023 08:27:48 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.849 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 21116.2623 | 9348.3911 | 1000.0352 | 14403.2575 | 7860.1619 | 13376.2608 | 6913.0064 | 14022.2122 | 18085.4444 | 51061.7606 | 34743.0385 | 74087.634 | 23102.6693 | 32397.3611 | 8137.6514 | 19568.0609 | 95953.1968 | 47119.5852 | 34179.6777 | 1316.7275 | 35105.7528 | 5175.604 | 787.3551 | 2475.5478 | 5288.3121 | 3549.4413 | 950.6983 | 92717.9315 | 73625.9209 | 173223.8998 | 167701.2983 | 220099.9823 | 182311.2691 | 150532.3489 | 17870.5041 | 44096.6201 | 105690.8924 | 185737.7312 | 389290.4658
201 | CCV 10 ppb MAJORS |  | 06/17/2023 08:34:07 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.850 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 21651.0637 | 9464.4681 | 967.3661 | 14654.1815 | 7915.5257 | 13238.131 | 7027.7286 | 14534.0575 | 18699.5642 | 52078.7688 | 36023.1528 | 76277.1038 | 23609.4933 | 32974.6847 | 8287.0713 | 19802.3819 | 99319.4055 | 48268.0846 | 34488.1741 | 1360.0648 | 35958.5326 | 5173.6034 | 850.0253 | 2493.5514 | 5368.3422 | 3790.558 | 965.366 | 94388.161 | 75659.8401 | 186395.9508 | 171068.8862 | 224956.7505 | 185295.31 | 153548.121 | 18033.3754 | 44861.6633 | 107945.6474 | 186870.0585 | 392885.9204
202 | 10_15_21_FB |  | 06/17/2023 08:40:08 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_FB.851 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 639.3477 | 6 | 991.3678 | 2704.2562 | 2098.8208 | 51.3334 | 8.6667 | 2 | 194.0013 | 418.0061 | 7560.8562 | 428.0064 | 9940.791 | 8 | 25.3334 | 264.0024 | 1406.7361 | 38059.2979 | -3119.5877 | 0.66667 | 48.6668 | 1057.3725 | 0 | 1.3333 | 2.6667 | 36.857 | 0 | 121.3338 | 78.6669 | 183467.2601 | 164794.7923 | 237451.8646 | 197293.7003 | 161769.4543 | 19349.0966 | 3.3333 | 24 | 1124.684 | 2238.0689
203 | 10_15_21_FB |  | 06/17/2023 08:46:26 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_FB.852 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 608.6796 | 6.6667 | 988.7009 | 2557.5623 | 2017.4759 | 62.0001 | 9.3333 | 5.3333 | 187.3346 | 427.3397 | 7442.7542 | 406.0058 | 9609.8982 | 6 | 29.3334 | 244.0021 | 1393.4014 | 37457.7111 | -3006.7617 | 0.66667 | 50.6668 | 1058.0392 | 0.66667 | 1.3333 | 2 | 48.8842 | 0 | 106.6671 | 68.0002 | 180872.5465 | 179665.4286 | 233010.9727 | 195649.997 | 158807.1457 | 19061.3754 | 1.3333 | 17.3333 | 1054.0153 | 2169.3972
204 | 10_15_21_BB |  | 06/17/2023 08:52:26 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_BB.853 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 6 | 1.3333 | 67.3335 | 554.6774 | 1912.128 | 29.3334 | 4 | 1.3333 | 64.6668 | 133.334 | 1135.0426 | 36.6667 | 6194.0097 | 6 | 26.6667 | 194.0013 | 1062.0395 | 1137.3786 | -2884.5118 | 2.6667 | 22.6667 | 1044.705 | 0.66667 | 1.3333 | 2.6667 | 45.5292 | 0 | 34.6667 | 11.3333 | 183020.9228 | 180738.1611 | 234684.1056 | 196004.1531 | 156876.0691 | 18827.7354 | 0.66667 | 10.6667 | 444.0026 | 924.0116
205 | 10_15_21_BB |  | 06/17/2023 08:58:45 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_BB.854 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 6 | 0 | 84.0002 | 554.0107 | 1940.1317 | 25.3334 | 5.3333 | 0 | 62.0001 | 138.0007 | 1182.0221 | 47.3334 | 6330.0692 | 0.66667 | 20.6667 | 212.0016 | 1041.3713 | 1141.379 | -2801.8911 | 0 | 15.3333 | 985.3673 | 0 | 2 | 3.3333 | 36.1092 | 0.66667 | 46.6667 | 11.3333 | 180644.9835 | 180014.5071 | 236686.8142 | 194032.2516 | 156898.9173 | 18842.4185 | 0 | 8.6667 | 434.0026 | 916.0115
206 | 10_15_21_LAB_A |  | 06/17/2023 09:04:45 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_LAB_A.855 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 8 | 0 | 64.6668 | 500.0088 | 1831.4508 | 24 | 2 | 0 | 109.3338 | 268.0025 | 1798.3993 | 68.0002 | 6488.1402 | 0.66667 | 16 | 220.6684 | 1145.3793 | 1290.7251 | -2928.597 | 0 | 14.6667 | 992.7012 | 0 | 0 | 0.66667 | 55.5184 | 0 | 84.0002 | 40.6667 | 180190.6348 | 178764.8737 | 245485.3192 | 193558.2882 | 156125.2729 | 18809.0419 | 1.3333 | 9.3333 | 485.3366 | 959.3461
207 | 10_15_21_LAB_A |  | 06/17/2023 09:11:03 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_LAB_A.856 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 6.6667 | 0 | 66.0002 | 512.0092 | 2029.4776 | 24 | 2.6667 | 2 | 113.3338 | 276.0027 | 1804.2871 | 61.3335 | 6578.1811 | 2 | 18 | 242.6687 | 1144.7126 | 1308.0599 | -2915.8412 | 0.66667 | 12.6667 | 1061.3728 | 0 | 0.66667 | 2 | 35.5021 | 0 | 96.0003 | 47.3334 | 179282.6427 | 179396.7392 | 235154.9706 | 192936.165 | 156928.5754 | 18933.2047 | 1.3333 | 9.3333 | 508.6703 | 953.346
208 | 10_15_21_LAB_B |  | 06/17/2023 09:17:03 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_LAB_B.857 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 5.3333 | 0.66667 | 70.6668 | 594.0124 | 1904.7937 | 26 | 2 | 0 | 60.0001 | 114.0005 | 1064.4209 | 34.6667 | 6244.6983 | 2 | 22.6667 | 239.3353 | 1664.1132 | 1865.4552 | -2891.5559 | 1.3333 | 8.6667 | 1016.0363 | 0 | 1.3333 | 2 | 13.464 | 0.66667 | 22 | 8 | 177619.3811 | 176058.2782 | 230011.567 | 190914.5945 | 153092.0135 | 18637.4837 | 0.66667 | 5.3333 | 487.3365 | 942.6787
209 | 10_15_21_LAB_B |  | 06/17/2023 09:23:20 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_LAB_B.858 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.3333 | 0 | 84.6669 | 640.0144 | 1766.1092 | 32.6667 | 2 | 2 | 53.3334 | 113.3338 | 1085.5677 | 29.3334 | 6130.6487 | 2.6667 | 17.3333 | 266.6692 | 1207.3845 | 1877.4567 | -2688.0459 | 0 | 9.3333 | 997.3681 | 0.66667 | 0 | 0 | 14.7162 | 0.66667 | 24.6667 | 7.3333 | 177216.4264 | 174632.2434 | 232158.6703 | 192046.9696 | 152817.022 | 18630.809 | 1.3333 | 11.3333 | 501.3366 | 966.0134
210 | 5_18_23_DB |  | 06/17/2023 09:29:20 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_DB.859 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.6667 | 0 | 45.3334 | 5667.1242 | 1856.7873 | 22.6667 | 2.6667 | 0 | 24.6667 | 24.6667 | 546.8799 | 18 | 5908.555 | 2.6667 | 10.6667 | 197.3347 | 1014.036 | 732.6855 | -2936.0429 | 0 | 11.3333 | 968.6997 | 0 | 0 | 1.3333 | 27.4262 | 0 | 16.6667 | 2.6667 | 177957.0123 | 175151.9136 | 229255.5684 | 190128.9253 | 151968.1117 | 18532.0128 | 1.3333 | 2.6667 | 516.6703 | 1020.6808
211 | 5_18_23_DB |  | 06/17/2023 09:35:39 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_DB.860 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2 | 0 | 44.0001 | 7436.6026 | 1770.7765 | 30.6667 | 2.6667 | 2 | 30 | 25.3334 | 582.5345 | 15.3333 | 5689.7996 | 2 | 14 | 216.0016 | 1006.0355 | 726.0185 | -2680.3595 | 0.66667 | 12 | 977.3668 | 0 | 0.66667 | 0.66667 | 24.1524 | 1.3333 | 21.3334 | 1.3333 | 175624.477 | 174104.6837 | 230217.568 | 192173.9529 | 153037.4124 | 18353.1159 | 2 | 8.6667 | 514.6702 | 1040.015
212 | LBLANK |  | 06/17/2023 09:41:39 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.861 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.66667 | 0 | 3.3333 | 1984.8046 | 1802.1138 | 24 | 1.3333 | 1.3333 | 21.3334 | 20.6667 | 306.0228 | 10 | 5720.4785 | 2 | 7.3333 | 13.3333 | 69.3335 | 306.0033 | -2685.6989 | 0.66667 | 14.6667 | 881.3606 | 0 | 1.3333 | 1.3333 | 42.7976 | 0.66667 | 32 | 8.6667 | 166693.3693 | 173101.7956 | 226292.9182 | 187967.1943 | 150968.8673 | 18109.4714 | 1.3333 | 5.3333 | 16 | 42.6667
213 | LBLANK |  | 06/17/2023 09:47:59 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.862 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.66667 | 0 | 6 | 482.0081 | 1727.4378 | 35.3334 | 0.66667 | 0 | 26 | 28 | 320.7168 | 6.6667 | 5735.8181 | 1.3333 | 11.3333 | 6.6667 | 73.3335 | 313.3368 | -2873.1371 | 0 | 8.6667 | 977.3668 | 0.66667 | 0 | 0 | 2.1144 | 0 | 24.6667 | 9.3333 | 181900.4921 | 173570.7439 | 227585.8613 | 188469.7772 | 150483.1032 | 18361.7964 | 0.66667 | 8 | 24 | 45.3334
214 | CCV 1ppb TRACE |  | 06/17/2023 09:53:59 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.863 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2068.1498 | 956.032 | 178.0011 | 1734.1056 | 2263.5127 | 1382.7336 | 703.3507 | 1468.7422 | 1883.4577 | 5367.6749 | 5179.7455 | 7742.7644 | 8164.3324 | 3319.7191 | 856.0257 | 2098.821 | 10570.5767 | 5164.2668 | 346.2282 | 120.6672 | 3484.4249 | 1441.4061 | 79.3336 | 266.6692 | 519.3428 | 349.9773 | 90.667 | 9287.019 | 7489.9633 | 172047.1632 | 170954.214 | 223223.9629 | 185156.9016 | 149846.4885 | 18022.6984 | 4486.7046 | 10873.4705 | 20734.5589 | 42430.6713
215 | CCV 1ppb TRACE |  | 06/17/2023 10:00:19 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.864 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2152.1621 | 960.0323 | 192.668 | 1640.761 | 2525.5566 | 1428.0716 | 746.0196 | 1531.4154 | 1958.801 | 5340.9984 | 5271.5802 | 7805.4652 | 8239.042 | 3362.3957 | 881.3606 | 2052.1474 | 10495.1876 | 5014.2133 | 606.6303 | 136.6673 | 3620.4589 | 1413.4033 | 83.3336 | 234.6686 | 500.0088 | 367.9133 | 104.0004 | 9364.4053 | 7539.9901 | 173312.9954 | 171369.0982 | 220720.6849 | 186393.4711 | 149362.7978 | 18163.5404 | 4544.0563 | 10834.7755 | 20723.8725 | 42312.5683
216 | LFB |  | 06/17/2023 10:06:18 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.865 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 19151.4972 | 8373.7875 | 951.3651 | 16484.8392 | 7298.531 | 12286.6151 | 6404.103 | 13242.8019 | 16772.5104 | 45405.3778 | 31306.6071 | 66040.9677 | 21204.3927 | 29594.6301 | 7538.6566 | 17963.2906 | 88069.3105 | 41795.0504 | 31559.9239 | 1202.7173 | 32154.1466 | 4753.4585 | 692.6836 | 2197.5024 | 4917.5132 | 3218.0783 | 895.3615 | 83823.8963 | 66765.6668 | 167212.9349 | 159085.5608 | 211967.1015 | 175396.1828 | 143686.9967 | 17333.1771 | 40840.2965 | 98109.1052 | 175425.8761 | 365306.4011
217 | LFB |  | 06/17/2023 10:12:36 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.866 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 19428.537 | 8627.2712 | 940.0311 | 17288.4571 | 7386.5778 | 12333.9889 | 6442.119 | 13256.1495 | 16705.0951 | 47084.807 | 32013.3449 | 68436.8718 | 21838.0134 | 30209.2411 | 7764.1094 | 18512.655 | 89978.5254 | 42539.2462 | 32162.135 | 1328.0618 | 33145.0751 | 4864.1615 | 713.3512 | 2271.5141 | 4957.527 | 3387.8097 | 931.3637 | 85711.7242 | 68894.4005 | 168531.8239 | 161963.7092 | 211751.6132 | 178718.3251 | 144880.3177 | 17822.444 | 41462.7741 | 100148.6044 | 176755.4279 | 367864.4209
218 | CCV 10 ppb MAJORS |  | 06/17/2023 10:18:36 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.867 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 20814.4891 | 9208.3005 | 1002.7019 | 14470.6587 | 7845.4876 | 13262.1534 | 6936.3503 | 14279.1335 | 18230.2913 | 50493.0773 | 35067.1556 | 73321.0188 | 23117.3619 | 32355.936 | 8032.9247 | 19447.2281 | 96578.7195 | 46060.8048 | 34314.8328 | 1255.3885 | 34501.6128 | 5079.5697 | 770.6875 | 2450.8769 | 5136.9238 | 3500.8797 | 931.3637 | 91359.889 | 73158.8638 | 173663.9072 | 165362.9325 | 218880.2782 | 183155.9974 | 148057.3072 | 17829.122 | 43136.3827 | 104069.0446 | 187456.1681 | 390820.1082
219 | CCV 10 ppb MAJORS |  | 06/17/2023 10:24:55 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.868 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 21392.6742 | 9298.3586 | 988.0342 | 14432.6221 | 8007.5772 | 13371.5888 | 7097.096 | 14576.0999 | 18265.6707 | 51241.7507 | 34826.047 | 74120.4555 | 23622.85 | 33024.799 | 8358.4446 | 19728.9471 | 97439.858 | 46848.0273 | 34645.7306 | 1310.7268 | 34680.0471 | 5083.5712 | 780.0213 | 2406.2028 | 5127.5868 | 3668.6782 | 929.3636 | 92428.7364 | 74370.4178 | 175004.7591 | 167546.7906 | 219496.7395 | 187891.5718 | 151398.7252 | 18228.2921 | 43593.0804 | 105533.7075 | 186538.7853 | 391210.6402
220 | 5_18_23_DB |  | 06/17/2023 10:30:55 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_DB.869 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 6 | 0 | 70.6668 | 6248.7001 | 1968.1358 | 26 | 2.6667 | 2 | 38.6667 | 68.6668 | 6339.3871 | 73.3335 | 9187.6208 | 6.6667 | 12 | 211.3349 | 980.7003 | 778.0212 | -2893.152 | 0.66667 | 12 | 1036.7044 | 0 | 2.6667 | 2.6667 | 42.1633 | 2 | 148.0008 | 20.6667 | 180432.4421 | 181122.9688 | 240309.8697 | 203881.6095 | 158121.6578 | 19149.4946 | 0 | 12 | 520.6704 | 1034.681
221 | 5_18_23_DB |  | 06/17/2023 10:37:14 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_DB.870 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4.6667 | 0.66667 | 59.3335 | 8010.2454 | 1927.4634 | 30 | 1.3333 | 0.66667 | 40.0001 | 56.0001 | 6225.0551 | 76.0002 | 9100.234 | 2 | 11.3333 | 212.0016 | 1022.0366 | 750.0197 | -3032.2225 | 1.3333 | 14.6667 | 1035.3709 | 0.66667 | 0.66667 | 0.66667 | 21.4695 | 1.3333 | 104.6671 | 6.6667 | 183166.8723 | 181032.5228 | 237361.1006 | 196856.3968 | 156323.9866 | 19152.8313 | 0.66667 | 5.3333 | 497.3369 | 1049.3486
222 | 5_18_23_DB |  | 06/17/2023 10:43:14 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_DB.871 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.3333 | 0 | 63.3335 | 8279.066 | 1985.4713 | 36 | 5.3333 | 1.3333 | 52.0001 | 100.667 | 1039.0336 | 40.0001 | 6516.8213 | 266.6692 | 21.3334 | 227.3352 | 1088.0415 | 977.3668 | -2889.1573 | 1.3333 | 14.6667 | 1044.7051 | 0 | 1.3333 | 2 | 41.3776 | 0.66667 | 70.6668 | 9.3333 | 162678.0653 | 177862.465 | 234938.1038 | 193957.8415 | 155689.7924 | 19054.6997 | 1.3333 | 8 | 587.3379 | 1129.3508
223 | 5_18_23_DB |  | 06/17/2023 10:49:33 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_DB.872 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4.6667 | 0 | 50.0001 | 7869.5005 | 1966.1356 | 22 | 2 | 1.3333 | 46.0001 | 78.6669 | 1115.6897 | 36.6667 | 6419.4421 | 246.6688 | 10.6667 | 216.0016 | 1106.7096 | 927.3634 | -2719.9645 | 0 | 13.3333 | 1023.37 | 2.6667 | 0.66667 | 2.6667 | 34.8895 | 0 | 56.6668 | 10.6667 | 173898.7665 | 176212.8421 | 233088.8592 | 215892.0064 | 153733.5011 | 18778.3345 | 2 | 3.3333 | 556.0043 | 1040.6818
224 | 5_18_23_DB |  | 06/17/2023 10:55:33 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_DB.873 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 6.6667 | 0 | 54.0001 | 7527.9832 | 1971.4699 | 26 | 2 | 0.66667 | 34.6667 | 40.0001 | 735.2706 | 55.3334 | 6354.7468 | 2.6667 | 16 | 171.3344 | 788.0219 | 958.6989 | -2796.6614 | 0 | 7.3333 | 986.034 | 0 | 0 | 3.3333 | 30.7812 | 0 | 58.0001 | 5.3333 | 177857.0979 | 172906.0988 | 229777.2186 | 190276.1405 | 151666.2296 | 18581.4104 | 2.6667 | 8.6667 | 374.002 | 729.3405
225 | 5_18_23_DB |  | 06/17/2023 11:01:51 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_DB.874 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 6 | 0 | 46.0001 | 7191.8108 | 1926.1299 | 26.6667 | 3.3333 | 0 | 29.3334 | 32 | 717.8235 | 46.6667 | 6232.6936 | 2 | 20 | 156.6675 | 796.0222 | 936.0308 | -2930.4349 | 2 | 3.3333 | 956.6988 | 0 | 3.3333 | 3.3333 | 30.7975 | 0 | 49.3334 | 4.6667 | 177178.6385 | 173933.1065 | 227774.3647 | 189868.7608 | 149604.6124 | 18706.2423 | 0.66667 | 9.3333 | 344.6683 | 781.3417
226 | 5_18_23_MFB |  | 06/17/2023 11:07:51 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_MFB.875 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10006.8377 | 4347.9954 | 490.0084 | 15027.2332 | 4748.1224 | 6332.7371 | 3281.0435 | 6847.6412 | 8962.8115 | 24446.2327 | 17757.3192 | 35810.8287 | 14649.5077 | 15679.2667 | 3925.8728 | 9723.3084 | 47906.8617 | 21846.6929 | 14665.2103 | 630.6806 | 16814.5573 | 2966.3081 | 397.3389 | 1198.717 | 2508.8871 | 1682.1826 | 438.0067 | 43236.6663 | 35148.5264 | 169606.937 | 161533.7043 | 216413.0629 | 180530.2078 | 145646.8429 | 17403.261 | 20588.8279 | 49845.4765 | 95126.7499 | 196394.3189
227 | 5_18_23_MFB |  | 06/17/2023 11:14:09 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_MFB.876 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10237.0009 | 4620.7472 | 542.0103 | 15258.1446 | 4690.1034 | 6430.7813 | 3427.0777 | 6982.3729 | 8887.4305 | 24840.9136 | 18040.0828 | 36235.8996 | 14795.6583 | 15988.2776 | 3952.5473 | 9861.4053 | 48229.2768 | 22213.258 | 14818.1332 | 639.3477 | 16900.6599 | 3006.9832 | 409.3393 | 1160.0471 | 2551.5612 | 1789.7734 | 441.3402 | 43861.2318 | 35156.5439 | 173987.1948 | 167143.5182 | 219026.2928 | 183715.8288 | 148100.3887 | 17964.6224 | 21249.7956 | 50891.1542 | 95821.5281 | 199152.9935
228 | 5_18_23_MFB |  | 06/17/2023 11:20:08 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_MFB.877 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10367.761 | 4436.0222 | 547.3439 | 16117.0868 | 4955.5264 | 6503.481 | 3391.7363 | 7015.0565 | 9104.2339 | 24498.9933 | 17198.462 | 36001.3067 | 14498.0213 | 16038.999 | 4075.9147 | 9978.8177 | 48353.7001 | 21966.8766 | 14662.7426 | 598.6792 | 16748.4818 | 3052.9933 | 408.6725 | 1218.0519 | 2515.555 | 1712.8222 | 454.6739 | 43215.9367 | 34938.674 | 172329.8817 | 170626.9733 | 223347.0644 | 168623.2992 | 152571.0843 | 18334.425 | 21246.4559 | 50863.7281 | 95332.6319 | 200339.4856
229 | 5_18_23_MFB |  | 06/17/2023 11:26:26 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_MFB.878 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10150.2722 | 4599.4082 | 521.3429 | 16036.3311 | 4830.8168 | 6447.455 | 3359.7284 | 6981.0391 | 9140.9237 | 24943.0922 | 17232.0432 | 35727.9561 | 14498.0199 | 15982.2692 | 4144.6017 | 9975.4818 | 48689.4999 | 22050.3379 | 14355.623 | 642.6812 | 16896.6566 | 3008.317 | 376.6716 | 1174.0483 | 2464.2133 | 1678.8114 | 434.6733 | 44152.7951 | 35080.3539 | 185179.2427 | 167839.5484 | 220168.9896 | 183578.0273 | 150692.6801 | 18120.1515 | 21047.4938 | 50401.4243 | 93207.5701 | 195034.5698
230 | LBLANK |  | 06/17/2023 11:32:26 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.879 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2 | 0.66667 | 7.3333 | 2259.5121 | 1905.4606 | 36 | 0.66667 | 2.6667 | 28 | 25.3334 | 343.0453 | 9.3333 | 6268.709 | 1.3333 | 20.6667 | 14.6667 | 76.6669 | 298.6698 | -2946.8734 | 0 | 15.3333 | 1023.3701 | 0 | 1.3333 | 3.3333 | 45.5562 | 0.66667 | 130.0006 | 14 | 183096.8862 | 176878.9632 | 232224.4978 | 191770.5901 | 153207.1672 | 18596.7635 | 0.66667 | 10.6667 | 18.6667 | 41.3334
231 | LBLANK |  | 06/17/2023 11:38:51 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.880 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.6667 | 0 | 7.3333 | 644.0145 | 1923.463 | 28.6667 | 0.66667 | 1.3333 | 28.6667 | 22 | 289.2322 | 8 | 6091.9657 | 1.3333 | 14.6667 | 10 | 56.0001 | 315.3368 | -2848.7817 | 0 | 8.6667 | 1037.371 | 0 | 0 | 2.6667 | 38.1741 | 0 | 93.3336 | 8.6667 | 181653.9068 | 176141.264 | 231901.7594 | 192738.8369 | 154062.961 | 18837.7455 | 0 | 6 | 18.6667 | 40
232 | CCV 1ppb TRACE |  | 06/17/2023 11:44:51 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.881 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2196.1688 | 1020.0365 | 234.0019 | 1941.4653 | 2528.2238 | 1379.4 | 754.0199 | 1510.0798 | 1999.4738 | 5516.3985 | 5385.2783 | 7896.8487 | 8683.9722 | 3385.7345 | 907.3622 | 2145.4945 | 10373.7654 | 5205.6152 | 619.1368 | 140.0007 | 3645.1325 | 1422.0709 | 70.6668 | 251.3355 | 566.6779 | 421.3404 | 97.3337 | 9543.8537 | 7524.648 | 174742.9395 | 172655.0787 | 227959.8665 | 189577.6165 | 150964.8472 | 18596.7634 | 4510.7121 | 10957.5341 | 21108.7969 | 42511.3674
233 | CCV 1ppb TRACE |  | 06/17/2023 11:51:11 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.882 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2192.1682 | 1021.3699 | 220.6684 | 1927.4634 | 2609.5718 | 1442.0729 | 717.3514 | 1537.4161 | 1925.4632 | 5286.3117 | 5258.1488 | 8020.2509 | 8553.228 | 3439.0807 | 928.6969 | 2163.4973 | 10475.1726 | 5077.5696 | 642.0823 | 148.0008 | 3511.0982 | 1520.0809 | 78.0002 | 238.6687 | 532.6766 | 394.7212 | 98.0003 | 9393.7546 | 7582.0118 | 173668.5873 | 171473.6434 | 226133.1703 | 187341.7412 | 151458.6845 | 18484.618 | 4432.0211 | 10916.1691 | 20808.6578 | 42264.3755
234 | CCV 10 ppb MAJORS |  | 06/17/2023 11:57:11 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.883 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 20408.5684 | 8919.4508 | 933.3639 | 13656.5246 | 7539.3236 | 13031.9415 | 6834.9681 | 13998.8585 | 17620.1934 | 49413.3187 | 34184.7655 | 71436.1824 | 22543.7742 | 31369.7397 | 7971.5576 | 18999.2947 | 93633.8564 | 44962.6508 | 32918.0626 | 1328.0618 | 33476.5147 | 4970.1982 | 748.6863 | 2300.1854 | 5058.8959 | 3484.9849 | 884.0274 | 89588.0482 | 70910.2254 | 171142.452 | 160408.2368 | 212534.0175 | 177545.2355 | 143608.8938 | 17579.4768 | 42019.0631 | 101259.6275 | 181981.006 | 380403.4477
235 | CAL BLANK |  | 06/17/2023 12:25:05 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CAL BLANK.884 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0

## Raw Intensities RSDs
| Sample Id | R | Acquisition Time | QC Status | Dataset File | Method File | Na KED 23/23 Helium KED_LowLow | Mg KED 24/24 Helium KED_LowLow | Al KED 27/27 Helium KED_LowLow | P Oshift 31/47 Oxygen DRC | S (34) -1 34/34 Helium KED_LowLow | K KED 39/39 Helium KED_LowLow | Ca KED 43/43 Helium KED_LowLow | V KED 51/51 Helium KED_LowLow | Cr (52) 52/52 Helium KED_LowLow | Cr (53) NH3 53/53 Ammonia DRC | Fe (54) NH3 54/54 Ammonia DRC | Mn Oshift 55/71 Oxygen DRC | Fe (56) KED 56/56 Helium KED_LowLow | Co KED 59/59 Helium KED_LowLow | Ni KED 60/60 Helium KED_LowLow | Cu (63) KED 63/63 Helium KED_LowLow | Cu (65) NH3 65/65 Ammonia DRC | Zn NH3 66/66 Ammonia DRC | As 75/75 | As KED 75/75 Helium KED_LowLow | As Oshift 75/91 Oxygen DRC | Se (77) 77/77 | Se (77) Oshift 77/93 Oxygen DRC | Se (78) Oshift 78/94 Oxygen DRC | Se (80) Oshift 80/96 Oxygen DRC | Se (82) 82/82 | Se (82) Oshift 82/98 Oxygen DRC | Mo (95) NH3 95/95 Ammonia DRC | Mo (95) OShift 95/127 Oxygen DRC | Rh (IS1) 103/103 (IS) | Rh (IS2) 103/103 (IS) Oxygen DRC | Rh (IS4) 103/103 (IS) Ammonia DRC | Rh (IS5) 103/103 (IS) Ammonia DRC | Rh (IS6) 103/103 (IS) Oxygen DRC | Rh (IS7) 103/103 (IS) Helium KED_LowLow | Cd (111) NH3 111/111 Ammonia DRC | Cd (114) NH3 114/114 Ammonia DRC | Pb 208/208 | Pb NH3 208/208 Ammonia DRC
1 | sac blank |  | 6/16/2023 11:58:56 AM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\sac blank .650 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 21.65% |  | 14.66% | 5.852% | 2.415% | 15.06% | 173.2% | 173.2% | 29.4% | 11.76% | 5.216% | 14.78% | 0.6935% | 100% | 7.692% | 25.7% | 6.664% | 13.45% | 3.92% | 86.6% | 66.14% | 3.544% |  | 43.3% | 45.83% | 13.94% |  | 6.063% | 17.32% | 0.8211% | 1.123% | 1.118% | 1.066% | 2.031% | 1.658% | 86.6% | 26.96% | 8.813% | 14.39%
2 | sac blank cal |  | 6/16/2023 12:04:57 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\sac blank cal.651 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 24.74% | 173.2% | 25.71% | 1.582% | 1.172% | 32.87% |  | 91.65% | 14.52% | 28.64% | 3.108% | 43.48% | 3.453% | 65.47% | 18.9% | 41.66% | 16.88% | 3.737% | 2.311% | 100% | 49.49% | 3.487% | 86.6% | 49.49% | 124.9% | 55.01% | 0% | 8.53% | 56.77% | 0.9443% | 0.8027% | 1.274% | 0.301% | 2.06% | 1.339% | 173.2% | 22.71% | 16.38% | 12.74%
3 | sac blank cal |  | 6/16/2023 12:10:57 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\sac blank cal.652 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 65.47% |  | 13.58% | 3.715% | 3.289% | 7.524% |  | 132.3% | 15.75% | 52.85% | 4.832% | 33.33% | 1.864% | 33.33% | 13.29% | 30.14% | 9.524% | 10.17% | 1.563% |  | 47.19% | 4.239% |  | 173.2% | 78.06% | 25.17% |  | 2.1% | 43.3% | 1.708% | 0.8716% | 0.9613% | 0.6692% | 0.5924% | 1.444% | 86.6% | 50.33% | 30.06% | 21.68%
4 | CAL BLANK |  | 6/16/2023 12:16:57 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CAL BLANK.653 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 49.49% |  | 12.06% | 7.112% | 6.257% | 15.72% | 173.2% | 66.67% | 29.57% | 10% | 9.291% | 18.23% | 1.93% | 94.37% | 15.61% | 25% | 16.35% | 4.092% | 4.003% | 173.2% | 42.86% | 4.405% |  | 173.2% | 66.67% | 14.17% | 86.6% | 14.06% | 100% | 0.6452% | 0.1482% | 0.7484% | 0.4378% | 0.8227% | 1.556% | 114.6% | 34.99% | 14.75% | 20.38%
5 | LOW 0.01 |  | 6/16/2023 12:22:58 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LOW 0.01.654 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 6.186% | 41.81% | 6.329% | 4.12% | 1.459% | 1.506% | 28.64% | 17.56% | 4.967% | 13.85% | 5.478% | 6.613% | 7.338% | 9.042% | 4.545% | 8.517% | 7.303% | 7.031% | 2.233% | 57.28% | 10.7% | 6.437% | 173.2% | 58.08% | 10.83% | 38.39% | 34.64% | 3.417% | 4.721% | 1.137% | 0.8423% | 0.9212% | 0.873% | 0.507% | 1.977% | 19.06% | 11.54% | 0.9155% | 3.686%
6 | LOW 0.05 |  | 6/16/2023 12:28:58 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LOW 0.05.655 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 6.068% | 7.862% | 3.38% | 8.193% | 0.822% | 24.37% | 16.82% | 11.02% | 7.873% | 3.086% | 1.179% | 6.381% | 0.3454% | 5.112% | 13.35% | 11.03% | 3.004% | 2.321% | 3.633% | 16.67% | 7.985% | 2.475% | 43.3% | 29.57% | 10.92% | 14.15% | 47.89% | 2% | 3.811% | 1.348% | 0.5856% | 1.247% | 0.3364% | 0.4161% | 0.9849% | 13.34% | 3.592% | 4.791% | 0.4946%
7 | LOW 0.1 |  | 6/16/2023 12:34:59 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LOW 0.1.656 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 5.376% | 5.223% | 15.75% | 1.094% | 3.75% | 13.67% | 2.225% | 6.893% | 13.69% | 3.915% | 1.8% | 1.055% | 1.866% | 12.83% | 10.69% | 9.013% | 2.809% | 0.9616% | 2.961% | 22.91% | 5.817% | 1.002% | 66.67% | 9.184% | 10.67% | 3.258% | 52.07% | 0.404% | 3.696% | 0.7358% | 0.9433% | 1.578% | 0.8387% | 1.27% | 0.9627% | 2.189% | 6.189% | 4.205% | 0.9051%
8 | LOW 0.5 |  | 6/16/2023 12:40:59 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LOW 0.5.657 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 7.76% | 7.101% | 6.785% | 2.335% | 2.365% | 8.121% | 9.364% | 2.477% | 5.661% | 3.699% | 1.798% | 1.929% | 1.437% | 1.633% | 2.761% | 3.592% | 2.382% | 3.49% | 12.65% | 11.51% | 2.129% | 2.157% | 10.38% | 7.127% | 7.013% | 29.36% | 11.09% | 0.8855% | 1.402% | 0.7953% | 0.8605% | 0.2624% | 1.017% | 2.533% | 0.759% | 2.667% | 1.055% | 0.5295% | 1.811%
9 | STD 1 |  | 6/16/2023 12:46:58 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 1.658 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.407% | 2.006% | 4.138% | 1.007% | 1.574% | 2.544% | 7.791% | 4.608% | 0.8014% | 0.1714% | 0.8917% | 1.962% | 0.9238% | 0.5696% | 4.306% | 3.284% | 1.81% | 1.344% | 30.12% | 16.38% | 2.041% | 4.219% | 17.76% | 6.389% | 1.156% | 7.962% | 12.62% | 1.118% | 1.289% | 0.3183% | 1.265% | 1.929% | 0.8632% | 1.834% | 0.8067% | 1.389% | 1.472% | 0.6484% | 1.901%
10 | STD 5 |  | 6/16/2023 12:52:58 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 5.659 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.3709% | 2.021% | 6.144% | 1.303% | 1.915% | 1.807% | 1.878% | 1.055% | 0.4953% | 0.4888% | 2.021% | 0.8186% | 2.426% | 1.557% | 1.96% | 2.939% | 0.2715% | 0.9822% | 1.252% | 3.188% | 0.7391% | 2.168% | 5.92% | 2.275% | 1.398% | 4.722% | 2.143% | 0.637% | 0.6187% | 0.536% | 1.333% | 1.775% | 1.024% | 1.027% | 1.792% | 1.057% | 0.4838% | 1.294% | 1.303%
11 | STD 10 |  | 6/16/2023 12:58:58 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 10.660 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.013% | 2.128% | 3.779% | 1.192% | 2.259% | 1.859% | 0.2539% | 0.5421% | 1.026% | 0.3921% | 0.5965% | 0.5249% | 1.34% | 1.021% | 2.432% | 0.8251% | 0.7075% | 0.4397% | 1.02% | 1.321% | 0.4354% | 2.186% | 3.907% | 2.178% | 1.943% | 2.676% | 3.9% | 0.8215% | 0.5698% | 1.287% | 0.3345% | 0.3715% | 0.8411% | 0.951% | 0.05589% | 0.6778% | 0.2377% | 2.116% | 0.9095%
12 | STD 25 |  | 6/16/2023 1:04:58 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 25.661 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.6144% | 0.5622% | 1.93% | 3.428% | 1.67% | 1.158% | 0.3335% | 0.2411% | 1.201% | 0.3558% | 1.309% | 0.7137% | 0.7084% | 1.083% | 0.6031% | 0.7026% | 0.5183% | 0.5254% | 0.9052% | 1.77% | 0.1718% | 0.8716% | 1.927% | 1.035% | 0.6833% | 0.5434% | 0.7288% | 0.579% | 0.9508% | 0.4345% | 0.873% | 1.241% | 0.8751% | 12.49% | 0.2745% | 0.8127% | 0.6348% | 1.918% | 1.573%
13 | STD 50 |  | 6/16/2023 1:10:58 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 50.662 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.528% | 1.127% | 3.068% | 0.8858% | 1.793% | 1.439% | 1.683% | 1.179% | 1.956% | 0.6015% | 0.9329% | 0.8915% | 1.108% | 0.8592% | 0.9856% | 0.4448% | 0.8733% | 0.7997% | 0.07423% | 2.796% | 0.8947% | 0.7812% | 2.828% | 1.157% | 1.55% | 0.5652% | 4.098% | 0.9425% | 0.9259% | 0.3426% | 1.221% | 0.7681% | 0.7598% | 0.9539% | 1.873% | 0.2559% | 1.01% | 2.551% | 2.356%
14 | STD 100 |  | 6/16/2023 1:16:57 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 100.663 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.447% | 1.154% | 0.9832% | 0.9428% | 0.1803% | 2.456% | 1.684% | 0.6813% | 0.6526% | 0.418% | 0.8576% | 0.9417% | 1.674% | 1.476% | 0.5807% | 0.1436% | 0.2506% | 0.2675% | 0.9421% | 1.665% | 0.6133% | 1.24% | 2.183% | 0.5315% | 0.3645% | 1.803% | 0.6788% | 0.5008% | 0.6708% | 1.263% | 0.6764% | 0.9956% | 0.7512% | 0.532% | 0.3539% | 0.8082% | 1.212% | 4.09% | 0.4215%
16 | LBLANK |  | 6/16/2023 1:22:58 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.664 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 32.99% | 68.63% | 10.83% | 1.735% | 5.115% | 20.35% | 50% | 78.06% | 17.63% | 30.77% | 1.665% | 12.37% | 1.404% | 47.78% | 12.94% | 31.17% | 5.449% | 8.367% | 2.758% |  | 39.59% | 1.392% | 0% | 28.87% | 120.2% | 53.3% |  | 0.8662% | 19.79% | 0.9482% | 0.9773% | 0.6451% | 0.2886% | 1.639% | 1.666% | 33.33% | 33.48% | 27.49% | 25.63%
17 | LBLANK |  | 6/16/2023 1:41:50 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.665 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 20% | 173.2% | 76.71% | 3.252% | 18.83% | 22.05% | 173.2% | 33.33% | 11.8% | 27.6% | 8.582% | 34.44% | 7.684% | 33.33% | 21.65% | 34.44% | 13.42% | 40.39% | 3.257% |  | 5.973% | 2.945% | 86.6% | 66.14% | 89.21% | 28.9% |  | 2.718% | 5.945% | 0.5292% | 0.7508% | 0.6203% | 0.286% | 0.6941% | 7.71% | 86.6% | 31.91% | 24.74% | 29.24%
18 | ICV PE |  | 6/16/2023 1:47:50 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\ICV PE.666 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.814% | 0.5307% | 1.866% | 0.5423% | 1.108% | 1.438% | 0.7858% | 2.114% | 0.1002% | 1.19% | 1.722% | 0.5268% | 1.62% | 0.437% | 2.094% | 0.5399% | 1.98% | 1.091% | 1.291% | 3.311% | 0.9846% | 2.575% | 4.557% | 0.4415% | 1.597% | 1.124% | 3.502% | 1.085% | 0.7685% | 0.2963% | 0.3461% | 1.092% | 0.656% | 0.4973% | 0.4802% | 0.5884% | 0.9202% | 0.092% | 0.9294%
19 | CCV 1ppb TRACE |  | 6/16/2023 1:53:50 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.667 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 12.37% | 173.2% | 48.04% | 12.93% | 9.866% | 37.08% | 0% | 86.6% | 22.05% | 23.94% | 4.153% | 25% | 7.995% | 48.04% | 19.52% | 21.65% | 8.985% | 9.902% | 141.6% | 173.2% | 45.83% | 12.95% |  | 57.28% | 69.28% | 119.1% |  | 3.464% | 89.21% | 2.49% | 2.482% | 2.464% | 1.203% | 7.071% | 1.19% | 100% | 33.07% | 41.93% | 37.8%
20 | CCV 1ppb TRACE |  | 6/16/2023 2:00:07 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.668 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.611% | 4.717% | 3.891% | 3.149% | 2.732% | 2.7% | 3.659% | 1.912% | 1.923% | 0.7763% | 2.432% | 1.067% | 1.207% | 0.7357% | 1.44% | 2.418% | 0.2869% | 1.577% | 22.34% | 3.56% | 1.478% | 4.461% | 8.533% | 3.268% | 1.894% | 4.255% | 5.854% | 0.4352% | 1.337% | 1.609% | 2.152% | 0.7165% | 0.008299% | 6.066% | 1.123% | 2.379% | 1.507% | 3.006% | 0.9444%
21 | LFB |  | 6/16/2023 2:06:07 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.669 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.082% | 2.497% | 3.392% | 0.5833% | 2.276% | 2.363% | 1.189% | 2.958% | 3.745% | 1.86% | 2.161% | 0.6769% | 2.48% | 3.005% | 3.521% | 3.123% | 0.8469% | 1.914% | 2.043% | 7.533% | 0.4926% | 2.043% | 2.16% | 3.487% | 1.035% | 3.112% | 2.785% | 2.796% | 0.7952% | 0.4745% | 1.575% | 2.537% | 0.2438% | 1.024% | 7.511% | 2.239% | 2.087% | 1.012% | 1.9%
22 | LLOQ 0.01 |  | 6/16/2023 2:12:05 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LLOQ 0.01.670 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 5.207% | 5.973% | 46.98% | 1.685% | 3.38% | 19.05% | 50% | 43.72% | 12.62% | 9.081% | 5.311% | 10.12% | 0.8094% | 6.122% | 18.41% | 15.07% | 5.277% | 2.862% | 6.413% | 69.28% | 6.977% | 3.382% | 100% | 26.96% | 0% | 39.14% | 132.3% | 2.32% | 9.117% | 0.9267% | 0.7248% | 0.4548% | 0.4669% | 0.7163% | 0.5086% | 14.43% | 12.48% | 7.612% | 2.722%
23 | CCV 10 ppb MAJORS |  | 6/16/2023 2:18:05 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.671 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.307% | 1.705% | 1.33% | 0.9977% | 2.171% | 0.1152% | 2.027% | 0.3891% | 0.8371% | 0.8658% | 0.7046% | 0.405% | 0.9712% | 1.089% | 1.207% | 1.042% | 0.6043% | 0.6032% | 2.41% | 5.645% | 0.3151% | 2.409% | 5.767% | 2.747% | 1.983% | 3.642% | 3.564% | 1.065% | 0.595% | 0.5598% | 1.158% | 0.2376% | 0.9476% | 0.9511% | 0.4046% | 0.2291% | 0.4316% | 0.8733% | 0.4753%
24 | LLOQ 0.1 |  | 6/16/2023 2:24:05 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LLOQ 0.1.672 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 8.718% | 5.036% | 9.491% | 8.902% | 4.808% | 5.517% | 11.93% | 8.896% | 3.646% | 3.628% | 3.911% | 5.554% | 1.811% | 8.859% | 5.105% | 6.715% | 1.566% | 5.897% | 8.384% | 4.949% | 4.216% | 5.311% | 43.3% | 14.42% | 18.7% | 14.35% | 15.61% | 3.056% | 2.205% | 0.03196% | 0.558% | 1.709% | 0.535% | 0.9518% | 0.9947% | 3.533% | 5.029% | 2.429% | 1.594%
25 | 9_13_21_F_11_DUP |  | 6/16/2023 2:30:04 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_F_11_DUP.673 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.9349% | 3.113% | 3.744% | 1.794% | 0.8876% | 2.077% | 0.6281% | 4.014% | 5.344% | 2.303% | 1.559% | 1.21% | 1% | 5.815% | 3.215% | 1.454% | 1.273% | 1.584% | 1.573% | 2.884% | 0.5843% | 4.039% | 63.81% | 25.82% | 16.44% | 7.423% | 8.66% | 1.992% | 0.6518% | 0.9009% | 1.324% | 2.934% | 0.765% | 0.9075% | 1.778% | 7.575% | 4.15% | 1.477% | 1.63%
26 | LBLANK |  | 6/16/2023 2:47:08 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.675 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 41.66% | 173.2% | 59.59% | 4.976% | 2.297% | 30.32% | 100% | 107.9% | 11.86% | 31.21% | 4.927% | 51.63% | 0.9851% | 48.04% | 34.64% | 25% | 11.35% | 16.55% | 7.566% | 173.2% | 25% | 5.378% | 173.2% | 114.6% | 137.8% | 54.86% | 173.2% | 4.485% | 36.74% | 0.3744% | 0.8687% | 1.274% | 1.54% | 2.42% | 1.249% | 173.2% | 20.38% | 14.32% | 27.15%
27 | 0.5 STD |  | 6/16/2023 2:53:09 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\0.5 STD.676 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4.42% | 2.158% | 15.16% | 3.142% | 1.712% | 2.286% | 1.982% | 1.781% | 2.578% | 2.79% | 1.296% | 3.852% | 0.7132% | 2.372% | 2.709% | 2.812% | 3.694% | 1.928% | 13.98% | 15.64% | 3.934% | 3.757% | 26.27% | 13.67% | 12.39% | 3.722% | 10.13% | 1.586% | 0.7161% | 0.8426% | 0.909% | 0.3572% | 0.3128% | 2.212% | 1.801% | 1.165% | 2.061% | 0.5071% | 2.504%
28 | 0.1 STD |  | 6/16/2023 2:59:12 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\0.1 STD.677 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.309% | 2.564% | 17.29% | 5.146% | 2.462% | 6.655% | 12.07% | 3.606% | 4.21% | 4.855% | 3.53% | 2.205% | 3.662% | 3.249% | 13.97% | 16.37% | 3.026% | 2.685% | 2.06% | 16.11% | 4.4% | 4.669% | 19.25% | 37.15% | 18.83% | 27.07% | 43.3% | 0.5705% | 6.471% | 2.147% | 0.2205% | 0.8267% | 0.4524% | 0.9052% | 0.3296% | 0.4811% | 2.27% | 2.192% | 2.679%
29 | CCV 10 ppb MAJORS |  | 6/16/2023 3:05:12 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.678 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.6399% | 1.756% | 8.908% | 1.185% | 1.292% | 0.4353% | 1.637% | 1.014% | 0.1327% | 0.6725% | 0.2598% | 0.4859% | 1.402% | 2.323% | 2.426% | 2.423% | 0.4767% | 1.052% | 1.391% | 2.679% | 1.246% | 3.367% | 3.092% | 2.123% | 0.7944% | 4.192% | 1.412% | 0.7215% | 0.3644% | 1.111% | 0.4492% | 1.105% | 1.027% | 0.6158% | 0.921% | 1.221% | 1.069% | 0.7077% | 0.9558%
30 | 9_13_21_F_11_DUP |  | 6/16/2023 3:11:13 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_F_11_DUP.679 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.8259% | 1.486% | 6.872% | 3.485% | 1.086% | 1.71% | 0.5569% | 3.377% | 6.736% | 4.71% | 1.513% | 0.8904% | 0.4428% | 9.231% | 5.098% | 1.796% | 1.323% | 0.103% | 1.191% | 1.497% | 0.7689% | 5.116% | 44.1% | 24.74% | 2.703% | 4.662% | 26.65% | 2.068% | 0.3846% | 0.5127% | 1.097% | 0.2767% | 0.5001% | 0.9798% | 0.2759% | 16.86% | 5.88% | 3.175% | 2.46%
31 | 9_13_21_F_11_DUP_LSP |  | 6/16/2023 3:17:13 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_F_11_DUP_LSP.680 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.2855% | 0.5089% | 2.436% | 1.011% | 0.5193% | 1.973% | 0.5157% | 0.6393% | 0.624% | 0.643% | 0.9958% | 0.4493% | 0.8731% | 1.755% | 1.187% | 0.9497% | 1.195% | 0.8808% | 0.7639% | 6.281% | 0.3092% | 1.292% | 4.912% | 0.7139% | 1.199% | 2.917% | 4.022% | 1.744% | 0.2841% | 1.517% | 0.4821% | 0.4831% | 1.866% | 1.838% | 1.934% | 1.55% | 0.78% | 1.986% | 0.965%
32 | 9_13_21_UF_1 |  | 6/16/2023 3:23:13 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_UF_1.681 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.097% | 0.9033% | 4.147% | 1.895% | 2.586% | 1.621% | 0.9797% | 2.613% | 14.88% | 6.046% | 1.785% | 0.5463% | 2.462% | 8.858% | 4.334% | 1.513% | 1.639% | 3.891% | 2.741% | 9.12% | 1.131% | 5.471% | 9.116% | 7.873% | 9.608% | 2.767% | 26.96% | 0.6532% | 1.109% | 0.281% | 0.7427% | 0.2023% | 1.003% | 2.241% | 0.5457% | 22.09% | 14.01% | 4.143% | 0.701%
33 | 9_13_21_UF_6 |  | 6/16/2023 3:29:13 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_UF_6.682 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.686% | 3.09% | 8.182% | 4.363% | 2.207% | 1.973% | 1.836% | 1.618% | 11.01% | 5.334% | 1.798% | 0.6391% | 0.362% | 5.149% | 7.777% | 1.932% | 0.8777% | 1.016% | 2.06% | 1.314% | 1.129% | 3.139% | 22.91% | 25.32% | 22.12% | 10.11% | 24.74% | 0.9947% | 0.6772% | 1.183% | 0.7426% | 0.9056% | 2.44% | 0.7638% | 0.704% | 13.04% | 7.265% | 5.617% | 1.99%
34 | 9_13_21_UF_6_DUP |  | 6/16/2023 3:35:13 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_UF_6_DUP.683 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.194% | 2.009% | 9.334% | 0.826% | 2.204% | 0.729% | 0.672% | 2.669% | 1.32% | 2.862% | 3.426% | 1.32% | 0.9644% | 7.784% | 10.49% | 0.9839% | 1.43% | 0.1648% | 1.112% | 5.247% | 0.5827% | 0.3961% | 35.25% | 19.87% | 0% | 4.946% | 28.87% | 0.5594% | 1.09% | 1.278% | 0.5713% | 0.4614% | 0.5448% | 1.12% | 1.845% | 24.35% | 7.899% | 2.255% | 0.2782%
35 | 9_13_21_UF_8 |  | 6/16/2023 3:41:13 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_UF_8.684 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.125% | 2.917% | 6.41% | 1.421% | 2.587% | 1.557% | 0.9953% | 4.386% | 8.882% | 1.675% | 3.66% | 1.194% | 1.099% | 4.348% | 14.63% | 2.106% | 2.096% | 1.121% | 1.404% | 6.442% | 1.298% | 3.638% | 58.08% | 16.41% | 17.8% | 5.01% | 50.94% | 1.842% | 0.1976% | 1.747% | 0.2641% | 1.159% | 0.9784% | 1.742% | 0.5843% | 27.61% | 17.81% | 5.633% | 1.108%
36 | 9_13_21_UF_9 |  | 6/16/2023 3:47:13 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_UF_9.685 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.1735% | 1.4% | 17.35% | 1.634% | 2.606% | 1.695% | 0.1871% | 3.623% | 4.558% | 12.15% | 0.9276% | 1.021% | 2.8% | 4.996% | 6.766% | 2.319% | 0.3139% | 1.432% | 3.187% | 3.846% | 2.275% | 6.157% | 34.64% | 8.449% | 34.63% | 7.966% | 34.32% | 0.4123% | 2.036% | 0.7681% | 1.131% | 1.681% | 1.3% | 1.13% | 1.956% | 31.07% | 16.89% | 3.797% | 1.298%
37 | LBLANK |  | 6/16/2023 3:53:13 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.686 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 43.3% |  | 48.24% | 2.779% | 6.391% | 8.248% | 86.6% | 100% | 22.05% | 30.93% | 9.234% | 43.99% | 0.957% | 86.6% | 13.48% | 43.59% | 7.58% | 5.862% | 5.134% | 32.73% | 7.805% | 2.781% |  | 34.64% | 24.74% | 17.4% |  | 14.07% | 13.48% | 1.583% | 0.9188% | 0.8614% | 1.071% | 0.3249% | 1.125% | 173.2% | 43.3% | 36.09% | 35%
38 | CCV 1ppb TRACE |  | 6/16/2023 3:59:13 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.687 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.423% | 4.06% | 6.648% | 1.651% | 4.877% | 2.279% | 6.337% | 3.137% | 4.964% | 1.358% | 3.553% | 0.7541% | 1.162% | 5.399% | 6.913% | 1.59% | 0.497% | 2.651% | 14.34% | 13.64% | 1.342% | 3.454% | 15.42% | 3.693% | 2.57% | 3.957% | 5.588% | 1.083% | 2.464% | 0.2891% | 1.007% | 1.056% | 1.099% | 1.01% | 7.83% | 1.959% | 2.684% | 1.253% | 0.9793%
39 | CCV 1ppb TRACE |  | 6/16/2023 4:05:32 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.688 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.035% | 4.616% | 3.457% | 0% | 1.311% | 1.226% | 1.863% | 2.7% | 3.099% | 0.749% | 3.291% | 0.3512% | 0.2228% | 0.8736% | 0.5698% | 2.393% | 0.4363% | 1.303% | 12.47% | 8.498% | 1.098% | 2.376% | 13.43% | 2.703% | 0.4673% | 5.613% | 7.536% | 0.05939% | 2.628% | 2.002% | 0.2311% | 1.067% | 0.8012% | 1.069% | 2.357% | 2.751% | 1.206% | 2.814% | 1.417%
40 | CCV 10 ppb MAJORS |  | 6/16/2023 4:11:33 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.689 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.353% | 1.377% | 3.229% | 0.8019% | 2.841% | 3.392% | 0.8688% | 1.294% | 0.8108% | 0.9762% | 1.295% | 1.41% | 0.9205% | 0.6731% | 1.244% | 0.4652% | 0.1749% | 0.3626% | 1.776% | 2.994% | 1.139% | 1.133% | 5.921% | 2.647% | 1.227% | 5.362% | 2.477% | 0.6149% | 0.6086% | 0.9888% | 0.5536% | 0.3194% | 0.3864% | 0.835% | 1.568% | 1.138% | 0.525% | 0.9009% | 1.195%
41 | CCV 10 ppb MAJORS |  | 6/16/2023 4:17:52 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.690 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.8619% | 0.8903% | 5.34% | 1.447% | 3.096% | 1.329% | 0.5439% | 1.863% | 0.4469% | 0.9728% | 0.5535% | 0.2168% | 0.421% | 1.226% | 0.6656% | 0.9004% | 0.6748% | 0.7637% | 1.197% | 2.925% | 1.203% | 1.301% | 7.329% | 1.331% | 2.188% | 2.118% | 3.674% | 0.9167% | 0.4915% | 1.178% | 0.76% | 0.7053% | 1.42% | 0.6062% | 0.9535% | 0.6939% | 0.5624% | 1.725% | 0.813%
42 | 9_13_21_UF_10 |  | 6/16/2023 4:23:53 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_UF_10.691 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.492% | 0.9292% | 6.667% | 1.109% | 1.687% | 2.834% | 0.5106% | 3.964% | 12.08% | 6.713% | 1.993% | 0.5991% | 0.6333% | 12.01% | 5.707% | 2.708% | 1.842% | 2.634% | 3.262% | 5.045% | 0.6401% | 3.334% | 43.3% | 22.43% | 17.8% | 9.537% | 36.74% | 0.819% | 0.7551% | 1.462% | 1.066% | 0.7566% | 0.5826% | 0.2239% | 1.46% | 18.9% | 16.47% | 3.642% | 2.574%
43 | 9_13_21_UF_11 |  | 6/16/2023 4:29:52 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_UF_11.692 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.8586% | 1.638% | 13.93% | 0.08532% | 0.5063% | 0.7132% | 0.5203% | 2.608% | 5.634% | 0.9993% | 4.382% | 0.1375% | 2.993% | 23.51% | 8.804% | 1.695% | 1.151% | 1.474% | 1.04% | 3.198% | 0.9961% | 8.847% | 25% | 26.75% | 11.76% | 10.46% | 14.29% | 1.162% | 2.556% | 1.227% | 1.29% | 0.6789% | 3.544% | 0.4046% | 0.4105% | 12.74% | 12.72% | 4.152% | 2.666%
44 | 9_13_21_LAB_A |  | 6/16/2023 4:35:53 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_LAB_A.693 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 16.67% |  | 27.87% | 2.195% | 0.7858% | 15.75% | 69.28% | 173.2% | 18.01% | 10.7% | 1.989% | 2.79% | 2.456% | 62.45% | 13.33% | 1.627% | 1.078% | 2.644% | 8.674% | 49.49% | 12.59% | 6.911% | 100% | 173.2% | 21.65% | 10.05% |  | 18.62% | 44.61% | 1.869% | 0.8749% | 0.654% | 1.593% | 0.6699% | 0.573% | 173.2% | 36.66% | 2.697% | 0.9852%
45 | 9_13_21_BB |  | 6/16/2023 4:41:52 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_BB.694 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 52.92% |  | 15.86% | 2.392% | 5.836% | 21.09% | 62.45% | 49.49% | 13.89% | 24.05% | 2.323% | 9.805% | 1.026% | 86.6% | 7.531% | 8.425% | 5.035% | 4.995% | 4.182% | 124.9% | 2.9% | 7.932% |  | 86.6% | 78.06% | 5.718% | 173.2% | 14.65% | 27.55% | 0.2672% | 0.4443% | 0.6803% | 1.461% | 0.6236% | 1.267% | 100% | 16.88% | 4.935% | 5.652%
46 | 9_13_21_FB |  | 6/16/2023 4:47:52 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_FB.695 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.544% | 20.32% | 1.666% | 2.035% | 2.286% | 17.93% | 7.671% | 22.53% | 13.21% | 8.8% | 3.083% | 2.455% | 1.705% | 10.66% | 7.278% | 2.301% | 2.662% | 0.5509% | 2.706% | 50.76% | 6.476% | 2.047% | 86.6% | 34.64% | 33.07% | 22.06% | 86.6% | 3.616% | 4.356% | 0.5812% | 0.8864% | 0.3722% | 0.5201% | 2.576% | 4.93% | 17.3% | 16.85% | 1.534% | 1.939%
47 | 5_18_23_MFB |  | 6/16/2023 4:53:52 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_MFB.696 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.74% | 2.405% | 2.627% | 1.841% | 2.375% | 0.7868% | 6.248% | 0.8223% | 2.722% | 1.328% | 1.975% | 0.8631% | 1.2% | 0.6271% | 1.163% | 1.906% | 0.9051% | 0.9815% | 4.82% | 2.22% | 1.239% | 3.351% | 3.842% | 4.144% | 1.431% | 4.989% | 3.928% | 0.8566% | 0.929% | 10.33% | 0.4279% | 0.3683% | 0.6024% | 0.7247% | 0.5646% | 0.6274% | 0.4324% | 4.066% | 1.092%
48 | 5_18_23_DB |  | 6/16/2023 4:59:52 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_DB.697 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.704% | 33.33% | 4.225% | 1.637% | 2.138% | 8.646% | 24.35% | 36.3% | 14.58% | 8.032% | 4.174% | 19.72% | 0.07664% | 4.558% | 27.99% | 6.138% | 3.111% | 4.819% | 1.6% | 89.21% | 12.33% | 5.418% | 120.2% | 34.64% | 11.17% | 37.33% | 124.9% | 10.5% | 2.68% | 1.234% | 1.118% | 1.078% | 1.644% | 0.7032% | 0.7146% | 9.984% | 14.34% | 4.233% | 3.621%
49 | 5_18_23_T_239 |  | 6/16/2023 5:05:52 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_T_239.698 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.893% | 1.012% | 1.573% | 1.015% | 3.3% | 2.226% | 1.316% | 1.757% | 5.654% | 1.464% | 0.2584% | 0.4467% | 0.8583% | 3.07% | 1.018% | 4.648% | 2.854% | 0.8492% | 7.386% | 7.088% | 1.955% | 2.285% | 9.005% | 5.266% | 1.635% | 9.622% | 5.879% | 1.554% | 1.025% | 6.404% | 0.5446% | 1.433% | 1.146% | 0.595% | 2.293% | 0.988% | 1.145% | 1.979% | 1.438%
50 | 5_18_23_T_239 |  | 6/16/2023 5:11:52 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_T_239.699 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.4455% | 1.12% | 1.56% | 1.077% | 1.629% | 3.372% | 1.06% | 3.303% | 1.869% | 2.108% | 0.4223% | 0.08072% | 0.8778% | 3.106% | 0.4578% | 2.178% | 1.469% | 0.4611% | 3.352% | 8.21% | 0.5631% | 0.5802% | 12.72% | 5.277% | 3.074% | 4.564% | 14.22% | 1.561% | 2.324% | 1.134% | 1.015% | 0.3858% | 0.5237% | 0.1556% | 4.989% | 0.3251% | 0.6606% | 0.8017% | 0.1804%
51 | 5_18_23_M_224 |  | 6/16/2023 5:17:52 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_M_224.700 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.9704% | 8.515% | 1.167% | 1.091% | 0.525% | 3.303% | 2.188% | 2.636% | 5.707% | 8.622% | 2.6% | 1.082% | 2.314% | 15.89% | 22.65% | 8.969% | 5.145% | 0.9671% | 7.106% | 69.28% | 10.97% | 4.322% | 15.75% | 95.78% | 18.37% | 5.491% | 25% | 1.01% | 1.937% | 0.8471% | 0.2931% | 0.3809% | 2.088% | 0.4269% | 1.043% | 13.86% | 28.31% | 10.92% | 2.149%
52 | LBLANK |  | 6/16/2023 5:23:52 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.701 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 31.49% | 173.2% | 86.6% | 1.972% | 8.358% | 0% | 100% | 50% | 26.36% | 38.21% | 6.513% | 48.04% | 1.863% | 91.65% | 12.37% | 39.79% | 16.46% | 6.531% | 3.752% | 173.2% | 12.37% | 2.949% |  | 86.6% | 69.28% | 0.2064% |  | 5.075% | 28.64% | 7.484% | 1.298% | 3.372% | 1.101% | 0.909% | 1.297% |  | 20.83% | 19.52% | 9.758%
53 | CCV 1ppb TRACE |  | 6/16/2023 5:29:53 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.702 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.569% | 2.087% | 9.024% | 4.532% | 2.115% | 1.382% | 5.184% | 5.567% | 2.758% | 0.5538% | 3.014% | 1.597% | 1.727% | 0.8285% | 10.06% | 2.392% | 1.36% | 1.903% | 3.027% | 3.608% | 1.596% | 1.502% | 15.85% | 7.402% | 1.834% | 6.338% | 11.36% | 0.8238% | 1.335% | 0.6882% | 0.8524% | 0.7412% | 1.292% | 1.373% | 0.2669% | 1.379% | 2.525% | 2.048% | 0.5782%
54 | CCV 1ppb TRACE |  | 6/16/2023 5:36:12 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.703 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.5861% | 2.424% | 18.58% | 2.507% | 3.535% | 2.293% | 5.226% | 5.486% | 3.132% | 2.998% | 1.632% | 0.2542% | 1.754% | 1.772% | 2.305% | 3.23% | 2.04% | 1.747% | 0.1315% | 14.45% | 2.177% | 7.541% | 4.149% | 5.128% | 6.664% | 5.446% | 14.59% | 2.545% | 2.189% | 1.197% | 0.7841% | 0.5756% | 0.9823% | 0.9679% | 2.215% | 1.819% | 1.43% | 2.121% | 0.8298%
55 | LFB |  | 6/16/2023 5:42:12 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.704 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.02% | 1.46% | 2.658% | 1.752% | 1.455% | 1.468% | 0.7191% | 1.27% | 1.3% | 0.2134% | 0.1579% | 0.3534% | 0.6186% | 0.7509% | 1.402% | 1.658% | 0.3235% | 0.3315% | 1.3% | 3.589% | 0.4156% | 3.309% | 3.331% | 3.322% | 1.107% | 3.872% | 3.773% | 1.356% | 0.3466% | 0.8486% | 0.7544% | 0.7418% | 0.4742% | 1.207% | 1.048% | 1.016% | 0.8539% | 0.7253% | 1.824%
56 | CCV 10 ppb MAJORS |  | 6/16/2023 5:48:11 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.705 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.4% | 0.2635% | 3.535% | 1.47% | 1.164% | 1.224% | 1.659% | 1.203% | 1.494% | 0.7949% | 0.7825% | 0.4705% | 0.4078% | 0.9236% | 1.179% | 1.149% | 0.1726% | 0.6819% | 1.992% | 0.9638% | 0.3294% | 2.199% | 2.26% | 2.139% | 0.5579% | 2.594% | 1.959% | 0.978% | 0.2458% | 1.081% | 0.4757% | 0.6096% | 0.7384% | 0.7501% | 1.019% | 0.8589% | 0.3417% | 1.315% | 1.502%
57 | CCV 10 ppb MAJORS |  | 6/16/2023 5:54:30 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.706 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.443% | 1.129% | 5.324% | 2.367% | 2.272% | 1.452% | 1.708% | 1.741% | 1.497% | 1.028% | 1.256% | 1.348% | 1.339% | 1.677% | 1.535% | 1.247% | 0.6493% | 0.7102% | 2.286% | 1.751% | 0.4447% | 2.57% | 4.723% | 2.228% | 0.8794% | 1.196% | 3.103% | 0.9959% | 0.8041% | 1.002% | 1.19% | 0.2564% | 1.13% | 1.691% | 1.553% | 1.233% | 0.5932% | 1.244% | 1.383%
58 | 5_18_23_MFB |  | 6/16/2023 6:00:31 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_MFB.707 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.9593% | 3.025% | 1.449% | 0.8699% | 1.468% | 1.803% | 2.331% | 0.1615% | 1.925% | 2.492% | 2.087% | 0.815% | 0.6648% | 1.101% | 2.267% | 1.738% | 2.208% | 2.397% | 1.673% | 9.742% | 1.886% | 1.218% | 3.626% | 1.819% | 2.523% | 3.055% | 8.333% | 2.076% | 0.47% | 1.004% | 0.5323% | 10.78% | 1.046% | 1.368% | 1.773% | 2.356% | 2.775% | 1.545% | 0.6968%
59 | 10_1_21_W_1 |  | 6/16/2023 6:06:31 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_W_1.708 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.841% | 3.851% | 7.54% | 2.439% | 1.606% | 5.55% | 0.8551% | 18.22% | 11.75% | 7.048% | 1.51% | 0.2596% | 1.829% | 18.37% | 1.872% | 0.5398% | 2.255% | 3.008% | 304.1% | 1.741% | 1.434% | 5.324% | 91.65% | 81.04% | 66.14% | 12.83% | 34.64% | 0.4729% | 0.261% | 0.8712% | 1.276% | 0.6219% | 0.2237% | 0.9858% | 2.059% | 34.64% | 37.9% | 3.705% | 7.213%
60 | 10_1_21_W_6 |  | 6/16/2023 6:12:31 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_W_6.709 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.654% | 0.3827% | 4.059% | 0.5794% | 2.111% | 4.514% | 2.275% | 5.457% | 13.91% | 12.27% | 0.916% | 1.409% | 1.509% | 7.16% | 14.78% | 3.932% | 2.075% | 3.316% | 74.54% | 17.03% | 2.388% | 5.989% | 98.97% | 21.65% | 29.57% | 18.07% | 86.6% | 1.116% | 0.3793% | 0.9056% | 0.8356% | 2.591% | 1.434% | 1.621% | 0.7154% | 6.928% | 11.66% | 2.613% | 3.919%
61 | 10_1_21_W_8 |  | 6/16/2023 6:18:32 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_W_8.710 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.988% | 3.275% | 4.635% | 1.83% | 3.645% | 3.795% | 0.6077% | 14.36% | 19.86% | 3.485% | 1.064% | 0.09441% | 1.01% | 14.78% | 23.67% | 2.612% | 1.305% | 1.643% | 65.02% | 8.303% | 1.118% | 10.02% | 69.28% | 52.92% | 28.39% | 36.57% | 100% | 0.5455% | 1.274% | 1.643% | 0.409% | 0.09719% | 0.5937% | 0.6828% | 1.014% | 30.12% | 25.46% | 9.079% | 4.467%
62 | 10_1_21_W_9 |  | 6/16/2023 6:24:32 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_W_9.711 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.607% | 1.449% | 5.881% | 0.9485% | 1.805% | 3.145% | 1.664% | 13.71% | 8.642% | 5.803% | 1.336% | 0.9455% | 2.305% | 14.43% | 7.37% | 5.147% | 1.113% | 3.675% | 123.6% | 8.969% | 2.193% | 0.9555% | 86.6% | 57.28% | 14.29% | 18.29% | 91.65% | 2.476% | 1.434% | 0.8161% | 0.1883% | 1.016% | 0.9502% | 1.183% | 1.215% | 52.04% | 24.4% | 4.856% | 5.133%
63 | 10_1_21_W_10 |  | 6/16/2023 6:30:32 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_W_10.712 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.187% | 0.9406% | 3.785% | 1.299% | 1.389% | 14.35% | 0.5005% | 7.476% | 2.56% | 9.01% | 0.2785% | 0.322% | 0.5315% | 11.9% | 13.01% | 6.032% | 0.9048% | 3.569% | 110.7% | 4.41% | 0.6455% | 2.001% | 86.6% | 20.38% | 44.1% | 49.73% | 100% | 1.473% | 2.486% | 1.056% | 1.327% | 1.253% | 0.1956% | 1.093% | 0.7761% | 32.78% | 10.5% | 1.768% | 0.9129%
64 | 10_1_21_W_11 |  | 6/16/2023 6:36:31 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_W_11.713 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.092% | 1.908% | 3.98% | 0.4071% | 2.121% | 3.949% | 2.704% | 8.578% | 6.085% | 1.11% | 0.7811% | 0.5285% | 0.9656% | 5.728% | 15.88% | 5.262% | 1.243% | 3.189% | 71.76% | 11.83% | 1.784% | 0.9638% | 114.6% | 44.61% | 28.87% | 24.9% | 49.49% | 2.119% | 1.33% | 0.8303% | 0.9175% | 0.5402% | 0.6778% | 0.5215% | 1.719% | 27.74% | 17.75% | 2.339% | 1.76%
65 | 10_1_21_W_11_LD |  | 6/16/2023 6:42:31 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_W_11_LD.714 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 69.28% |  | 114.6% | 5.98% | 8.606% | 12.37% | 173.2% | 173.2% | 55.11% | 36.32% | 9.846% | 39.03% | 5.444% | 173.2% | 14.29% | 17.32% | 36.07% | 10.21% | 47.32% | 173.2% | 105.4% | 4.081% |  |  | 100% | 91.15% |  | 4.581% | 50% | 3.855% | 2.406% | 1.592% | 7.767% | 3.211% | 5.794% | 173.2% | 107.9% | 43.3% | 16.43%
66 | 10_1_21_W_11_LD |  | 6/16/2023 6:48:48 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_W_11_LD.715 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.452% | 3.021% | 2.193% | 2.359% | 1.585% | 1.982% | 6.077% | 12.92% | 12.42% | 1.671% | 1.296% | 0.7061% | 1.194% | 33.93% | 19.25% | 2.229% | 2.704% | 1.682% | 15.66% | 10.38% | 2.856% | 4.676% | 91.65% | 50% | 12.37% | 3.159% | 86.6% | 3.533% | 1.251% | 1.642% | 0.2073% | 0.7535% | 0.251% | 1.196% | 0.2052% | 39.36% | 16.56% | 2.781% | 2.02%
67 | 10_1_21_W_11_DUP |  | 6/16/2023 6:54:48 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_W_11_DUP.716 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.39% | 4.368% | 2.172% | 3.564% | 1.907% | 6.637% | 2.035% | 8.191% | 16.11% | 3.731% | 0.7343% | 1.26% | 1.163% | 17.32% | 6.186% | 5.331% | 2.515% | 3.402% | 61.14% | 8.816% | 0.9167% | 2.118% | 91.65% | 50.94% | 30% | 6.288% | 43.3% | 2.595% | 2.458% | 0.7506% | 0.3305% | 0.0735% | 0.6337% | 1.047% | 1.247% | 18.87% | 10.73% | 0.9111% | 1.202%
68 | LBLANK |  | 6/16/2023 7:00:48 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.717 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 86.6% | 173.2% | 66.14% | 2.958% | 1.524% | 29.04% |  | 50% | 6.25% | 44.06% | 5.174% | 9.116% | 0.2273% | 173.2% | 4.681% | 64.16% | 20.16% | 4.057% | 7.872% |  | 81.04% | 6.15% |  | 173.2% | 173.2% | 36.45% | 86.6% | 10.58% | 49.49% | 1.322% | 1.149% | 1.42% | 0.7713% | 0.2646% | 1.73% | 173.2% | 14.29% | 16.37% | 22.88%
69 | CCV 1ppb TRACE |  | 6/16/2023 7:06:48 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.718 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 6.101% | 3.89% | 6.367% | 3.157% | 2.6% | 2.779% | 7.004% | 3.398% | 1.523% | 4.387% | 1.257% | 0.9959% | 2.409% | 3.727% | 1.415% | 2.086% | 2.413% | 1.009% | 23.06% | 12.41% | 1.11% | 3.408% | 17.6% | 2.735% | 6.101% | 9.393% | 24.27% | 1.905% | 0.9901% | 0.2421% | 0.6822% | 0.3343% | 0.3424% | 1.82% | 0.8832% | 2.301% | 1.14% | 1.4% | 1.038%
70 | CCV 1ppb TRACE |  | 6/16/2023 7:13:07 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.719 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.889% | 2.939% | 0.7715% | 3.09% | 2.189% | 4.581% | 3.03% | 2.01% | 3.478% | 3.684% | 1.811% | 1.417% | 1.509% | 1.713% | 4.786% | 1.381% | 1.127% | 0.9127% | 9.141% | 11.73% | 1.47% | 2.728% | 19% | 7.127% | 8.437% | 10.54% | 14.11% | 3.058% | 1.189% | 0.07654% | 0.387% | 1.265% | 1.57% | 1.173% | 0.4386% | 1.838% | 0.7936% | 0.861% | 2.464%
71 | CCV 10 ppb MAJORS |  | 6/16/2023 7:19:08 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.720 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.5006% | 0.9946% | 2.872% | 1.263% | 2.262% | 1.261% | 1.575% | 0.6466% | 1.228% | 0.8885% | 1.339% | 0.2854% | 1.468% | 1.135% | 1.385% | 1.613% | 0.9199% | 1.074% | 1.294% | 2.76% | 0.426% | 4.508% | 3.832% | 2.658% | 2.491% | 3.327% | 2.996% | 0.4722% | 0.5943% | 0.9196% | 0.458% | 1.032% | 0.6228% | 1.176% | 0.2518% | 0.3936% | 1.206% | 1.482% | 1.405%
72 | CCV 10 ppb MAJORS |  | 6/16/2023 7:25:27 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.721 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.908% | 1.719% | 5.61% | 0.8532% | 1.306% | 0.8378% | 1.627% | 0.8906% | 0.8751% | 1.099% | 0.8994% | 1.551% | 0.6306% | 0.6546% | 0.4947% | 0.438% | 0.6993% | 0.3151% | 1.298% | 1.798% | 0.3952% | 0.5737% | 2.415% | 1.49% | 0.9994% | 2.143% | 3.464% | 0.7036% | 0.8285% | 0.5563% | 0.8365% | 0.739% | 1.099% | 0.7696% | 1.348% | 0.8942% | 0.4094% | 1.895% | 0.3041%
73 | 10_1_21_W_11_DUP_LSP |  | 6/16/2023 7:31:27 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_W_11_DUP_LSP.722 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.812% | 0.7274% | 1.567% | 1.496% | 0.3172% | 1.181% | 0.6549% | 1.239% | 0.678% | 1.903% | 0.7534% | 0.7531% | 0.2556% | 0.7446% | 0.821% | 0.6892% | 0.7336% | 0.9596% | 2.126% | 2.11% | 0.4818% | 0.5867% | 4.471% | 2.19% | 1.411% | 2.695% | 1.25% | 0.8071% | 0.219% | 0.5045% | 0.874% | 0.3329% | 0.8878% | 0.4816% | 0.6117% | 0.244% | 1.077% | 1.18% | 1.209%
74 | 10_1_21_F_1 |  | 6/16/2023 7:37:27 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_F_1.723 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.7567% | 1.542% | 1.878% | 2.06% | 0.7038% | 2.73% | 1.078% | 3.416% | 7.697% | 4.702% | 0.7095% | 0.6113% | 0.6705% | 2.585% | 4.715% | 5.139% | 1.972% | 0.3758% | 2.582% | 8.949% | 0.3461% | 3.719% | 78.73% | 45.81% | 38.77% | 3.337% | 25% | 0.4468% | 0.6117% | 2.205% | 0.382% | 1.085% | 0.3892% | 0.6383% | 0.8666% | 23.21% | 5.634% | 1.404% | 3.494%
75 | 10_1_21_F_1_DUP |  | 6/16/2023 7:43:27 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_F_1_DUP.724 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.296% | 0.3468% | 3.694% | 1.309% | 1.409% | 3.5% | 1.155% | 3.368% | 6.25% | 4.2% | 1.608% | 0.4087% | 1.744% | 14.37% | 6.183% | 3.286% | 1.467% | 0.3315% | 1.743% | 6.727% | 0.4932% | 5.321% | 78.06% | 14.32% | 11.95% | 15.21% | 21.65% | 0.9509% | 0.7073% | 1.078% | 0.9386% | 0.4764% | 0.7746% | 0.9433% | 0.8129% | 16.5% | 22.63% | 1.192% | 1.171%
76 | 10_1_21_F_6 |  | 6/16/2023 7:49:26 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_F_6.725 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.22% | 1.403% | 0.4662% | 2.967% | 1.216% | 0.9836% | 0.9406% | 3.564% | 11.2% | 6.145% | 1.336% | 0.5125% | 0.823% | 16.15% | 8.848% | 3.095% | 0.6012% | 0.5524% | 0.6452% | 7.827% | 0.5361% | 7.957% | 43.3% | 10.83% | 30.46% | 4.777% | 52.68% | 1.189% | 1.926% | 0.4811% | 0.5873% | 1.195% | 0.7977% | 1.239% | 0.628% | 14.84% | 10.61% | 6.87% | 1.633%
77 | 10_1_21_F_6_DUP |  | 6/16/2023 7:55:26 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_F_6_DUP.726 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.389% | 2.163% | 9.751% | 0.9588% | 1.548% | 1.589% | 0.4743% | 6.479% | 8.04% | 4.203% | 0.317% | 0.6769% | 0.7569% | 18.03% | 7.914% | 3.106% | 0.7442% | 0.1552% | 1.339% | 3.586% | 0.6093% | 5.255% | 101.3% | 26.19% | 29.17% | 5.37% | 52.92% | 0.5524% | 0.8327% | 1.425% | 0.2286% | 0.498% | 0.5496% | 0.5948% | 0.9555% | 18.42% | 7.809% | 5.905% | 3.351%
78 | 10_1_21_F_8 |  | 6/16/2023 8:01:27 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_F_8.727 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.032% | 2.758% | 4.766% | 3.12% | 2.314% | 2.997% | 1.649% | 1.467% | 8.628% | 7.173% | 0.6681% | 0.7053% | 0.4307% | 13.33% | 6.555% | 1.267% | 1.468% | 0.8764% | 1.724% | 4.819% | 2.032% | 2.628% | 53.93% | 16.41% | 11.44% | 12.66% | 21.65% | 1.657% | 0.8916% | 0.8553% | 0.4319% | 0.2814% | 0.9658% | 0.8528% | 2.239% | 45.08% | 3.593% | 2.558% | 2.282%
79 | 10_1_21_F_9 |  | 6/16/2023 8:07:27 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_F_9.728 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.3885% | 0.4573% | 8.951% | 2.298% | 1.129% | 1.587% | 1.743% | 5.354% | 8.77% | 4.556% | 0.1562% | 0.6995% | 0.9637% | 14.32% | 11.07% | 1.888% | 0.5872% | 0.9376% | 2.691% | 8.199% | 2.235% | 1.792% | 32.73% | 6.662% | 20.82% | 5.594% | 38.57% | 0.5619% | 0.6564% | 1.202% | 0.5019% | 0.6072% | 0.4501% | 1.3% | 1.393% | 13.78% | 7.997% | 1.819% | 1.549%
80 | 10_1_21_F_9_DUP |  | 6/16/2023 8:13:27 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_F_9_DUP.729 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.001% | 0.7551% | 4.21% | 2.348% | 1.633% | 5.492% | 1.292% | 4.103% | 9.663% | 3.652% | 0.7765% | 0.1026% | 0.9765% | 7.376% | 5.155% | 3.251% | 1.233% | 0.7719% | 0.6806% | 7.5% | 0.2266% | 2.155% | 57.28% | 36.06% | 22.58% | 8.148% | 72.65% | 1.212% | 1.136% | 1.132% | 0.1082% | 0.6324% | 0.2692% | 1.257% | 0.7% | 17.76% | 4.052% | 2.952% | 2.922%
81 | 10_1_21_F_10 |  | 6/16/2023 8:19:27 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_F_10.730 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.344% | 1.903% | 0.9799% | 1.847% | 2.035% | 1.772% | 0.5105% | 4.586% | 5.83% | 8.604% | 0.67% | 1.048% | 2.164% | 13.04% | 5.303% | 4.589% | 1.547% | 0.5195% | 1.126% | 8.437% | 0.3611% | 3.098% | 41.66% | 30.14% | 24.12% | 9.022% | 53.29% | 0.558% | 0.6393% | 0.8002% | 0.7045% | 0.2221% | 0.7774% | 0.7656% | 1.287% | 5.413% | 12.8% | 1.926% | 1.348%
82 | 10_1_21_F_11 |  | 6/16/2023 8:25:27 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_F_11.731 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.493% | 2.605% | 2.99% | 2.02% | 1.581% | 1.684% | 1.519% | 3.77% | 5.931% | 2.172% | 0.6909% | 0.5447% | 1.35% | 11.26% | 10.48% | 2.249% | 0.6332% | 0.3085% | 2.008% | 1.051% | 0.8064% | 4.301% | 32.73% | 19.68% | 24.24% | 11.66% | 26.65% | 1.263% | 0.4393% | 1.547% | 0.4645% | 1.041% | 0.7159% | 0.54% | 1.854% | 7.531% | 8.541% | 1.211% | 3.396%
83 | LBLANK |  | 6/16/2023 8:31:27 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.732 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 57.28% | 173.2% | 78.06% | 7.708% | 2.279% | 22.18% | 173.2% | 173.2% | 12.74% | 25.64% | 0.7062% | 22.91% | 2.117% |  | 20.38% | 45.07% | 18.16% | 9.747% | 7.428% |  | 31.49% | 6.233% |  | 100% |  | 42.69% | 173.2% | 24.9% | 6.186% | 1.747% | 1.257% | 0.8401% | 0.8553% | 0.4416% | 1.881% |  | 20% | 53.91% | 27.26%
84 | CCV 1ppb TRACE |  | 6/16/2023 8:37:27 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.733 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 5.821% | 4.015% | 9.324% | 4.212% | 2.99% | 3.925% | 5.089% | 2.153% | 1.919% | 4.066% | 2.272% | 0.5385% | 1.946% | 1.211% | 3.237% | 0.7652% | 1.032% | 3.657% | 38.95% | 6.831% | 1.802% | 3.926% | 3.039% | 4.669% | 1.389% | 8.833% | 5.634% | 0.4867% | 0.2943% | 1.783% | 0.6754% | 0.5846% | 0.6097% | 1.136% | 0.8885% | 0.7964% | 1.496% | 0.2717% | 0.9107%
85 | CCV 1ppb TRACE |  | 6/16/2023 8:43:46 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.734 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.032% | 1.366% | 6.287% | 2.4% | 1.949% | 1.834% | 7.531% | 4.719% | 2.398% | 1.821% | 0.9124% | 0.4056% | 1.267% | 1.317% | 1.03% | 4.838% | 1.478% | 0.7363% | 32.6% | 5.208% | 1.735% | 4.335% | 6.811% | 11.19% | 2.309% | 12.79% | 9.818% | 0.9569% | 1.649% | 2.358% | 0.458% | 0.7415% | 1.106% | 0.4806% | 1.06% | 0.8177% | 0.5784% | 1.258% | 0.3843%
86 | LFB |  | 6/16/2023 8:49:46 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.735 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.7596% | 1.543% | 5.496% | 0.4657% | 2.188% | 2.86% | 0.7745% | 0.7244% | 0.847% | 0.1532% | 1.564% | 0.9491% | 1.176% | 1.441% | 0.5469% | 0.6667% | 0.3794% | 0.2771% | 2.976% | 1.105% | 1.032% | 2.838% | 3.725% | 2.416% | 2.965% | 3.376% | 2.502% | 1.034% | 0.4947% | 2.448% | 0.8675% | 0.9842% | 0.7748% | 0.3062% | 0.6025% | 0.7177% | 0.5109% | 1.033% | 0.9764%
87 | CCV 10 ppb MAJORS |  | 6/16/2023 8:55:46 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.736 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.821% | 1.883% | 5.04% | 0.1765% | 1.594% | 1.253% | 1.51% | 0.4522% | 0.6925% | 1.591% | 0.3637% | 0.3956% | 0.5802% | 0.3809% | 0.9072% | 0.3507% | 0.5547% | 0.1734% | 1.051% | 1.386% | 0.7525% | 3.755% | 2.873% | 2.161% | 0.889% | 2.521% | 3.193% | 1.262% | 0.6124% | 0.9861% | 0.4187% | 0.4929% | 1.466% | 0.1154% | 0.2436% | 0.7061% | 0.2039% | 0.4783% | 1.832%
88 | 10_1_21_UF_1 |  | 6/16/2023 9:01:46 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_UF_1.737 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.6245% | 1.885% | 12.01% | 1.082% | 0.3166% | 4.41% | 1.291% | 10.32% | 11% | 12.15% | 3.012% | 1.492% | 2.753% | 9.841% | 3.939% | 5.005% | 2.14% | 0.5622% | 3.597% | 5.825% | 1.992% | 2.673% | 88.19% | 12.5% | 25.46% | 1.552% | 58.08% | 0.2457% | 0.8079% | 0.6437% | 0.8076% | 0.3625% | 0.9473% | 1.633% | 1.092% | 5.774% | 14.17% | 4.481% | 3.883%
89 | 10_1_21_UF_6 |  | 6/16/2023 9:07:46 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_UF_6.738 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.4759% | 0.9175% | 7.37% | 0.4421% | 2.052% | 3.366% | 1.453% | 4.45% | 8.465% | 8.595% | 1.805% | 0.6709% | 1.802% | 4.367% | 1.349% | 3.321% | 0.5575% | 2.544% | 2.958% | 2.105% | 1.769% | 6.217% | 34.64% | 23.08% | 14.8% | 2.302% | 57.28% | 0.6365% | 1.163% | 3.668% | 0.9077% | 0.2122% | 0.3631% | 0.7168% | 1.548% | 22.15% | 9.916% | 6.083% | 5.11%
90 | 10_1_21_UF_8 |  | 6/16/2023 9:13:46 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_UF_8.739 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.046% | 0.9426% | 11.59% | 0.4601% | 2.746% | 3.573% | 1.193% | 2.428% | 9.375% | 2.776% | 0.5139% | 0.4213% | 1.838% | 11.39% | 5.303% | 4.109% | 0.5412% | 1.089% | 1.094% | 4.507% | 0.5897% | 0.8225% | 114.6% | 20.35% | 13.18% | 18.62% | 26.96% | 0.3703% | 0.6672% | 1.618% | 0.6212% | 0.3365% | 1.295% | 0.3257% | 1.388% | 18.56% | 12.28% | 9.645% | 1.957%
91 | 10_1_21_UF_8_LD |  | 6/16/2023 9:19:45 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_UF_8_LD.740 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.4928% | 1.272% | 1.4% | 0.6936% | 1.498% | 1.49% | 1.424% | 6.517% | 4.302% | 1.326% | 2.654% | 0.6052% | 0.2316% | 23.31% | 8.442% | 4.684% | 1.97% | 0.6861% | 2.905% | 6.204% | 0.9784% | 3.461% | 66.67% | 26.36% | 20.82% | 11.67% | 31.23% | 0.7091% | 0.2189% | 0.1703% | 0.3394% | 0.2906% | 2.277% | 0.7436% | 0.7012% | 22.08% | 10.3% | 5.148% | 3.148%
92 | 10_1_21_UF_8_LD |  | 6/16/2023 9:26:03 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_UF_8_LD.741 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.072% | 2.467% | 11.7% | 2.875% | 1.349% | 2.246% | 0.7489% | 3.273% | 3.604% | 0.1227% | 2.574% | 1.149% | 0.3747% | 6.153% | 3.584% | 1.972% | 1.838% | 0.1293% | 1.162% | 6.775% | 1.792% | 2.468% | 78.06% | 28.39% | 19.62% | 7.689% | 58.08% | 0.5136% | 0.3982% | 0.7377% | 0.7907% | 2.171% | 2.481% | 1.714% | 0.2614% | 16.26% | 8.047% | 3.079% | 2.703%
93 | 10_1_21_UF_8_DUP |  | 6/16/2023 9:32:02 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_UF_8_DUP.742 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.371% | 1.606% | 7.371% | 1.257% | 0.2498% | 3.224% | 0.9917% | 6.396% | 2.097% | 1.071% | 1.056% | 0.63% | 1.463% | 20.71% | 10.03% | 0.9685% | 2.36% | 1.171% | 2.572% | 4.48% | 1.006% | 7.843% | 40% | 24.74% | 16.34% | 5.175% | 39.74% | 0.9878% | 0.6122% | 0.5614% | 0.3978% | 0.9743% | 1.165% | 1.111% | 0.7187% | 16.38% | 9.934% | 5.579% | 3.054%
94 | 10_1_21_UF_8_DUP_LSP |  | 6/16/2023 9:38:02 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_UF_8_DUP_LSP.743 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.505% | 0.9396% | 5.498% | 0.5813% | 0.9457% | 0.659% | 2.063% | 0.3405% | 0.8211% | 1.342% | 1.403% | 0.07567% | 0.3725% | 1.187% | 2.275% | 1.173% | 0.1622% | 0.7312% | 0.681% | 3.868% | 0.8806% | 1.161% | 6.243% | 3.567% | 1.107% | 1.803% | 2.414% | 0.3851% | 0.7405% | 0.5751% | 0.3661% | 0.8009% | 0.1717% | 1.2% | 0.8372% | 1.961% | 0.5206% | 0.901% | 1.257%
95 | 10_1_21_UF_9 |  | 6/16/2023 9:44:02 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_UF_9.744 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.617% | 1.147% | 10.04% | 1.314% | 0.09481% | 1.726% | 1.011% | 1.813% | 6.623% | 11.73% | 1.52% | 1.289% | 0.4905% | 7.593% | 2.585% | 3.817% | 1.05% | 1.475% | 2.611% | 3.089% | 2.192% | 6.508% | 31.49% | 24.25% | 31.33% | 6.283% | 18.23% | 0.1871% | 1.268% | 1.32% | 0.8741% | 0.6536% | 0.8234% | 0.3847% | 1.231% | 7.692% | 15.96% | 5.678% | 7.141%
96 | 10_1_21_UF_10 |  | 6/16/2023 9:50:03 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_UF_10.745 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.255% | 0.7278% | 9.655% | 1.171% | 1.468% | 3.061% | 0.7029% | 5.621% | 6.399% | 3.106% | 2.424% | 0.8832% | 0.5122% | 16.43% | 19.78% | 4.656% | 0.2298% | 0.2626% | 4.035% | 6.052% | 0.4519% | 4.276% | 12.37% | 35.15% | 26.65% | 9.391% | 24.12% | 1.275% | 0.602% | 3.721% | 1.277% | 0.9107% | 1.123% | 0.3988% | 0.4589% | 41.72% | 12.65% | 6.444% | 2.055%
97 | 10_1_21_UF_11 |  | 6/16/2023 9:56:03 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_UF_11.746 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.8984% | 0.6378% | 10.66% | 2.69% | 0.7955% | 2.94% | 1.598% | 2.327% | 11.62% | 13.62% | 1.114% | 0.4066% | 0.8705% | 6.565% | 5.815% | 2.827% | 1.059% | 0.7308% | 3.293% | 6.974% | 0.7543% | 2.171% | 44.41% | 7.531% | 11.66% | 14.16% | 16.67% | 0.7668% | 1.212% | 4.674% | 0.5039% | 0.4325% | 1.185% | 0.9599% | 1.006% | 22.94% | 9.14% | 4.862% | 1.263%
98 | LBLANK |  | 6/16/2023 10:02:04 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.747 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 57.28% | 86.6% | 57.74% | 4.21% | 5.943% | 22.04% | 173.2% | 86.6% | 10.83% | 12.49% | 8.067% | 20% | 0.3729% | 173.2% | 26.67% | 8.66% | 10.21% | 4.657% | 8.06% | 173.2% | 8.935% | 6.132% | 173.2% | 34.64% | 34.64% | 29.6% |  | 25.46% | 51.96% | 2.235% | 0.6251% | 0.5635% | 0.4907% | 0.4222% | 0.8836% | 100% | 59.78% | 9.116% | 6.928%
99 | CCV 1ppb TRACE |  | 6/16/2023 10:08:04 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.748 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.878% | 3.948% | 7.199% | 2.947% | 0.8208% | 0.5089% | 0.5527% | 5.329% | 3.447% | 2.232% | 3.095% | 1.559% | 1.938% | 3.477% | 4.726% | 1.806% | 0.19% | 1.242% | 24% | 8.599% | 0.07951% | 2.177% | 23.26% | 1.637% | 5.018% | 9.46% | 21.54% | 1.082% | 0.9362% | 1.193% | 0.5393% | 0.7116% | 0.9686% | 1.285% | 1.066% | 1.441% | 0.2615% | 0.9507% | 1.318%
100 | CCV 1ppb TRACE |  | 6/16/2023 10:14:23 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.749 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.977% | 3.021% | 9.598% | 1.616% | 3.219% | 2.24% | 6.022% | 4.048% | 0.9813% | 1.115% | 2.545% | 2.087% | 0.8751% | 1.81% | 5.339% | 1.566% | 1.602% | 0.4377% | 34.5% | 12.78% | 2.18% | 2.195% | 8.772% | 2.959% | 4.259% | 5.758% | 11.6% | 1.131% | 0.1679% | 0.8387% | 0.3328% | 0.6684% | 3.154% | 0.9213% | 1.786% | 2.358% | 0.6083% | 0.6789% | 0.6268%
101 | CCV 10 ppb MAJORS |  | 6/16/2023 10:20:23 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.750 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.305% | 2.41% | 11.04% | 0.8715% | 1.046% | 1.697% | 1.131% | 2.596% | 1.582% | 0.419% | 1.808% | 0.9077% | 1.567% | 1.196% | 0.3645% | 1.09% | 0.1252% | 0.7558% | 1.082% | 3.977% | 1.225% | 0.2452% | 2.872% | 2.698% | 2.324% | 1.98% | 5.286% | 0.3356% | 1.018% | 0.3927% | 0.8377% | 0.3682% | 0.5189% | 0.7274% | 1.204% | 0.519% | 0.1463% | 1.919% | 1.302%
102 | 10_1_21_BB |  | 6/16/2023 10:26:24 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_BB.751 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 49.49% | 173.2% | 16.43% | 4.852% | 3.176% | 11.11% | 50% | 100% | 16.15% | 8.679% | 0.6468% | 12.88% | 0.1418% | 68.63% | 22.05% | 2.674% | 5.87% | 1.946% | 2.991% | 69.28% | 9.096% | 2.261% | 173.2% | 114.6% | 173.2% | 38.79% | 173.2% | 33.48% | 12.06% | 1.103% | 0.5277% | 1.434% | 1.109% | 0.3852% | 1.471% | 86.6% | 6.186% | 3.049% | 4.16%
103 | 10_1_21_FB |  | 6/16/2023 10:32:24 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_FB.752 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 8.773% | 86.6% | 2.793% | 1.284% | 4.246% | 6.739% | 76.38% | 57.74% | 15.89% | 1% | 1.195% | 4.115% | 0.9391% | 53.91% | 11.72% | 6.243% | 4.836% | 2.495% | 2.298% | 86.6% | 9.799% | 8.823% | 173.2% | 100% | 65.47% | 44.66% |  | 18.52% | 22.49% | 1.537% | 0.631% | 0.6416% | 0.3558% | 0.3348% | 1.151% | 100% | 56.97% | 4.479% | 5.774%
104 | 10_1_21_LAB_A |  | 6/16/2023 10:38:24 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_LAB_A.753 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 35.25% |  | 18.94% | 6.916% | 2.85% | 37.13% | 57.74% | 86.6% | 3.43% | 15.56% | 4.675% | 9.984% | 1.635% | 86.6% | 18.41% | 12.02% | 1.655% | 0.6533% | 8.654% | 100% | 14.74% | 1.463% | 173.2% | 100% | 114.6% | 13.64% | 173.2% | 9.091% | 11.95% | 0.9147% | 0.2723% | 0.6468% | 0.4915% | 2.048% | 0.8414% |  | 50.76% | 4.453% | 1.465%
105 | 10_1_21_LAB_B |  | 6/16/2023 10:44:24 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_LAB_B.754 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 53.93% | 173.2% | 10.41% | 7.769% | 0.6693% | 21.43% | 43.3% | 86.6% | 12.39% | 2.468% | 1.097% | 15.04% | 2.184% | 50% | 30.74% | 6.287% | 1.716% | 2.35% | 2.067% | 173.2% | 29.05% | 4.472% |  | 173.2% | 69.28% | 28.53% | 173.2% | 19.88% | 28.64% | 1.494% | 0.3412% | 0.3601% | 0.1852% | 1.299% | 2.03% | 173.2% | 37.5% | 3.439% | 1.143%
106 | 5_18_23_DB |  | 6/16/2023 10:50:24 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_DB.755 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.365% | 1.609% | 1.858% | 0.1756% | 4.878% | 1.406% | 0.6971% | 1.822% | 5.454% | 1.718% | 1.041% | 1.321% | 1.365% | 2.038% | 2.478% | 6.567% | 1.196% | 1.432% | 2.615% | 10.83% | 1.099% | 2.428% | 8.75% | 6.404% | 5.459% | 11.11% | 4.075% | 0.8% | 0.9428% | 1.37% | 0.3016% | 0.7725% | 0.9521% | 1.032% | 0.8361% | 1.519% | 1.705% | 1.181% | 0.7222%
107 | 5_18_23_DB |  | 6/16/2023 10:56:42 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_DB.756 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.8849% | 1.563% | 1.792% | 1.058% | 1.05% | 1.326% | 0.9648% | 3.385% | 1.303% | 1.144% | 1.732% | 1.266% | 0.5652% | 2.747% | 3.257% | 0.7686% | 2.32% | 1.153% | 7.782% | 5.944% | 1.604% | 1.302% | 16.09% | 7.739% | 1.988% | 3.538% | 0.9524% | 0.5519% | 1.847% | 0.2917% | 0.6116% | 0.6478% | 0.4366% | 0.3877% | 2.18% | 0.9919% | 0.2165% | 2.125% | 0.8605%
108 | 5_18_23_T_239 |  | 6/16/2023 11:02:42 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_T_239.757 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.326% | 1.72% | 1.459% | 1.321% | 4.315% | 0.3731% | 0.8887% | 1.799% | 0.7238% | 1.217% | 0.4681% | 0.5453% | 0.9811% | 1.971% | 1.458% | 1.803% | 2.46% | 1.067% | 1.373% | 5.81% | 0.635% | 1.294% | 3.41% | 2.236% | 5.24% | 6.274% | 16.2% | 0.5449% | 1.006% | 0.3745% | 0.618% | 0.6688% | 0.9669% | 0.4168% | 0.5022% | 0.2368% | 1.282% | 1.184% | 0.8429%
109 | 5_18_23_T_239 |  | 6/16/2023 11:09:00 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_T_239.758 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.8047% | 2.162% | 0.5004% | 0.5215% | 2.489% | 4.036% | 1.318% | 1.837% | 1.656% | 1.427% | 0.3191% | 0.779% | 0.2452% | 2.649% | 2.28% | 1.537% | 2.078% | 0.6483% | 4.745% | 12.22% | 0.9412% | 2.632% | 9.314% | 7.727% | 5.471% | 1.986% | 1.151% | 2.756% | 1.25% | 0.8346% | 1.095% | 0.3458% | 1.506% | 0.8793% | 0.7705% | 0.6748% | 0.3406% | 1.093% | 0.5374%
110 | 5_18_23_T_239 |  | 6/16/2023 11:15:00 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_T_239.759 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.427% | 8.077% | 2.712% | 2.214% | 3.579% | 2.322% | 2.314% | 1.524% | 12.06% | 7.986% | 1.947% | 0.08943% | 0.2414% | 6.005% | 15.91% | 13.91% | 4.59% | 2.051% | 9.331% | 56.77% | 15.34% | 2.186% | 86.6% | 33.33% | 59.19% | 2.878% | 107.9% | 1.573% | 3.05% | 0.7488% | 0.4336% | 0.58% | 4.1% | 1.307% | 3.672% | 55.3% | 21.65% | 6.569% | 8.442%
111 | 5_18_23_M_224 |  | 6/16/2023 11:20:59 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_M_224.760 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.3115% | 3.796% | 7.084% | 1.292% | 1.2% | 1.353% | 0.5429% | 1.229% | 1.096% | 0.754% | 1.426% | 0.5771% | 1.736% | 1.893% | 0.9847% | 2.169% | 0.8217% | 1.281% | 3.735% | 0.797% | 1.035% | 3.355% | 7.38% | 2.557% | 2.017% | 5.515% | 5.997% | 0.9506% | 0.9305% | 6.87% | 0.5318% | 0.1772% | 0.921% | 0.99% | 0.4839% | 1.688% | 0.4052% | 0.6875% | 1.256%
112 | LBLANK |  | 6/16/2023 11:27:00 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.761 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 86.6% | 173.2% | 86.6% | 0.6511% | 1.481% | 15.41% | 173.2% | 98.97% | 36.64% | 38.65% | 1.217% | 43.3% | 1.76% | 48.04% | 6.928% | 100.3% | 18.94% | 5.724% | 3.771% |  | 11.11% | 4.162% | 173.2% | 100% | 173.2% | 23.41% | 173.2% | 4.237% | 14.32% | 0.3356% | 0.8375% | 0.8353% | 0.9035% | 0.728% | 1.007% | 173.2% | 6.662% | 39.16% | 3.268%
113 | LBLANK |  | 6/16/2023 11:33:18 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.762 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 88.19% | 173.2% | 17.32% | 3.395% | 2.627% | 21.5% | 86.6% | 114.6% | 38.57% | 14.19% | 4.264% | 34.32% | 1.625% | 86.6% | 6.928% | 26.65% | 5.43% | 13.47% | 2.414% | 173.2% | 36.74% | 5.366% |  | 173.2% | 173.2% | 23.1% |  | 13.32% | 32.87% | 1.577% | 0.4858% | 2.63% | 0.5922% | 1.335% | 0.9195% | 173.2% | 78.73% | 19.92% | 6.278%
114 | CCV 1ppb TRACE |  | 6/16/2023 11:39:19 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.763 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4.615% | 2.385% | 9.693% | 0.4714% | 2.03% | 5.179% | 1.656% | 1.196% | 2.244% | 3.84% | 2.208% | 0.857% | 0.9866% | 2.008% | 6.623% | 2.237% | 1.493% | 1.015% | 11.95% | 11.38% | 5.803% | 1.332% | 8.986% | 4.124% | 6.523% | 11.6% | 3.431% | 0.9991% | 1.557% | 1.072% | 1.182% | 1.366% | 0.5358% | 0.6714% | 0.9509% | 0.02053% | 1.894% | 1.563% | 0.5008%
115 | CCV 1ppb TRACE |  | 6/16/2023 11:45:38 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.764 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.5526% | 5.994% | 7.569% | 2.008% | 3.278% | 4.367% | 11.55% | 1.687% | 2.007% | 0.5081% | 1.777% | 2.076% | 1.166% | 1.663% | 1.713% | 2.83% | 1.042% | 0.9943% | 6.108% | 20.63% | 1.811% | 2.73% | 32.78% | 4.983% | 5.62% | 4.03% | 20.6% | 0.9423% | 1.911% | 5.296% | 0.9594% | 0.7673% | 0.5262% | 0.5698% | 1.376% | 4.057% | 0.7109% | 0.5745% | 1.386%
116 | LFB |  | 6/16/2023 11:51:38 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.765 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.878% | 1.723% | 10.83% | 0.5883% | 0.2072% | 1.489% | 1.593% | 1.532% | 1.736% | 0.3597% | 2.182% | 1.063% | 2.307% | 0.5761% | 0.1927% | 0.2004% | 0.7421% | 0.2724% | 1.648% | 3.364% | 1.174% | 0.2982% | 4.651% | 2.013% | 1.351% | 4.335% | 2.915% | 0.693% | 0.5347% | 2.44% | 0.4442% | 0.8114% | 0.965% | 1.074% | 1.004% | 0.6879% | 0.7773% | 0.6934% | 0.6664%
117 | LFB |  | 6/16/2023 11:57:55 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.766 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.2228% | 1.533% | 3.498% | 0.7292% | 0.7079% | 0.9644% | 0.3514% | 1.441% | 0.6082% | 0.3677% | 0.4922% | 1.123% | 0.5188% | 0.948% | 2.651% | 0.7541% | 0.8591% | 0.7187% | 0.819% | 3.226% | 0.9936% | 2.842% | 3.444% | 1.335% | 1.698% | 4.808% | 0.5072% | 1.487% | 0.913% | 0.6494% | 0.9099% | 0.7759% | 0.3577% | 0.5826% | 1.404% | 1.106% | 0.6514% | 0.9473% | 1.003%
118 | CCV 10 ppb MAJORS |  | 6/17/2023 12:03:55 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.767 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.188% | 0.6584% | 6.042% | 0.9536% | 2.731% | 1.023% | 1.413% | 1.19% | 0.9271% | 0.3914% | 1.433% | 0.4244% | 0.4184% | 1.701% | 1.889% | 1.99% | 0.8397% | 0.4183% | 2.004% | 3.372% | 0.7564% | 1.152% | 2.767% | 0.8987% | 2.367% | 2.782% | 4.174% | 0.6758% | 0.7222% | 1.176% | 0.3713% | 0.6955% | 0.888% | 0.2153% | 0.2875% | 0.3566% | 0.2628% | 0.4405% | 0.3665%
119 | 5_18_23_MFB |  | 6/17/2023 12:09:55 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_MFB.768 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 12.37% |  | 28.23% | 1.68% | 4.727% | 35.95% | 56.77% | 50% | 8.449% | 11.5% | 1.491% | 12.38% | 1.719% | 21.65% | 27.06% | 9.147% | 5.558% | 1.841% | 5.29% | 173.2% | 41.66% | 0.9356% | 173.2% | 91.65% | 173.2% | 59.85% |  | 10.02% | 17.63% | 0.5741% | 0.343% | 0.7403% | 0.3774% | 0.4238% | 0.3288% | 114.6% | 83.33% | 5.538% | 4.277%
120 | 10_15_21_W_1 |  | 6/17/2023 12:15:55 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_W_1.769 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.741% | 4.2% | 6.797% | 1.525% | 2.374% | 4.788% | 0.2938% | 8.54% | 18.44% | 1.446% | 0.6747% | 0.2059% | 0.5425% | 22.54% | 15.16% | 1.376% | 0.4202% | 2.371% | 15.34% | 10.14% | 2.79% | 6.439% | 86.6% | 57.74% | 36.74% | 29.76% | 86.6% | 1.109% | 1.307% | 0.902% | 0.4892% | 0.5454% | 0.07775% | 0.8708% | 1.325% | 21.53% | 19.88% | 5.098% | 1.98%
121 | 10_15_21_W_6 |  | 6/17/2023 12:21:56 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_W_6.770 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.637% | 0.4205% | 3.099% | 2.303% | 1.947% | 1.481% | 0.6925% | 4.402% | 22.78% | 52.14% | 0.2916% | 0.5323% | 1.128% | 16.95% | 6.614% | 4.568% | 1.886% | 2.286% | 79.42% | 14.95% | 0.6716% | 5.834% | 34.64% | 80% | 7.873% | 42.08% | 173.2% | 1.264% | 0.2606% | 1.057% | 1.391% | 1.007% | 1.223% | 0.1285% | 1.332% | 28.97% | 34.33% | 1.901% | 0.1714%
122 | 10_15_21_W_8 |  | 6/17/2023 12:27:56 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_W_8.771 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.833% | 8.107% | 1.887% | 2.812% | 0.6882% | 1.028% | 1.449% | 7.756% | 3.851% | 3.002% | 1.732% | 1.239% | 1.231% | 5.882% | 5.43% | 5.502% | 0.8496% | 1.505% | 118.8% | 16.11% | 2.054% | 6.05% | 43.3% | 100% | 50.76% | 6.861% | 0% | 2.346% | 1.122% | 0.5462% | 0.5638% | 0.9321% | 0.6423% | 0.1904% | 0.6107% | 45.83% | 23.01% | 3.309% | 2.027%
123 | 10_15_21_W_9 |  | 6/17/2023 12:33:56 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_W_9.772 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.6368% | 1.362% | 0.4659% | 1.949% | 1.762% | 2.919% | 1.741% | 18.22% | 8.882% | 6.612% | 0.1207% | 2.627% | 0.9434% | 30.31% | 9.014% | 4.895% | 2.401% | 2.925% | 59.61% | 16.08% | 1.602% | 5.283% | 86.6% | 0% | 44.1% | 50.34% | 91.65% | 1.642% | 1.391% | 15% | 0.7616% | 0.54% | 0.269% | 0.9237% | 1.057% | 18.18% | 30.43% | 4.422% | 1.498%
124 | 10_15_21_W_9 |  | 6/17/2023 12:40:15 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_W_9.773 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.048% | 2.78% | 5.422% | 1.397% | 3.94% | 6.61% | 2.366% | 17.4% | 12.96% | 5.803% | 0.9063% | 1.081% | 1.189% | 12.37% | 8.056% | 2.242% | 0.4925% | 1.338% | 61.49% | 27.04% | 7.436% | 2.422% | 173.2% | 50% | 96.22% | 38.19% | 69.28% | 2.097% | 1.302% | 0.2768% | 1.192% | 1.315% | 1.345% | 0.6338% | 1.182% | 41.93% | 25.82% | 1.377% | 2.334%
125 | 10_15_21_W_10 |  | 6/17/2023 12:46:15 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_W_10.774 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.644% | 2.365% | 1.31% | 2.885% | 1.316% | 3.447% | 3.397% | 3.897% | 7.22% | 20.05% | 1.131% | 1.502% | 1.165% | 11.58% | 11.66% | 3.796% | 2.059% | 0.9211% | 36.36% | 15.81% | 4.622% | 3.125% | 43.3% | 41.66% | 37.75% | 17.83% | 173.2% | 2.627% | 0.4762% | 1.148% | 0.2914% | 0.4235% | 0.5708% | 0.6537% | 0.5292% | 46.69% | 11.06% | 3.137% | 2.726%
126 | 10_15_21_W_10 |  | 6/17/2023 12:52:33 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_W_10.775 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.62% | 2.775% | 1.712% | 2.122% | 3.275% | 7.042% | 1.887% | 16.51% | 13.5% | 12.05% | 0.2253% | 0.361% | 0.6648% | 7.613% | 8.206% | 2.452% | 1.028% | 2.073% | 64.94% | 2.863% | 3.376% | 2.534% | 86.6% | 57.28% | 24.74% | 13.25% | 100% | 1.056% | 2.227% | 0.7711% | 0.5148% | 0.3204% | 0.5286% | 0.2867% | 1.01% | 24.74% | 7.099% | 1.674% | 1.739%
127 | 10_15_21_W_10_DUP |  | 6/17/2023 12:58:33 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_W_10_DUP.776 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.22% | 2.556% | 1.942% | 1.7% | 4.212% | 9.63% | 2.2% | 5.941% | 0% | 5.785% | 2.31% | 0.701% | 1.425% | 9.449% | 19.44% | 0.7643% | 1.331% | 0.4733% | 73.88% | 18.37% | 1.699% | 2.351% |  | 69.28% | 56.77% | 7.331% | 173.2% | 1.764% | 1.672% | 0.5889% | 0.9428% | 1.56% | 0.8012% | 0.5468% | 1.441% | 6.186% | 15.25% | 0.1433% | 0.613%
128 | 10_15_21_W_10_DUP |  | 6/17/2023 1:04:52 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_W_10_DUP.777 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.883% | 3.177% | 3.165% | 1.423% | 1.273% | 4.492% | 0.8207% | 7.812% | 9.587% | 5.241% | 2.954% | 1.106% | 0.3716% | 24.9% | 22.21% | 2.366% | 1.55% | 0.8503% | 52.48% | 16.7% | 3.974% | 3.653% |  | 33.33% | 13.86% | 9.427% | 173.2% | 1.772% | 0.8831% | 0.5219% | 0.3854% | 0.5241% | 1.382% | 0.4935% | 1.045% | 9.116% | 23.07% | 2.064% | 5.081%
129 | LBLANK |  | 6/17/2023 1:10:52 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.778 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 100% | 173.2% | 34.64% | 2.2% | 3.964% | 23.57% | 173.2% | 114.6% | 57.41% | 30.05% | 1.52% | 61.86% | 0.7551% | 173.2% | 29.4% | 15.06% | 10.07% | 6.261% | 7.098% |  | 13.32% | 3.891% |  | 173.2% | 173.2% | 35.02% | 173.2% | 20.78% | 9.116% | 0.3605% | 0.9115% | 0.5611% | 0.5531% | 0.133% | 1.395% |  | 36.74% | 27.06% | 30.79%
130 | LBLANK |  | 6/17/2023 1:17:12 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.779 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 34.64% |  | 26.96% | 4.517% | 3.562% | 8.183% | 173.2% | 173.2% | 35.44% | 18.23% | 4.899% | 69.28% | 2.238% | 100% | 36.06% | 49.49% | 9.791% | 3.862% | 5.554% |  | 34.64% | 6.072% |  | 173.2% | 50% | 72.81% |  | 7.922% | 9.116% | 2.965% | 0.5999% | 0.3541% | 0.7981% | 0.9525% | 1.186% | 173.2% | 12.37% | 6.662% | 10.19%
131 | CCV 1ppb TRACE |  | 6/17/2023 1:23:12 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.780 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.293% | 10.15% | 7.009% | 2.995% | 4.525% | 5.893% | 6.228% | 1.885% | 3.135% | 2.088% | 0.8487% | 1.984% | 0.417% | 0.9431% | 0.7098% | 3.397% | 1.176% | 3.538% | 13.83% | 23.39% | 2.397% | 5.051% | 14.24% | 4.912% | 6.485% | 9.754% | 12.2% | 1.839% | 1.265% | 0.5928% | 0.4277% | 0.8454% | 0.6498% | 0.3416% | 0.6929% | 0.951% | 1.722% | 1.217% | 0.8061%
132 | CCV 1ppb TRACE |  | 6/17/2023 1:29:31 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.781 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.867% | 5.623% | 13.9% | 3.316% | 4.218% | 3.311% | 5.236% | 0.7555% | 3.036% | 2.307% | 1.625% | 0.2737% | 0.6869% | 2.135% | 6.967% | 4.933% | 1.835% | 1.137% | 24.74% | 13.16% | 2.593% | 2.948% | 16.23% | 6.358% | 1.613% | 19.04% | 5.516% | 0.7903% | 2.002% | 0.7887% | 1.005% | 0.8591% | 0.5559% | 0.7531% | 1.07% | 2.025% | 1.053% | 1.06% | 1.072%
133 | CCV 10 ppb MAJORS |  | 6/17/2023 1:35:32 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.782 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.476% | 1.91% | 6.633% | 0.9336% | 1.446% | 1.09% | 2.41% | 1.523% | 1.188% | 0.7444% | 1.977% | 1.489% | 0.8546% | 0.4947% | 2.797% | 2.03% | 0.8708% | 0.2734% | 2.205% | 1.504% | 0.5698% | 2.056% | 4.563% | 4.499% | 1.398% | 0.773% | 2.986% | 0.9434% | 0.1174% | 1.054% | 0.3696% | 1.044% | 0.2332% | 0.4327% | 0.5378% | 1.083% | 0.2848% | 2.272% | 0.7783%
134 | CCV 10 ppb MAJORS |  | 6/17/2023 1:41:51 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.783 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.981% | 0.2596% | 4.614% | 0.5986% | 0.6326% | 0.4508% | 1.452% | 0.3366% | 0.3872% | 0.7643% | 0.9509% | 0.24% | 0.534% | 0.3722% | 1.579% | 1.078% | 1.309% | 1.501% | 1.957% | 5.452% | 0.1815% | 2.407% | 4.664% | 4.163% | 1.06% | 3.914% | 6.075% | 0.9538% | 0.3485% | 0.593% | 0.9523% | 2.859% | 0.9469% | 0.4514% | 0.6567% | 2.178% | 1.176% | 1.25% | 0.6052%
135 | 10_15_21_W_11 |  | 6/17/2023 1:47:51 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_W_11.784 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.214% | 2.646% | 2.859% | 1.354% | 3.49% | 8% | 1.56% | 4.082% | 21.31% | 4.062% | 1.241% | 1.628% | 0.4155% | 6.963% | 22.9% | 4.966% | 1.105% | 0.7915% | 46.81% | 26.28% | 4.727% | 0.2595% | 173.2% | 100% | 39.85% | 26.66% | 100% | 3.075% | 0.3588% | 2.839% | 0.7473% | 0.2272% | 0.6943% | 1.043% | 0.6431% | 30.12% | 6.554% | 3.286% | 1.475%
136 | 10_15_21_W_11 |  | 6/17/2023 1:54:10 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_W_11.785 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.782% | 2.251% | 1.829% | 2.468% | 1.016% | 4.538% | 0.4993% | 3.304% | 4.419% | 4.489% | 1.965% | 0.8613% | 0.7359% | 14% | 38.13% | 4.472% | 1.183% | 1.754% | 111.2% | 13.5% | 1.25% | 3.284% | 86.6% | 56.77% | 35.25% | 17.85% | 100% | 0.9736% | 0.3829% | 0.6382% | 0.3568% | 0.8426% | 1.27% | 1.042% | 0.5285% | 17.08% | 15.44% | 2.511% | 1.026%
137 | 10_15_21_F_1 |  | 6/17/2023 2:00:09 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_1.786 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.016% | 2.376% | 6.354% | 1.215% | 1.031% | 1.154% | 0.7416% | 9.486% | 17.08% | 6.421% | 0.4236% | 0.2033% | 0.7147% | 23.1% | 2.549% | 3.24% | 0.3968% | 0.9877% | 1.8% | 2.444% | 1.382% | 1.833% | 25% | 10% | 5.556% | 14.25% | 68.63% | 0.5048% | 0.2831% | 0.5382% | 0.6462% | 0.6477% | 3.372% | 1.247% | 0.7562% | 43.91% | 12.47% | 1.017% | 1.386%
138 | 10_15_21_F_1 |  | 6/17/2023 2:06:28 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_1.787 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.21% | 0.2466% | 8.788% | 2.194% | 0.4085% | 3.249% | 1.62% | 3.237% | 13.33% | 2.462% | 0.4854% | 0.116% | 0.3491% | 15.07% | 13.01% | 1.149% | 1.663% | 0.9288% | 2.18% | 2.024% | 0.7375% | 2.719% | 41.66% | 28.64% | 10.53% | 6.113% | 109.9% | 0.545% | 0.8017% | 1.082% | 0.1814% | 0.2782% | 0.1876% | 1.107% | 0.8423% | 20.78% | 23.06% | 1.377% | 0.7599%
139 | 10_15_21_F_1_DUP |  | 6/17/2023 2:12:28 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_1_DUP.788 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.23% | 1.26% | 11.48% | 0.748% | 0.6455% | 3.262% | 1.003% | 9.755% | 14.45% | 8.776% | 0.2995% | 0.8709% | 0.7695% | 6.484% | 2.066% | 4.596% | 2.26% | 1.689% | 2.001% | 8.314% | 0.5524% | 4.076% | 56.77% | 29.04% | 14.2% | 8.015% | 47.19% | 0.5948% | 0.7345% | 0.8147% | 0.7142% | 0.5063% | 0.2845% | 0.3679% | 0.535% | 3.704% | 15.21% | 1.897% | 1.752%
140 | 10_15_21_F_1_DUP |  | 6/17/2023 2:18:46 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_1_DUP.789 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.9133% | 3.105% | 11.82% | 2.536% | 1.136% | 1.138% | 0.8106% | 4.498% | 10.79% | 7.566% | 1.553% | 0.6995% | 1.111% | 13.09% | 6.558% | 1.899% | 0.769% | 0.7496% | 2.537% | 7.554% | 0.5902% | 6.04% | 49.49% | 15.38% | 13.86% | 16.79% | 15.75% | 0.2465% | 0.686% | 1.767% | 1.071% | 0.5785% | 0.5509% | 0.4484% | 1.734% | 43.3% | 13.33% | 2.891% | 1.111%
141 | 10_15_21_F_6 |  | 6/17/2023 2:24:45 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_6.790 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.9552% | 0.9655% | 6.504% | 2.098% | 1.973% | 1.364% | 0.5833% | 9.459% | 9.759% | 4.04% | 0.1228% | 0.5628% | 0.4332% | 9.476% | 4.551% | 4.22% | 0.9964% | 0.8226% | 2.314% | 2.932% | 0.9135% | 10.16% | 39.74% | 32.83% | 17.91% | 6.135% | 33.33% | 1.19% | 0.3793% | 0.6614% | 0.647% | 1.256% | 4.023% | 0.812% | 1.467% | 22.78% | 3.749% | 2.165% | 5.479%
142 | 10_15_21_F_6 |  | 6/17/2023 2:31:03 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_6.791 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.6745% | 2.2% | 0.9156% | 1.913% | 1.262% | 1.152% | 1.098% | 2.629% | 4.872% | 5.838% | 0.7223% | 0.8646% | 0.1737% | 9.024% | 8.784% | 1.813% | 1.063% | 1.804% | 1.31% | 8.718% | 1.124% | 3.121% | 19.92% | 42.28% | 22.15% | 4.276% | 22.91% | 0.4948% | 0.3948% | 0.7932% | 0.2577% | 0.2251% | 0.2746% | 0.8431% | 0.5364% | 14.83% | 6.616% | 3.547% | 2.508%
143 | 10_15_21_F_6_LD |  | 6/17/2023 2:37:02 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_6_LD.792 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.8451% | 2.408% | 0.2135% | 2.584% | 1.925% | 2.25% | 0.7507% | 1.25% | 7.449% | 6.618% | 1.418% | 0.3836% | 0.9641% | 16.65% | 0.697% | 1.71% | 0.6195% | 1.464% | 0.7642% | 6.612% | 1.707% | 2.456% | 66.14% | 10.41% | 6.278% | 14.26% | 16.67% | 0.7381% | 0.3648% | 0.7848% | 1.222% | 0.9164% | 0.7052% | 0.2116% | 1.097% | 19.58% | 2.758% | 4.366% | 2.093%
144 | 10_15_21_F_6_LD |  | 6/17/2023 2:43:20 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_6_LD.793 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.9201% | 2% | 4.301% | 3.254% | 0.8691% | 1.34% | 0.7432% | 1.993% | 11.22% | 2.474% | 1.146% | 1.004% | 1.576% | 10.93% | 12.3% | 1.838% | 1.83% | 0.925% | 3.632% | 13.92% | 0.721% | 4.835% | 43.3% | 13.32% | 3.535% | 2.248% | 43.3% | 0.5948% | 0.2389% | 5.774% | 0.3839% | 3.872% | 0.1897% | 1.387% | 1.006% | 6.153% | 7.042% | 3.244% | 0.5111%
145 | LBLANK |  | 6/17/2023 2:49:20 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.794 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 173.2% |  | 56.77% | 1.774% | 2.277% | 8.646% | 0% | 173.2% | 7.143% | 14.52% | 7.418% | 22.91% | 2.034% | 43.3% | 32.45% | 62.42% | 10.27% | 8.813% | 6.549% | 86.6% | 40% | 6.331% |  | 100% | 86.6% | 11.02% | 173.2% | 9.213% | 22.42% | 5.908% | 0.6351% | 0.924% | 1.533% | 0.5081% | 1.933% |  | 57.74% | 18.33% | 9.116%
146 | CCV 1ppb TRACE |  | 6/17/2023 2:55:21 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.795 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.471% | 2.243% | 5.342% | 2.822% | 3.122% | 4.44% | 4.612% | 5.88% | 3.492% | 0.7601% | 3.245% | 1.383% | 1.316% | 2.254% | 3.271% | 0.4906% | 1.324% | 0.8686% | 18.79% | 8.988% | 2.133% | 4.672% | 10.88% | 13.24% | 2.537% | 13.61% | 2.619% | 2.034% | 1.188% | 1.788% | 0.1805% | 0.9617% | 0.7834% | 0.5332% | 0.4923% | 2.724% | 1.693% | 1.237% | 1.073%
147 | CCV 1ppb TRACE |  | 6/17/2023 3:01:39 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.796 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.806% | 4.903% | 6.25% | 4.591% | 4% | 1.887% | 3.434% | 1.848% | 2.506% | 1.587% | 2.312% | 0.8811% | 1.478% | 2.823% | 6.393% | 4.823% | 1.35% | 2.982% | 4.953% | 21.46% | 2.451% | 2.386% | 33.38% | 3.729% | 4.059% | 4.441% | 2.014e-06% | 1.353% | 1.894% | 7.55% | 0.8115% | 0.5285% | 0.6279% | 0.6499% | 0.6774% | 0.6172% | 1.547% | 1.9% | 1.003%
148 | LFB |  | 6/17/2023 3:07:39 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.797 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.8573% | 1.75% | 7.682% | 0.4322% | 1.743% | 1.582% | 2.192% | 1.218% | 2.03% | 0.8613% | 0.5647% | 0.7633% | 0.9246% | 1.275% | 2.59% | 0.9547% | 1.09% | 0.2998% | 0.2534% | 3.077% | 0.5864% | 2.126% | 5.795% | 1.783% | 2.816% | 2.084% | 2.442% | 0.3956% | 0.3881% | 1.999% | 1.025% | 0.5075% | 3.364% | 0.3256% | 1.357% | 0.8914% | 0.6983% | 0.8172% | 0.8756%
149 | LFB |  | 6/17/2023 3:13:57 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.798 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.8449% | 2.339% | 2.941% | 0.7431% | 1.142% | 0.3152% | 2.103% | 2.17% | 1.018% | 0.9895% | 0.6813% | 1.315% | 1.755% | 0.6489% | 1.228% | 1.226% | 0.1382% | 0.3377% | 1.8% | 3.567% | 0.06925% | 0.4851% | 7.112% | 0.8017% | 0.5355% | 5.157% | 4.811% | 0.784% | 0.8825% | 0.6623% | 0.3639% | 0.5003% | 0.7924% | 0.8585% | 0.9356% | 1.376% | 0.6956% | 1.094% | 1.609%
150 | CCV 10 ppb MAJORS |  | 6/17/2023 3:19:56 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.799 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.544% | 0.729% | 4.933% | 1.393% | 2.366% | 2.489% | 3.211% | 1.724% | 0.9795% | 0.8264% | 1.099% | 0.4592% | 1.335% | 1.307% | 1.024% | 0.9304% | 0.9867% | 0.672% | 1.777% | 6.242% | 0.4018% | 1.159% | 8.951% | 2.31% | 2.21% | 2.386% | 0.5876% | 0.8365% | 0.2727% | 0.2922% | 0.5046% | 0.4959% | 0.7286% | 0.1791% | 2.148% | 0.3156% | 1.447% | 0.56% | 1.889%
151 | CCV 10 ppb MAJORS |  | 6/17/2023 3:26:14 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.800 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.565% | 3.362% | 3.146% | 2.413% | 1.669% | 0.6749% | 2.711% | 1.397% | 0.7347% | 0.5064% | 1.016% | 1.114% | 0.7778% | 2.026% | 0.9371% | 0.3095% | 0.4813% | 0.8374% | 0.9856% | 1.828% | 1.223% | 2.693% | 5.486% | 2.716% | 3.087% | 3.232% | 0.6038% | 0.2847% | 0.7846% | 3.224% | 0.144% | 0.6114% | 0.323% | 0.3116% | 1.771% | 1.096% | 0.6598% | 1.747% | 0.6731%
152 | 10_15_21_F_6_DUP |  | 6/17/2023 3:32:14 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_6_DUP.801 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.768% | 1.048% | 2.711% | 1.375% | 1.348% | 4.739% | 1.877% | 7.864% | 6.796% | 9.959% | 0.3494% | 0.2862% | 0.5839% | 15.25% | 9.525% | 3.461% | 0.518% | 0.7927% | 3.414% | 3.552% | 0.8916% | 5.272% | 12.37% | 48.89% | 50% | 11.86% | 66.14% | 0.8945% | 1.06% | 1.558% | 0.292% | 0.5471% | 0.4767% | 0.8061% | 0.7152% | 2.128% | 9.352% | 2.163% | 1.136%
153 | 10_15_21_F_6_DUP |  | 6/17/2023 3:38:32 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_6_DUP.802 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.754% | 0.7158% | 6.769% | 1.27% | 1.239% | 0.7972% | 1.002% | 9.544% | 16.18% | 7.707% | 1.521% | 0.1859% | 1.52% | 11.89% | 11.22% | 1.656% | 1.314% | 0.3813% | 1.711% | 8.617% | 1.845% | 1.881% | 89.21% | 10.19% | 19.88% | 21.82% | 53.93% | 0.8052% | 0.8674% | 2.836% | 1.254% | 0.4233% | 2.772% | 0.4706% | 0.6008% | 7.698% | 4.243% | 0.3275% | 1.399%
154 | 10_15_21_F_6_DUP_LSP |  | 6/17/2023 3:44:32 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_6_DUP_LSP.803 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.2% | 1.198% | 0.6591% | 1.188% | 1.311% | 0.716% | 0.5254% | 0.7072% | 1.177% | 0.6719% | 0.3257% | 0.4813% | 0.9511% | 0.8001% | 1.4% | 0.7232% | 0.7805% | 1.008% | 0.1628% | 5.195% | 0.7278% | 2.383% | 1.478e-06% | 0.2188% | 2.655% | 2.724% | 7.211% | 0.4473% | 0.4725% | 0.8807% | 0.1632% | 0.4564% | 0.9774% | 1.092% | 1.435% | 0.8117% | 0.06545% | 1.547% | 0.5396%
155 | 10_15_21_F_6_DUP_LSP |  | 6/17/2023 3:50:52 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_6_DUP_LSP.804 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.03% | 0.9699% | 1.754% | 1.265% | 0.2756% | 1.438% | 0.8042% | 0.7146% | 0.5601% | 1.618% | 0.9626% | 0.5001% | 0.3691% | 0.2824% | 1.363% | 1.798% | 1.389% | 1.449% | 1.633% | 4.358% | 0.474% | 3.15% | 5.618% | 2.349% | 1.446% | 2.072% | 5.316% | 0.7254% | 0.4578% | 0.6434% | 0.7323% | 0.6684% | 0.6743% | 0.5631% | 3.299% | 0.7771% | 0.9314% | 1.049% | 0.8082%
156 | 10_15_21_F_9 |  | 6/17/2023 3:56:52 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_9.805 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.243% | 1.518% | 3.834% | 0.1586% | 2.057% | 1.776% | 1.863% | 3.848% | 6.373% | 5.196% | 0.07379% | 0.5157% | 0.3028% | 6.784% | 8.968% | 2.536% | 0.5613% | 0.895% | 5.508% | 9.988% | 1.252% | 6.357% | 57.74% | 38.19% | 18.16% | 17.88% | 43.3% | 1.186% | 0.9172% | 1.257% | 0.7266% | 0.6754% | 0.232% | 0.7402% | 0.3964% | 14.06% | 13.82% | 4.121% | 1.45%
157 | 10_15_21_F_9 |  | 6/17/2023 4:03:12 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_9.806 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.029% | 1.976% | 4.305% | 1.021% | 1.537% | 6.365% | 1.354% | 4.489% | 4.956% | 6.486% | 0.9152% | 0.656% | 0.48% | 10.75% | 13.49% | 3.174% | 1.452% | 0.8777% | 2.335% | 8.399% | 0.7007% | 0.8482% | 33.33% | 20.15% | 22.04% | 11.46% | 10.83% | 0.8854% | 1.258% | 0.6497% | 0.4503% | 0.2038% | 0.6271% | 0.5051% | 1.061% | 12.6% | 14.21% | 3.514% | 2.111%
158 | 10_15_21_F_10 |  | 6/17/2023 4:09:13 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_10.807 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.841% | 0.5235% | 4.371% | 0.3437% | 2.663% | 1.296% | 0.5436% | 4.11% | 10.18% | 5.011% | 0.4854% | 0.5362% | 1.347% | 16.22% | 0.8898% | 1.774% | 1.825% | 1.621% | 0.9817% | 6.752% | 2.236% | 3.099% | 28.87% | 24.05% | 21.74% | 4.185% | 86.6% | 0.4688% | 2.073% | 2.2% | 4.762% | 0.5503% | 0.9162% | 0.5534% | 0.5655% | 36.78% | 20.96% | 1.841% | 2.041%
159 | 10_15_21_F_10 |  | 6/17/2023 4:15:32 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_10.808 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.3236% | 0.727% | 4.52% | 0.8997% | 1.932% | 3.318% | 0.5971% | 3.882% | 12.08% | 4.576% | 3.432% | 1.476% | 0.3494% | 30.77% | 9.718% | 2.423% | 5.157% | 2.686% | 2.57% | 4.295% | 2.477% | 3.748% | 48.04% | 18.18% | 13.87% | 2.998% | 52.92% | 4.177% | 1.203% | 0.1382% | 0.7617% | 13.17% | 0.7009% | 0.663% | 2.296% | 26.57% | 10.35% | 1.88% | 1.121%
160 | 10_15_21_F_10_DUP |  | 6/17/2023 4:21:32 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_10_DUP.809 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.6559% | 2.401% | 6.087% | 1.211% | 2.549% | 0.8681% | 1.037% | 6.349% | 13.47% | 4.158% | 0.952% | 1.164% | 2.615% | 11.02% | 2.794% | 3.676% | 1.527% | 0.2842% | 2.652% | 10.81% | 1.442% | 4.102% | 86.6% | 42.14% | 26.96% | 10.87% | 114.6% | 0.679% | 0.7687% | 1.174% | 1.045% | 1.064% | 0.9918% | 0.909% | 0.6035% | 36.91% | 19.08% | 2.531% | 1.087%
161 | 10_15_21_F_10_DUP |  | 6/17/2023 4:27:51 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_10_DUP.810 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.57% | 1.155% | 8.298% | 1.525% | 0.4435% | 1.821% | 2.103% | 3.808% | 20.89% | 6.052% | 1.076% | 0.4141% | 1.52% | 32.73% | 5.83% | 1.804% | 1.238% | 1.257% | 1.258% | 4.335% | 0.2277% | 3.856% | 86.6% | 27.71% | 29.04% | 12.08% | 62.45% | 1.025% | 0.2935% | 2.058% | 0.6494% | 0.7846% | 0.9728% | 1.412% | 0.5801% | 27.84% | 21.36% | 1.358% | 0.2494%
162 | LBLANK |  | 6/17/2023 4:33:51 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.811 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 50% | 173.2% | 33.33% | 2.563% | 2.849% | 15.72% | 91.65% |  | 20.4% | 9.75% | 9.506% | 57.14% | 1.062% | 173.2% | 45.4% | 57.28% | 46.33% | 7.217% | 7.022% |  | 57.74% | 6.015% | 173.2% | 114.6% | 43.3% | 22.75% |  | 13.62% | 43.3% | 1.042% | 0.6218% | 0.2723% | 1.28% | 0.4278% | 2.298% | 173.2% | 53.29% | 29.57% | 15.81%
163 | LBLANK |  | 6/17/2023 4:40:11 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.812 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 86.6% | 173.2% | 100% | 10.93% | 0.715% | 23.94% | 173.2% | 173.2% | 27.74% | 7.07% | 6.711% | 43.3% | 1.866% | 173.2% | 21.53% | 45.07% | 25.68% | 3.093% | 3.472% | 173.2% | 44.41% | 2.492% |  |  |  | 18.52% |  | 29.98% | 52.92% | 3.627% | 1.124% | 0.9879% | 0.4514% | 0.4492% | 2.821% |  | 33.33% | 1.46e-06% | 2.547%
164 | CCV 1ppb TRACE |  | 6/17/2023 4:46:11 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.813 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.548% | 2.563% | 6.397% | 2.974% | 2.826% | 6.673% | 4.969% | 1.112% | 1.226% | 1.043% | 5.733% | 1.347% | 2.056% | 4.182% | 8.551% | 3.141% | 1.483% | 0.5341% | 23.96% | 9.805% | 1.604% | 1.305% | 26.41% | 6.489% | 7.378% | 15.28% | 15.34% | 0.7006% | 1.751% | 1.316% | 0.5856% | 0.763% | 0.7265% | 0.7338% | 1.232% | 2.39% | 2.143% | 1.432% | 1.848%
165 | CCV 1ppb TRACE |  | 6/17/2023 4:52:31 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.814 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.765% | 3.857% | 10.22% | 2.655% | 1.732% | 5.531% | 5.969% | 3.497% | 4.679% | 1.08% | 1.359% | 0.5412% | 0.9355% | 4.815% | 5.598% | 3.497% | 0.7587% | 1.495% | 14.04% | 17.63% | 1.034% | 4.552% | 13.82% | 10.43% | 4.616% | 18.22% | 5.391% | 1.768% | 0.9704% | 7.182% | 0.6419% | 0.9373% | 0.7019% | 0.02218% | 2.09% | 2.08% | 0.4727% | 3.174% | 2.168%
166 | CCV 10 ppb MAJORS |  | 6/17/2023 4:58:31 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.815 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.743% | 0.6689% | 7.387% | 1.692% | 3.462% | 1.462% | 3.616% | 1.668% | 1.856% | 1.007% | 1.152% | 0.1713% | 1.024% | 0.8842% | 2.325% | 0.8004% | 1.428% | 1.273% | 1.507% | 3.291% | 0.6699% | 0.3917% | 3.415% | 4.549% | 3.11% | 1.786% | 6.271% | 0.4146% | 0.552% | 0.9511% | 0.3233% | 0.7959% | 0.6926% | 0.3525% | 0.6979% | 1.28% | 1.529% | 1.008% | 1.484%
167 | CCV 10 ppb MAJORS |  | 6/17/2023 5:04:50 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.816 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.4181% | 1.307% | 0.5383% | 0.9595% | 1.651% | 1.277% | 1.581% | 1.234% | 1.08% | 0.8527% | 1.545% | 1.101% | 0.5604% | 1.646% | 2.751% | 1.336% | 0.1899% | 0.7218% | 1.487% | 5.967% | 0.6937% | 2.213% | 7.717% | 0.9354% | 1.496% | 3.137% | 4.726% | 0.8049% | 0.338% | 2.251% | 0.5143% | 0.8808% | 0.516% | 0.8302% | 2.247% | 0.2643% | 0.4343% | 1.099% | 0.9951%
168 | 10_15_21_F_11 |  | 6/17/2023 5:10:50 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_11.817 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.575% | 2.125% | 8.038% | 0.8641% | 2.313% | 1.198% | 0.9033% | 3.447% | 16.31% | 9.091% | 2.404% | 0.5013% | 1.288% | 16.04% | 18.69% | 2.558% | 1.328% | 0.7237% | 3.494% | 9.739% | 1.5% | 4.393% | 48.04% | 37.63% | 18.65% | 14.99% | 54.13% | 1.109% | 0.5402% | 0.07403% | 0.4769% | 0.5942% | 0.1247% | 0.2647% | 1.043% | 16.11% | 5.292% | 2.728% | 1.173%
169 | 10_15_21_F_11 |  | 6/17/2023 5:17:09 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_11.818 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.4989% | 0.8141% | 3.244% | 2.289% | 1.167% | 2.255% | 1.067% | 1.954% | 9.358% | 5.382% | 0.9845% | 0.6351% | 0.7513% | 19.87% | 12.14% | 2.432% | 2.046% | 0.5216% | 4.596% | 6.983% | 0.5865% | 4.299% | 24.74% | 44.89% | 31.77% | 7.358% | 17.32% | 1.945% | 0.5084% | 0.4129% | 0.9654% | 0.5893% | 3.357% | 0.7831% | 0.9304% | 19.16% | 2.585% | 3.096% | 2.307%
170 | 10_15_21_UF_1 |  | 6/17/2023 5:23:09 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_1.819 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.7054% | 1.558% | 14.58% | 1.011% | 1.364% | 2.631% | 1.294% | 6.967% | 5.957% | 4.783% | 2.579% | 0.4434% | 1.385% | 15.62% | 9.126% | 1.818% | 2.803% | 0.6381% | 1.891% | 10.86% | 0.78% | 4.661% | 86.6% | 48.24% | 28.64% | 10.41% | 86.6% | 1.388% | 0.4803% | 0.9167% | 0.4293% | 0.508% | 1.344% | 0.3541% | 0.7725% | 21.14% | 15.33% | 4.391% | 6.71%
171 | 10_15_21_UF_1 |  | 6/17/2023 5:29:28 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_1.820 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.15% | 0.7173% | 9.024% | 1.522% | 0.4896% | 0.8666% | 0.1318% | 5.407% | 4.528% | 3.693% | 2.038% | 0.3893% | 0.698% | 6.926% | 3.382% | 4.826% | 1.708% | 1.506% | 3.336% | 7.434% | 1.632% | 3.147% | 91.65% | 14.29% | 10.66% | 7.324% | 94.37% | 0.9684% | 0.7225% | 0.1832% | 1.218% | 0.2944% | 0.4764% | 0.9472% | 0.7133% | 12.18% | 19.7% | 5.46% | 7.246%
172 | 10_15_21_UF_6 |  | 6/17/2023 5:35:27 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_6.821 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.6635% | 3.656% | 8.183% | 1.527% | 2.482% | 1.69% | 1.716% | 1.01% | 4.822% | 12.66% | 3.319% | 0.7421% | 2.963% | 22.48% | 13.78% | 1.517% | 2.256% | 1.448% | 2.49% | 3.705% | 1.792% | 2.237% | 25% | 35.66% | 20.8% | 15.68% | 25% | 0.8075% | 0.799% | 1.196% | 1.086% | 0.1733% | 0.6111% | 0.8935% | 1.225% | 12.76% | 5.376% | 4.441% | 3.456%
173 | 10_15_21_UF_6 |  | 6/17/2023 5:41:46 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_6.822 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.215% | 0.8277% | 10.75% | 2.994% | 1.845% | 4.846% | 0.2418% | 5.962% | 10.76% | 5.717% | 1.89% | 0.4396% | 1.004% | 20.03% | 4.137% | 2.53% | 1.025% | 0.2818% | 4.357% | 1.776% | 0.9998% | 5.138% | 0% | 42.87% | 13.03% | 6.79% | 88.19% | 1.476% | 0.4597% | 2.371% | 0.2845% | 0.7375% | 0.683% | 0.8663% | 0.3766% | 7.05% | 7.344% | 2.528% | 0.6994%
174 | 10_15_21_UF_8 |  | 6/17/2023 5:47:46 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_8.823 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.344% | 1.965% | 8.57% | 0.8306% | 0.2494% | 1.109% | 0.9854% | 1.959% | 3.494% | 10.65% | 0.5866% | 0.1702% | 2.699% | 18.49% | 9.856% | 4.554% | 1.087% | 0.7174% | 2.704% | 17.9% | 1.182% | 5.915% | 34.64% | 36.36% | 26.59% | 9.983% | 49.49% | 1.49% | 0.1354% | 0.1926% | 0.3072% | 0.5153% | 0.5438% | 0.6148% | 0.7879% | 12.74% | 4.41% | 1.799% | 4.112%
175 | 10_15_21_UF_8 |  | 6/17/2023 5:54:04 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_8.824 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.601% | 1.026% | 3.637% | 2.553% | 1.258% | 2.286% | 0.9205% | 2.597% | 10.22% | 1.23% | 0.7226% | 0.6479% | 1.146% | 13.23% | 11.07% | 2.877% | 0.3777% | 1.577% | 3.804% | 4.826% | 0.395% | 7.144% | 41.66% | 48.71% | 13.68% | 3.145% | 17.32% | 1.034% | 0.7107% | 1.203% | 0.6943% | 0.8069% | 2.57% | 0.7964% | 1.393% | 16.5% | 20.31% | 3.71% | 0.1719%
176 | 10_15_21_UF_8_LD |  | 6/17/2023 6:00:03 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_8_LD.825 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.4073% | 1.399% | 4.798% | 1.578% | 0.9767% | 3.261% | 0.5137% | 2.553% | 9.148% | 3.617% | 2.077% | 1.985% | 0.8228% | 4.461% | 3.804% | 3.209% | 1.643% | 0.5739% | 0.3512% | 3.791% | 0.5205% | 2.08% | 34.64% | 43.59% | 11.62% | 9.869% | 13.32% | 0.6914% | 1.893% | 0.5062% | 5.726% | 0.41% | 0.1099% | 0.5577% | 1.264% | 14.84% | 12.22% | 1.331% | 2.45%
177 | 10_15_21_UF_8_LD |  | 6/17/2023 6:06:21 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_8_LD.826 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.6722% | 1.59% | 2.164% | 0.9965% | 0.9691% | 1.076% | 1.619% | 1.434% | 4.995% | 8.63% | 1.281% | 0.2836% | 1.651% | 19.08% | 11.65% | 0.9532% | 1.801% | 0.3146% | 2.99% | 10.19% | 0.5239% | 12.9% | 21.65% | 29.57% | 17.8% | 2.331% | 24.74% | 1.564% | 0.4399% | 0.6671% | 0.3272% | 1.419% | 0.8862% | 0.7562% | 0.8365% | 19.92% | 9.719% | 3.188% | 2.805%
178 | LBLANK |  | 6/17/2023 6:12:21 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.827 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 86.6% | 86.6% | 33.33% | 3.358% | 5.837% | 29.89% | 100% | 100% | 22.21% | 15.75% | 8.344% | 31.49% | 1.837% |  | 20.15% | 50.92% | 5.094% | 16.66% | 7.371% | 86.6% | 19.52% | 1.409% |  | 173.2% | 50% | 27.41% | 173.2% | 22.63% | 18.9% | 4.406% | 0.3825% | 0.1172% | 0.5691% | 0.8644% | 0.7997% | 100% | 57.74% | 45.58% | 25.52%
179 | LBLANK |  | 6/17/2023 6:18:40 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.828 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 50% | 173.2% | 27.15% | 12.15% | 7.832% | 34.42% | 86.6% | 173.2% | 20.38% | 12.37% | 4.801% | 24.12% | 1.051% | 86.6% | 33.07% | 26.65% | 7.559% | 2.567% | 4.098% | 173.2% | 26.96% | 1.423% |  | 173.2% | 57.28% | 22.38% |  | 31.25% | 24.98% | 1.018% | 0.2777% | 0.4676% | 2.53% | 0.6979% | 1.457% | 173.2% | 33.33% | 32.73% | 45.42%
180 | CCV 1ppb TRACE |  | 6/17/2023 6:24:40 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.829 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4.28% | 7.067% | 2.606% | 1.117% | 4.062% | 2.359% | 4.912% | 3.193% | 2.358% | 1.76% | 3.521% | 1.171% | 3.611% | 3.46% | 2.168% | 3.024% | 2.003% | 1.969% | 36.34% | 6.191% | 1.247% | 2.717% | 5% | 7.2% | 6.477% | 10.05% | 18.7% | 0.5729% | 2.173% | 1.844% | 0.3839% | 0.2344% | 0.6877% | 0.6185% | 0.492% | 1.6% | 1.025% | 1.404% | 0.9239%
181 | CCV 1ppb TRACE |  | 6/17/2023 6:31:00 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.830 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.929% | 5.1% | 8.579% | 2.281% | 0.9041% | 2.708% | 6.207% | 3.056% | 4.561% | 1.637% | 2.772% | 1.731% | 1.767% | 0.464% | 0.8468% | 1.121% | 1.186% | 2.819% | 28.99% | 5.957% | 1.786% | 4.552% | 8.314% | 12.62% | 5.067% | 9.173% | 15.7% | 1.408% | 1.712% | 1.818% | 0.9932% | 0.4423% | 0.5217% | 1.183% | 1.402% | 1.548% | 2.217% | 1.937% | 1.29%
182 | LFB |  | 6/17/2023 6:37:00 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.831 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.314% | 0.5129% | 7.703% | 1.673% | 2.143% | 0.6784% | 2.048% | 2.515% | 1.668% | 0.8353% | 0.6004% | 0.6325% | 1.778% | 1.908% | 0.5198% | 0.755% | 0.4202% | 1.867% | 1.841% | 4.128% | 1.126% | 1.227% | 3.438% | 1.393% | 0.7817% | 2.468% | 2.469% | 0.8037% | 0.9142% | 0.2936% | 0.09303% | 1.796% | 0.6008% | 0.2231% | 1.394% | 0.315% | 0.1738% | 1.344% | 1.36%
183 | LFB |  | 6/17/2023 6:43:18 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.832 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.378% | 0.4363% | 1.837% | 0.1608% | 1.226% | 1.925% | 1.06% | 2.17% | 0.9598% | 0.714% | 0.8152% | 0.2485% | 1.928% | 1.059% | 2.323% | 1.488% | 0.1904% | 1.16% | 1.349% | 5.124% | 0.7472% | 1.985% | 4.676% | 4.024% | 1.237% | 1.556% | 3.245% | 0.8438% | 0.09945% | 0.2805% | 0.7076% | 0.7562% | 3.879% | 1.099% | 1.651% | 0.2986% | 0.3437% | 0.06979% | 1.389%
184 | CCV 10 ppb MAJORS |  | 6/17/2023 6:49:17 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.833 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.087% | 1.587% | 4.7% | 0.796% | 0.6774% | 0.6267% | 1.95% | 0.5208% | 0.9044% | 1.751% | 8.041% | 0.9295% | 2.521% | 0.5416% | 2.222% | 0.226% | 1.109% | 0.4008% | 1.297% | 1.994% | 1.049% | 1.755% | 5.67% | 1.638% | 0.5224% | 1.536% | 2.946% | 0.5542% | 0.2284% | 2.554% | 0.4529% | 0.4295% | 0.9437% | 0.3084% | 0.8975% | 1.465% | 0.5326% | 1.041% | 1.432%
185 | CCV 10 ppb MAJORS |  | 6/17/2023 6:55:37 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.834 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.191% | 4.003% | 5.414% | 1.534% | 2.094% | 1.765% | 1.149% | 0.9959% | 0.8604% | 1.246% | 1.502% | 1.223% | 1.542% | 0.5401% | 1.666% | 0.8141% | 0.7669% | 0.631% | 3.214% | 6.673% | 0.9869% | 1.948% | 4.61% | 2.104% | 0.3664% | 2.156% | 5.185% | 1.561% | 0.5336% | 0.6864% | 0.1965% | 1.142% | 0.3485% | 0.7907% | 1.439% | 1.069% | 0.336% | 1.126% | 1.19%
186 | 10_15_21_UF_8_DUP |  | 6/17/2023 7:01:37 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_8_DUP.835 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.02% | 1.901% | 11.98% | 2.192% | 0.3037% | 1.8% | 1.14% | 1.556% | 11.63% | 1.939% | 0.9808% | 0.661% | 0.3404% | 16.82% | 6.398% | 3.631% | 1.183% | 1.347% | 4.14% | 7.195% | 2.496% | 6.733% | 66.67% | 31.13% | 12.88% | 14.09% | 13.32% | 1.186% | 0.9849% | 0.1749% | 0.2182% | 1.236% | 1.021% | 0.4684% | 2.809% | 35.18% | 6.347% | 1.63% | 2.035%
187 | 10_15_21_UF_8_DUP |  | 6/17/2023 7:07:56 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_8_DUP.836 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.708% | 1.227% | 3.16% | 3.318% | 2.777% | 1.694% | 1.309% | 7.516% | 9.997% | 5.166% | 2.438% | 1.554% | 1.495% | 11.36% | 8.586% | 2.493% | 0.3009% | 2.571% | 1.514% | 7.184% | 2.033% | 6.277% | 0% | 40.29% | 19.6% | 21.27% | 48.04% | 0.8529% | 1.06% | 0.5484% | 0.5125% | 1.34% | 0.9152% | 0.5244% | 1.625% | 11.35% | 21.1% | 3.975% | 2.401%
188 | 10_15_21_UF_8_DUP_LSP |  | 6/17/2023 7:13:55 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_8_DUP_LSP.837 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.8982% | 1.339% | 4.469% | 4.598% | 1.195% | 1.619% | 1.595% | 0.7169% | 1.195% | 0.674% | 0.215% | 1.06% | 0.1841% | 1.135% | 1.581% | 0.9806% | 0.5566% | 0.3955% | 0.4507% | 1.956% | 0.7605% | 1.439% | 7.623% | 1.984% | 2.196% | 1.841% | 2.167% | 0.8565% | 0.8053% | 0.4237% | 0.3315% | 0.4178% | 1.069% | 0.8426% | 1.13% | 0.5317% | 0.4931% | 1.378% | 0.7458%
189 | 10_15_21_UF_8_DUP_LSP |  | 6/17/2023 7:20:13 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_8_DUP_LSP.838 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.8742% | 0.8342% | 2.293% | 1.284% | 0.1497% | 0.7012% | 0.7736% | 1.571% | 0.9286% | 0.7037% | 0.4128% | 0.8838% | 0.8169% | 0.3932% | 1.294% | 0.4259% | 0.2792% | 0.4943% | 1.991% | 4.194% | 0.8281% | 2.763% | 3.558% | 3.355% | 1.033% | 1.742% | 2.876% | 0.3431% | 0.5189% | 2.649% | 0.4814% | 0.8073% | 0.515% | 0.4411% | 1.224% | 0.4252% | 0.4373% | 1.555% | 1.198%
190 | 10_15_21_UF_9 |  | 6/17/2023 7:26:13 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_9.839 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.5192% | 2.412% | 2.621% | 0.3978% | 0.6659% | 3.15% | 1.123% | 3.58% | 3.696% | 1.691% | 1.916% | 0.5275% | 0.08778% | 23.85% | 3.028% | 0.3251% | 2.233% | 0.4431% | 0.874% | 13.95% | 1.828% | 5.495% | 21.65% | 30.14% | 26.65% | 17.59% | 57.74% | 0.8398% | 0.8184% | 0.2182% | 0.6003% | 0.7352% | 0.3603% | 1.181% | 0.8849% | 29.63% | 17.58% | 1.297% | 4.049%
191 | 10_15_21_UF_9 |  | 6/17/2023 7:32:31 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_9.840 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.3781% | 1.134% | 1.309% | 2.638% | 0.9255% | 3.082% | 0.6422% | 8.708% | 13.43% | 3.011% | 0.9871% | 1.514% | 0.805% | 17.16% | 16.17% | 2.775% | 0.5838% | 0.514% | 4.121% | 7.563% | 1.826% | 3.711% | 69.28% | 21.65% | 23.32% | 21.34% | 65.47% | 1.154% | 0.72% | 0.9742% | 0.5266% | 1.202% | 0.3178% | 0.5594% | 1.763% | 13.04% | 7.146% | 8.352% | 3.97%
192 | 10_15_21_UF_10 |  | 6/17/2023 7:38:31 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_10.841 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.9488% | 1.21% | 13.01% | 1.911% | 1.279% | 0.1331% | 0.9675% | 4.541% | 8.126% | 1.8% | 3.2% | 0.9367% | 2.069% | 21.46% | 7.073% | 3.568% | 1.508% | 0.9421% | 2.813% | 6.097% | 1.249% | 4.263% | 56.73% | 36.06% | 4.348% | 1.301% | 26.96% | 1.696% | 1.078% | 0.1584% | 1.186% | 0.9719% | 14.82% | 0.1394% | 0.3658% | 17.65% | 5.249% | 2.345% | 6.388%
193 | 10_15_21_UF_10 |  | 6/17/2023 7:44:50 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_10.842 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.591% | 1.302% | 18.01% | 2.794% | 3.024% | 3.302% | 1.583% | 4.544% | 9.107% | 6.643% | 1.499% | 0.4643% | 0.7483% | 9.362% | 7.081% | 2.089% | 0.7427% | 1.167% | 1.413% | 6.16% | 2.38% | 2.942% | 21.65% | 19.87% | 33.86% | 33.62% | 49.49% | 1.307% | 1.06% | 0.7467% | 0.8777% | 0.2577% | 2.528% | 0.6488% | 1.17% | 7.07% | 20.59% | 1.867% | 2.441%
194 | 10_15_21_UF_11 |  | 6/17/2023 7:50:50 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_11.843 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.4957% | 0.9083% | 2.373% | 2.313% | 2.085% | 3.886% | 0.3198% | 2.204% | 1.176% | 5.365% | 4.085% | 0.9362% | 1.078% | 18.23% | 8.98% | 0.4909% | 0.2011% | 0.2097% | 1.393% | 2.77% | 1.437% | 3.407% | 124.9% | 26.96% | 11.11% | 14.21% | 71.32% | 0.9966% | 1.543% | 0.505% | 0.2873% | 0.1994% | 3.137% | 0.8055% | 0.9374% | 51.7% | 9.835% | 4.549% | 1.901%
195 | 10_15_21_UF_11 |  | 6/17/2023 7:57:09 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_11.844 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.9205% | 2.15% | 5.313% | 1.512% | 1.5% | 0.743% | 0.9911% | 6.569% | 3.485% | 16.37% | 0.6619% | 0.05568% | 0.5897% | 10.18% | 17.09% | 1.454% | 1.799% | 1.187% | 2.537% | 5.963% | 0.7147% | 2.426% | 78.06% | 27.73% | 18.94% | 22.83% | 74.18% | 1.143% | 0.9792% | 0.1615% | 0.3486% | 0.1205% | 0.1712% | 1.083% | 1.883% | 18.9% | 11.46% | 1.873% | 1.742%
196 | LBLANK |  | 6/17/2023 8:03:10 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.845 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 34.64% | 173.2% | 86.6% | 5.794% | 3.761% | 14.19% |  | 173.2% | 21.43% | 74.41% | 11.87% | 28.87% | 1.313% |  | 18.33% | 0% | 10.13% | 5.443% | 3.383% | 173.2% | 23.82% | 0.7996% |  | 91.65% | 86.6% | 18.99% | 86.6% | 14.18% | 45.83% | 3.274% | 0.9434% | 4.597% | 1.051% | 0.982% | 0.5099% |  | 96.44% | 13.86% | 21.53%
197 | LBLANK |  | 6/17/2023 8:09:28 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.846 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 173.2% |  | 88.19% | 7.097% | 2.375% | 22.88% | 43.3% | 173.2% | 33.33% | 18.37% | 10.6% | 33.33% | 1.834% | 86.6% | 33.33% | 57.28% | 24.41% | 2.505% | 8.291% | 86.6% | 12.85% | 4.559% | 173.2% | 100% | 69.28% | 24.55% | 173.2% | 17.56% | 10.19% | 11.6% | 0.3508% | 3.078% | 0.775% | 1.091% | 0.5998% | 43.3% | 26.65% | 30.1% | 23.32%
198 | CCV 1ppb TRACE |  | 6/17/2023 8:15:29 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.847 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.492% | 5.083% | 1.805% | 7.076% | 4.448% | 3.156% | 2.319% | 1.264% | 2.897% | 0.719% | 2.714% | 2.815% | 2.084% | 1.799% | 8.169% | 2.905% | 3.611% | 0.1694% | 30.83% | 18.98% | 2.745% | 2.493% | 15.29% | 8.31% | 1.815% | 18.89% | 14.74% | 2.633% | 0.9337% | 0.4053% | 0.1566% | 0.143% | 0.3961% | 0.2567% | 0.4616% | 1.267% | 1.589% | 0.5074% | 1.516%
199 | CCV 1ppb TRACE |  | 6/17/2023 8:21:48 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.848 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.672% | 3.342% | 15.23% | 3.882% | 5.16% | 4.645% | 2.3% | 4.629% | 3.07% | 1.707% | 1.732% | 1.171% | 0.2454% | 1.715% | 1.848% | 2.815% | 1.302% | 2.257% | 13.44% | 4.545% | 1.604% | 3.733% | 25.65% | 6.767% | 9.019% | 7.055% | 19.47% | 0.8493% | 1.296% | 0.7409% | 0.5037% | 0.6126% | 0.645% | 0.4516% | 0.822% | 0.2957% | 1.929% | 1.024% | 1.088%
200 | CCV 10 ppb MAJORS |  | 6/17/2023 8:27:48 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.849 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.8399% | 0.3685% | 9.515% | 0.7599% | 0.7277% | 1.849% | 2.798% | 1% | 2.207% | 0.4212% | 0.5995% | 1.143% | 1.736% | 0.3944% | 2.558% | 1.348% | 1.185% | 0.9063% | 1.982% | 5.775% | 1.259% | 0.4133% | 5.795% | 1.641% | 1.193% | 1.656% | 0.6763% | 1.081% | 1.081% | 0.09716% | 0.8362% | 0.7998% | 0.1872% | 1.181% | 0.7117% | 0.2852% | 1.083% | 1.43% | 1.012%
201 | CCV 10 ppb MAJORS |  | 6/17/2023 8:34:07 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.850 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.544% | 1.702% | 4.653% | 2.302% | 0.6347% | 0.687% | 1.776% | 1.605% | 0.5952% | 1.521% | 2.513% | 0.8212% | 0.2447% | 1.455% | 3.208% | 0.3088% | 0.8406% | 1.501% | 3.21% | 3.824% | 0.4302% | 1.448% | 2.939% | 5.577% | 2.433% | 1.957% | 3.662% | 1.304% | 1.059% | 2.453% | 0.9137% | 0.7203% | 0.4736% | 0.2292% | 0.7842% | 0.3294% | 0.8592% | 2.668% | 0.941%
202 | 10_15_21_FB |  | 6/17/2023 8:40:08 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_FB.851 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9.72% | 33.33% | 5.668% | 3.78% | 1.117% | 31.73% | 53.29% | 100% | 4.724% | 5.328% | 1.34% | 7.056% | 0.9301% | 43.3% | 16.43% | 7.982% | 5.787% | 0.5003% | 8.169% | 173.2% | 12.56% | 4.905% |  | 86.6% | 114.6% | 90.78% |  | 5.036% | 11.74% | 0.1586% | 11.15% | 0.5913% | 0.3092% | 0.5303% | 1.478% | 91.65% | 25% | 3.978% | 4.787%
203 | 10_15_21_FB |  | 6/17/2023 8:46:26 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_FB.852 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4.187% | 34.64% | 1.038% | 2.031% | 3.949% | 29.57% | 75.25% | 43.3% | 8.629% | 7.508% | 2.1% | 6.627% | 1.344% | 100% | 28.39% | 17.21% | 5.678% | 0.5302% | 4.326% | 173.2% | 9.116% | 4.172% | 173.2% | 173.2% | 100% | 80.92% |  | 5.728% | 23.53% | 0.9108% | 0.6744% | 0.6422% | 0.3913% | 0.1658% | 0.6935% | 86.6% | 35.25% | 6.271% | 2.398%
204 | 10_15_21_BB |  | 6/17/2023 8:52:26 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_BB.853 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 66.67% | 86.6% | 6.183% | 2.203% | 2.116% | 32.22% | 50% | 173.2% | 22.8% | 10.85% | 9.34% | 43.75% | 2.51% | 33.33% | 18.87% | 15.19% | 3.213% | 1.708% | 6.237% | 173.2% | 10.19% | 7.105% | 173.2% | 86.6% | 86.6% | 17.41% |  | 14.52% | 73.47% | 0.1579% | 0.6144% | 1.012% | 0.5925% | 1.232% | 1.993% | 173.2% | 28.64% | 5.908% | 7.037%
205 | 10_15_21_BB |  | 6/17/2023 8:58:45 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_BB.854 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 33.33% |  | 13.26% | 0.722% | 0.7435% | 4.558% | 78.06% |  | 20.15% | 11.59% | 0.6649% | 28.13% | 2.105% | 173.2% | 24.35% | 6.604% | 2.275% | 6.081% | 7.724% |  | 7.531% | 3.633% |  | 100% | 69.28% | 84.85% | 173.2% | 15.05% | 61.97% | 0.4912% | 1.114% | 1.255% | 1.123% | 0.5711% | 0.7145% |  | 13.32% | 4.539% | 2.574%
206 | 10_15_21_LAB_A |  | 6/17/2023 9:04:45 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_LAB_A.855 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 43.3% |  | 12.5% | 3.02% | 2.794% | 8.333% | 100% |  | 16.6% | 2.585% | 4.052% | 12.82% | 2.213% | 173.2% | 45.07% | 4.087% | 5.092% | 4.67% | 4.407% |  | 28.39% | 2.866% |  |  | 173.2% | 34.41% |  | 8.585% | 41.83% | 0.5217% | 0.2406% | 4.992% | 2.44% | 1.429% | 1.157% | 173.2% | 53.93% | 6.7% | 3.708%
207 | 10_15_21_LAB_A |  | 6/17/2023 9:11:03 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_LAB_A.856 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 62.45% |  | 10.93% | 4.437% | 3.943% | 43.3% | 43.3% | 100% | 7.132% | 1.449% | 0.5104% | 9.962% | 1.54% | 100% | 33.33% | 10.76% | 6.242% | 3.859% | 4.948% | 173.2% | 24.12% | 1.925% |  | 173.2% | 0% | 16.81% |  | 10.42% | 12.2% | 0.4016% | 0.5544% | 0.3027% | 0.7753% | 0.5973% | 0.2759% | 173.2% | 81.13% | 7.808% | 4.245%
208 | 10_15_21_LAB_B |  | 6/17/2023 9:17:03 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_LAB_B.857 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 57.28% | 173.2% | 18.84% | 3.97% | 1.433% | 35.25% | 100% |  | 6.667% | 13.25% | 4.413% | 29.04% | 1.996% | 100% | 26.96% | 8.244% | 50.18% | 2.999% | 3.933% | 86.6% | 48.04% | 8.635% |  | 173.2% | 100% | 132.7% | 173.2% | 32.78% | 66.14% | 0.09203% | 0.7893% | 0.1261% | 0.6512% | 1.254% | 1.026% | 173.2% | 57.28% | 2.26% | 4.941%
209 | 10_15_21_LAB_B |  | 6/17/2023 9:23:20 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_LAB_B.858 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 173.2% |  | 10.65% | 6.768% | 3.177% | 7.07% | 173.2% | 0% | 8.66% | 8.882% | 3.725% | 14.19% | 1.557% | 86.6% | 24.02% | 21.55% | 5.801% | 2.32% | 4.767% |  | 12.37% | 0.3063% | 173.2% |  |  | 33.9% | 173.2% | 4.681% | 68.63% | 0.238% | 0.5297% | 1.005% | 0.773% | 0.5101% | 1.359% | 173.2% | 26.96% | 6.041% | 2.586%
210 | 5_18_23_DB |  | 6/17/2023 9:29:20 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_DB.859 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 173.2% |  | 24.3% | 2.148% | 0.8228% | 18.37% | 43.3% |  | 16.88% | 9.362% | 4.005% | 58.79% | 0.8974% | 43.3% | 43.3% | 5.582% | 0.3416% | 0.1576% | 6.02% |  | 26.96% | 9.606% |  |  | 86.6% | 65% |  | 24.98% | 114.6% | 1.367% | 0.8639% | 0.2967% | 0.9445% | 1.284% | 0.6405% | 173.2% | 43.3% | 6.258% | 1.471%
211 | 5_18_23_DB |  | 6/17/2023 9:35:39 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_DB.860 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 100% |  | 35.5% | 2.463% | 3.9% | 3.765% | 114.6% | 0% | 17.64% | 18.23% | 1.636% | 7.531% | 1.099% | 100% | 28.57% | 10.02% | 5.928% | 4.33% | 1.526% | 173.2% | 76.38% | 2.867% |  | 173.2% | 173.2% | 9.413% | 173.2% | 37.89% | 86.6% | 1.684% | 1.583% | 0.691% | 0.3408% | 1.09% | 1.047% | 100% | 35.25% | 6.413% | 0.9615%
212 | LBLANK |  | 6/17/2023 9:41:39 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.861 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 173.2% |  | 124.9% | 2.029% | 4.852% | 50.69% | 86.6% | 173.2% | 27.06% | 49.66% | 12.12% | 52.92% | 0.9123% | 173.2% | 56.77% | 31.23% | 16.4% | 11.56% | 1.045% | 173.2% | 28.39% | 6.741% |  | 173.2% | 86.6% | 56.06% | 173.2% | 18.75% | 35.25% | 6.235% | 0.6584% | 0.5743% | 0.1421% | 0.631% | 0.6248% | 86.6% | 43.3% | 25% | 26.65%
213 | LBLANK |  | 6/17/2023 9:47:59 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.862 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 173.2% |  | 33.33% | 3.32% | 3.084% | 25.52% | 173.2% |  | 26.65% | 24.74% | 4.375% | 62.45% | 1.706% | 86.6% | 50.94% | 62.45% | 20.65% | 6.2% | 4.09% |  | 74.18% | 5.621% | 173.2% |  |  | 673.8% |  | 24.77% | 12.37% | 3.775% | 0.2978% | 0.3872% | 0.5834% | 0.4001% | 2.111% | 173.2% | 25% | 36.32% | 13.48%
214 | CCV 1ppb TRACE |  | 6/17/2023 9:53:59 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.863 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.843% | 0.4184% | 18.08% | 6.828% | 1.024% | 3.331% | 5.465% | 3.229% | 4.897% | 0.317% | 3.049% | 0.504% | 0.5989% | 1.921% | 6.195% | 4.347% | 1.576% | 1.883% | 68.71% | 15.67% | 0.7194% | 1.914% | 16.4% | 13.53% | 5.128% | 6.03% | 24.4% | 2.512% | 1.655% | 0.5838% | 1.076% | 1.098% | 0.8227% | 0.8006% | 2.185% | 1.46% | 1.545% | 1.527% | 0.8447%
215 | CCV 1ppb TRACE |  | 6/17/2023 10:00:19 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.864 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.105% | 2.366% | 11.2% | 4.241% | 2.355% | 6.262% | 7.899% | 1.503% | 1.839% | 1.558% | 3.535% | 0.7994% | 1.045% | 0.8805% | 4.437% | 2.075% | 1.384% | 1.437% | 11.49% | 11.83% | 2.313% | 4.385% | 14.66% | 19.37% | 3.666% | 17.82% | 13.46% | 4.266% | 2.546% | 0.6072% | 0.7237% | 7.333% | 0.6219% | 0.8248% | 1.009% | 2.968% | 2.429% | 1.306% | 0.2552%
216 | LFB |  | 6/17/2023 10:06:18 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.865 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.485% | 1.999% | 6.28% | 0.3453% | 1.882% | 1.131% | 3.592% | 0.3753% | 2.276% | 0.5112% | 1.709% | 0.6044% | 0.8796% | 1.893% | 3.189% | 2.282% | 0.7807% | 0.4379% | 2.382% | 2.93% | 0.6958% | 4.715% | 9.241% | 1.785% | 2.178% | 1.367% | 6.955% | 1.201% | 1.106% | 0.5999% | 0.4423% | 1.127% | 0.5146% | 0.2978% | 1.436% | 0.7639% | 1.27% | 0.577% | 0.9344%
217 | LFB |  | 6/17/2023 10:12:36 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.866 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.22% | 1.152% | 8.713% | 1.813% | 3.891% | 0.5798% | 1.057% | 2.167% | 0.7929% | 1.167% | 2.587% | 0.6039% | 1.086% | 0.5015% | 0.9835% | 1.08% | 1.639% | 1.066% | 1.387% | 3.174% | 0.9217% | 2.057% | 8.662% | 4.181% | 2.176% | 2.555% | 4.885% | 1.382% | 0.8564% | 1.821% | 1.218% | 2.141% | 0.6679% | 0.7912% | 0.5973% | 2.449% | 2.304% | 1.153% | 0.9175%
218 | CCV 10 ppb MAJORS |  | 6/17/2023 10:18:36 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.867 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.801% | 1.326% | 1.937% | 0.4753% | 2.034% | 0.9092% | 0.9018% | 1.253% | 0.1164% | 0.2653% | 0.9806% | 0.8075% | 2.174% | 0.9861% | 1.241% | 0.53% | 1.083% | 0.5845% | 1.445% | 3.585% | 0.3603% | 1.347% | 2.997% | 1.452% | 2.606% | 2.147% | 3.055% | 1.055% | 1.085% | 0.8282% | 0.7479% | 1.438% | 0.6097% | 0.7542% | 2.072% | 2.223% | 0.9836% | 0.8509% | 1.499%
219 | CCV 10 ppb MAJORS |  | 6/17/2023 10:24:55 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.868 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.319% | 1.019% | 2.281% | 1.97% | 1.317% | 1.036% | 1.409% | 1.306% | 1.103% | 1.575% | 1.021% | 0.3824% | 1.054% | 1.269% | 0.88% | 0.3483% | 0.7501% | 0.8768% | 1.667% | 1.482% | 1.11% | 1.6% | 5.354% | 3.103% | 1.21% | 1.075% | 6.104% | 1.157% | 0.519% | 0.5097% | 0.1366% | 0.1719% | 0.6183% | 0.7821% | 2.089% | 0.7671% | 0.1902% | 0.9224% | 0.911%
220 | 5_18_23_DB |  | 6/17/2023 10:30:55 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_DB.869 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 57.74% |  | 14.52% | 2.303% | 5.139% | 33.53% | 114.6% | 100% | 10.77% | 20.46% | 1.561% | 21.18% | 1.786% | 17.32% | 16.67% | 13.79% | 2.784% | 6.738% | 2.253% | 173.2% | 28.87% | 5.351% |  | 114.6% | 86.6% | 14.67% | 100% | 9.362% | 31.11% | 3.284% | 0.3298% | 1.188% | 2.003% | 0.9762% | 1.486% |  | 28.87% | 7.138% | 1.837%
221 | 5_18_23_DB |  | 6/17/2023 10:37:14 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_DB.870 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 98.97% | 173.2% | 1.946% | 1.4% | 0.765% | 34.64% | 86.6% | 173.2% | 5% | 16.37% | 2.684% | 25.92% | 3.996% | 173.2% | 26.96% | 14.28% | 5.871% | 5.413% | 8.66% | 86.6% | 41.66% | 4.733% | 173.2% | 173.2% | 173.2% | 33.7% | 86.6% | 9.03% | 34.64% | 1.134% | 0.6438% | 0.9004% | 0.3227% | 0.3216% | 0.9927% | 173.2% | 43.3% | 8.923% | 6.476%
222 | 5_18_23_DB |  | 6/17/2023 10:43:14 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_DB.871 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 69.28% |  | 13.15% | 2.528% | 2.222% | 14.7% | 94.37% | 173.2% | 13.32% | 8.272% | 7.382% | 17.32% | 4.391% | 9.073% | 42.27% | 12.76% | 4.755% | 2.354% | 9.568% | 86.6% | 20.83% | 9.744% |  | 86.6% | 100% | 63.88% | 173.2% | 9.939% | 44.61% | 15.75% | 1.103% | 0.1436% | 0.6102% | 0.5082% | 0.5605% | 86.6% | 25% | 2.896% | 1.683%
223 | 5_18_23_DB |  | 6/17/2023 10:49:33 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_DB.872 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 49.49% |  | 13.86% | 1.418% | 5.653% | 18.18% | 0% | 173.2% | 28.51% | 15.32% | 1.889% | 30.04% | 0.9551% | 8.517% | 47.19% | 10.68% | 4.977% | 2.848% | 2.976% |  | 52.68% | 3.342% | 86.6% | 173.2% | 43.3% | 78.17% |  | 17.76% | 28.64% | 1.045% | 0.8654% | 0.355% | 11.84% | 1.151% | 0.9395% | 100% | 91.65% | 0.3597% | 6.541%
224 | 5_18_23_DB |  | 6/17/2023 10:55:33 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_DB.873 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 62.45% |  | 25.93% | 1.257% | 7.766% | 20.35% | 173.2% | 173.2% | 20.26% | 5% | 9.74% | 26.64% | 2.079% | 43.3% | 12.5% | 2.696% | 9.774% | 3.011% | 6.272% |  | 56.77% | 1.267% |  |  | 124.9% | 133.9% |  | 20.69% | 57.28% | 1.29% | 0.4664% | 0.6213% | 0.4251% | 1.018% | 0.8267% | 86.6% | 48.04% | 8.932% | 8.913%
225 | 5_18_23_DB |  | 6/17/2023 11:01:51 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_DB.874 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 57.74% |  | 32.83% | 2.856% | 3.004% | 22.91% | 34.64% |  | 21.92% | 22.53% | 9.479% | 8.921% | 3.163% | 173.2% | 45.83% | 10.63% | 5.318% | 6.717% | 3.595% | 100% | 69.28% | 5.677% |  | 124.9% | 91.65% | 68.29% |  | 18.28% | 65.47% | 0.4389% | 0.1387% | 1.361% | 0.1499% | 0.3278% | 1.918% | 173.2% | 49.49% | 3.35% | 5.633%
226 | 5_18_23_MFB |  | 6/17/2023 11:07:51 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_MFB.875 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.867% | 3.432% | 5.714% | 0.8365% | 1.59% | 2.18% | 1.548% | 2.052% | 2.01% | 0.6876% | 0.9924% | 0.3283% | 0.7715% | 0.7673% | 1.697% | 1.482% | 0.6386% | 0.9191% | 2.808% | 7.385% | 1.199% | 3.141% | 3.076% | 3.141% | 3.018% | 2.108% | 2.994% | 0.7578% | 1.505% | 0.8757% | 2.025% | 1.277% | 0.4159% | 1.774% | 0.576% | 1.441% | 0.4671% | 0.3097% | 1.32%
227 | 5_18_23_MFB |  | 6/17/2023 11:14:09 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_MFB.876 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.062% | 0.8649% | 6.273% | 0.9782% | 2.55% | 3.04% | 0.243% | 1.654% | 1.059% | 1.048% | 1.322% | 0.6723% | 0.8321% | 1.995% | 3.83% | 3.534% | 0.5722% | 1.093% | 1.638% | 5.788% | 1.744% | 2.648% | 15.71% | 1.473% | 1.345% | 6.16% | 4.211% | 0.6243% | 1.233% | 1.223% | 1.105% | 0.5077% | 0.9751% | 0.09416% | 0.9342% | 1.676% | 0.5583% | 1.419% | 0.5399%
228 | 5_18_23_MFB |  | 6/17/2023 11:20:08 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_MFB.877 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.8798% | 2.466% | 7.42% | 0.7461% | 2.449% | 3.163% | 3.634% | 3.292% | 1.485% | 1.756% | 1.32% | 0.3955% | 1.834% | 0.7289% | 0.6536% | 1.377% | 1.014% | 0.712% | 1.279% | 6.634% | 2.128% | 4.339% | 5.279% | 1.998% | 3.264% | 5.36% | 7.824% | 0.6986% | 0.6683% | 0.6761% | 0.8451% | 0.1682% | 5.476% | 0.5712% | 0.8766% | 1.216% | 1.043% | 0.5449% | 0.08641%
229 | 5_18_23_MFB |  | 6/17/2023 11:26:26 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_MFB.878 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.876% | 4.959% | 4.735% | 1.861% | 1.573% | 2.044% | 1.397% | 1.951% | 0.9685% | 1.258% | 1.476% | 0.7283% | 0.6424% | 1.122% | 3.672% | 0.8093% | 0.4351% | 0.2168% | 3.28% | 5.957% | 2.287% | 3.745% | 3.985% | 5.8% | 7.307% | 4.664% | 7.867% | 0.7064% | 0.6611% | 3.141% | 0.5436% | 0.4914% | 0.5853% | 1.073% | 0.5426% | 0.5166% | 0.7804% | 0.3506% | 0.6752%
230 | LBLANK |  | 6/17/2023 11:32:26 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.879 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0% | 173.2% | 31.49% | 2.91% | 4.178% | 0% | 173.2% | 43.3% | 7.143% | 24.12% | 10.25% | 61.86% | 2.463% | 86.6% | 31.11% | 43.83% | 26.13% | 14.4% | 9.065% |  | 32.83% | 5.766% |  | 86.6% | 34.64% | 37.65% | 173.2% | 9.231% | 24.74% | 2.088% | 0.5003% | 1.362% | 0.3981% | 0.5041% | 0.6416% | 173.2% | 21.65% | 24.74% | 30.73%
231 | LBLANK |  | 6/17/2023 11:38:51 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.880 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 43.3% |  | 15.75% | 4.821% | 5.206% | 8.056% | 173.2% | 173.2% | 14.52% | 0% | 3.097% | 25% | 2.135% | 86.6% | 15.75% | 52.92% | 9.449% | 8.517% | 4.351% |  | 26.65% | 0.9511% |  |  | 114.6% | 62.75% |  | 18.22% | 13.32% | 1.885% | 0.4801% | 0.1343% | 0.5602% | 0.4477% | 0.5455% |  | 33.33% | 16.37% | 18.03%
232 | CCV 1ppb TRACE |  | 6/17/2023 11:44:51 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.881 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.977% | 5.49% | 8.418% | 2.533% | 2.947% | 3.481% | 0.9564% | 1.528% | 7.546% | 2.031% | 3.37% | 0.344% | 1.167% | 0.6848% | 2.478% | 2.045% | 1.054% | 1.834% | 45.13% | 13.78% | 5.183% | 5.03% | 21.43% | 6.025% | 3.26% | 2.453% | 6.605% | 0.7533% | 1.194% | 0.7664% | 0.2276% | 1.049% | 0.4844% | 0.9052% | 0.5254% | 1.014% | 0.385% | 0.5501% | 1.027%
233 | CCV 1ppb TRACE |  | 6/17/2023 11:51:11 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.882 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.6889% | 2.556% | 1.887% | 1.678% | 3.429% | 4.584% | 5.716% | 1.647% | 3.051% | 2.685% | 1.386% | 1.136% | 3.012% | 1.848% | 4.745% | 3.405% | 0.8241% | 3.539% | 33.79% | 9.362% | 1.54% | 0.7895% | 13.57% | 18.06% | 4.886% | 2.201% | 19.47% | 1.408% | 1.531% | 0.2555% | 0.2396% | 1.031% | 0.4663% | 0.7229% | 0.5741% | 3.029% | 0.09527% | 2.059% | 1.14%
234 | CCV 10 ppb MAJORS |  | 6/17/2023 11:57:11 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.883 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.976% | 1.629% | 4.034% | 0.624% | 3.135% | 0.5903% | 0.7477% | 2.561% | 0.6259% | 1.156% | 2.887% | 1.247% | 0.6309% | 0.9467% | 2.344% | 1.682% | 0.4044% | 0.9856% | 2.538% | 4.458% | 1.166% | 2.566% | 6.854% | 4.382% | 2.232% | 3.181% | 7.065% | 0.9185% | 0.6792% | 1.104% | 0.5806% | 0.6938% | 1.186% | 0.5447% | 0.7936% | 2.443% | 0.9632% | 0.8073% | 0.5036%
235 | CAL BLANK |  | 6/17/2023 12:25:05 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CAL BLANK.884 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth

## Net Intensities
| Sample Id | R | Acquisition Time | QC Status | Dataset File | Method File | Na KED 23/23 Helium KED_LowLow (cps) | Mg KED 24/24 Helium KED_LowLow (cps) | Al KED 27/27 Helium KED_LowLow (cps) | P Oshift 31/47 Oxygen DRC (cps) | S (34) -1 34/34 Helium KED_LowLow (cps) | K KED 39/39 Helium KED_LowLow (cps) | Ca KED 43/43 Helium KED_LowLow (cps) | V KED 51/51 Helium KED_LowLow (cps) | Cr (52) 52/52 Helium KED_LowLow (cps) | Cr (53) NH3 53/53 Ammonia DRC (cps) | Fe (54) NH3 54/54 Ammonia DRC (cps) | Mn Oshift 55/71 Oxygen DRC (cps) | Fe (56) KED 56/56 Helium KED_LowLow (cps) | Co KED 59/59 Helium KED_LowLow (cps) | Ni KED 60/60 Helium KED_LowLow (cps) | Cu (63) KED 63/63 Helium KED_LowLow (cps) | Cu (65) NH3 65/65 Ammonia DRC (cps) | Zn NH3 66/66 Ammonia DRC (cps) | As 75/75 (cps) | As KED 75/75 Helium KED_LowLow (cps) | As Oshift 75/91 Oxygen DRC (cps) | Se (77) 77/77 (cps) | Se (77) Oshift 77/93 Oxygen DRC (cps) | Se (78) Oshift 78/94 Oxygen DRC (cps) | Se (80) Oshift 80/96 Oxygen DRC (cps) | Se (82) 82/82 (cps) | Se (82) Oshift 82/98 Oxygen DRC (cps) | Mo (95) NH3 95/95 Ammonia DRC (cps) | Mo (95) OShift 95/127 Oxygen DRC (cps) | Rh (IS1) 103/103 (IS) (cps) | Rh (IS2) 103/103 (IS) Oxygen DRC (cps) | Rh (IS4) 103/103 (IS) Ammonia DRC (cps) | Rh (IS5) 103/103 (IS) Ammonia DRC (cps) | Rh (IS6) 103/103 (IS) Oxygen DRC (cps) | Rh (IS7) 103/103 (IS) Helium KED_LowLow (cps) | Cd (111) NH3 111/111 Ammonia DRC (cps) | Cd (114) NH3 114/114 Ammonia DRC (cps) | Pb 208/208 (cps) | Pb NH3 208/208 Ammonia DRC (cps)
1 | sac blank |  | 06/16/2023 11:58:56 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\sac blank .650 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.00015326 | 0 | 0.0019774 | 0.0034592 | 0.083565 | 0.0013243 | 1.8815e-05 | 1.9289e-05 | 0.0010384 | 8.2823e-05 | 0.0015736 | 5.6696e-05 | 0.2878 | 5.7641e-05 | 0.00074761 | 0.00078633 | 0.00032985 | 0.0054219 | -0.015115 | 3.8702e-05 | 4.4003e-05 | 0.0051147 | 0 | 7.315e-06 | 1.8363e-05 | 0.00022394 | 0 | 0.00066934 | 5.4913e-05 | 271255.1708 | 364338.5337 | 410502.873 | 325134.5776 | 318004.0787 | 34778.2898 | 3.2611e-06 | 5.5112e-05 | 0.00012787 | 0.00019261
2 | sac blank cal |  | 06/16/2023 12:04:57 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\sac blank cal.651 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.00013886 | 1.9615e-05 | 0.0017099 | 0.0034908 | 0.082253 | 0.0015147 | 0 | 9.967e-05 | 0.00085439 | 5.1918e-05 | 0.0015911 | 6.9615e-05 | 0.29362 | 0.00013861 | 0.00083506 | 0.00065739 | 0.00028919 | 0.0044323 | -0.014102 | 5.9873e-05 | 3.8599e-05 | 0.0050015 | 3.6656e-06 | 1.284e-05 | 9.2279e-06 | 0.00018485 | 5.5084e-06 | 0.00063155 | 2.0207e-05 | 273640.0501 | 363100.0755 | 410914.6013 | 323638.2545 | 316582.3718 | 33547.3484 | 1.5994e-06 | 8.907e-05 | 0.00018048 | 0.00034406
3 | sac blank cal |  | 06/16/2023 12:10:57 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\sac blank cal.652 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.00013986 | 0 | 0.00092119 | 0.0033041 | 0.083645 | 0.0016635 | 0 | 0.00011893 | 0.00066062 | 4.6048e-05 | 0.0016176 | 4.9982e-05 | 0.29121 | 0.00018006 | 0.00094321 | 0.00064151 | 0.00031186 | 0.0088872 | -0.014707 | 0 | 2.9607e-05 | 0.0051655 | 0 | 1.8438e-06 | 1.48e-05 | 0.00025887 | 0 | 0.00062372 | 1.479e-05 | 269150.2831 | 360078.3844 | 404036.9377 | 319318.0668 | 309692.239 | 33259.3435 | 6.6368e-06 | 7.4209e-05 | 0.00021481 | 0.00027572
4 | CAL BLANK |  | 06/16/2023 12:16:57 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CAL BLANK.653 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.00014284 | 0 | 0.0015478 | 0.0030216 | 0.088881 | 0.0013652 | 2.013e-05 | 0.00018206 | 0.00083563 | 5.0112e-05 | 0.0016622 | 7.1531e-05 | 0.29844 | 0.00016278 | 0.00081479 | 0.00048943 | 0.00031879 | 0.0036221 | -0.014614 | 2.0712e-05 | 3.9539e-05 | 0.0050956 | 0 | 1.8852e-06 | 1.693e-05 | 0.00023662 | 3.762e-06 | 0.00058002 | 1.1289e-05 | 269907.9681 | 354160.9715 | 399133.1355 | 321765.3432 | 314307.4835 | 32774.2249 | 6.7171e-06 | 7.3371e-05 | 0.00019026 | 0.00029627
5 | LOW 0.01 |  | 06/16/2023 12:22:58 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LOW 0.01.654 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.00097601 | 0.00058268 | 0.0031893 | 0.015847 | -0.003125 | 0.00093561 | 0.00061745 | 0.00067436 | 0.0022045 | 0.0003117 | 0.0065782 | 0.0012664 | 0.0526 | 0.003177 | 0.00050591 | 0.0033893 | 0.0011897 | 0.0028771 | 0.00077511 | 0.00013903 | 0.00027799 | 0.00014187 | 1.8863e-06 | 2.2467e-05 | 1.3101e-05 | 2.0893e-05 | 5.6474e-06 | 0.00069869 | 0.0007083 | 270980.1273 | 354851.3171 | 400108.6326 | 321453.915 | 308628.908 | 33364.9273 | 0.00018655 | 0.0005551 | 0.001943 | 0.0032856
6 | LOW 0.05 |  | 06/16/2023 12:28:58 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LOW 0.05.655 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.0062915 | 0.0027456 | 0.0026568 | 0.017774 | 0.0012711 | 0.0041736 | 0.0017853 | 0.0043245 | 0.0054782 | 0.0011921 | 0.0027365 | 0.0021696 | 0.018411 | 0.010207 | 0.0025096 | 0.0070187 | 0.0026637 | 0.0015237 | 0.0011009 | 0.000349 | 0.001147 | 0.00037706 | 2.2978e-05 | 7.6512e-05 | 0.00018217 | 0.00011641 | 3.8425e-05 | 0.0020797 | 0.0021002 | 268088.4116 | 348270.1943 | 395087.317 | 318408.0119 | 299948.2338 | 32527.6586 | 0.0010253 | 0.0026461 | 0.0056527 | 0.01028
7 | LOW 0.1 |  | 06/16/2023 12:34:59 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LOW 0.1.656 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.011315 | 0.005318 | 3.296e-05 | 0.0066461 | -0.0031621 | 0.0072383 | 0.0042084 | 0.0073088 | 0.0098994 | 0.0021786 | 0.0043935 | 0.004309 | 0.031486 | 0.018757 | 0.0050156 | 0.012306 | 0.0047635 | 0.0022327 | 0.0010272 | 0.00080189 | 0.0020896 | 0.0005956 | 3.4434e-05 | 0.00012872 | 0.00031167 | 0.00016698 | 5.2056e-05 | 0.0037627 | 0.0040577 | 265582.7134 | 347157.0686 | 400478.0263 | 321727.2865 | 302713.9605 | 32481.5538 | 0.002021 | 0.0048288 | 0.010951 | 0.019766
8 | LOW 0.5 |  | 06/16/2023 12:40:59 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LOW 0.5.657 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.058833 | 0.027711 | 0.006558 | 0.022789 | 0.0098971 | 0.035091 | 0.018226 | 0.037728 | 0.050102 | 0.011417 | 0.015853 | 0.021026 | 0.10603 | 0.094145 | 0.024968 | 0.060912 | 0.025039 | 0.013185 | 0.010875 | 0.0040861 | 0.010411 | 0.0014055 | 0.00024747 | 0.0007016 | 0.0015607 | 0.0010219 | 0.00028897 | 0.02106 | 0.020825 | 260297.4376 | 339285.2492 | 385459.7385 | 314549.4697 | 295286.5087 | 32322.5262 | 0.010458 | 0.024781 | 0.055695 | 0.10113
9 | STD 1 |  | 06/16/2023 12:46:58 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 1.658 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.11869 | 0.05312 | 0.0075838 | 0.016916 | 0.030514 | 0.072723 | 0.033982 | 0.075328 | 0.10031 | 0.021955 | 0.025356 | 0.041587 | 0.1717 | 0.18331 | 0.048217 | 0.11249 | 0.045779 | 0.021057 | 0.019565 | 0.0071876 | 0.02034 | 0.0027363 | 0.00043924 | 0.0014385 | 0.0030934 | 0.0018416 | 0.00051476 | 0.040926 | 0.041541 | 256927.841 | 338303.9383 | 382744.0335 | 310102.4133 | 296139.2927 | 32209.6051 | 0.020093 | 0.048151 | 0.10655 | 0.20064
10 | STD 5 |  | 06/16/2023 12:52:58 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 5.659 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.61018 | 0.26913 | 0.029782 | 0.060928 | 0.15812 | 0.36189 | 0.18298 | 0.39091 | 0.50964 | 0.1072 | 0.077711 | 0.21156 | 0.54588 | 0.93455 | 0.24072 | 0.55809 | 0.21919 | 0.10317 | 0.1086 | 0.037138 | 0.10279 | 0.011588 | 0.002322 | 0.0073472 | 0.01503 | 0.010239 | 0.0027905 | 0.20564 | 0.20584 | 253318.9374 | 322069.7942 | 368756.3281 | 298152.7386 | 287394.9296 | 30266.7021 | 0.10162 | 0.24148 | 0.51213 | 0.98303
11 | STD 10 |  | 06/16/2023 12:58:58 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 10.660 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.2174 | 0.54058 | 0.053162 | 0.097289 | 0.30962 | 0.7272 | 0.3699 | 0.79198 | 1.0138 | 0.21491 | 0.15077 | 0.43267 | 1.0321 | 1.8608 | 0.47291 | 1.0912 | 0.43077 | 0.20678 | 0.21938 | 0.071948 | 0.20515 | 0.024771 | 0.0044771 | 0.014455 | 0.030518 | 0.019952 | 0.0054652 | 0.41416 | 0.41835 | 251310.7966 | 316131.6182 | 361552.0644 | 294804.2408 | 281081.1205 | 29857.8363 | 0.20014 | 0.47714 | 0.95655 | 1.9055
12 | STD 25 |  | 06/16/2023 13:04:58 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 25.661 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.2264 | 1.3706 | 0.13324 | 0.23055 | 0.79785 | 1.9125 | 0.93303 | 2.0488 | 2.6012 | 0.57185 | 0.39123 | 1.1195 | 2.5803 | 4.8279 | 1.2138 | 2.7829 | 1.116 | 0.53291 | 0.57722 | 0.19151 | 0.54918 | 0.060263 | 0.012355 | 0.038482 | 0.082784 | 0.052295 | 0.014861 | 1.0987 | 1.1024 | 250484.9366 | 298535.2123 | 336580.588 | 276834.7038 | 292829.6889 | 28656.0456 | 0.51969 | 1.2361 | 2.1293 | 4.7414
13 | STD 50 |  | 06/16/2023 13:10:58 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 50.662 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 6.5472 | 2.6981 | 0.26479 | 0.4855 | 1.5907 | 3.9769 | 1.9325 | 4.236 | 5.3228 | 1.1536 | 0.76323 | 2.2277 | 5.4289 | 9.6863 | 2.387 | 5.4271 | 2.1613 | 1.0394 | 1.1642 | 0.37233 | 1.1003 | 0.11987 | 0.024728 | 0.077895 | 0.16471 | 0.10442 | 0.029742 | 2.2076 | 2.2272 | 237649.7514 | 280575.4282 | 316455.4454 | 261750.9622 | 253532.0934 | 27131.7462 | 1.0079 | 2.38 | 3.3409 | 8.8664
14 | STD 100 |  | 06/16/2023 13:16:57 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 100.663 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 14.073 | 5.7027 | 0.56008 | 0.9938 | 3.2881 | 8.6676 | 4.1216 | 9.2496 | 11.4347 | 2.4814 | 1.5486 | 4.581 | 10.7233 | 20.3691 | 5.0008 | 11.1144 | 4.373 | 2.0791 | 2.4977 | 0.77695 | 2.2166 | 0.25093 | 0.050138 | 0.15973 | 0.33782 | 0.22662 | 0.06148 | 4.6608 | 4.6166 | 210906.948 | 255858.9654 | 277990.8731 | 230754.8096 | 230674.1552 | 24521.027 | 2.0146 | 4.728 | 5.4842 | 19.4866
15 | Calibration Curves |  | 06/16/2023 13:16:57
16 | LBLANK |  | 06/16/2023 13:22:58 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.664 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.00073315 | 0.00023313 | -0.00086582 | 0.0027282 | 0.0088172 | 0.0011228 | 0.00010802 | -1.4038e-05 | 0.0008224 | 0.00011162 | -0.00020267 | 3.6728e-05 | 0.027724 | 0.0004569 | 0.0017149 | 0.00064236 | 9.8536e-05 | 0.010025 | 0.00050372 | -2.0341e-05 | 2.818e-05 | 0.00017261 | 5.8036e-06 | 3.2996e-05 | 5.0805e-07 | 1.5022e-05 | -3.7648e-06 | 0.0084483 | 0.0004629 | 264900.2734 | 344637.1589 | 383179.046 | 314055.9082 | 301031.8543 | 31338.3419 | 8.9842e-06 | 3.2771e-05 | -4.9328e-05 | -9.7558e-06
17 | LBLANK |  | 06/16/2023 13:41:50 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.665 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.00016209 | 3.7449e-05 | -0.00016659 | -0.00037494 | -0.016364 | 0.00075122 | 1.1478e-06 | 1.4042e-06 | 0.00046895 | 8.5442e-05 | -0.00013327 | -1.3499e-05 | -0.0060099 | 1.8665e-05 | 0.0003444 | 7.5035e-05 | 1.5772e-05 | 0.00025614 | -0.00015491 | -2.0341e-05 | 2.0476e-05 | -1.6455e-05 | 4.1217e-06 | 2.2878e-05 | -2.435e-06 | 2.7766e-05 | -3.7648e-06 | 0.0022228 | 0.0002515 | 254503.0387 | 322216.0874 | 364269.8879 | 305105.7248 | 279979.7305 | 32704.8747 | -3.0319e-06 | -3.8905e-06 | -0.00011686 | -0.00012798
18 | ICV PE |  | 06/16/2023 13:47:50 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\ICV PE.666 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.2927 | 0.56794 | 0.05538 | 0.18675 | 0.31383 | 0.75969 | 0.37562 | 0.79687 | 1.0353 | 0.2235 | 0.15748 | 0.43993 | 1.0698 | 1.872 | 0.47619 | 1.1213 | 0.44064 | 0.20133 | 0.21814 | 0.067806 | 0.21079 | 0.016644 | 0.0046077 | 0.014841 | 0.031644 | 0.019503 | 0.0054935 | 0.40062 | 0.40507 | 245191.604 | 300675.269 | 345722.3099 | 284714.6029 | 265426.4691 | 28577.8896 | 0.19562 | 0.46796 | 0.91864 | 1.8296
19 | CCV 1ppb TRACE |  | 06/16/2023 13:53:50 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.667 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.00011338 | 3.7073e-05 | -0.0013078 | -0.0020276 | -0.078357 | -0.00051962 | 3.452e-05 | -0.00010975 | -0.00017655 | 2.237e-05 | -0.00059188 | -4.9075e-05 | -0.17409 | 7.4166e-05 | -0.00022763 | -4.9034e-05 | -6.6772e-05 | -0.0018753 | 0.014914 | -2.215e-06 | -2.0776e-05 | -0.0041437 | 0 | 1.2932e-05 | -7.5686e-06 | -0.00017207 | -3.7648e-06 | 0.0013881 | 1.9409e-06 | 275520.5968 | 357683.5542 | 404966.6031 | 332299.3953 | 314707.8476 | 36459.1362 | -1.6705e-06 | -3.4032e-05 | -0.00013413 | -0.00021227
20 | CCV 1ppb TRACE |  | 06/16/2023 14:00:07 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.668 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.12878 | 0.058786 | 0.010601 | 0.0088833 | 0.039081 | 0.081633 | 0.03613 | 0.08346 | 0.10841 | 0.023169 | 0.021039 | 0.044081 | 0.17401 | 0.19574 | 0.050765 | 0.1214 | 0.046965 | 0.019393 | 0.020535 | 0.0070597 | 0.0207 | 0.0031232 | 0.00042928 | 0.001448 | 0.0032369 | 0.0019104 | 0.00054435 | 0.042938 | 0.041782 | 249367.3099 | 319820.7585 | 366417.1397 | 299374.3473 | 284383.4349 | 30032.8719 | 0.019934 | 0.047304 | 0.10287 | 0.19047
21 | LFB |  | 06/16/2023 14:06:07 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.669 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.3669 | 0.59729 | 0.059735 | 0.11789 | 0.34704 | 0.82297 | 0.39406 | 0.85773 | 1.1164 | 0.22011 | 0.1498 | 0.4309 | 1.1484 | 2.0392 | 0.51491 | 1.2067 | 0.43565 | 0.20904 | 0.22212 | 0.083715 | 0.20993 | 0.024464 | 0.0047196 | 0.014826 | 0.031065 | 0.020721 | 0.0058014 | 0.40829 | 0.41541 | 240822.2342 | 295243.5241 | 340267.0087 | 278862.1119 | 260573.1265 | 26507.9964 | 0.19758 | 0.46476 | 0.90592 | 1.8115
22 | LLOQ 0.01 |  | 06/16/2023 14:12:05 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LLOQ 0.01.670 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.0018074 | 0.00064232 | 0.0056029 | 0.019175 | 0.011403 | 0.0013603 | 0.00037895 | 0.0008134 | 0.0018694 | 0.00038317 | 0.0074402 | 0.0012848 | 0.081759 | 0.0030939 | 0.0010037 | 0.003609 | 0.0013401 | 0.0023927 | 5.4857e-05 | 9.0334e-05 | 0.00023214 | 0.00042432 | 6.3345e-06 | 3.3951e-05 | 2.0953e-05 | 6.9082e-05 | 8.9461e-06 | 0.0020135 | 0.00085421 | 251360.2066 | 316677.6067 | 367645.3383 | 296732.1147 | 276145.5118 | 30093.6642 | 0.00018918 | 0.00050317 | 0.0017117 | 0.0033119
23 | CCV 10 ppb MAJORS |  | 06/16/2023 14:18:05 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.671 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.3013 | 0.56114 | 0.054744 | 0.095291 | 0.32632 | 0.77287 | 0.36742 | 0.80676 | 1.0391 | 0.23051 | 0.15675 | 0.4509 | 1.0608 | 1.9367 | 0.50004 | 1.149 | 0.45378 | 0.21596 | 0.22369 | 0.076494 | 0.2158 | 0.024853 | 0.0046444 | 0.014625 | 0.031797 | 0.020232 | 0.0056733 | 0.4289 | 0.435 | 238342.5087 | 287577.2647 | 332470.9726 | 276625.3969 | 258561.7143 | 27893.2048 | 0.20112 | 0.47929 | 0.92632 | 1.8586
24 | LLOQ 0.1 |  | 06/16/2023 14:24:05 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LLOQ 0.1.672 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.013286 | 0.0055771 | 9.6724e-05 | 0.0025393 | 0.01325 | 0.0084186 | 0.0034203 | 0.0082108 | 0.0098494 | 0.002338 | 0.0045553 | 0.0046302 | 0.067983 | 0.018998 | 0.0051484 | 0.013596 | 0.005146 | 0.0027643 | 0.0018895 | 0.00076662 | 0.0020672 | 0.00068058 | 4.2305e-05 | 0.00015658 | 0.00030654 | 0.00012003 | 8.0708e-05 | 0.0056032 | 0.0042234 | 251609.074 | 315533.9998 | 362521.6662 | 297856.501 | 271411.4699 | 29644.0604 | 0.0019841 | 0.0046129 | 0.0097939 | 0.018673
25 | 9_13_21_F_11_DUP |  | 06/16/2023 14:30:04 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_F_11_DUP.673 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.86523 | 0.47421 | 0.049783 | 0.033846 | 0.46865 | 0.1456 | 1.2544 | 0.060724 | 0.012722 | 0.0026441 | 0.074549 | 0.21764 | 0.57543 | 0.0044091 | 0.011828 | 0.24965 | 0.10111 | 0.075253 | 0.093049 | 0.029942 | 0.088045 | 0.0010495 | 4.5532e-05 | 0.00016796 | 0.00031097 | 0.00086364 | 4.4139e-05 | 0.098069 | 0.097409 | 222221.8074 | 278349.2177 | 318240.2778 | 264312.8247 | 250563.0459 | 27121.0586 | 0.00024723 | 0.00061029 | 0.012656 | 0.023342
26 | LBLANK |  | 06/16/2023 14:47:08 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.675 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.00012206 | 7.2268e-05 | -0.00055674 | -0.00084493 | 0.0062447 | 0.00059412 | 5.1678e-05 | -1.3741e-05 | 0.00092847 | 7.6437e-05 | -0.00015445 | -2.1987e-05 | 0.026392 | 0.00014964 | 2.9562e-05 | 9.1488e-05 | 7.0779e-06 | -0.0014029 | -0.00055953 | 5.1181e-05 | 1.4628e-05 | 0.00011353 | 2.2606e-06 | 7.1753e-06 | -1.2432e-06 | -2.7638e-05 | -1.4888e-06 | 0.00074356 | 0.00010383 | 232895.7495 | 295279.9247 | 337579.0948 | 277993.1585 | 251865.0368 | 27638.0454 | -4.7349e-06 | -6.2821e-06 | -9.8618e-05 | -0.00013078
27 | 0.5 STD |  | 06/16/2023 14:53:09 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\0.5 STD.676 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.066216 | 0.029646 | 0.0064951 | 0.021176 | 0.024754 | 0.041414 | 0.019578 | 0.040407 | 0.054111 | 0.011697 | 0.016898 | 0.022186 | 0.13079 | 0.096627 | 0.025309 | 0.064201 | 0.026036 | 0.013924 | 0.011018 | 0.0036212 | 0.010751 | 0.001352 | 0.00023662 | 0.00078242 | 0.0015947 | 0.00091083 | 0.00031753 | 0.021673 | 0.021324 | 230199.9828 | 290615.7396 | 333024.3242 | 273633.7973 | 252893.5011 | 27255.314 | 0.010164 | 0.024427 | 0.052281 | 0.0988
28 | 0.1 STD |  | 06/16/2023 14:59:12 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\0.1 STD.677 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.013067 | 0.0055802 | -0.00028203 | 0.0073596 | 0.011508 | 0.0080098 | 0.0041299 | 0.0069701 | 0.010732 | 0.0022799 | 0.0046015 | 0.0044409 | 0.051593 | 0.019654 | 0.0050768 | 0.012266 | 0.0049644 | 0.0018083 | 0.0026271 | 0.0010054 | 0.0021719 | 0.00015881 | 6.1887e-05 | 0.00015628 | 0.00032215 | 0.00014384 | 3.2879e-05 | 0.0045485 | 0.004288 | 232108.9808 | 290921.0448 | 337273.0849 | 275960.6031 | 251580.6985 | 27957.9979 | 0.0021282 | 0.0049482 | 0.010452 | 0.018764
29 | CCV 10 ppb MAJORS |  | 06/16/2023 15:05:12 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.678 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.2851 | 0.56051 | 0.050102 | 0.094247 | 0.32646 | 0.77938 | 0.3763 | 0.80963 | 1.033 | 0.22935 | 0.15702 | 0.44746 | 1.06 | 1.9046 | 0.49099 | 1.134 | 0.45343 | 0.21435 | 0.22453 | 0.07688 | 0.21598 | 0.024 | 0.0048444 | 0.014932 | 0.032117 | 0.020484 | 0.0057079 | 0.42768 | 0.43786 | 226803.0185 | 265194.2076 | 312687.2907 | 258218.9534 | 241226.3203 | 26232.0635 | 0.20215 | 0.48255 | 0.9284 | 1.8784
30 | 9_13_21_F_11_DUP |  | 06/16/2023 15:11:13 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_F_11_DUP.679 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.85387 | 0.48519 | 0.051966 | 0.03266 | 0.46637 | 0.14378 | 1.2277 | 0.057199 | 0.012726 | 0.0026223 | 0.074798 | 0.21579 | 0.56981 | 0.0043317 | 0.012068 | 0.24544 | 0.1006 | 0.074336 | 0.09012 | 0.028005 | 0.086969 | 0.00083126 | 4.473e-05 | 0.00013738 | 0.0002586 | 0.00074316 | 2.8442e-05 | 0.09732 | 0.097504 | 218252.5864 | 268588.916 | 306604.6407 | 253331.8242 | 241789.1264 | 26547.3105 | 0.00020647 | 0.00053747 | 0.012192 | 0.023315
31 | 9_13_21_F_11_DUP_LSP |  | 06/16/2023 15:17:13 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_F_11_DUP_LSP.680 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.109 | 0.99505 | 0.09933 | 0.13703 | 0.75132 | 0.92757 | 1.5373 | 0.8658 | 1.038 | 0.23088 | 0.2285 | 0.63772 | 1.5705 | 1.8923 | 0.48406 | 1.3208 | 0.52269 | 0.2724 | 0.30924 | 0.10308 | 0.28887 | 0.023964 | 0.0045217 | 0.014475 | 0.029852 | 0.021298 | 0.0053545 | 0.51282 | 0.51764 | 215057.5576 | 259936.8998 | 295256.0017 | 244398.65 | 230444.1962 | 25108.7185 | 0.19944 | 0.47347 | 0.92962 | 1.8516
32 | 9_13_21_UF_1 |  | 06/16/2023 15:23:13 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_UF_1.681 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.1046 | 0.33707 | 0.0099062 | 0.081951 | 0.67084 | 0.16043 | 1.2778 | 0.025241 | 0.0081623 | 0.0013443 | 0.014306 | 0.5669 | 0.19439 | 0.0061841 | 0.02432 | 0.067354 | 0.028726 | 0.026762 | 0.087503 | 0.027526 | 0.082872 | 0.0011275 | 4.737e-05 | 5.2969e-05 | 0.00014523 | 0.00072975 | 3.8665e-05 | 0.65039 | 0.66558 | 217805.1482 | 267313.8922 | 306000.5528 | 256214.671 | 241074.402 | 25727.1456 | 0.0001393 | 0.00032514 | 0.0050656 | 0.0094331
33 | 9_13_21_UF_6 |  | 06/16/2023 15:29:13 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_UF_6.682 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.1716 | 0.39144 | 0.0041546 | 0.030788 | 0.4882 | 0.1444 | 1.2796 | 0.051401 | 0.0093994 | 0.0014449 | 0.0102 | 0.51928 | 0.17465 | 0.0044271 | 0.012575 | 0.20239 | 0.082453 | 0.047565 | 0.089253 | 0.029021 | 0.08659 | 0.0012678 | 5.0054e-05 | 0.00013848 | 0.00026093 | 0.00095205 | 4.884e-05 | 0.25193 | 0.25497 | 217645.5528 | 266102.2501 | 299295.5547 | 255242.904 | 241211.4325 | 25851.3699 | 0.00030064 | 0.0007281 | 0.0067744 | 0.012766
34 | 9_13_21_UF_6_DUP |  | 06/16/2023 15:35:13 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_UF_6_DUP.683 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.1632 | 0.38887 | 0.0040815 | 0.030818 | 0.49115 | 0.1414 | 1.2589 | 0.052291 | 0.011592 | 0.0021144 | 0.012553 | 0.51385 | 0.17491 | 0.0042469 | 0.012908 | 0.20318 | 0.082109 | 0.039028 | 0.092294 | 0.029308 | 0.087864 | 0.0012241 | 3.329e-05 | 9.5323e-05 | 0.00025175 | 0.00092316 | 4.2288e-05 | 0.24866 | 0.25437 | 211424.5194 | 260531.6015 | 299978.8753 | 245936.9483 | 232790.1873 | 25386.5418 | 0.00032202 | 0.00064451 | 0.0071609 | 0.012884
35 | 9_13_21_UF_8 |  | 06/16/2023 15:41:13 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_UF_8.684 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.2096 | 0.38156 | 0.0038961 | 0.035383 | 0.44283 | 0.15015 | 1.2294 | 0.067164 | 0.0085282 | 0.0013522 | 0.0088187 | 0.45874 | 0.14559 | 0.0035182 | 0.012635 | 0.1799 | 0.071145 | 0.022677 | 0.10137 | 0.02981 | 0.095233 | 0.00083919 | 3.3495e-05 | 0.00011651 | 0.00024821 | 0.00072054 | 4.0024e-05 | 0.20822 | 0.21384 | 210079.0281 | 259007.5555 | 299086.0065 | 247261.6142 | 232864.9443 | 24993.1778 | 0.00021188 | 0.00041405 | 0.0058797 | 0.010657
36 | 9_13_21_UF_9 |  | 06/16/2023 15:47:13 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_UF_9.685 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.1139 | 0.39062 | 0.0053228 | 0.033237 | 0.41022 | 0.14211 | 1.1224 | 0.058521 | 0.011067 | 0.0017988 | 0.011302 | 0.16609 | 0.16889 | 0.0031583 | 0.010126 | 0.16689 | 0.065374 | 0.017568 | 0.083821 | 0.027098 | 0.081705 | 0.0013191 | 3.873e-05 | 0.00010377 | 0.00021731 | 0.00077346 | 5.3034e-05 | 0.15797 | 0.16304 | 212480.5758 | 258893.0249 | 297856.341 | 244576.0713 | 231758.2798 | 25092.6904 | 0.00012568 | 0.00024155 | 0.0059251 | 0.010555
37 | LBLANK |  | 06/16/2023 15:53:13 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.686 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | -4.3144e-05 | 0 | -0.0010735 | -0.0008858 | 0.0060172 | 0.00020417 | 2.9563e-05 | -0.00010855 | 6.2385e-05 | 6.065e-05 | -0.00010407 | 1.2266e-05 | 0.013104 | -0.00011283 | 3.1287e-05 | 0.00025874 | -3.4127e-05 | -0.0021298 | 0.00099081 | 0.00032684 | 0.00088244 | 0.00030853 | 0 | 1.0083e-05 | -1.6005e-07 | -1.6648e-06 | -3.7648e-06 | -0.0002517 | 7.02e-05 | 226720.5403 | 278322.546 | 325386.1147 | 268735.6278 | 243668.5438 | 26821.1566 | -4.6467e-06 | -4.8921e-05 | -6.9162e-05 | -0.00014774
38 | CCV 1ppb TRACE |  | 06/16/2023 15:59:13 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.687 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.1468 | 0.063866 | 0.010376 | 0.0081463 | 0.053548 | 0.085028 | 0.041115 | 0.088819 | 0.11348 | 0.023079 | 0.021005 | 0.043579 | 0.19646 | 0.20685 | 0.055062 | 0.1273 | 0.047274 | 0.019748 | 0.020635 | 0.0083805 | 0.021224 | 0.002876 | 0.00045809 | 0.0015106 | 0.0030009 | 0.0020343 | 0.00055721 | 0.041094 | 0.042441 | 224402.0785 | 278187.3128 | 318651.5832 | 262640.4992 | 245157.8257 | 25306.4862 | 0.020515 | 0.047627 | 0.10493 | 0.19584
39 | CCV 1ppb TRACE |  | 06/16/2023 16:05:32 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.688 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.12962 | 0.058602 | 0.010307 | 0.0084115 | 0.044092 | 0.081016 | 0.038057 | 0.083859 | 0.10633 | 0.023751 | 0.022141 | 0.04447 | 0.18149 | 0.19411 | 0.05353 | 0.11768 | 0.047334 | 0.019617 | 0.019716 | 0.0081224 | 0.021466 | 0.003049 | 0.00043639 | 0.0014797 | 0.0030612 | 0.0019084 | 0.00061706 | 0.041666 | 0.042799 | 224690.3416 | 278094.7005 | 322519.6244 | 265604.4616 | 245624.4224 | 26899.9785 | 0.020363 | 0.047814 | 0.10632 | 0.19405
40 | CCV 10 ppb MAJORS |  | 06/16/2023 16:11:33 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.689 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.2804 | 0.57399 | 0.05476 | 0.092292 | 0.33678 | 0.78786 | 0.37659 | 0.8087 | 1.0299 | 0.22756 | 0.16008 | 0.44087 | 1.0676 | 1.9082 | 0.48282 | 1.1278 | 0.44517 | 0.21429 | 0.21787 | 0.076705 | 0.21067 | 0.024244 | 0.0046429 | 0.014876 | 0.031482 | 0.019954 | 0.0056124 | 0.41684 | 0.42798 | 223053.9793 | 261302.1483 | 306860.724 | 252052.8155 | 233908.1594 | 25695.0913 | 0.20096 | 0.48129 | 0.9569 | 1.9005
41 | CCV 10 ppb MAJORS |  | 06/16/2023 16:17:52 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.690 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.3025 | 0.58586 | 0.053293 | 0.092932 | 0.33822 | 0.78763 | 0.3754 | 0.81383 | 1.0469 | 0.2305 | 0.15774 | 0.44682 | 1.0807 | 1.9185 | 0.48253 | 1.1415 | 0.45424 | 0.21795 | 0.2226 | 0.074413 | 0.2133 | 0.023986 | 0.0046243 | 0.014712 | 0.030884 | 0.020364 | 0.0055894 | 0.423 | 0.43034 | 224935.1818 | 263575.3844 | 308180.7727 | 254476.1549 | 238497.7411 | 25829.9986 | 0.19934 | 0.47775 | 0.94246 | 1.8841
42 | 9_13_21_UF_10 |  | 06/16/2023 16:23:53 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_UF_10.691 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.0257 | 0.48149 | 0.0066934 | 0.042338 | 0.47318 | 0.15211 | 1.278 | 0.061668 | 0.0086208 | 0.001326 | 0.006845 | 0.18958 | 0.15904 | 0.002923 | 0.0095943 | 0.13282 | 0.054048 | 0.020036 | 0.093607 | 0.029646 | 0.088737 | 0.00069148 | 3.0707e-05 | 0.00010821 | 0.00024618 | 0.00090592 | 3.9798e-05 | 0.10336 | 0.1049 | 212735.826 | 260791.7426 | 295509.905 | 249160.768 | 234713.7987 | 25505.418 | 0.00018266 | 0.00028081 | 0.0062173 | 0.011703
43 | 9_13_21_UF_11 |  | 06/16/2023 16:29:52 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_UF_11.692 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.0262 | 0.49452 | 0.0045105 | 0.041714 | 0.46539 | 0.15289 | 1.2911 | 0.060968 | 0.017637 | 0.0032974 | 0.020542 | 0.20984 | 0.23057 | 0.0026787 | 0.0093888 | 0.16001 | 0.063824 | 0.026126 | 0.092702 | 0.027773 | 0.087823 | 0.00082981 | 3.1434e-05 | 0.00013749 | 0.00025093 | 0.00080476 | 5.138e-05 | 0.096557 | 0.098727 | 208468.1962 | 254132.0577 | 294973.6357 | 240084.9302 | 228577.5079 | 25088.678 | 0.00014698 | 0.0004099 | 0.0073338 | 0.013416
44 | 9_13_21_LAB_A |  | 06/16/2023 16:35:53 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_LAB_A.693 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.001168 | 0 | 0.0033071 | 0.0015102 | 0.0075756 | 0.00024061 | 0.00010138 | -0.00011047 | 0.0022021 | 0.00040481 | 0.0032771 | 0.0016633 | 0.028535 | 8.0477e-05 | 0.0013706 | 0.066579 | 0.026585 | 0.0049861 | 0.00089803 | 0.00014978 | 0.00056108 | 0.00021588 | 7.2633e-06 | 5.4328e-06 | 2.4352e-06 | 7.2508e-05 | -3.7648e-06 | -0.0002948 | 2.2642e-05 | 226541.1767 | 275132.9634 | 320898.8992 | 265179.6474 | 239875.8733 | 27457.696 | -4.6187e-06 | -2.1548e-05 | 0.0027445 | 0.0052259
45 | 9_13_21_BB |  | 06/16/2023 16:41:52 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_BB.694 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.00023798 | 0 | 0.0044371 | 0.001441 | 0.0057982 | 0.00048323 | 0.00023257 | -7.1001e-06 | 0.0018922 | 0.00044324 | 0.0030046 | 0.00011627 | 0.032752 | -6.2493e-05 | 0.00034793 | 0.014998 | 0.0058449 | 0.0015806 | 6.0648e-05 | 0.00010503 | 0.00034577 | 4.5494e-05 | 0 | 2.9824e-06 | 2.5434e-06 | -6.0945e-05 | -1.314e-06 | -0.0003421 | 7.3985e-05 | 227480.8662 | 273392.5949 | 320221.2874 | 263843.1728 | 237195.6658 | 26411.0608 | -4.088e-07 | 3.4852e-06 | 0.0028666 | 0.0054092
46 | 9_13_21_FB |  | 06/16/2023 16:47:52 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_FB.695 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.02351 | 0.0023741 | 0.079385 | 0.013226 | 0.018846 | 0.0054608 | 0.0087703 | 0.0011139 | 0.0085915 | 0.0018706 | 0.019737 | 0.013211 | 0.17525 | 0.0021543 | 0.0023756 | 0.025321 | 0.011688 | 0.38443 | -0.00082098 | 0.00048625 | 0.0011104 | 0.00053292 | 4.9594e-06 | 1.0523e-05 | 4.2629e-05 | 1.2172e-05 | 1.1658e-06 | 0.0039313 | 0.0044124 | 219852.1767 | 269126.1513 | 315132.5781 | 259127.0575 | 237578.0783 | 24873.6706 | 0.00018168 | 0.00040039 | 0.0097896 | 0.018469
47 | 5_18_23_MFB |  | 06/16/2023 16:53:52 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_MFB.696 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.61798 | 0.27753 | 0.028119 | 0.089639 | 0.16419 | 0.3726 | 0.18421 | 0.39242 | 0.51902 | 0.11475 | 0.076683 | 0.2184 | 0.53412 | 0.93461 | 0.24321 | 0.57154 | 0.22699 | 0.10196 | 0.097745 | 0.037448 | 0.10289 | 0.010626 | 0.0023849 | 0.007166 | 0.015443 | 0.009092 | 0.0028565 | 0.20252 | 0.20664 | 238030.4131 | 262239.7433 | 308394.5686 | 255513.517 | 232536.0075 | 25712.4526 | 0.098762 | 0.23584 | 0.46193 | 0.9682
48 | 5_18_23_DB |  | 06/16/2023 16:59:52 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_DB.697 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.0018675 | 0.00067095 | 0.0025238 | 0.047612 | 0.013937 | 0.0012663 | 0.00074799 | 0.00098312 | 0.0022406 | 0.00049058 | 0.0037527 | 0.00071751 | 0.051981 | 0.0026664 | 0.00055012 | 0.0096103 | 0.0037026 | 0.00019753 | -0.00089069 | 0.00015398 | 0.00038814 | 0.00051999 | 2.1673e-05 | 2.247e-05 | 5.8325e-05 | 1.5879e-05 | 8.2717e-06 | 0.00014918 | 0.00055455 | 230308.4754 | 274486.113 | 326597.392 | 265475.7653 | 240996.5053 | 26869.9135 | 0.00028321 | 0.00061054 | 0.0033467 | 0.0065449
49 | 5_18_23_T_239 |  | 06/16/2023 17:05:52 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_T_239.698 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 5.9911 | 0.71336 | 0.60842 | 0.12508 | 0.035967 | 0.30292 | 1.8955 | 0.15264 | 0.14418 | 0.032615 | 0.64932 | 0.1591 | 4.3335 | 0.26548 | 0.083334 | 0.087576 | 0.037941 | 0.12723 | 0.055244 | 0.016734 | 0.051616 | 0.012986 | 0.00096431 | 0.0028822 | 0.0059551 | 0.0045793 | 0.0010923 | 0.045859 | 0.046835 | 194973.0541 | 232993.3409 | 278031.1657 | 232347.7448 | 205432.5785 | 23385.799 | 0.068965 | 0.16145 | 0.13991 | 0.24366
50 | 5_18_23_T_239 |  | 06/16/2023 17:11:52 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_T_239.699 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 6.2059 | 0.73754 | 0.63533 | 0.13097 | 0.042007 | 0.31219 | 2.0154 | 0.16018 | 0.15878 | 0.032017 | 0.63907 | 0.15617 | 4.5697 | 0.27959 | 0.088172 | 0.096696 | 0.03885 | 0.12635 | 0.053668 | 0.018182 | 0.050018 | 0.011824 | 0.00088277 | 0.0026319 | 0.005944 | 0.0041882 | 0.0010756 | 0.04674 | 0.046485 | 204143.7208 | 231693.2066 | 276709.5398 | 231500.0199 | 205744.9947 | 22539.1292 | 0.068581 | 0.15892 | 0.12869 | 0.2381
51 | 5_18_23_M_224 |  | 06/16/2023 17:17:52 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_M_224.700 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.51663 | 0.035098 | 0.042486 | 0.17095 | 0.027759 | 0.036044 | 0.11122 | 0.0813 | 0.0042825 | 0.00079407 | 0.003776 | 0.2377 | 0.045218 | 0.0025128 | 0.0025258 | 0.008613 | 0.0031991 | 0.016704 | 0.00408 | 0.00036692 | 0.0013204 | 0.00090088 | 2.9037e-05 | 2.7108e-05 | 7.2816e-05 | 0.0015173 | 2.793e-05 | 0.010051 | 0.0087333 | 219624.0834 | 252525.7153 | 306452.7819 | 254962.3906 | 226682.3356 | 25922.1656 | 4.7693e-05 | 0.00011143 | 0.0018324 | 0.0030971
52 | LBLANK |  | 06/16/2023 17:23:52 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.701 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.0001369 | 2.5812e-05 | -0.0013419 | 0.0056996 | 0.0035044 | 0.00062104 | 5.5452e-05 | -3.035e-05 | 0.00033507 | 4.23e-05 | -0.00025365 | -3.9048e-05 | 0.015618 | -3.5537e-05 | 0.00025421 | 0.00037691 | -5.4323e-06 | -0.0020888 | -0.0016804 | 5.4713e-06 | 6.5741e-05 | 0.00086648 | 0 | 3.1072e-06 | -4.4782e-06 | -2.1004e-05 | -3.7648e-06 | 0.00063337 | 2.8908e-05 | 201553.7404 | 265923.1425 | 330118.1227 | 264892.8544 | 235002.586 | 26214.0318 | -6.6811e-06 | -2.9181e-05 | -8.3402e-05 | -0.0001353
53 | CCV 1ppb TRACE |  | 06/16/2023 17:29:53 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.702 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.12059 | 0.055187 | 0.010426 | 0.0096847 | 0.033441 | 0.077153 | 0.035433 | 0.081133 | 0.10473 | 0.023497 | 0.021732 | 0.043676 | 0.16049 | 0.18791 | 0.047026 | 0.12056 | 0.046872 | 0.019627 | 0.02064 | 0.0073621 | 0.020576 | 0.0027969 | 0.00055333 | 0.0015285 | 0.0030554 | 0.0019689 | 0.00053442 | 0.04202 | 0.042707 | 223263.1519 | 266509.0913 | 315013.4182 | 257338.7071 | 230993.4473 | 26008.9883 | 0.019604 | 0.047703 | 0.10371 | 0.19413
54 | CCV 1ppb TRACE |  | 06/16/2023 17:36:12 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.703 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.12341 | 0.056393 | 0.0095952 | 0.0087633 | 0.037428 | 0.076304 | 0.036401 | 0.079502 | 0.10232 | 0.02339 | 0.021029 | 0.043633 | 0.17814 | 0.18752 | 0.050586 | 0.1207 | 0.046727 | 0.019555 | 0.019125 | 0.0075546 | 0.020967 | 0.0030391 | 0.00041631 | 0.0013843 | 0.0031321 | 0.0020187 | 0.00053933 | 0.04181 | 0.04298 | 224805.8413 | 267475.5532 | 315418.1545 | 262800.0217 | 234862.9989 | 26553.3296 | 0.019828 | 0.047821 | 0.10598 | 0.19425
55 | LFB |  | 06/16/2023 17:42:12 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.704 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.1973 | 0.52909 | 0.054684 | 0.11701 | 0.32497 | 0.73216 | 0.36473 | 0.76616 | 0.98783 | 0.21661 | 0.15182 | 0.4246 | 1.0144 | 1.8222 | 0.46425 | 1.0984 | 0.43165 | 0.20144 | 0.21348 | 0.074968 | 0.20734 | 0.024536 | 0.0048766 | 0.014428 | 0.031163 | 0.020051 | 0.0054368 | 0.4024 | 0.4157 | 220002.5068 | 251315.3867 | 302674.9263 | 248180.5763 | 224792.9774 | 25123.407 | 0.19405 | 0.46058 | 0.93438 | 1.8491
56 | CCV 10 ppb MAJORS |  | 06/16/2023 17:48:11 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.705 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.2392 | 0.5553 | 0.050939 | 0.093523 | 0.32584 | 0.76974 | 0.37135 | 0.78857 | 1.0124 | 0.23065 | 0.1602 | 0.44354 | 1.0542 | 1.8942 | 0.4737 | 1.1179 | 0.4509 | 0.21628 | 0.21372 | 0.073832 | 0.20983 | 0.02442 | 0.0046467 | 0.014506 | 0.031569 | 0.019317 | 0.0054027 | 0.42257 | 0.43296 | 223989.8079 | 253676.5238 | 304676.7752 | 249380.5831 | 227631.292 | 25757.8687 | 0.19982 | 0.47715 | 0.94695 | 1.9044
57 | CCV 10 ppb MAJORS |  | 06/16/2023 17:54:30 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.706 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.265 | 0.56282 | 0.054004 | 0.094035 | 0.34273 | 0.76878 | 0.37904 | 0.81415 | 1.0291 | 0.2329 | 0.15919 | 0.44175 | 1.0782 | 1.9362 | 0.49367 | 1.1567 | 0.4534 | 0.21509 | 0.22197 | 0.0744 | 0.20749 | 0.024807 | 0.0044946 | 0.014724 | 0.030913 | 0.020422 | 0.0057345 | 0.42439 | 0.43135 | 225456.5945 | 259229.737 | 306434.3765 | 250498.6629 | 231380.4208 | 25567.5293 | 0.19815 | 0.47598 | 0.95145 | 1.8894
58 | 5_18_23_MFB |  | 06/16/2023 18:00:31 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_MFB.707 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.63612 | 0.2811 | 0.029954 | 0.089283 | 0.17285 | 0.37864 | 0.19056 | 0.40277 | 0.52015 | 0.12421 | 0.08699 | 0.22189 | 0.56333 | 0.97765 | 0.24932 | 0.58604 | 0.24519 | 0.10979 | 0.10565 | 0.038658 | 0.10455 | 0.01186 | 0.0022211 | 0.0071899 | 0.015049 | 0.010053 | 0.0027695 | 0.22021 | 0.20937 | 227168.7348 | 263210.706 | 299416.7669 | 259611.1886 | 234635.4219 | 26291.5097 | 0.1056 | 0.25259 | 0.49077 | 0.93348
59 | 10_1_21_W_1 |  | 06/16/2023 18:06:31 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_W_1.708 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.22682 | 0.074611 | 0.030504 | 0.038095 | 0.1745 | 0.034764 | 0.29545 | 0.0041277 | 0.0021497 | 0.00046469 | 0.12025 | 0.2496 | 0.84484 | 0.0015719 | 0.0039125 | 0.053253 | 0.020885 | 0.0057595 | 0.014453 | 0.0050638 | 0.014354 | 0.00063978 | 1.2772e-05 | 3.1185e-05 | 1.3535e-05 | 0.00016188 | 8.908e-06 | 0.14839 | 0.15002 | 222341.6447 | 262292.9322 | 306837.6297 | 256352.9768 | 231197.2607 | 26107.8416 | 3.6718e-05 | 7.6628e-05 | 0.0031769 | 0.0057426
60 | 10_1_21_W_6 |  | 06/16/2023 18:12:31 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_W_6.709 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.22935 | 0.087743 | 0.050793 | 0.024634 | 0.13471 | 0.031116 | 0.28888 | 0.011371 | 0.0025042 | 0.00047644 | 0.14748 | 0.14197 | 1.0359 | 0.001481 | 0.0036799 | 0.095944 | 0.037307 | 0.015767 | 0.015776 | 0.004962 | 0.016611 | 0.00045207 | 1.8209e-05 | 1.8974e-05 | 6.3959e-05 | 0.00012009 | 6.6095e-06 | 0.066682 | 0.070646 | 220208.9965 | 255835.2487 | 310648.0455 | 230935.2044 | 227870.6403 | 25965.5768 | 0.00010067 | 0.00015019 | 0.0076254 | 0.015896
61 | 10_1_21_W_8 |  | 06/16/2023 18:18:32 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_W_8.710 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.23349 | 0.082946 | 0.021715 | 0.026446 | 0.12153 | 0.03132 | 0.27897 | 0.01266 | 0.0031378 | 0.0005385 | 0.08424 | 0.10558 | 0.59469 | 0.00066199 | 0.0022518 | 0.093115 | 0.036661 | 0.013019 | 0.016975 | 0.0052764 | 0.018113 | 0.00059656 | 1.3164e-05 | 3.7627e-05 | 4.1084e-05 | 0.00012879 | 4.1606e-06 | 0.06091 | 0.061372 | 217964.5017 | 252784.2394 | 297880.2667 | 251690.5409 | 224484.0621 | 25045.2699 | 4.4789e-05 | 7.6439e-05 | 0.0035911 | 0.0070029
62 | 10_1_21_W_9 |  | 06/16/2023 18:24:32 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_W_9.711 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.19563 | 0.075392 | 0.031009 | 0.023796 | 0.097713 | 0.028237 | 0.24232 | 0.010545 | 0.0054892 | 0.0010956 | 0.087457 | 0.077672 | 0.6146 | 0.00077391 | 0.001633 | 0.064057 | 0.025053 | 0.006285 | 0.013365 | 0.0050787 | 0.013994 | 0.00060284 | 5.2206e-06 | 1.9009e-05 | 3.7916e-05 | 0.00011107 | 9.2835e-06 | 0.041626 | 0.042519 | 220411.3985 | 255191.7939 | 304229.5321 | 251776.7868 | 228135.3368 | 25629.6391 | 7.1996e-05 | 0.00010178 | 0.0034571 | 0.0068382
63 | 10_1_21_W_10 |  | 06/16/2023 18:30:32 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_W_10.712 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.17252 | 0.09347 | 0.1612 | 0.040372 | 0.10666 | 0.028324 | 0.2623 | 0.011446 | 0.0038756 | 0.00079811 | 0.22982 | 0.16352 | 1.5347 | 0.0018625 | 0.0025293 | 0.070368 | 0.028228 | 0.016199 | 0.015425 | 0.0047159 | 0.016034 | 0.00080079 | 5.2909e-06 | 4.3318e-05 | 3.0808e-05 | 0.00017722 | 4.2685e-06 | 0.024525 | 0.025892 | 221263.8085 | 250960.0586 | 301984.7588 | 252930.7804 | 223834.7418 | 25336.4487 | 6.6355e-05 | 0.00018912 | 0.010885 | 0.020185
64 | 10_1_21_W_11 |  | 06/16/2023 18:36:31 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_W_11.713 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.16539 | 0.090279 | 0.11397 | 0.037355 | 0.10558 | 0.028302 | 0.25406 | 0.013218 | 0.0058773 | 0.00098131 | 0.26808 | 0.16777 | 1.7724 | 0.0019593 | 0.0027331 | 0.089233 | 0.036045 | 0.019751 | 0.016278 | 0.0054215 | 0.016707 | 0.00076529 | 1.0434e-05 | 3.4798e-05 | 3.0195e-05 | 7.3311e-06 | 1.4529e-05 | 0.023124 | 0.023872 | 221139.1867 | 254335.5104 | 302494.4431 | 252005.3675 | 226552.985 | 25152.7947 | 7.9357e-05 | 0.00020852 | 0.021915 | 0.040766
65 | 10_1_21_W_11_LD |  | 06/16/2023 18:42:31 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_W_11_LD.714 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | -4.3836e-05 | 0 | -0.001468 | -0.0013677 | -0.077666 | -0.00080373 | -1.1644e-06 | -0.00016348 | -0.00039583 | 1.5662e-05 | -0.00063844 | -3.8169e-05 | -0.19096 | -0.00011988 | -0.00039043 | -8.7809e-05 | -0.00013 | -0.002254 | 0.015351 | 2.2512e-05 | -1.9011e-05 | -0.0040183 | 0 | -1.8824e-06 | -1.0699e-05 | -0.0001172 | -3.7648e-06 | -0.00016022 | 1.1984e-06 | 259144.9361 | 320374.5675 | 365527.382 | 309786.5814 | 275561.4134 | 33300.1863 | -4.8894e-06 | -6.084e-05 | -0.00015916 | -0.00021416
66 | 10_1_21_W_11_LD |  | 06/16/2023 18:48:48 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_W_11_LD.715 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.166 | 0.092177 | 0.11905 | 0.031256 | 0.10218 | 0.028654 | 0.25961 | 0.012501 | 0.0058893 | 0.0012554 | 0.27304 | 0.17493 | 1.7749 | 0.0023845 | 0.0024685 | 0.08678 | 0.036151 | 0.020227 | 0.016829 | 0.0065971 | 0.017953 | 0.0007362 | 1.3394e-05 | 3.0259e-05 | 2.058e-05 | 6.9169e-05 | 1.594e-06 | 0.024062 | 0.024934 | 213700.7956 | 248734.0528 | 294694.6534 | 243946.5561 | 218436.1539 | 25387.2046 | 0.00010396 | 0.00023442 | 0.022557 | 0.041674
67 | 10_1_21_W_11_DUP |  | 06/16/2023 18:54:48 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_W_11_DUP.716 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.1728 | 0.092707 | 0.12422 | 0.030677 | 0.10332 | 0.026881 | 0.25667 | 0.012169 | 0.0037611 | 0.00068624 | 0.28176 | 0.18245 | 1.8782 | 0.0017401 | 0.00223 | 0.091061 | 0.036419 | 0.020392 | 0.016376 | 0.0055994 | 0.017142 | 0.00037556 | 1.3706e-05 | 4.465e-05 | 6.5226e-05 | 0.00023964 | 7.2036e-06 | 0.023048 | 0.024359 | 215824.816 | 243311.1972 | 294246.7851 | 245492.7531 | 217355.1481 | 24546.4058 | 8.3953e-05 | 0.00027541 | 0.013942 | 0.025243
68 | LBLANK |  | 06/16/2023 19:00:48 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.717 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | -8.7568e-05 | 2.7789e-05 | -0.0012201 | 0.0018786 | 0.01041 | 5.8008e-05 | -2.0341e-05 | -1.9723e-05 | 0.00047385 | 6.3584e-05 | -0.00032822 | -2.1433e-05 | 0.027961 | -0.00013582 | 0.00019406 | 0.00088417 | -4.848e-06 | -0.0021486 | 0.00053689 | -2.0341e-05 | -5.0208e-06 | 0.00029658 | 0 | 7.8545e-07 | -1.4274e-05 | 2.2366e-05 | 6.8505e-06 | -0.00032032 | 4.412e-05 | 220811.6288 | 252707.7376 | 303732.6637 | 248627.6351 | 221990.164 | 24474.952 | -4.483e-06 | -2.742e-05 | -0.00010576 | -0.00015414
69 | CCV 1ppb TRACE |  | 06/16/2023 19:06:48 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.718 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.12251 | 0.054844 | 0.011195 | 0.0090196 | 0.041228 | 0.074048 | 0.036043 | 0.081872 | 0.10587 | 0.0241 | 0.022091 | 0.042921 | 0.16402 | 0.19179 | 0.048732 | 0.11627 | 0.047117 | 0.0199 | 0.019353 | 0.0079995 | 0.02055 | 0.0027027 | 0.00047545 | 0.0014341 | 0.0030682 | 0.0018949 | 0.00061144 | 0.041989 | 0.042971 | 213475.1851 | 251145.0825 | 297457.6162 | 245559.1745 | 217306.6575 | 24708.0163 | 0.020072 | 0.047923 | 0.10853 | 0.19834
70 | CCV 1ppb TRACE |  | 06/16/2023 19:13:07 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.719 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.12397 | 0.058677 | 0.010369 | 0.0086191 | 0.043748 | 0.07543 | 0.036888 | 0.079566 | 0.10669 | 0.023857 | 0.022426 | 0.043747 | 0.1641 | 0.18826 | 0.048088 | 0.12047 | 0.047022 | 0.019852 | 0.020483 | 0.0070668 | 0.020888 | 0.0027113 | 0.00049088 | 0.0014118 | 0.00306 | 0.0020237 | 0.00058524 | 0.042113 | 0.042635 | 216313.3968 | 251321.4315 | 299332.4744 | 247438.8538 | 222111.5086 | 25124.0735 | 0.019883 | 0.04905 | 0.10927 | 0.20173
71 | CCV 10 ppb MAJORS |  | 06/16/2023 19:19:08 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.720 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.2382 | 0.55437 | 0.053098 | 0.09251 | 0.33716 | 0.7551 | 0.37539 | 0.78648 | 1.0042 | 0.2299 | 0.15755 | 0.44222 | 1.041 | 1.8552 | 0.47039 | 1.1179 | 0.446 | 0.21345 | 0.22012 | 0.074444 | 0.21044 | 0.025237 | 0.0046228 | 0.014655 | 0.031453 | 0.020792 | 0.0054217 | 0.41797 | 0.43321 | 211084.8035 | 234238.106 | 285234.9467 | 236019.0841 | 210393.4368 | 23924.6837 | 0.1973 | 0.47552 | 0.99557 | 1.9421
72 | CCV 10 ppb MAJORS |  | 06/16/2023 19:25:27 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.721 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.2452 | 0.54356 | 0.052947 | 0.093299 | 0.33754 | 0.75808 | 0.37847 | 0.80152 | 1.038 | 0.23509 | 0.16011 | 0.45181 | 1.0756 | 1.8958 | 0.47808 | 1.1318 | 0.44969 | 0.21538 | 0.21927 | 0.075524 | 0.21209 | 0.024969 | 0.0048262 | 0.014986 | 0.031716 | 0.02099 | 0.0058988 | 0.42229 | 0.43553 | 216719.5152 | 240521.9679 | 291033.3941 | 241381.0105 | 215514.8586 | 24525.704 | 0.19992 | 0.47452 | 0.96423 | 1.908
73 | 10_1_21_W_11_DUP_LSP |  | 06/16/2023 19:31:27 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_W_11_DUP_LSP.722 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.7199 | 1.1841 | 0.26746 | 0.23399 | 0.75975 | 1.5859 | 1.0292 | 1.6556 | 2.0948 | 0.45989 | 0.60489 | 1.0855 | 4.0941 | 3.8376 | 0.96691 | 2.3175 | 0.9002 | 0.43166 | 0.4677 | 0.15371 | 0.4454 | 0.047813 | 0.0095862 | 0.029274 | 0.063277 | 0.041775 | 0.011555 | 0.86502 | 0.89749 | 214672.9548 | 235448.4978 | 282882.0065 | 233593.0776 | 208420.0414 | 23512.6675 | 0.38879 | 0.92911 | 1.7299 | 3.6736
74 | 10_1_21_F_1 |  | 06/16/2023 19:37:27 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_F_1.723 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.1229 | 0.34978 | 0.020449 | 0.065609 | 0.82507 | 0.17538 | 1.4324 | 0.018323 | 0.0084627 | 0.0017005 | 0.17848 | 0.78336 | 1.2924 | 0.0054474 | 0.020798 | 0.095511 | 0.03997 | 0.30515 | 0.064487 | 0.020247 | 0.062032 | 0.0013841 | 3.0744e-05 | 6.2463e-05 | 0.00010606 | 0.00094583 | 2.9816e-05 | 0.71818 | 0.7491 | 205124.6338 | 238381.1234 | 279514.2026 | 232699.2542 | 214816.4371 | 23883.9498 | 0.00020341 | 0.00038467 | 0.0061788 | 0.011568
75 | 10_1_21_F_1_DUP |  | 06/16/2023 19:43:27 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_F_1_DUP.724 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.1902 | 0.37018 | 0.019225 | 0.061004 | 0.84937 | 0.18349 | 1.4828 | 0.017628 | 0.007284 | 0.001688 | 0.13089 | 0.80624 | 0.96979 | 0.0066531 | 0.023512 | 0.089756 | 0.039345 | 0.2979 | 0.066355 | 0.020727 | 0.064461 | 0.0014527 | 2.2619e-05 | 8.8214e-05 | 0.00014638 | 0.00082954 | 4.129e-05 | 0.74752 | 0.77152 | 204207.2745 | 236946.9859 | 275806.337 | 229754.8895 | 214827.3163 | 23660.2446 | 0.00019163 | 0.00035897 | 0.0060855 | 0.01125
76 | 10_1_21_F_6 |  | 06/16/2023 19:49:26 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_F_6.725 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.1948 | 0.42014 | 0.019423 | 0.028028 | 0.63377 | 0.16116 | 1.4906 | 0.048635 | 0.0078905 | 0.0016999 | 0.1165 | 0.48662 | 0.88507 | 0.0040395 | 0.013921 | 0.24056 | 0.096629 | 0.10983 | 0.077081 | 0.023481 | 0.07507 | 0.001108 | 3.3491e-05 | 0.00013199 | 0.00020312 | 0.00082779 | 5.2106e-05 | 0.33982 | 0.35171 | 205881.6289 | 239090.1854 | 277634.1111 | 230949.2563 | 213765.6342 | 23623.5168 | 0.00028837 | 0.00069728 | 0.010151 | 0.019291
77 | 10_1_21_F_6_DUP |  | 06/16/2023 19:55:26 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_F_6_DUP.726 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.1656 | 0.4174 | 0.020175 | 0.027236 | 0.63235 | 0.15648 | 1.4372 | 0.049667 | 0.0087986 | 0.001677 | 0.11691 | 0.4876 | 0.88388 | 0.0040844 | 0.012058 | 0.23793 | 0.097604 | 0.12427 | 0.077993 | 0.023888 | 0.07548 | 0.0012472 | 3.9668e-05 | 9.7168e-05 | 0.00025194 | 0.0010335 | 3.8719e-05 | 0.33481 | 0.35245 | 204082.9322 | 235478.2253 | 275983.7174 | 228008.5396 | 213804.7565 | 23679.6103 | 0.00026884 | 0.0005232 | 0.010916 | 0.020564
78 | 10_1_21_F_8 |  | 06/16/2023 20:01:27 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_F_8.727 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.2598 | 0.42998 | 0.019169 | 0.028343 | 0.58748 | 0.16291 | 1.4511 | 0.063834 | 0.007035 | 0.0013596 | 0.13449 | 0.46418 | 0.99727 | 0.0036832 | 0.010576 | 0.26521 | 0.10723 | 0.11894 | 0.089484 | 0.028384 | 0.086421 | 0.0016047 | 3.944e-05 | 0.0001276 | 0.00028146 | 0.0009059 | 4.1275e-05 | 0.30228 | 0.31659 | 204764.4579 | 236861.4969 | 277169.6527 | 232381.5749 | 214158.7051 | 23370.4403 | 0.00016894 | 0.00050614 | 0.02153 | 0.040064
79 | 10_1_21_F_9 |  | 06/16/2023 20:07:27 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_F_9.728 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.0619 | 0.40271 | 0.020013 | 0.034423 | 0.4974 | 0.14409 | 1.2854 | 0.051286 | 0.012513 | 0.0028029 | 0.095554 | 0.22847 | 0.73296 | 0.0034383 | 0.010621 | 0.2465 | 0.097812 | 0.14494 | 0.069148 | 0.021742 | 0.066844 | 0.0017012 | 3.9203e-05 | 0.00014373 | 0.00023494 | 0.00087817 | 6.6275e-05 | 0.20878 | 0.218 | 207947.347 | 238050.3564 | 279222.4364 | 233246.0702 | 215219.2342 | 23671.5984 | 0.00020352 | 0.00051369 | 0.057226 | 0.10513
80 | 10_1_21_F_9_DUP |  | 06/16/2023 20:13:27 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_F_9_DUP.729 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.0334 | 0.3974 | 0.024535 | 0.034752 | 0.48684 | 0.14203 | 1.2606 | 0.053225 | 0.010631 | 0.0024987 | 0.090896 | 0.22779 | 0.7015 | 0.0061927 | 0.012012 | 0.24838 | 0.096354 | 0.14459 | 0.067486 | 0.022006 | 0.066631 | 0.0015053 | 2.2479e-05 | 8.2466e-05 | 0.00024444 | 0.00080015 | 4.6829e-05 | 0.20831 | 0.21575 | 204967.39 | 237165.7943 | 279319.6481 | 231094.1699 | 213783.2324 | 23489.9637 | 0.00019605 | 0.00049448 | 0.0099171 | 0.018544
81 | 10_1_21_F_10 |  | 06/16/2023 20:19:27 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_F_10.730 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.92555 | 0.47891 | 0.041849 | 0.043202 | 0.4939 | 0.14666 | 1.3434 | 0.056719 | 0.0064565 | 0.0013857 | 0.089062 | 0.248 | 0.68595 | 0.0037236 | 0.0087887 | 0.18427 | 0.074493 | 0.1596 | 0.080843 | 0.024693 | 0.077903 | 0.0018225 | 3.1248e-05 | 8.9298e-05 | 0.00030766 | 0.00086579 | 3.3241e-05 | 0.12835 | 0.13386 | 216898.2326 | 234249.6623 | 273951.1177 | 229752.1803 | 213063.8334 | 23677.6079 | 0.00014907 | 0.00038393 | 1.8605 | 3.971
82 | 10_1_21_F_11 |  | 06/16/2023 20:25:27 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_F_11.731 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.91187 | 0.48181 | 0.03512 | 0.034206 | 0.50859 | 0.14458 | 1.3428 | 0.058501 | 0.010125 | 0.0022892 | 0.11006 | 0.21662 | 0.82169 | 0.003869 | 0.0094219 | 0.1771 | 0.074029 | 0.15137 | 0.081821 | 0.024936 | 0.081121 | 0.0014866 | 3.9755e-05 | 0.00012302 | 0.00026446 | 0.0006907 | 3.3192e-05 | 0.12265 | 0.12558 | 203191.7289 | 234705.0101 | 272726.1868 | 227456.5354 | 210235.6963 | 23301.6594 | 0.00016208 | 0.00033245 | 0.011278 | 0.020831
83 | LBLANK |  | 06/16/2023 20:31:27 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.732 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 7.4438e-05 | 2.6841e-05 | -0.0013295 | -0.001297 | 0.01277 | 0.00029973 | 6.4999e-06 | -0.00015623 | 0.00050159 | 6.2752e-05 | -0.00012654 | -1.9361e-05 | 0.027302 | -0.00016273 | 0.00011417 | 0.0001654 | 2.8945e-05 | -0.0022308 | -0.0015913 | -2.0341e-05 | -1.0744e-05 | 0.00036749 | 0 | 5.9646e-06 | -1.6941e-05 | -1.9113e-05 | -1.1289e-06 | -0.00039826 | 6.1875e-05 | 217429.1389 | 255193.3839 | 301075.6912 | 244986.8911 | 222540.6439 | 24425.535 | -6.6811e-06 | -4.0279e-05 | -8.557e-05 | -0.00015491
84 | CCV 1ppb TRACE |  | 06/16/2023 20:37:27 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.733 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.12694 | 0.05886 | 0.010464 | 0.0086721 | 0.039217 | 0.078449 | 0.034742 | 0.084124 | 0.10652 | 0.024432 | 0.022576 | 0.043788 | 0.16948 | 0.19128 | 0.047382 | 0.11939 | 0.047222 | 0.020007 | 0.019815 | 0.007722 | 0.020887 | 0.002842 | 0.00045285 | 0.0014152 | 0.0030101 | 0.0019709 | 0.00050456 | 0.040872 | 0.042094 | 213637.239 | 251752.975 | 293376.5585 | 243083.4035 | 221706.6148 | 24801.5115 | 0.019641 | 0.048254 | 0.10706 | 0.19582
85 | CCV 1ppb TRACE |  | 06/16/2023 20:43:46 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.734 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.1274 | 0.060274 | 0.010161 | 0.0085236 | 0.044625 | 0.078387 | 0.039305 | 0.082672 | 0.10673 | 0.024065 | 0.022647 | 0.044823 | 0.17911 | 0.19196 | 0.051639 | 0.1216 | 0.046846 | 0.019824 | 0.019454 | 0.0071194 | 0.021096 | 0.0029214 | 0.00051191 | 0.0013842 | 0.0029828 | 0.0019335 | 0.00055096 | 0.041218 | 0.043232 | 218104.9314 | 250040.5897 | 295752.8513 | 244913.0413 | 218465.805 | 24658.5982 | 0.019992 | 0.047564 | 0.10425 | 0.19671
86 | LFB |  | 06/16/2023 20:49:46 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.735 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.2164 | 0.54091 | 0.053334 | 0.12185 | 0.331 | 0.73884 | 0.36523 | 0.76353 | 0.99256 | 0.22198 | 0.15337 | 0.42903 | 1.003 | 1.8062 | 0.45899 | 1.0818 | 0.43192 | 0.20256 | 0.21979 | 0.074941 | 0.20524 | 0.025129 | 0.0046061 | 0.014515 | 0.031154 | 0.020461 | 0.0054853 | 0.40062 | 0.41601 | 205684.7127 | 234097.2258 | 282494.6079 | 231977.0383 | 208758.9221 | 23774.4335 | 0.19209 | 0.46016 | 0.96803 | 1.8516
87 | CCV 10 ppb MAJORS |  | 06/16/2023 20:55:46 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.736 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.2734 | 0.56527 | 0.055832 | 0.097269 | 0.33997 | 0.77074 | 0.38337 | 0.81377 | 1.0383 | 0.23695 | 0.16098 | 0.45022 | 1.0483 | 1.8844 | 0.47637 | 1.1156 | 0.45275 | 0.21814 | 0.2185 | 0.076555 | 0.21067 | 0.024413 | 0.0046867 | 0.014833 | 0.031636 | 0.02098 | 0.0056278 | 0.41979 | 0.43365 | 214245.3488 | 234992.3377 | 280397.7682 | 234565.6582 | 211960.8813 | 23856.5699 | 0.19933 | 0.47592 | 0.96187 | 1.9146
88 | 10_1_21_UF_1 |  | 06/16/2023 21:01:46 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_UF_1.737 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.2293 | 0.34712 | 0.0057673 | 0.055224 | 0.7994 | 0.17563 | 1.4411 | 0.018286 | 0.0091792 | 0.0015146 | 0.014017 | 0.77037 | 0.2005 | 0.0053443 | 0.018615 | 0.065139 | 0.028063 | 0.25605 | 0.06033 | 0.019742 | 0.060011 | 0.0013792 | 2.5782e-05 | 6.6839e-05 | 0.00013748 | 0.00084234 | 3.3491e-05 | 0.71699 | 0.73797 | 206878.5613 | 232866.6779 | 273077.9848 | 226776.564 | 209068.9974 | 23713.6671 | 0.00021302 | 0.00033659 | 0.003162 | 0.0064053
89 | 10_1_21_UF_6 |  | 06/16/2023 21:07:46 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_UF_6.738 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.3588 | 0.43618 | 0.0066675 | 0.037341 | 0.65445 | 0.16668 | 1.5296 | 0.049748 | 0.010587 | 0.0018872 | 0.012561 | 0.49219 | 0.20316 | 0.004004 | 0.01268 | 0.16219 | 0.068428 | 0.092834 | 0.073333 | 0.025242 | 0.074596 | 0.0013536 | 2.88e-05 | 0.0001107 | 0.00027148 | 0.00085348 | 6.5264e-05 | 0.34226 | 0.35064 | 212243.7496 | 231230.4399 | 271488.7743 | 228232.0462 | 208406.5412 | 22879.6433 | 0.0002339 | 0.00050355 | 0.005535 | 0.01122
90 | 10_1_21_UF_8 |  | 06/16/2023 21:13:46 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_UF_8.739 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.3225 | 0.41812 | 0.0066959 | 0.051268 | 0.57523 | 0.16187 | 1.4401 | 0.064309 | 0.017408 | 0.0036341 | 0.021519 | 0.44505 | 0.26033 | 0.0045052 | 0.010661 | 0.14551 | 0.060511 | 0.09245 | 0.085288 | 0.029125 | 0.083979 | 0.0015834 | 3.4952e-05 | 0.00011217 | 0.00025214 | 0.00093588 | 4.6035e-05 | 0.2943 | 0.30362 | 202752.9135 | 227831.0851 | 271251.6785 | 224406.3274 | 208813.039 | 22999.8355 | 0.00021455 | 0.00034668 | 0.0055356 | 0.010688
91 | 10_1_21_UF_8_LD |  | 06/16/2023 21:19:45 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_UF_8_LD.740 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.3402 | 0.4245 | 0.011345 | 0.08191 | 0.55642 | 0.16062 | 1.4039 | 0.063037 | 0.017226 | 0.0033598 | 0.021558 | 0.43855 | 0.24246 | 0.003511 | 0.01135 | 0.14513 | 0.059852 | 0.091568 | 0.085935 | 0.027227 | 0.08348 | 0.0020091 | 2.6145e-05 | 0.00013198 | 0.00024486 | 0.0009042 | 5.4383e-05 | 0.29472 | 0.30499 | 202584.9798 | 229231.1851 | 270205.4583 | 219655.0453 | 204412.7913 | 23067.9435 | 0.00019817 | 0.00036568 | 0.0053381 | 0.01084
92 | 10_1_21_UF_8_LD |  | 06/16/2023 21:26:03 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_UF_8_LD.741 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.3207 | 0.41003 | 0.011313 | 0.085766 | 0.56219 | 0.16431 | 1.4045 | 0.065313 | 0.018365 | 0.0034399 | 0.021179 | 0.44218 | 0.23451 | 0.0035563 | 0.011726 | 0.14246 | 0.060154 | 0.091292 | 0.084766 | 0.026355 | 0.083228 | 0.0015937 | 2.3468e-05 | 0.00012753 | 0.00025655 | 0.00096525 | 3.4386e-05 | 0.29323 | 0.3064 | 201739.1843 | 226874.062 | 269808.8308 | 225407.0857 | 204114.8025 | 23126.0368 | 0.00021814 | 0.00045112 | 0.0054517 | 0.01045
93 | 10_1_21_UF_8_DUP |  | 06/16/2023 21:32:02 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_UF_8_DUP.742 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.3366 | 0.42115 | 0.0065237 | 0.058895 | 0.57132 | 0.16387 | 1.4288 | 0.063097 | 0.029381 | 0.0059905 | 0.033549 | 0.44261 | 0.32636 | 0.0042125 | 0.011418 | 0.15722 | 0.062173 | 0.10208 | 0.087073 | 0.026 | 0.083904 | 0.0013508 | 4.4593e-05 | 0.00012301 | 0.00029851 | 0.0010368 | 5.2702e-05 | 0.29782 | 0.30656 | 198818.6442 | 224087.8791 | 264249.7311 | 221903.6464 | 201488.4129 | 22394.206 | 0.00017986 | 0.00050156 | 0.006218 | 0.011394
94 | 10_1_21_UF_8_DUP_LSP |  | 06/16/2023 21:38:02 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_UF_8_DUP_LSP.743 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.4997 | 0.92287 | 0.062213 | 0.16503 | 0.84304 | 0.93354 | 1.6566 | 0.88324 | 1.0504 | 0.23735 | 0.19667 | 0.84206 | 1.35 | 1.8874 | 0.47141 | 1.2159 | 0.48796 | 0.29154 | 0.29534 | 0.10006 | 0.28144 | 0.025338 | 0.0046059 | 0.014932 | 0.031143 | 0.021697 | 0.0054012 | 0.68948 | 0.7027 | 201124.5892 | 221769.4493 | 260233.6254 | 215018.7401 | 198589.5924 | 22183.8785 | 0.19417 | 0.46334 | 0.93406 | 1.874
95 | 10_1_21_UF_9 |  | 06/16/2023 21:44:02 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_UF_9.744 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.1795 | 0.39675 | 0.0044159 | 0.060007 | 0.49072 | 0.142 | 1.2941 | 0.052351 | 0.009018 | 0.0014032 | 0.016806 | 0.22171 | 0.21653 | 0.0034958 | 0.0099349 | 0.15136 | 0.061658 | 0.12553 | 0.066791 | 0.022604 | 0.066599 | 0.0013739 | 3.2292e-05 | 0.00014528 | 0.00025006 | 0.0007648 | 5.2099e-05 | 0.20978 | 0.21681 | 206863.8133 | 226839.5365 | 269885.5523 | 221371.4075 | 207318.1731 | 23136.7225 | 0.00018606 | 0.00050724 | 0.0049661 | 0.0098489
96 | 10_1_21_UF_10 |  | 06/16/2023 21:50:03 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_UF_10.745 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.0638 | 0.47895 | 0.0040087 | 0.05114 | 0.4997 | 0.14664 | 1.3614 | 0.058296 | 0.030221 | 0.0063729 | 0.033321 | 0.23439 | 0.31772 | 0.0031867 | 0.0094748 | 0.11614 | 0.047695 | 0.10593 | 0.072652 | 0.024218 | 0.077266 | 0.0011588 | 4.1438e-05 | 0.00013687 | 0.00025282 | 0.00077803 | 5.257e-05 | 0.12769 | 0.1335 | 218643.3073 | 225333.413 | 269321.7942 | 224555.9367 | 203461.9909 | 22688.002 | 0.00012419 | 0.00028075 | 0.006666 | 0.012996
97 | 10_1_21_UF_11 |  | 06/16/2023 21:56:03 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_UF_11.746 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.0006 | 0.47771 | 0.0037049 | 0.044299 | 0.4967 | 0.1447 | 1.3456 | 0.056918 | 0.0078073 | 0.0014191 | 0.032281 | 0.19864 | 0.3073 | 0.003171 | 0.0082005 | 0.10613 | 0.044381 | 0.10914 | 0.073769 | 0.025217 | 0.078841 | 0.0012501 | 5.0028e-05 | 0.00013326 | 0.00028873 | 0.00062197 | 4.9118e-05 | 0.11849 | 0.12181 | 215727.6239 | 226880.1134 | 271269.3551 | 226255.1187 | 204109.8743 | 22993.8248 | 0.00013331 | 0.00029014 | 0.0054803 | 0.011218
98 | LBLANK |  | 06/16/2023 22:02:04 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.747 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 8.0294e-05 | 0.00011202 | -0.00079356 | -0.00083832 | 0.021181 | 0.00017198 | 9.0677e-05 | -0.00012751 | 0.00050659 | 6.378e-05 | -0.00015026 | -3.0259e-05 | 0.021239 | -0.00010712 | 0.0004412 | 7.0403e-05 | 1.1033e-05 | -0.002212 | -0.00036465 | 3.6071e-05 | 0.00042568 | 0.0003694 | 2.7337e-06 | 1.1905e-05 | -3.1827e-06 | -7.2032e-05 | -3.7648e-06 | -0.00042703 | 7.1232e-05 | 215915.3063 | 242124.6624 | 292825.4174 | 241198.5049 | 214592.4169 | 23877.9398 | 1.7019e-07 | -3.0332e-05 | -0.00013156 | -0.00015806
99 | CCV 1ppb TRACE |  | 06/16/2023 22:08:04 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.748 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.12548 | 0.058448 | 0.0098949 | 0.0083409 | 0.054687 | 0.076548 | 0.038179 | 0.083276 | 0.10599 | 0.024349 | 0.022557 | 0.043659 | 0.15689 | 0.1928 | 0.050221 | 0.12031 | 0.047969 | 0.019632 | 0.019989 | 0.0075536 | 0.020774 | 0.0031205 | 0.00051253 | 0.0014567 | 0.0030247 | 0.001856 | 0.00055281 | 0.040742 | 0.042948 | 212561.1809 | 241807.3418 | 283916.2958 | 236973.4264 | 209015.4544 | 23843.2157 | 0.020058 | 0.047117 | 0.11029 | 0.20014
100 | CCV 1ppb TRACE |  | 06/16/2023 22:14:23 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.749 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.12772 | 0.058811 | 0.0096479 | 0.0080092 | 0.049246 | 0.077127 | 0.039342 | 0.084487 | 0.10697 | 0.023904 | 0.021861 | 0.044189 | 0.17212 | 0.20203 | 0.051551 | 0.12002 | 0.047171 | 0.019287 | 0.01901 | 0.007643 | 0.021116 | 0.0032529 | 0.00047579 | 0.001409 | 0.0030635 | 0.0018927 | 0.00061668 | 0.040721 | 0.043083 | 212830.4044 | 239561.9603 | 287831.474 | 255690.0581 | 207811.3303 | 23647.5602 | 0.019888 | 0.04795 | 0.10869 | 0.18118
101 | CCV 10 ppb MAJORS |  | 06/16/2023 22:20:23 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.750 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.2505 | 0.55423 | 0.053477 | 0.094394 | 0.34645 | 0.77347 | 0.37765 | 0.80912 | 1.0489 | 0.235 | 0.16093 | 0.45474 | 1.0511 | 1.9072 | 0.47494 | 1.132 | 0.44943 | 0.21599 | 0.21746 | 0.073852 | 0.21425 | 0.024658 | 0.0048272 | 0.014758 | 0.030974 | 0.02029 | 0.005732 | 0.42037 | 0.44016 | 207076.0329 | 218468.5647 | 270299.8949 | 220860.8839 | 197825.6871 | 22507.7187 | 0.19964 | 0.47716 | 0.97828 | 1.9578
102 | 10_1_21_BB |  | 06/16/2023 22:26:24 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_BB.751 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4.8632e-05 | 5.393e-05 | 0.002624 | 0.0015681 | 0.020581 | 0.00011701 | 0.00014353 | -0.00010169 | 0.0024032 | 0.00063462 | 0.0095867 | 0.00019722 | 0.087654 | 0.00013664 | 0.00017393 | 0.010303 | 0.0040298 | 0.0018284 | -0.00068504 | 0.00011577 | 0.00019037 | 0.0001818 | 2.7558e-06 | 9.2281e-06 | -8.6387e-06 | -1.0643e-05 | -9.7988e-07 | -0.00030562 | 9.3911e-05 | 218195.1558 | 240728.0062 | 294017.1978 | 239892.9251 | 209727.6287 | 24340.054 | 2.4534e-06 | -9.9623e-06 | 0.0029381 | 0.0057024
103 | 10_1_21_FB |  | 06/16/2023 22:32:24 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_FB.752 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.025009 | 5.6446e-05 | 0.026605 | 0.010586 | 0.013995 | 0.00055474 | 0.00049032 | 7.1756e-05 | 0.0058313 | 0.0013453 | 0.015456 | 0.001763 | 0.12189 | 0.00031824 | 0.0025994 | 0.026917 | 0.010715 | 0.087227 | -0.001921 | 9.3123e-05 | 0.00018977 | 0.00057017 | 2.8215e-06 | 6.6436e-06 | 2.8386e-06 | -2.3843e-05 | -3.7648e-06 | -0.00014441 | 0.00029144 | 213267.6927 | 235447.7732 | 286657.5471 | 236200.0046 | 206355.0413 | 23634.2024 | 2.6648e-07 | -6.2105e-06 | 0.0028626 | 0.0053862
104 | 10_1_21_LAB_A |  | 06/16/2023 22:38:24 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_LAB_A.753 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.00022339 | 0 | 0.0028466 | 0.0010096 | 0.013864 | 0.00018983 | 0.0002325 | -0.00012674 | 0.0020119 | 0.00041974 | 0.0066284 | 0.00033319 | 0.047003 | 7.003e-06 | 0.00034237 | 0.010726 | 0.0038866 | 0.0050409 | 0.00018196 | 6.3918e-05 | 9.4433e-05 | 0.00036677 | 2.8579e-06 | 1.5228e-05 | -5.5533e-06 | 1.3852e-05 | -9.2224e-07 | -0.0004275 | 7.1332e-05 | 212866.2801 | 233946.7053 | 289338.1168 | 236296.2581 | 203284.539 | 23661.5802 | -6.6811e-06 | -2.9648e-05 | 0.0027422 | 0.0049146
105 | 10_1_21_LAB_B |  | 06/16/2023 22:44:24 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_LAB_B.754 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.00025497 | 2.8587e-05 | 0.0021721 | 0.0016865 | 0.0098421 | -0.00017859 | 9.1828e-05 | -0.00012601 | 0.005437 | 0.0011191 | 0.0073057 | 0.00044685 | 0.058169 | 5.4685e-06 | 0.00043215 | 0.011732 | 0.0044663 | 0.0042972 | 7.5931e-06 | 8.1328e-06 | 8.357e-05 | -6.6728e-05 | 0 | 9.7112e-07 | -2.6037e-06 | -6.1973e-05 | -8.926e-07 | -0.00045719 | 3.455e-05 | 214285.4084 | 232737.1938 | 288530.6415 | 237085.8194 | 202243.0071 | 23642.887 | -4.373e-06 | -1.8026e-05 | 0.0049802 | 0.0089244
106 | 5_18_23_DB |  | 06/16/2023 22:50:24 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_DB.755 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 5.65 | 0.67272 | 0.59365 | 0.11148 | 0.037068 | 0.28529 | 1.8742 | 0.15549 | 0.14416 | 0.032571 | 0.64009 | 0.15451 | 4.1074 | 0.24951 | 0.079342 | 0.087496 | 0.038301 | 0.122 | 0.049005 | 0.016243 | 0.049781 | 0.011589 | 0.000949 | 0.0026699 | 0.0057718 | 0.0038111 | 0.001142 | 0.043833 | 0.045754 | 189543.9762 | 193157.7526 | 244004.9463 | 206489.0341 | 171780.7009 | 19993.9824 | 0.066748 | 0.15581 | 0.12979 | 0.24451
107 | 5_18_23_DB |  | 06/16/2023 22:56:42 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_DB.756 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 5.5974 | 0.67774 | 0.57102 | 0.12057 | 0.029198 | 0.28258 | 1.8825 | 0.14682 | 0.13893 | 0.031394 | 0.62779 | 0.15406 | 4.0318 | 0.25274 | 0.075629 | 0.084373 | 0.03817 | 0.12111 | 0.05034 | 0.015803 | 0.048883 | 0.01221 | 0.00089563 | 0.002746 | 0.0057994 | 0.0043402 | 0.0010765 | 0.044295 | 0.046244 | 184970.4632 | 194397.7609 | 248972.8791 | 205198.5688 | 171766.4695 | 20424.5949 | 0.066083 | 0.15474 | 0.13537 | 0.24836
108 | 5_18_23_T_239 |  | 06/16/2023 23:02:42 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_T_239.757 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 5.6563 | 0.67728 | 0.59056 | 0.12419 | 0.031998 | 0.29071 | 1.912 | 0.14975 | 0.13933 | 0.032684 | 0.63818 | 0.15605 | 4.1238 | 0.25391 | 0.077735 | 0.085084 | 0.037612 | 0.12371 | 0.049375 | 0.016134 | 0.049586 | 0.012443 | 0.00087101 | 0.0027959 | 0.0058353 | 0.0041871 | 0.00098012 | 0.046217 | 0.045638 | 186275.9332 | 194429.5228 | 246050.9063 | 208871.9738 | 170610.0704 | 20140.8548 | 0.06628 | 0.156 | 0.13382 | 0.24633
109 | 5_18_23_T_239 |  | 06/16/2023 23:09:00 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_T_239.758 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 5.633 | 0.6727 | 0.60038 | 0.11964 | 0.0385 | 0.2921 | 1.9266 | 0.15005 | 0.14164 | 0.031437 | 0.6274 | 0.1488 | 4.1287 | 0.25205 | 0.083146 | 0.084714 | 0.037037 | 0.1204 | 0.050143 | 0.016088 | 0.04689 | 0.012112 | 0.00077959 | 0.0026606 | 0.0055046 | 0.0041975 | 0.0010028 | 0.046017 | 0.043898 | 183970.3873 | 199381.8258 | 247156.3664 | 203305.3163 | 171066.1834 | 19937.904 | 0.065854 | 0.15488 | 0.13436 | 0.25135
110 | 5_18_23_T_239 |  | 06/16/2023 23:15:00 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_T_239.759 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.51163 | 0.035631 | 0.044976 | 0.16874 | 0.041668 | 0.036494 | 0.12204 | 0.081769 | 0.0040426 | 0.00091325 | 0.0040353 | 0.23438 | 0.049333 | 0.0030388 | 0.0024337 | 0.0072184 | 0.003565 | 0.014605 | 0.003434 | 0.00031419 | 0.0011884 | 0.00097573 | 6.2e-06 | 2.5981e-05 | 6.0498e-05 | 0.001647 | 1.7952e-05 | 0.0098721 | 0.0083592 | 203896.8299 | 215545.2695 | 270531.5858 | 237180.181 | 190534.3413 | 21644.3991 | 7.4719e-05 | 6.4419e-05 | 0.0017592 | 0.0032997
111 | 5_18_23_M_224 |  | 06/16/2023 23:20:59 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_M_224.760 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.56381 | 0.25032 | 0.026532 | 0.097741 | 0.16628 | 0.35498 | 0.18347 | 0.38306 | 0.49064 | 0.11317 | 0.078595 | 0.21738 | 0.5087 | 0.90103 | 0.22744 | 0.55127 | 0.2227 | 0.10304 | 0.10979 | 0.035624 | 0.10158 | 0.014285 | 0.0023493 | 0.0067093 | 0.014922 | 0.010601 | 0.0027692 | 0.20222 | 0.20958 | 187241.5425 | 216581.442 | 273923.2546 | 225922.5181 | 191519.9975 | 22631.9132 | 0.097315 | 0.23282 | 0.59466 | 1.0089
112 | LBLANK |  | 06/16/2023 23:27:00 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.761 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.8206e-05 | 2.8872e-05 | -0.0012039 | 0.010143 | 0.012945 | 3.5499e-05 | 3.7402e-05 | 1.6402e-05 | 0.00065153 | 3.0635e-05 | -0.00016606 | -1.981e-06 | 0.025276 | 0.00020942 | -0.00010003 | -4.5672e-06 | -2.3488e-05 | -0.0021568 | -0.00089379 | -2.0341e-05 | 3.8502e-05 | 0.00012705 | 5.8325e-06 | 6.8417e-06 | -1.4074e-05 | 2.1065e-06 | 1.9708e-06 | 0.0005343 | 8.1271e-05 | 215594.6652 | 230541.4221 | 288532.7914 | 236399.3492 | 198474.5954 | 23353.7413 | -4.3866e-06 | -1.341e-05 | -0.00010059 | -0.00014684
113 | LBLANK |  | 06/16/2023 23:33:18 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.762 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.00011737 | 2.8528e-05 | -0.0012589 | 0.00071561 | 0.015452 | 4.569e-05 | 3.7204e-05 | -6.8754e-05 | 0.00063041 | 4.5973e-05 | -0.00016381 | -7.2215e-06 | 0.025204 | -4.8481e-05 | -9.5917e-05 | -0.00011552 | -6.7366e-05 | -0.0020843 | 1.4342e-05 | 8.6765e-06 | 1.0069e-05 | 0.00016164 | 0 | 3.9499e-06 | -1.4025e-05 | -8.2655e-05 | -3.7648e-06 | 0.00022383 | 4.4247e-05 | 213159.4411 | 228269.3137 | 304774.8757 | 236837.7964 | 197036.3344 | 23218.8548 | -2.2689e-06 | -4.9739e-05 | -4.6555e-05 | -9.0788e-05
114 | CCV 1ppb TRACE |  | 06/16/2023 23:39:19 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.763 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.11874 | 0.054386 | 0.009462 | 0.0089037 | 0.041416 | 0.073009 | 0.036999 | 0.080392 | 0.10197 | 0.023299 | 0.021058 | 0.043123 | 0.16227 | 0.18253 | 0.045951 | 0.11439 | 0.045354 | 0.018449 | 0.01972 | 0.00747 | 0.020127 | 0.0033219 | 0.00040943 | 0.0014024 | 0.0030528 | 0.001895 | 0.00053207 | 0.040975 | 0.042049 | 211419.7332 | 226471.1679 | 287143.4317 | 236891.3317 | 195948.0669 | 23145.4025 | 0.019588 | 0.047094 | 0.11381 | 0.20889
115 | CCV 1ppb TRACE |  | 06/16/2023 23:45:38 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.764 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.12069 | 0.051524 | 0.010027 | 0.0083341 | 0.048501 | 0.072776 | 0.039467 | 0.07916 | 0.10456 | 0.023564 | 0.022171 | 0.042841 | 0.16338 | 0.18572 | 0.048389 | 0.11336 | 0.046885 | 0.019459 | 0.021122 | 0.0067856 | 0.020707 | 0.0040993 | 0.00049727 | 0.0015563 | 0.0030547 | 0.0022828 | 0.00052401 | 0.041971 | 0.042603 | 184000.8965 | 225039.417 | 279711.9124 | 232508.1963 | 195322.2392 | 22886.3201 | 0.019546 | 0.047602 | 0.13036 | 0.20462
116 | LFB |  | 06/16/2023 23:51:38 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.765 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.1376 | 0.50424 | 0.050992 | 0.11763 | 0.32538 | 0.69486 | 0.36076 | 0.74955 | 0.9697 | 0.2209 | 0.15176 | 0.42094 | 0.94625 | 1.7438 | 0.45232 | 1.0661 | 0.42932 | 0.20331 | 0.20384 | 0.071784 | 0.20402 | 0.02376 | 0.0044141 | 0.014092 | 0.031008 | 0.01907 | 0.0054483 | 0.40587 | 0.41793 | 206742.4862 | 209474.6286 | 261882.0107 | 217709.1266 | 184799.0158 | 21661.0774 | 0.19493 | 0.46616 | 0.93771 | 1.9283
117 | LFB |  | 06/16/2023 23:57:55 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.766 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.1464 | 0.51901 | 0.051709 | 0.11823 | 0.33 | 0.71261 | 0.36784 | 0.76146 | 0.99155 | 0.2214 | 0.15299 | 0.42536 | 0.95077 | 1.794 | 0.45698 | 1.0798 | 0.43571 | 0.20298 | 0.20809 | 0.071184 | 0.20784 | 0.024306 | 0.0046384 | 0.014914 | 0.031409 | 0.019493 | 0.0056834 | 0.40588 | 0.41962 | 203689.2196 | 211835.7965 | 265950.8103 | 220853.4208 | 190234.2682 | 21916.1337 | 0.19504 | 0.46745 | 0.96017 | 1.9109
118 | CCV 10 ppb MAJORS |  | 06/17/2023 00:03:55 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.767 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.1751 | 0.53149 | 0.051679 | 0.095821 | 0.32746 | 0.72974 | 0.37702 | 0.78317 | 1.0069 | 0.22931 | 0.15808 | 0.4407 | 1.014 | 1.8379 | 0.46226 | 1.1006 | 0.44394 | 0.21163 | 0.22185 | 0.069177 | 0.20879 | 0.025827 | 0.0044096 | 0.014343 | 0.031278 | 0.020387 | 0.005813 | 0.417 | 0.43312 | 199498.7918 | 215007.9239 | 271098.8805 | 222373.471 | 191990.8429 | 22623.2326 | 0.19688 | 0.47254 | 1.0351 | 1.9802
119 | 5_18_23_MFB |  | 06/17/2023 00:09:55 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_MFB.768 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.00024517 | 0 | 0.00086251 | 0.033417 | 0.011464 | 0.00038182 | 0.00028381 | -1.7016e-05 | 0.0014362 | 0.00025829 | 0.0018985 | 0.00010028 | 0.04689 | 5.8671e-05 | 7.2378e-05 | 0.0087847 | 0.0033095 | 0.0037474 | -6.0585e-05 | 7.4202e-06 | 2.2358e-05 | 0.00050703 | 2.8123e-06 | 1.2171e-05 | -2.8801e-06 | -7.2268e-05 | -3.7648e-06 | 0.00067426 | 6.1886e-05 | 223592.8997 | 236764.5677 | 298383.3793 | 243014.8884 | 208377.4212 | 24080.9459 | 2.232e-06 | -3.3169e-05 | 0.0017566 | 0.0033363
120 | 10_15_21_W_1 |  | 06/17/2023 00:15:55 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_W_1.769 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.25213 | 0.077177 | 0.026172 | 0.043246 | 0.20875 | 0.039341 | 0.33274 | 0.0037028 | 0.003919 | 0.00072234 | 0.17309 | 0.25707 | 1.1469 | 0.0017909 | 0.0039123 | 0.034593 | 0.014852 | 0.0057369 | 0.012569 | 0.0047363 | 0.01422 | 0.0010329 | 1.2264e-05 | 2.5679e-05 | 3.5246e-05 | 0.00024204 | 2.367e-06 | 0.17467 | 0.18172 | 204898.3655 | 217310.3825 | 273595.1005 | 228692.6433 | 194188.3075 | 22828.227 | 6.4028e-05 | 7.5216e-05 | 0.0036882 | 0.0064495
121 | 10_15_21_W_6 |  | 06/17/2023 00:21:56 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_W_6.770 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.22818 | 0.080209 | 0.080657 | 0.026953 | 0.15485 | 0.033882 | 0.30629 | 0.0099985 | 0.003552 | 0.0012179 | 0.26244 | 0.16456 | 1.6919 | 0.001828 | 0.0027489 | 0.12608 | 0.050768 | 0.017471 | 0.013286 | 0.0047584 | 0.014823 | 0.00060531 | 1.5572e-05 | 4.4719e-05 | 5.1548e-05 | 0.00019168 | -6.0434e-07 | 0.096715 | 0.099651 | 203600.0275 | 214328.0093 | 269876.8198 | 223530.7375 | 192750.2981 | 22452.9663 | 0.00014866 | 0.00021821 | 0.013089 | 0.024379
122 | 10_15_21_W_8 |  | 06/17/2023 00:27:56 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_W_8.771 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.2299 | 0.075863 | 0.034852 | 0.025758 | 0.13989 | 0.031765 | 0.29763 | 0.010876 | 0.0027336 | 0.00047234 | 0.12992 | 0.11152 | 0.84018 | 0.0013658 | 0.0026331 | 0.076457 | 0.032071 | 0.0081104 | 0.015145 | 0.0050476 | 0.015414 | 0.00083358 | 1.2495e-05 | 2.6344e-05 | 4.254e-05 | 0.00014338 | 1.4984e-05 | 0.086507 | 0.088728 | 203553.291 | 213357.473 | 265460.5259 | 224321.5326 | 190566.6289 | 22238.6298 | 4.3419e-05 | 0.00010005 | 0.0061547 | 0.011185
123 | 10_15_21_W_9 |  | 06/17/2023 00:33:56 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_W_9.772 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.19611 | 0.070586 | 0.053706 | 0.024806 | 0.11765 | 0.027023 | 0.26006 | 0.0091304 | 0.0042216 | 0.00084708 | 0.1427 | 0.099985 | 0.9151 | 0.0010286 | 0.0027525 | 0.067029 | 0.027688 | 0.0076664 | 0.012204 | 0.0041401 | 0.01246 | 0.0021893 | 6.286e-06 | 2.6499e-05 | 3.9712e-05 | 0.00033498 | 1.2013e-05 | 0.059638 | 0.062916 | 173400.8493 | 211414.2617 | 269675.6208 | 224335.0365 | 190470.0886 | 22432.9341 | 7.4938e-05 | 9.6975e-05 | 0.007402 | 0.011187
124 | 10_15_21_W_9 |  | 06/17/2023 00:40:15 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_W_9.773 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.19852 | 0.069601 | 0.04963 | 0.02519 | 0.11671 | 0.027554 | 0.26091 | 0.0092611 | 0.0040885 | 0.00092393 | 0.1431 | 0.098311 | 0.92712 | 0.0011126 | 0.0017985 | 0.070556 | 0.02835 | 0.0079518 | 0.011746 | 0.0036341 | 0.012653 | 0.00081765 | 3.1435e-06 | 1.711e-05 | 3.4078e-05 | 0.00022917 | 1.2019e-05 | 0.060599 | 0.062215 | 202042.7188 | 210628.1957 | 265529.2567 | 221470.3836 | 188652.8415 | 21938.8346 | 5.086e-05 | 8.7356e-05 | 0.0060923 | 0.011562
125 | 10_15_21_W_10 |  | 06/17/2023 00:46:15 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_W_10.774 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.15911 | 0.08222 | 0.16406 | 0.035472 | 0.10148 | 0.027014 | 0.25151 | 0.011988 | 0.0044966 | 0.00084603 | 0.389 | 0.22419 | 2.4763 | 0.0026401 | 0.00239 | 0.09821 | 0.039508 | 0.024576 | 0.013515 | 0.0048504 | 0.014095 | 0.00073301 | 1.2744e-05 | 3.313e-05 | 4.6753e-05 | 9.9613e-05 | -5.8239e-07 | 0.026811 | 0.028293 | 200447.9352 | 209412.3814 | 264989.7963 | 221720.7232 | 186431.2964 | 21636.3725 | 0.00012179 | 0.00028129 | 0.018692 | 0.03332
126 | 10_15_21_W_10 |  | 06/17/2023 00:52:33 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_W_10.775 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.15691 | 0.080439 | 0.16655 | 0.034123 | 0.10093 | 0.024199 | 0.25557 | 0.011865 | 0.0037054 | 0.0007909 | 0.38761 | 0.22631 | 2.4986 | 0.0026302 | 0.0020076 | 0.097394 | 0.039119 | 0.024836 | 0.01252 | 0.0036922 | 0.014478 | 0.00053342 | 1.2684e-05 | 2.3453e-05 | 4.9547e-05 | 0.00012805 | 5.6938e-06 | 0.026688 | 0.027956 | 203380.8894 | 210734.9695 | 267997.7885 | 224432.5744 | 189220.9215 | 21733.1868 | 0.00011525 | 0.00023 | 0.018149 | 0.033609
127 | 10_15_21_W_10_DUP |  | 06/17/2023 00:58:33 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_W_10_DUP.776 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.15183 | 0.077575 | 0.15386 | 0.034664 | 0.10026 | 0.025376 | 0.24846 | 0.011363 | 0.0068425 | 0.0014803 | 0.37823 | 0.21661 | 2.3821 | 0.0023967 | 0.0021763 | 0.096172 | 0.039174 | 0.022318 | 0.012281 | 0.0041238 | 0.014215 | 0.0009212 | 0 | 1.386e-05 | 1.769e-05 | 0.00014853 | 5.7049e-06 | 0.026128 | 0.027008 | 204563.5737 | 212257.3243 | 271447.7379 | 220441.8605 | 188169.1834 | 21888.0909 | 0.00013079 | 0.00024298 | 0.016976 | 0.032369
128 | 10_15_21_W_10_DUP |  | 06/17/2023 01:04:52 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_W_10_DUP.777 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.1541 | 0.079614 | 0.14798 | 0.034771 | 0.10105 | 0.025767 | 0.24232 | 0.011467 | 0.0066317 | 0.001348 | 0.38239 | 0.21682 | 2.4086 | 0.0023934 | 0.0023702 | 0.097241 | 0.038681 | 0.021998 | 0.012782 | 0.0045649 | 0.014611 | 0.00074653 | 0 | 2.6942e-05 | 6.3148e-05 | 0.0002506 | 2.6174e-06 | 0.025678 | 0.027569 | 198667.8981 | 208045.9729 | 265149.3989 | 219736.5918 | 184411.9914 | 21351.2783 | 0.00013664 | 0.00022341 | 0.016852 | 0.031249
129 | LBLANK |  | 06/17/2023 01:10:52 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.778 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | -5.1329e-05 | 3.0895e-05 | -0.0012404 | 0.0023375 | 0.013235 | 0.00025681 | 9.7823e-06 | -5.9573e-05 | 0.0002111 | 3.7189e-05 | -0.00029092 | -2.7855e-05 | 0.0077982 | -0.00013192 | 1.1537e-05 | 0.00021625 | -9.8356e-05 | -0.002336 | -0.0011133 | -2.0341e-05 | 1.0774e-06 | 0.000275 | 0 | 4.3511e-06 | -1.3825e-05 | -3.643e-05 | -6.6942e-07 | -0.00021548 | 4.8006e-05 | 203686.4868 | 213597.7438 | 274826.4025 | 222238.6765 | 186069.2562 | 21781.9287 | -6.6811e-06 | -3.2237e-05 | -8.5495e-05 | -0.00014627
130 | LBLANK |  | 06/17/2023 01:17:12 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.779 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.2056e-05 | 0 | -0.0010213 | -0.00026067 | 0.01465 | 0.00036392 | 1.0169e-05 | -0.00015233 | 0.00039902 | 4.3826e-05 | -0.00027872 | -2.4762e-05 | 0.018977 | -7.0743e-05 | 0.00011331 | 0.00015967 | -3.2332e-05 | -0.0021801 | 0.00028485 | -2.0341e-05 | -8.2825e-06 | -0.00014219 | 0 | 1.2279e-06 | 1.8307e-06 | 3.5858e-06 | -3.7648e-06 | -0.00025089 | 4.803e-05 | 212315.6613 | 213476.5702 | 269695.9722 | 221435.7216 | 184472.1457 | 21629.0293 | -4.2003e-06 | -3.8885e-05 | -0.00010858 | -0.00014281
131 | CCV 1ppb TRACE |  | 06/17/2023 01:23:12 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.780 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.1186 | 0.05639 | 0.010532 | 0.0082511 | 0.044961 | 0.077746 | 0.037908 | 0.080953 | 0.10275 | 0.023333 | 0.021679 | 0.042879 | 0.15284 | 0.1867 | 0.046868 | 0.11539 | 0.046235 | 0.018848 | 0.019 | 0.0080271 | 0.020276 | 0.0029345 | 0.00040143 | 0.0015062 | 0.0029964 | 0.0019045 | 0.00060304 | 0.041207 | 0.04315 | 197327.5187 | 210885.8297 | 266356.2786 | 218383.2445 | 183216.0742 | 21309.2149 | 0.019835 | 0.048555 | 0.11612 | 0.2096
132 | CCV 1ppb TRACE |  | 06/17/2023 01:29:31 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.781 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.1177 | 0.052527 | 0.0096455 | 0.0085314 | 0.047099 | 0.074855 | 0.037972 | 0.081256 | 0.10271 | 0.023609 | 0.022295 | 0.043406 | 0.15254 | 0.18398 | 0.047901 | 0.11331 | 0.04584 | 0.019464 | 0.019408 | 0.0076807 | 0.020522 | 0.0030954 | 0.00044395 | 0.001505 | 0.0031063 | 0.0019001 | 0.00060262 | 0.041376 | 0.042412 | 197167.3539 | 210078.1429 | 268014.1955 | 215480.9901 | 182430.8213 | 21646.3886 | 0.019903 | 0.047649 | 0.11532 | 0.21922
133 | CCV 10 ppb MAJORS |  | 06/17/2023 01:35:32 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.782 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.1818 | 0.52658 | 0.049841 | 0.091784 | 0.33631 | 0.73349 | 0.38044 | 0.79186 | 1.0139 | 0.22917 | 0.15988 | 0.44309 | 0.9963 | 1.8412 | 0.46038 | 1.0941 | 0.43911 | 0.21149 | 0.21322 | 0.074812 | 0.20738 | 0.02392 | 0.0046379 | 0.014451 | 0.031152 | 0.019869 | 0.0054777 | 0.41508 | 0.43772 | 193907.2248 | 196667.8291 | 251590.9151 | 208486.3137 | 175430.596 | 20501.3671 | 0.19914 | 0.4771 | 1.0293 | 2.0286
134 | CCV 10 ppb MAJORS |  | 06/17/2023 01:41:51 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.783 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.1834 | 0.52789 | 0.053631 | 0.0916 | 0.33517 | 0.72417 | 0.38219 | 0.79489 | 0.99915 | 0.23409 | 0.16048 | 0.44055 | 0.99865 | 1.8318 | 0.46559 | 1.0799 | 0.44563 | 0.21273 | 0.21288 | 0.072609 | 0.20737 | 0.024055 | 0.0046631 | 0.014498 | 0.031113 | 0.020438 | 0.0056021 | 0.41732 | 0.43391 | 200768.2984 | 203428.9291 | 259118.1023 | 213831.8813 | 180850.2197 | 21287.182 | 0.19904 | 0.47527 | 1.0123 | 2.0034
135 | 10_15_21_W_11 |  | 06/17/2023 01:47:51 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_W_11.784 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.14828 | 0.081907 | 0.22254 | 0.034513 | 0.10058 | 0.026635 | 0.25039 | 0.013614 | 0.0054511 | 0.0011252 | 0.5113 | 0.24651 | 3.2626 | 0.003264 | 0.0028262 | 0.11606 | 0.048462 | 0.037282 | 0.012897 | 0.0045514 | 0.014822 | 0.00024321 | 6.3038e-06 | 7.5004e-06 | 5.5403e-05 | 3.7088e-05 | 5.6987e-06 | 0.023562 | 0.023675 | 220652.5668 | 212295.18 | 266047.1578 | 222890.0609 | 188363.1115 | 22175.198 | 0.00010853 | 0.00039004 | 0.019239 | 0.040218
136 | 10_15_21_W_11 |  | 06/17/2023 01:54:10 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_W_11.785 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.14615 | 0.080666 | 0.21496 | 0.036725 | 0.093903 | 0.024539 | 0.24934 | 0.013742 | 0.005138 | 0.0010848 | 0.51519 | 0.24504 | 3.1979 | 0.0031907 | 0.002717 | 0.11234 | 0.048241 | 0.037139 | 0.015381 | 0.0056181 | 0.015214 | 0.00059025 | 1.2954e-05 | 3.3689e-05 | 2.5099e-05 | 9.6547e-05 | 5.9209e-06 | 0.023447 | 0.024427 | 198411.044 | 206185.2847 | 261352.452 | 218750.9363 | 184408.6813 | 21879.409 | 0.00017428 | 0.00033226 | 0.022177 | 0.039865
137 | 10_15_21_F_1 |  | 06/17/2023 02:00:09 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_1.786 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.3137 | 0.37396 | 0.013985 | 0.068898 | 1.0389 | 0.21281 | 1.7224 | 0.015452 | 0.0078219 | 0.0017547 | 0.1894 | 0.85905 | 1.299 | 0.0059213 | 0.019355 | 0.087919 | 0.037258 | 0.028098 | 0.061586 | 0.019345 | 0.059449 | 0.0017933 | 4.0724e-05 | 9.9792e-05 | 0.00016609 | 0.0010413 | 3.3474e-05 | 0.89761 | 0.92609 | 193546.9848 | 196659.7437 | 249778.0868 | 230545.4449 | 174882.702 | 19965.2755 | 0.00015881 | 0.00044139 | 0.35727 | 0.59677
138 | 10_15_21_F_1 |  | 06/17/2023 02:06:28 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_1.787 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.3354 | 0.38403 | 0.01408 | 0.067827 | 1.0394 | 0.21733 | 1.7303 | 0.016345 | 0.0078135 | 0.0017327 | 0.19288 | 0.86644 | 1.3299 | 0.0057685 | 0.016955 | 0.086871 | 0.037506 | 0.028844 | 0.060181 | 0.019166 | 0.059557 | 0.0019644 | 3.5959e-05 | 0.00010272 | 0.00016933 | 0.00090781 | 4.8588e-05 | 0.89328 | 0.9301 | 197274.196 | 204001.5929 | 250903.5038 | 208673.008 | 182555.111 | 20812.4836 | 0.00019253 | 0.0004579 | 0.34703 | 0.67246
139 | 10_15_21_F_1_DUP |  | 06/17/2023 02:12:28 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_1_DUP.788 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.2726 | 0.35973 | 0.0097016 | 0.056454 | 0.97444 | 0.20217 | 1.6206 | 0.016052 | 0.0064145 | 0.0012079 | 0.18027 | 0.83249 | 1.2521 | 0.0056295 | 0.017366 | 0.080126 | 0.034677 | 0.028139 | 0.059438 | 0.018372 | 0.05758 | 0.0019846 | 3.5306e-05 | 8.164e-05 | 0.00017907 | 0.00101 | 4.7717e-05 | 0.85473 | 0.88409 | 192971.307 | 207345.9406 | 253887.5612 | 208702.7761 | 185122.4165 | 21524.871 | 0.00020602 | 0.0003731 | 0.0059531 | 0.010996
140 | 10_15_21_F_1_DUP |  | 06/17/2023 02:18:46 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_1_DUP.789 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.3033 | 0.36977 | 0.0099725 | 0.057001 | 0.98552 | 0.20622 | 1.6448 | 0.01402 | 0.0057442 | 0.0011978 | 0.18074 | 0.82732 | 1.265 | 0.0052206 | 0.017096 | 0.082445 | 0.034276 | 0.027783 | 0.059556 | 0.01705 | 0.057053 | 0.0019001 | 2.2267e-05 | 0.00012242 | 0.00014259 | 0.00095283 | 3.1367e-05 | 0.84588 | 0.88111 | 193906.7242 | 208955.1989 | 255447.0057 | 211483.2433 | 186161.798 | 21309.2176 | 0.00017108 | 0.00039651 | 0.0056771 | 0.010281
141 | 10_15_21_F_6 |  | 06/17/2023 02:24:45 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_6.790 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.2375 | 0.41881 | 0.062597 | 0.029759 | 0.77029 | 0.17063 | 1.5868 | 0.045073 | 0.0090588 | 0.0019923 | 0.28639 | 0.35508 | 1.9767 | 0.004188 | 0.013785 | 0.22497 | 0.093876 | 0.10439 | 0.06357 | 0.020128 | 0.061001 | 0.0019689 | 6.0763e-05 | 0.00014492 | 0.00033157 | 0.0010011 | 2.4977e-05 | 0.48076 | 0.49654 | 193767.2872 | 208603.3735 | 252372.5941 | 233819.3169 | 187628.8413 | 21143.6375 | 0.00028977 | 0.000669 | 0.009957 | 0.015919
142 | 10_15_21_F_6 |  | 06/17/2023 02:31:03 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_6.791 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.2118 | 0.41565 | 0.061544 | 0.029761 | 0.77032 | 0.17445 | 1.5903 | 0.043009 | 0.009262 | 0.0020668 | 0.28457 | 0.35509 | 1.9178 | 0.0050745 | 0.015154 | 0.2269 | 0.09189 | 0.10667 | 0.061644 | 0.019051 | 0.062512 | 0.0016057 | 7.4995e-05 | 0.00011234 | 0.00023429 | 0.00092874 | 6.1485e-05 | 0.48038 | 0.50144 | 191007.8457 | 204389.0758 | 248821.558 | 206274.5545 | 183409.9169 | 20870.5678 | 0.00030137 | 0.00080271 | 0.010095 | 0.018508
143 | 10_15_21_F_6_LD |  | 06/17/2023 02:37:02 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_6_LD.792 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.1975 | 0.40128 | 0.066159 | 0.051779 | 0.73408 | 0.16476 | 1.5439 | 0.041501 | 0.009504 | 0.0021728 | 0.23832 | 0.35482 | 1.6084 | 0.0056443 | 0.014858 | 0.22347 | 0.091335 | 0.1044 | 0.062374 | 0.019407 | 0.061438 | 0.0018464 | 3.9165e-05 | 0.00014152 | 0.00022119 | 0.00097548 | 5.4874e-05 | 0.48191 | 0.49784 | 191104.4572 | 204396.6525 | 249191.3654 | 211886.5064 | 182950.7 | 21143.6366 | 0.00031136 | 0.00076648 | 0.0097962 | 0.018007
144 | 10_15_21_F_6_LD |  | 06/17/2023 02:43:20 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_6_LD.793 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.2135 | 0.40822 | 0.065554 | 0.058232 | 0.75932 | 0.16997 | 1.5936 | 0.043608 | 0.0093983 | 0.0020604 | 0.24292 | 0.35363 | 1.6451 | 0.0061812 | 0.015018 | 0.22718 | 0.094229 | 0.1059 | 0.059442 | 0.020096 | 0.061472 | 0.0013643 | 3.9278e-05 | 0.00012587 | 0.00030416 | 0.00081403 | 4.8652e-05 | 0.49425 | 0.49818 | 198174.0775 | 203468.0756 | 243402.743 | 204426.2661 | 181156.1891 | 20796.4606 | 0.00034682 | 0.00072261 | 0.0093654 | 0.0183
145 | LBLANK |  | 06/17/2023 02:49:20 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.794 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | -8.1048e-05 | 0 | -0.0012106 | 0.0062548 | 0.013338 | 0.00024451 | 7.0661e-05 | -0.00012174 | 0.00043904 | 7.7244e-05 | -0.00031068 | -1.1195e-05 | 0.030858 | -4.2063e-05 | 0.00025072 | 0.0004842 | -2.9791e-05 | -0.00222 | 4.8654e-05 | 3.9994e-05 | 5.7097e-06 | 0.00026359 | 0 | 7.1943e-06 | -1.0885e-05 | 2.2357e-05 | -7.2887e-07 | -0.00034934 | 0.00013042 | 224135.2586 | 220899.4888 | 272027.0046 | 222536.1512 | 192963.1635 | 21982.9049 | -6.6811e-06 | -5.1397e-05 | -0.00011535 | -0.00012554
146 | CCV 1ppb TRACE |  | 06/17/2023 02:55:21 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.795 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.12291 | 0.057242 | 0.010739 | 0.010809 | 0.053675 | 0.077811 | 0.038725 | 0.086353 | 0.10755 | 0.024425 | 0.022827 | 0.043604 | 0.16513 | 0.18489 | 0.051849 | 0.11735 | 0.047326 | 0.019608 | 0.018283 | 0.0076252 | 0.020755 | 0.0032171 | 0.00041396 | 0.0014075 | 0.0028928 | 0.0019232 | 0.00052885 | 0.041378 | 0.043411 | 199252.2171 | 219038.4525 | 270028.9737 | 226232.675 | 189364.1506 | 21976.8917 | 0.020404 | 0.048023 | 0.11289 | 0.20666
147 | CCV 1ppb TRACE |  | 06/17/2023 03:01:39 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.796 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.12723 | 0.057138 | 0.010066 | 0.0087893 | 0.052523 | 0.074668 | 0.038762 | 0.082065 | 0.10452 | 0.024651 | 0.022489 | 0.044021 | 0.17802 | 0.19573 | 0.050523 | 0.11803 | 0.047867 | 0.019287 | 0.019589 | 0.0073287 | 0.021069 | 0.0027359 | 0.00038453 | 0.0015043 | 0.0031165 | 0.0016334 | 0.00053665 | 0.041671 | 0.042685 | 209653.5902 | 214659.4635 | 265138.5341 | 219450.7477 | 189853.2564 | 21880.7445 | 0.019557 | 0.047134 | 0.105 | 0.20637
148 | LFB |  | 06/17/2023 03:07:39 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.797 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.1721 | 0.51039 | 0.052793 | 0.11926 | 0.34127 | 0.71304 | 0.37632 | 0.76542 | 0.99002 | 0.22574 | 0.15242 | 0.42999 | 0.9868 | 1.7881 | 0.45329 | 1.0757 | 0.43086 | 0.20344 | 0.21324 | 0.07314 | 0.20687 | 0.024267 | 0.0049873 | 0.014775 | 0.031499 | 0.020553 | 0.005687 | 0.40301 | 0.41895 | 192954.7021 | 197904.1426 | 249662.7403 | 197932.1121 | 177883.3084 | 20665.605 | 0.19456 | 0.46503 | 0.98243 | 2.0068
149 | LFB |  | 06/17/2023 03:13:57 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.798 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.1525 | 0.52732 | 0.05427 | 0.12234 | 0.33531 | 0.72342 | 0.36356 | 0.77857 | 0.98929 | 0.22508 | 0.15393 | 0.42946 | 0.99563 | 1.8009 | 0.44364 | 1.0739 | 0.43443 | 0.20336 | 0.20918 | 0.070561 | 0.2064 | 0.024135 | 0.0045595 | 0.014914 | 0.031553 | 0.019966 | 0.0055249 | 0.39996 | 0.42158 | 195712.8745 | 199577.7699 | 250319.4151 | 205886.4019 | 177171.2535 | 20783.7753 | 0.19432 | 0.46706 | 0.95784 | 1.9288
150 | CCV 10 ppb MAJORS |  | 06/17/2023 03:19:56 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.799 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.221 | 0.54875 | 0.052519 | 0.098462 | 0.3458 | 0.7634 | 0.39355 | 0.82161 | 1.0407 | 0.2345 | 0.16387 | 0.44785 | 1.0649 | 1.8838 | 0.47087 | 1.1249 | 0.44886 | 0.21512 | 0.21279 | 0.077304 | 0.209 | 0.025058 | 0.0046054 | 0.014668 | 0.031521 | 0.020316 | 0.0059322 | 0.41646 | 0.43845 | 198631.3787 | 201385.5535 | 254420.2451 | 209643.7831 | 180958.9035 | 20809.817 | 0.19844 | 0.47273 | 0.99367 | 1.9891
151 | CCV 10 ppb MAJORS |  | 06/17/2023 03:26:14 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.800 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.2152 | 0.5442 | 0.053611 | 0.096122 | 0.3439 | 0.75335 | 0.38808 | 0.82448 | 1.0354 | 0.23821 | 0.16351 | 0.45336 | 1.0377 | 1.8891 | 0.47118 | 1.1134 | 0.45043 | 0.21687 | 0.20944 | 0.073039 | 0.21009 | 0.023911 | 0.004556 | 0.014496 | 0.031688 | 0.019213 | 0.0057366 | 0.42255 | 0.43936 | 203321.6472 | 202666.1177 | 254713.4193 | 211442.6284 | 182760.9652 | 21309.8854 | 0.19888 | 0.47708 | 0.97023 | 1.9856
152 | 10_15_21_F_6_DUP |  | 06/17/2023 03:32:14 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_6_DUP.801 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.1317 | 0.38185 | 0.039384 | 0.030259 | 0.72803 | 0.16252 | 1.5142 | 0.039988 | 0.0092243 | 0.0021503 | 0.15946 | 0.33376 | 1.1162 | 0.0059222 | 0.015308 | 0.22287 | 0.091415 | 0.10244 | 0.061153 | 0.019805 | 0.059832 | 0.0017338 | 4.7305e-05 | 9.6037e-05 | 0.0002466 | 0.00092285 | 3.6827e-05 | 0.47281 | 0.4863 | 186482.1123 | 197330.8742 | 245336.76 | 203163.7983 | 175566.9552 | 20475.9977 | 0.00037644 | 0.00075778 | 0.011272 | 0.021501
153 | 10_15_21_F_6_DUP |  | 06/17/2023 03:38:32 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_6_DUP.802 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.1445 | 0.3851 | 0.041376 | 0.030066 | 0.72679 | 0.16237 | 1.5209 | 0.039885 | 0.010373 | 0.0022319 | 0.15552 | 0.33769 | 1.1287 | 0.0062921 | 0.014237 | 0.22308 | 0.089909 | 0.10215 | 0.066754 | 0.019845 | 0.060405 | 0.001954 | 2.4189e-05 | 0.00011477 | 0.00031567 | 0.0010645 | 4.4128e-05 | 0.46052 | 0.49073 | 173656.9988 | 194276.2278 | 246234.6617 | 195043.7801 | 176240.4757 | 20234.3202 | 0.0003588 | 0.0008037 | 0.012497 | 0.022786
154 | 10_15_21_F_6_DUP_LSP |  | 06/17/2023 03:44:32 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_6_DUP_LSP.803 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.2629 | 0.88873 | 0.092258 | 0.14451 | 1.0131 | 0.90338 | 1.7753 | 0.85125 | 1.0271 | 0.23293 | 0.30904 | 0.75515 | 2.0491 | 1.8508 | 0.46589 | 1.2594 | 0.50935 | 0.29719 | 0.27247 | 0.089518 | 0.26173 | 0.024987 | 0.0046897 | 0.014649 | 0.031256 | 0.02093 | 0.0052353 | 0.84819 | 0.88878 | 186851.4267 | 190635.5384 | 237305.4434 | 196237.3754 | 172860.2828 | 19682.2178 | 0.1912 | 0.45904 | 0.97594 | 1.9492
155 | 10_15_21_F_6_DUP_LSP |  | 06/17/2023 03:50:52 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_6_DUP_LSP.804 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.4162 | 0.93346 | 0.10047 | 0.15279 | 1.0715 | 0.9557 | 1.9001 | 0.90375 | 1.1039 | 0.23607 | 0.31317 | 0.76939 | 2.1806 | 1.9701 | 0.49286 | 1.3373 | 0.51608 | 0.29988 | 0.27092 | 0.092605 | 0.26084 | 0.024931 | 0.0048262 | 0.014731 | 0.03069 | 0.020941 | 0.0056429 | 0.85673 | 0.88728 | 188808.8838 | 192455.744 | 237971.7547 | 197266.0206 | 174389.98 | 19127.4725 | 0.19514 | 0.46475 | 0.97033 | 1.9519
156 | 10_15_21_F_9 |  | 06/17/2023 03:56:52 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_9.805 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.0402 | 0.37978 | 0.029139 | 0.03603 | 0.58088 | 0.14802 | 1.3837 | 0.045894 | 0.011974 | 0.0029017 | 0.12646 | 0.20398 | 0.90685 | 0.0036599 | 0.01147 | 0.26456 | 0.10631 | 0.086925 | 0.058033 | 0.019686 | 0.058113 | 0.0013418 | 3.0413e-05 | 0.00011972 | 0.00022601 | 0.00077357 | 3.6775e-05 | 0.31042 | 0.32083 | 188708.3298 | 197495.1606 | 242370.1145 | 201409.8741 | 177979.1903 | 20403.2266 | 0.00024925 | 0.0004683 | 0.0095369 | 0.018039
157 | 10_15_21_F_9 |  | 06/17/2023 04:03:12 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_9.806 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.0463 | 0.37915 | 0.029114 | 0.030742 | 0.57578 | 0.14531 | 1.379 | 0.044464 | 0.011653 | 0.0028393 | 0.12536 | 0.20558 | 0.87942 | 0.0030755 | 0.01094 | 0.26102 | 0.107 | 0.084963 | 0.058794 | 0.0194 | 0.058019 | 0.0018157 | 3.0602e-05 | 0.00010336 | 0.00016996 | 0.00086279 | 5.0602e-05 | 0.30838 | 0.32008 | 187910.4881 | 196260.9633 | 242056.1358 | 200537.7646 | 176582.665 | 20184.9177 | 0.00029624 | 0.00049107 | 0.0097391 | 0.01839
158 | 10_15_21_F_10 |  | 06/17/2023 04:09:13 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_10.807 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.80704 | 0.41738 | 0.031961 | 0.04533 | 0.46262 | 0.13096 | 1.294 | 0.050635 | 0.014109 | 0.0032133 | 0.11369 | 0.2325 | 0.83377 | 0.0035182 | 0.016256 | 0.15671 | 0.064417 | 0.068516 | 0.064387 | 0.020303 | 0.067022 | 0.0016047 | 6.4654e-05 | 0.00011651 | 0.00022917 | 0.00073223 | 1.7741e-05 | 0.13453 | 0.14895 | 186444.4845 | 187557.0127 | 242890.1485 | 203588.5054 | 174248.9405 | 20114.1506 | 0.00015523 | 0.00030503 | 0.024465 | 0.044543
159 | 10_15_21_F_10 |  | 06/17/2023 04:15:32 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_10.808 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.79365 | 0.39971 | 0.033564 | 0.043595 | 0.45416 | 0.12733 | 1.2684 | 0.048184 | 0.013661 | 0.0036417 | 0.13462 | 0.21923 | 0.80476 | 0.002923 | 0.0157 | 0.1569 | 0.076586 | 0.08062 | 0.064906 | 0.020105 | 0.063204 | 0.0013243 | 4.4747e-05 | 0.00011168 | 0.00025164 | 0.00065681 | 4.7809e-05 | 0.15878 | 0.14104 | 185670.1133 | 193683.5081 | 216102.7213 | 199739.3711 | 175076.9819 | 20168.8991 | 0.0001432 | 0.00035835 | 0.024317 | 0.045532
160 | 10_15_21_F_10_DUP |  | 06/17/2023 04:21:32 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_10_DUP.809 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.79657 | 0.40994 | 0.035945 | 0.041052 | 0.44723 | 0.13355 | 1.3129 | 0.050927 | 0.010926 | 0.0021983 | 0.11015 | 0.21528 | 0.81588 | 0.0035549 | 0.0075689 | 0.14796 | 0.058459 | 0.055332 | 0.05477 | 0.019881 | 0.056365 | 0.0018755 | 2.0717e-05 | 8.503e-05 | 0.00015977 | 0.00061773 | 2.4043e-05 | 0.12878 | 0.13903 | 189429.7684 | 192149.0368 | 246129.6734 | 204096.3558 | 173793.4784 | 19722.9391 | 0.00015865 | 0.00024572 | 0.038839 | 0.069613
161 | 10_15_21_F_10_DUP |  | 06/17/2023 04:27:51 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_10_DUP.810 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.77794 | 0.40241 | 0.033381 | 0.040924 | 0.43878 | 0.12993 | 1.2894 | 0.048989 | 0.011116 | 0.0022413 | 0.11252 | 0.21432 | 0.80273 | 0.0026526 | 0.0078748 | 0.14355 | 0.058876 | 0.056187 | 0.053374 | 0.017558 | 0.055863 | 0.0012434 | 2.0757e-05 | 8.4766e-05 | 0.00016337 | 0.00076912 | 3.0898e-05 | 0.13331 | 0.13857 | 192713.4984 | 192140.8715 | 240841.9472 | 201098.9556 | 171937.2976 | 19877.1526 | 0.00015924 | 0.00025842 | 0.036658 | 0.07158
162 | LBLANK |  | 06/17/2023 04:33:51 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.811 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4.8185e-05 | 3.2408e-05 | -0.0012605 | -0.0012562 | 0.013635 | 0.0007523 | 0.00013765 | -0.00018307 | 0.00033076 | 7.0555e-05 | -0.00022109 | -3.1452e-06 | 0.015784 | -0.0001315 | 0.00057202 | 0.00026895 | 0.00013048 | -0.0023274 | -0.0010069 | -2.0341e-05 | -1.0186e-05 | 0.00017199 | 3.2376e-06 | 1.1108e-05 | -3.886e-06 | 5.1056e-05 | -3.7648e-06 | -0.00035115 | 6.7039e-05 | 199352.0712 | 204548.7001 | 259672.8039 | 212781.108 | 178507.6969 | 21126.9495 | -4.1091e-06 | -4.0149e-05 | -8.6491e-05 | -0.0001049
163 | LBLANK |  | 06/17/2023 04:40:11 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.812 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | -7.8866e-05 | 3.3111e-05 | -0.0013563 | -0.0015161 | 0.014311 | 5.1629e-05 | 1.277e-05 | -8.3737e-05 | 0.00042034 | 7.6301e-05 | -0.00021662 | -4.5305e-05 | 0.012851 | -0.00012962 | 0.00012316 | 0.00028028 | -4.5364e-05 | -0.0022288 | -0.00048618 | 1.0955e-05 | 1.6226e-05 | 0.00014411 | 0 | -1.8824e-06 | -1.6941e-05 | -1.3465e-05 | -3.7648e-06 | -0.00039623 | 3.7848e-05 | 199958.9003 | 202728.4426 | 258474.8407 | 210818.1775 | 177057.1467 | 20707.672 | -6.6811e-06 | -5.024e-05 | -9.0082e-05 | -8.1239e-05
164 | CCV 1ppb TRACE |  | 06/17/2023 04:46:11 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.813 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.12006 | 0.058541 | 0.0099571 | 0.0080117 | 0.047621 | 0.075856 | 0.039563 | 0.08316 | 0.10491 | 0.023684 | 0.021682 | 0.044403 | 0.15378 | 0.18253 | 0.051427 | 0.11625 | 0.046357 | 0.019304 | 0.019591 | 0.007528 | 0.020427 | 0.0034408 | 0.00042633 | 0.0014567 | 0.0030933 | 0.0021152 | 0.00053875 | 0.041169 | 0.042587 | 190370.8516 | 201549.1784 | 256809.8192 | 210705.9054 | 174259.0792 | 20409.9041 | 0.019824 | 0.047244 | 0.11682 | 0.21369
165 | CCV 1ppb TRACE |  | 06/17/2023 04:52:31 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.814 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.1201 | 0.054382 | 0.0096059 | 0.0081354 | 0.05354 | 0.077521 | 0.038183 | 0.082131 | 0.10498 | 0.024127 | 0.022716 | 0.045021 | 0.1657 | 0.19155 | 0.047912 | 0.11616 | 0.047267 | 0.019528 | 0.018133 | 0.0067143 | 0.020846 | 0.003988 | 0.00043797 | 0.0014633 | 0.003124 | 0.0017337 | 0.00055991 | 0.041352 | 0.043447 | 185118.4009 | 201034.6777 | 255695.4727 | 210696.4303 | 176147.9965 | 20547.4372 | 0.019931 | 0.048198 | 0.12158 | 0.21364
166 | CCV 10 ppb MAJORS |  | 06/17/2023 04:58:31 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.815 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.1668 | 0.52321 | 0.050615 | 0.094023 | 0.34089 | 0.74047 | 0.38504 | 0.788 | 1.0172 | 0.23151 | 0.15959 | 0.43971 | 0.97252 | 1.8002 | 0.45578 | 1.086 | 0.44114 | 0.21072 | 0.21204 | 0.072852 | 0.20766 | 0.024586 | 0.0045954 | 0.014289 | 0.031071 | 0.020373 | 0.0055135 | 0.41635 | 0.43519 | 186159.8284 | 184408.604 | 237596.2725 | 194845.1367 | 164992.6608 | 19545.362 | 0.19891 | 0.4746 | 1.0454 | 2.0749
167 | CCV 10 ppb MAJORS |  | 06/17/2023 05:04:50 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.816 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.2055 | 0.53167 | 0.053543 | 0.09158 | 0.34868 | 0.74303 | 0.38552 | 0.80074 | 1.0277 | 0.23053 | 0.15826 | 0.44363 | 1.0222 | 1.8487 | 0.47025 | 1.0916 | 0.43959 | 0.21172 | 0.21424 | 0.075461 | 0.20894 | 0.024586 | 0.0048221 | 0.014328 | 0.030979 | 0.019913 | 0.0056301 | 0.41663 | 0.43346 | 188959.2321 | 187782.8129 | 241779.6095 | 199005.2106 | 168244.2933 | 19477.2732 | 0.19826 | 0.47591 | 1.0351 | 2.0405
168 | 10_15_21_F_11 |  | 06/17/2023 05:10:50 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_11.817 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.71579 | 0.39552 | 0.031986 | 0.035207 | 0.4292 | 0.1232 | 1.224 | 0.047583 | 0.0056515 | 0.0012421 | 0.10413 | 0.18925 | 0.73141 | 0.0033097 | 0.0078178 | 0.14812 | 0.061054 | 0.081487 | 0.065991 | 0.020429 | 0.064518 | 0.0015006 | 4.6943e-05 | 9.9108e-05 | 0.00025708 | 0.00096074 | 5.3997e-05 | 0.11361 | 0.11954 | 182129.6387 | 184794.8906 | 238429.9551 | 198138.5998 | 167672.2531 | 19754.3164 | 0.00011348 | 0.00025651 | 0.0083313 | 0.015219
169 | 10_15_21_F_11 |  | 06/17/2023 05:17:09 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_11.818 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.73082 | 0.39737 | 0.035467 | 0.036029 | 0.44721 | 0.12202 | 1.2314 | 0.050749 | 0.0059009 | 0.0013281 | 0.10398 | 0.18801 | 0.72468 | 0.0024586 | 0.0080913 | 0.14736 | 0.06138 | 0.082442 | 0.064553 | 0.021592 | 0.063932 | 0.0013475 | 2.546e-05 | 0.00013623 | 0.00022676 | 0.00089428 | 3.2621e-05 | 0.11494 | 0.11947 | 180979.1681 | 183052.0541 | 235529.0909 | 229846.293 | 164905.6958 | 19309.0413 | 0.00014891 | 0.00030583 | 0.0081823 | 0.012169
170 | 10_15_21_UF_1 |  | 06/17/2023 05:23:09 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_1.819 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.3729 | 0.35543 | 0.0034671 | 0.049831 | 0.97347 | 0.20574 | 1.6579 | 0.013987 | 0.015738 | 0.0031027 | 0.021839 | 0.81543 | 0.24639 | 0.0050564 | 0.015872 | 0.059328 | 0.026564 | 0.03789 | 0.055203 | 0.018634 | 0.053822 | 0.0019132 | 1.4546e-05 | 6.7316e-05 | 0.00015774 | 0.0010942 | 1.0816e-05 | 0.85023 | 0.87971 | 181252.6531 | 183301.8357 | 233677.7124 | 193795.1117 | 164006.0642 | 19147.4905 | 0.00017578 | 0.0003603 | 0.0036232 | 0.0068777
171 | 10_15_21_UF_1 |  | 06/17/2023 05:29:28 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_1.820 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.4034 | 0.35789 | 0.0041345 | 0.052245 | 0.98989 | 0.21204 | 1.6892 | 0.014754 | 0.016696 | 0.0030847 | 0.02146 | 0.82847 | 0.26406 | 0.0055194 | 0.017965 | 0.059861 | 0.026864 | 0.037589 | 0.053263 | 0.01769 | 0.054264 | 0.001918 | 1.7862e-05 | 7.3432e-05 | 0.0001375 | 0.00097704 | 2.5138e-05 | 0.85171 | 0.88956 | 186684.5401 | 185737.0912 | 235209.185 | 194777.5356 | 166703.7364 | 19241.6168 | 0.00016907 | 0.00037715 | 0.0035381 | 0.006596
172 | 10_15_21_UF_6 |  | 06/17/2023 05:35:27 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_6.821 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.2567 | 0.3772 | 0.0041841 | 0.035158 | 0.73323 | 0.16528 | 1.5432 | 0.040311 | 0.0079976 | 0.0015905 | 0.034318 | 0.32661 | 0.33472 | 0.0040675 | 0.012552 | 0.13805 | 0.058356 | 0.081263 | 0.057012 | 0.018085 | 0.05692 | 0.0018472 | 4.2439e-05 | 0.00011785 | 0.00024456 | 0.00088926 | 3.8543e-05 | 0.4564 | 0.47743 | 182740.1365 | 188840.6959 | 236473.7996 | 196665.1468 | 168618.6759 | 19559.3821 | 0.00026114 | 0.00071307 | 0.0053561 | 0.0092522
173 | 10_15_21_UF_6 |  | 06/17/2023 05:41:46 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_6.822 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.222 | 0.37563 | 0.0046471 | 0.034414 | 0.71423 | 0.16322 | 1.4978 | 0.037749 | 0.0082336 | 0.0015268 | 0.033593 | 0.32744 | 0.31771 | 0.0037987 | 0.011538 | 0.14218 | 0.0582 | 0.080951 | 0.058217 | 0.020589 | 0.055632 | 0.0018253 | 3.1966e-05 | 0.00014383 | 0.00024584 | 0.00102 | 2.8158e-05 | 0.45079 | 0.47516 | 178991.4257 | 187702.4185 | 239357.2954 | 195930.4964 | 169390.3892 | 19699.5733 | 0.00035553 | 0.00068419 | 0.0053301 | 0.0096401
174 | 10_15_21_UF_8 |  | 06/17/2023 05:47:46 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_8.823 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.307 | 0.38958 | 0.011326 | 0.033073 | 0.70851 | 0.17304 | 1.5251 | 0.052244 | 0.011204 | 0.0018454 | 0.019775 | 0.33887 | 0.24323 | 0.0038994 | 0.012474 | 0.12989 | 0.054878 | 0.07323 | 0.072563 | 0.021821 | 0.068463 | 0.0017428 | 5.4131e-05 | 0.00011728 | 0.00022862 | 0.00099773 | 4.6843e-05 | 0.42042 | 0.44127 | 179865.2216 | 184609.844 | 234867.6569 | 193577.3932 | 166714.4945 | 19214.2467 | 0.00018639 | 0.0004375 | 0.0048061 | 0.0087991
175 | 10_15_21_UF_8 |  | 06/17/2023 05:54:04 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_8.824 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.3253 | 0.39825 | 0.011754 | 0.033859 | 0.71351 | 0.16879 | 1.5368 | 0.055312 | 0.011725 | 0.0019634 | 0.019467 | 0.33689 | 0.24282 | 0.0040605 | 0.011884 | 0.13198 | 0.054554 | 0.073726 | 0.069873 | 0.022334 | 0.067803 | 0.0020013 | 3.9795e-05 | 0.00010987 | 0.00028292 | 0.000714 | 3.2338e-05 | 0.42116 | 0.43924 | 180036.1256 | 184530.8636 | 233117.384 | 195151.8135 | 164856.4671 | 18947.2249 | 0.00017365 | 0.00040132 | 0.0050176 | 0.0088153
176 | 10_15_21_UF_8_LD |  | 06/17/2023 06:00:03 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_8_LD.825 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.3215 | 0.39746 | 0.019593 | 0.056915 | 0.7166 | 0.17046 | 1.5352 | 0.056038 | 0.01297 | 0.0022294 | 0.025322 | 0.31889 | 0.28441 | 0.0047425 | 0.012823 | 0.13588 | 0.05779 | 0.15657 | 0.069559 | 0.021644 | 0.065944 | 0.0018017 | 1.7702e-05 | 0.00010429 | 0.00021043 | 0.0011304 | 4.1841e-05 | 0.41846 | 0.41918 | 181436.2512 | 191076.7761 | 233074.6351 | 193066.5251 | 165331.2197 | 19024.6607 | 0.00019632 | 0.00044425 | 0.005608 | 0.010302
177 | 10_15_21_UF_8_LD |  | 06/17/2023 06:06:21 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_8_LD.826 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.3007 | 0.38356 | 0.020112 | 0.061869 | 0.69923 | 0.17274 | 1.5184 | 0.055197 | 0.011881 | 0.0022885 | 0.025599 | 0.33467 | 0.26924 | 0.0050403 | 0.013086 | 0.14189 | 0.058296 | 0.15806 | 0.068589 | 0.02243 | 0.068962 | 0.0016515 | 5.7998e-05 | 0.00014676 | 0.00025832 | 0.0011932 | 4.6975e-05 | 0.41901 | 0.43756 | 180026.5989 | 184000.7286 | 231985.3448 | 192938.8837 | 163386.3846 | 19237.6116 | 0.00019157 | 0.00041467 | 0.0053457 | 0.010161
178 | LBLANK |  | 06/17/2023 06:12:21 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.827 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | -1.0696e-05 | 6.5846e-05 | -0.0012506 | 0.0027514 | 0.018374 | -0.00011498 | 7.8325e-05 | -8.4893e-05 | 0.00028227 | 3.6354e-05 | -0.00020822 | -3.4574e-05 | 0.023294 | -0.00016273 | 0.00020582 | 0.00040077 | -5.1747e-05 | -0.0022141 | -0.0007151 | 0.00011135 | 0.00028334 | 0.00021048 | 0 | 1.4925e-06 | 3.2048e-06 | 3.5754e-05 | -3.8989e-07 | -0.00038827 | 0.00012992 | 200966.8193 | 198296.7837 | 254416.8174 | 208455.225 | 173353.5086 | 20285.7266 | 1.1773e-06 | -4.9919e-05 | -0.00012717 | -0.00012689
179 | LBLANK |  | 06/17/2023 06:18:40 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.828 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 5.4887e-05 | 3.2733e-05 | -0.00078454 | -0.00094536 | 0.014947 | 5.4171e-05 | 4.5135e-05 | -8.4841e-05 | 0.0002875 | 6.1506e-05 | -0.0001855 | -7.3218e-06 | 0.021803 | -3.0105e-05 | -2.3721e-05 | -5.8398e-05 | -4.0011e-05 | -0.0022207 | -0.0022626 | 1.3238e-05 | 0.00014976 | 0.000658 | 0 | 1.5066e-06 | 1.009e-05 | -1.6152e-05 | -3.7648e-06 | -0.0004519 | 7.3259e-05 | 190866.0026 | 197206.5071 | 250830.2545 | 209042.9933 | 170652.5948 | 20193.5976 | -4.0375e-06 | -4.9592e-05 | -9.2599e-05 | -0.00012162
180 | CCV 1ppb TRACE |  | 06/17/2023 06:24:40 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.829 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.11804 | 0.052368 | 0.010693 | 0.0080484 | 0.046864 | 0.074397 | 0.036399 | 0.083813 | 0.10668 | 0.024254 | 0.022476 | 0.044299 | 0.15131 | 0.18324 | 0.047419 | 0.11127 | 0.047106 | 0.019815 | 0.01876 | 0.0075536 | 0.021249 | 0.0030663 | 0.00051914 | 0.0014935 | 0.0029024 | 0.0018076 | 0.0006429 | 0.042492 | 0.04216 | 191272.9224 | 193909.8438 | 246706.6842 | 203618.8993 | 168932.3219 | 20154.2069 | 0.020107 | 0.04789 | 0.11504 | 0.21522
181 | CCV 1ppb TRACE |  | 06/17/2023 06:31:00 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.830 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.12082 | 0.054864 | 0.010177 | 0.0082945 | 0.054744 | 0.073972 | 0.038658 | 0.082317 | 0.10877 | 0.024077 | 0.022069 | 0.044772 | 0.16025 | 0.18758 | 0.048304 | 0.11818 | 0.046878 | 0.019578 | 0.017612 | 0.007904 | 0.020548 | 0.0030508 | 0.00040242 | 0.0014821 | 0.0030837 | 0.0019405 | 0.00057585 | 0.041544 | 0.044267 | 194706.8485 | 192286.1883 | 246594.7966 | 204136.1845 | 168369.8052 | 20022.6905 | 0.019656 | 0.047253 | 0.11194 | 0.21744
182 | LFB |  | 06/17/2023 06:37:00 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.831 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.1523 | 0.5015 | 0.052403 | 0.11799 | 0.34636 | 0.71079 | 0.37018 | 0.77031 | 1.0012 | 0.22554 | 0.15427 | 0.42476 | 0.96143 | 1.7682 | 0.44583 | 1.0698 | 0.43605 | 0.20748 | 0.21234 | 0.073465 | 0.20715 | 0.023473 | 0.0045458 | 0.01449 | 0.030854 | 0.020081 | 0.0055349 | 0.40517 | 0.42335 | 181138.4955 | 177307.5431 | 228563.6534 | 191457.111 | 160398.107 | 18649.5001 | 0.19643 | 0.47096 | 1.0053 | 1.9924
183 | LFB |  | 06/17/2023 06:43:18 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.832 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.139 | 0.4983 | 0.05121 | 0.11889 | 0.32963 | 0.70767 | 0.37099 | 0.76829 | 0.9832 | 0.22047 | 0.14951 | 0.42355 | 0.95976 | 1.7725 | 0.44797 | 1.0654 | 0.42431 | 0.19614 | 0.20233 | 0.072676 | 0.20377 | 0.022645 | 0.0047138 | 0.014458 | 0.031587 | 0.019486 | 0.0055011 | 0.39444 | 0.41404 | 187461.283 | 180219.6815 | 236882.5577 | 180644.0781 | 160326.0397 | 18918.5208 | 0.19167 | 0.45875 | 0.97604 | 2.1556
184 | CCV 10 ppb MAJORS |  | 06/17/2023 06:49:17 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.833 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.1958 | 0.53045 | 0.05157 | 0.095811 | 0.33888 | 0.7415 | 0.39113 | 0.81356 | 1.0255 | 0.23221 | 0.16434 | 0.44355 | 1.0138 | 1.833 | 0.45999 | 1.1072 | 0.44486 | 0.21367 | 0.20795 | 0.07342 | 0.20932 | 0.024047 | 0.0045768 | 0.014939 | 0.030391 | 0.019904 | 0.0056996 | 0.41747 | 0.43836 | 186718.8349 | 182251.8645 | 236351.1477 | 197875.063 | 163652.7113 | 19117.4505 | 0.19936 | 0.48017 | 1.0215 | 2.032
185 | CCV 10 ppb MAJORS |  | 06/17/2023 06:55:37 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.834 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.1977 | 0.52543 | 0.054389 | 0.094187 | 0.34479 | 0.74786 | 0.38186 | 0.80779 | 1.0369 | 0.2332 | 0.15961 | 0.45283 | 1.0355 | 1.8423 | 0.4635 | 1.1011 | 0.44476 | 0.21075 | 0.21409 | 0.075011 | 0.20853 | 0.023924 | 0.0048134 | 0.014896 | 0.030978 | 0.020283 | 0.0056796 | 0.42045 | 0.44083 | 187258.0078 | 183939.9481 | 240665.0947 | 197206.5111 | 164958.973 | 19572.0666 | 0.19772 | 0.47532 | 1.0385 | 2.075
186 | 10_15_21_UF_8_DUP |  | 06/17/2023 07:01:37 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_8_DUP.835 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.335 | 0.39825 | 0.0040037 | 0.032698 | 0.72975 | 0.17401 | 1.5668 | 0.055201 | 0.01396 | 0.0023113 | 0.032024 | 0.33081 | 0.33962 | 0.0038535 | 0.012217 | 0.14013 | 0.055693 | 0.056674 | 0.069544 | 0.021891 | 0.069101 | 0.0017949 | 3.3778e-05 | 0.00018945 | 0.0002984 | 0.0010047 | 4.5034e-05 | 0.42352 | 0.43796 | 175619.5521 | 177620.059 | 226994.0516 | 189489.8634 | 160156.7652 | 18140.8502 | 0.00018174 | 0.00041385 | 0.0051017 | 0.0097112
187 | 10_15_21_UF_8_DUP |  | 06/17/2023 07:07:56 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_8_DUP.836 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.266 | 0.37611 | 0.0036895 | 0.03444 | 0.70332 | 0.16843 | 1.5135 | 0.054274 | 0.012876 | 0.0023783 | 0.032241 | 0.3304 | 0.3111 | 0.0034513 | 0.011839 | 0.13354 | 0.054344 | 0.056679 | 0.069421 | 0.022283 | 0.066867 | 0.0016608 | 3.4069e-05 | 0.00011541 | 0.0002898 | 0.00078852 | 4.5496e-05 | 0.41636 | 0.43485 | 174848.8704 | 176115.6196 | 226146.1251 | 187589.0043 | 157829.0767 | 18461.9239 | 0.00024381 | 0.0004335 | 0.0052768 | 0.0095392
188 | 10_15_21_UF_8_DUP_LSP |  | 06/17/2023 07:13:55 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_8_DUP_LSP.837 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.3527 | 0.87143 | 0.061505 | 0.15835 | 0.9685 | 0.90457 | 1.766 | 0.85661 | 1.0372 | 0.2294 | 0.1873 | 0.7329 | 1.2793 | 1.8164 | 0.4579 | 1.1746 | 0.46932 | 0.23843 | 0.27015 | 0.095096 | 0.26486 | 0.024085 | 0.0047414 | 0.014367 | 0.030348 | 0.020549 | 0.0053565 | 0.79399 | 0.82957 | 175805.8329 | 172510.6873 | 221580.5565 | 182451.1219 | 153653.2783 | 17739.6751 | 0.19243 | 0.45839 | 1.0141 | 2.0002
189 | 10_15_21_UF_8_DUP_LSP |  | 06/17/2023 07:20:13 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_8_DUP_LSP.838 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.3568 | 0.84407 | 0.060607 | 0.15741 | 0.9499 | 0.88705 | 1.7489 | 0.85723 | 1.0198 | 0.23155 | 0.18562 | 0.73191 | 1.2579 | 1.8064 | 0.45557 | 1.1373 | 0.46853 | 0.23847 | 0.27732 | 0.090987 | 0.26278 | 0.024989 | 0.0047096 | 0.014255 | 0.031221 | 0.021258 | 0.0054918 | 0.79844 | 0.82867 | 174802.7817 | 175803.1371 | 224809.1359 | 184127.0192 | 158617.061 | 18410.5234 | 0.19391 | 0.46331 | 1.0263 | 2.0074
190 | 10_15_21_UF_9 |  | 06/17/2023 07:26:13 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_9.839 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.1149 | 0.35574 | 0.070648 | 0.072515 | 0.56771 | 0.14643 | 1.3625 | 0.043911 | 0.012357 | 0.0025551 | 0.016098 | 0.1884 | 0.20636 | 0.0030676 | 0.0093256 | 0.11442 | 0.048063 | 0.060578 | 0.05615 | 0.018288 | 0.05476 | 0.0019792 | 2.981e-05 | 0.00011718 | 0.00021366 | 0.00085967 | 6.3348e-05 | 0.29586 | 0.30742 | 176957.9175 | 179070.0154 | 228511.8481 | 189262.1326 | 161040.6892 | 18805.7037 | 0.00021535 | 0.00050099 | 0.0037393 | 0.0062304
191 | 10_15_21_UF_9 |  | 06/17/2023 07:32:31 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_9.840 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.1042 | 0.35896 | 0.068472 | 0.065877 | 0.57471 | 0.14229 | 1.3538 | 0.04235 | 0.01291 | 0.0024842 | 0.015567 | 0.18979 | 0.19875 | 0.003005 | 0.0085395 | 0.11187 | 0.048144 | 0.061249 | 0.054097 | 0.017794 | 0.055518 | 0.0015414 | 1.8842e-05 | 0.00011884 | 0.00017934 | 0.00072427 | 4.9103e-05 | 0.29608 | 0.30854 | 176461.9026 | 176592.1157 | 225989.6366 | 187774.694 | 157785.9396 | 18560.0513 | 0.00019672 | 0.00046611 | 0.0035771 | 0.0065874
192 | 10_15_21_UF_10 |  | 06/17/2023 07:38:31 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_10.841 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.88577 | 0.3974 | 0.0036982 | 0.048841 | 0.45512 | 0.12396 | 1.2801 | 0.048895 | 0.011361 | 0.0019905 | 0.012654 | 0.19736 | 0.18656 | 0.0026421 | 0.0085809 | 0.090941 | 0.0382 | 0.045842 | 0.061481 | 0.021351 | 0.06039 | 0.0016303 | 6.3489e-05 | 0.00011052 | 0.00024197 | 0.00073266 | 6.0066e-05 | 0.13075 | 0.13534 | 176045.412 | 177743.0057 | 226734.5776 | 195773.579 | 158987.1306 | 18311.7288 | 0.00014341 | 0.00021755 | 0.0037748 | 0.0070657
193 | 10_15_21_UF_10 |  | 06/17/2023 07:44:50 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_10.842 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.87387 | 0.38229 | 0.0029634 | 0.046599 | 0.44666 | 0.12776 | 1.2584 | 0.047475 | 0.010886 | 0.0020271 | 0.012385 | 0.19798 | 0.17218 | 0.002506 | 0.0067306 | 0.090024 | 0.038065 | 0.04658 | 0.060347 | 0.020931 | 0.060871 | 0.0017988 | 3.044e-05 | 0.00014246 | 0.00020697 | 0.00080097 | 2.2893e-05 | 0.13027 | 0.13594 | 176471.9983 | 175324.6821 | 225631.2105 | 210859.4073 | 156111.7018 | 18480.6137 | 0.00013812 | 0.00022782 | 0.0039428 | 0.0060344
194 | 10_15_21_UF_11 |  | 06/17/2023 07:50:50 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_11.843 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.84125 | 0.39427 | 0.0037298 | 0.039401 | 0.44462 | 0.12016 | 1.2504 | 0.046575 | 0.0083787 | 0.0013318 | 0.029749 | 0.17259 | 0.27602 | 0.002586 | 0.0059057 | 0.088757 | 0.037771 | 0.065107 | 0.06437 | 0.020682 | 0.063049 | 0.0011554 | 1.894e-05 | 0.00010417 | 0.00018746 | 0.00077752 | 6.0604e-05 | 0.11075 | 0.11731 | 173534.3168 | 176096.0343 | 226270.534 | 195551.5092 | 157579.0269 | 18452.5771 | 9.3563e-05 | 0.00030074 | 0.0041316 | 0.0070166
195 | 10_15_21_UF_11 |  | 06/17/2023 07:57:09 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_11.844 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.85819 | 0.39145 | 0.0044408 | 0.040997 | 0.44913 | 0.11769 | 1.2328 | 0.046496 | 0.0088309 | 0.0014809 | 0.029962 | 0.17527 | 0.27528 | 0.0027067 | 0.0079892 | 0.091483 | 0.038587 | 0.065756 | 0.062125 | 0.020834 | 0.063985 | 0.0015926 | 3.0936e-05 | 0.00014522 | 0.0002309 | 0.00077614 | 4.6493e-05 | 0.11048 | 0.11704 | 174433.8234 | 172161.851 | 223393.1077 | 183754.9223 | 153532.6838 | 18143.517 | 0.00011864 | 0.00028461 | 0.0042242 | 0.0076419
196 | LBLANK |  | 06/17/2023 08:03:10 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.845 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.3162e-05 | 3.5013e-05 | -0.001336 | -0.00091884 | 0.01664 | 0.0001815 | -2.0341e-05 | -0.00014806 | 0.00063869 | 4.2046e-05 | -0.00014865 | -6.9397e-06 | 0.022364 | -0.00016273 | 6.3918e-05 | 0.00045907 | -3.5263e-05 | -0.002269 | -0.00088444 | 1.4672e-05 | 0.00024784 | 0.00041067 | 0 | 1.6014e-05 | -2.5194e-06 | -1.0596e-05 | 3.4071e-06 | -0.00039281 | 6.0412e-05 | 189282.6021 | 185827.5427 | 244047.783 | 198986.36 | 161653.5274 | 19002.6303 | -6.6811e-06 | -4.5884e-05 | -0.00010217 | -0.00010197
197 | LBLANK |  | 06/17/2023 08:09:28 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.846 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | -0.00010797 | 0 | -0.0012356 | -0.0015513 | 0.013002 | 0.00046456 | 0.00011786 | -0.00014874 | 0.0004089 | 4.2386e-05 | -0.0002002 | -6.5186e-06 | 0.016337 | -2.5231e-05 | 0.00042963 | 6.3616e-05 | -3.4577e-05 | -0.0023041 | -0.002734 | 4.8408e-05 | 0.00025637 | 0.00090806 | 3.6228e-06 | 8.9533e-06 | 1.1311e-06 | -2.9651e-05 | -1.5832e-07 | -0.00046276 | 5.0045e-05 | 170330.3624 | 184724.6494 | 245022.0128 | 197140.9919 | 159579.0776 | 19329.0679 | 4.2946e-06 | -3.7935e-05 | -5.0998e-05 | -0.0001203
198 | CCV 1ppb TRACE |  | 06/17/2023 08:15:29 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.847 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.12247 | 0.05442 | 0.010778 | 0.0080841 | 0.057132 | 0.075643 | 0.038629 | 0.085341 | 0.10825 | 0.023671 | 0.022533 | 0.043771 | 0.16235 | 0.1942 | 0.048879 | 0.1173 | 0.046859 | 0.019074 | 0.017965 | 0.0081038 | 0.020943 | 0.0033587 | 0.00042017 | 0.0014082 | 0.0030809 | 0.0018524 | 0.000577 | 0.041634 | 0.043176 | 179175.3089 | 182495.5995 | 234603.3359 | 195131.6463 | 158877.9284 | 18716.9201 | 0.020633 | 0.048289 | 0.11989 | 0.22174
199 | CCV 1ppb TRACE |  | 06/17/2023 08:21:48 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.848 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.12179 | 0.05433 | 0.011287 | 0.0079781 | 0.053438 | 0.078289 | 0.035718 | 0.084027 | 0.10757 | 0.023835 | 0.022219 | 0.04482 | 0.163 | 0.18842 | 0.046071 | 0.11483 | 0.047192 | 0.019194 | 0.01909 | 0.0069661 | 0.02073 | 0.0033162 | 0.00049425 | 0.0014812 | 0.0032333 | 0.0018975 | 0.00054291 | 0.041563 | 0.043482 | 176383.5824 | 179315.0525 | 232796.8527 | 192144.925 | 158625.1506 | 18897.1576 | 0.020407 | 0.048291 | 0.11917 | 0.22086
200 | CCV 10 ppb MAJORS |  | 06/17/2023 08:27:48 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.849 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.1815 | 0.52314 | 0.054402 | 0.092669 | 0.35095 | 0.74714 | 0.38688 | 0.78448 | 1.0111 | 0.23195 | 0.1562 | 0.4417 | 0.99431 | 1.8128 | 0.45453 | 1.0945 | 0.43562 | 0.21046 | 0.21193 | 0.073674 | 0.2093 | 0.024782 | 0.0046942 | 0.014759 | 0.031521 | 0.020254 | 0.0056653 | 0.42067 | 0.43906 | 173223.8998 | 167701.2983 | 220099.9823 | 182311.2691 | 150532.3489 | 17870.5041 | 0.20035 | 0.48011 | 1.0721 | 2.135
201 | CCV 10 ppb MAJORS |  | 06/17/2023 08:34:07 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.850 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.2005 | 0.52486 | 0.052088 | 0.092418 | 0.35003 | 0.73276 | 0.38968 | 0.80578 | 1.0362 | 0.23145 | 0.15847 | 0.44586 | 1.0108 | 1.8286 | 0.45882 | 1.0976 | 0.4412 | 0.21094 | 0.19967 | 0.075388 | 0.21017 | 0.022676 | 0.0049692 | 0.014576 | 0.031363 | 0.02011 | 0.0056388 | 0.41899 | 0.44232 | 186395.9508 | 171068.8862 | 224956.7505 | 185295.31 | 153548.121 | 18033.3754 | 0.19942 | 0.47979 | 1.0031 | 2.1201
202 | 10_15_21_FB |  | 06/17/2023 08:40:08 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_FB.851 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.03293 | 0.00031003 | 0.049725 | 0.013694 | 0.019575 | 0.0012862 | 0.00042625 | -8.0629e-05 | 0.0091952 | 0.0017105 | 0.030179 | 0.0025436 | 0.21536 | 0.00024902 | 0.00049783 | 0.013166 | 0.0056067 | 0.15666 | -0.0023894 | 1.3771e-05 | 0.0002595 | 0.00066763 | 0 | 6.3356e-06 | -1.4231e-06 | -3.6029e-05 | -3.7648e-06 | -6.8499e-05 | 0.0004662 | 183467.2601 | 164794.7923 | 237451.8646 | 197293.7003 | 161769.4543 | 19349.0966 | 7.3072e-06 | 2.7532e-05 | 0.0059399 | 0.011047
203 | 10_15_21_FB |  | 06/17/2023 08:46:26 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_FB.852 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.031796 | 0.00034975 | 0.050324 | 0.013082 | 0.016911 | 0.0018943 | 0.00046763 | 9.6744e-05 | 0.0089943 | 0.0017834 | 0.030278 | 0.0021884 | 0.20576 | 0.00015351 | 0.00072336 | 0.012308 | 0.0056626 | 0.15714 | -0.0020152 | 1.4875e-05 | 0.00024242 | 0.00075385 | 3.7396e-06 | 5.5052e-06 | -5.8043e-06 | 3.3709e-05 | -3.7648e-06 | -0.00012188 | 0.00036687 | 180872.5465 | 179665.4286 | 233010.9727 | 195649.997 | 158807.1457 | 19061.3754 | -9.6731e-07 | 7.8881e-07 | 0.005639 | 0.010792
204 | 10_15_21_BB |  | 06/17/2023 08:52:26 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_BB.853 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.00017362 | 7.068e-05 | 0.0020339 | 0.00051385 | 0.012641 | 0.00019541 | 0.00019125 | -0.00011101 | 0.0025953 | 0.00051838 | 0.0031727 | 0.00013164 | 0.030616 | 0.00015696 | 0.00060508 | 0.0098025 | 0.0042063 | 0.0012251 | -0.0011464 | 0.00011827 | 8.5936e-05 | 0.00061243 | 3.697e-06 | 5.5205e-06 | -2.2212e-06 | 1.2067e-05 | -3.7648e-06 | -0.00043179 | 5.1601e-05 | 183020.9228 | 180738.1611 | 234684.1056 | 196004.1531 | 156876.0691 | 18827.7354 | -3.8142e-06 | -2.7956e-05 | 0.0022357 | 0.004417
205 | 10_15_21_BB |  | 06/17/2023 08:58:45 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_BB.854 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.00017559 | 0 | 0.0029103 | 0.00050841 | 0.014048 | -1.8042e-05 | 0.00026334 | -0.00018307 | 0.0024541 | 0.00053245 | 0.0033325 | 0.00019196 | 0.037531 | -0.00012724 | 0.00028212 | 0.01076 | 0.0040806 | 0.0012005 | -0.00089578 | -2.0341e-05 | 4.5679e-05 | 0.00035854 | 0 | 9.3007e-06 | 1.5291e-06 | -3.6534e-05 | -1.309e-08 | -0.00038246 | 5.1405e-05 | 180644.9835 | 180014.5071 | 236686.8142 | 194032.2516 | 156898.9173 | 18842.4185 | -6.6811e-06 | -3.6842e-05 | 0.002212 | 0.0044256
206 | 10_15_21_LAB_A |  | 06/17/2023 09:04:45 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_LAB_A.855 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.0002822 | 0 | 0.0018927 | 0.0001801 | 0.0084474 | -8.6719e-05 | 8.6045e-05 | -0.00018307 | 0.0049856 | 0.0010437 | 0.0056693 | 0.00030892 | 0.046612 | -0.00012754 | 3.6213e-05 | 0.011244 | 0.0043493 | 0.0016524 | -0.0016375 | -2.0341e-05 | 4.2549e-05 | 0.00041335 | 0 | -1.8824e-06 | -1.3203e-05 | 7.1314e-05 | -3.7648e-06 | -0.00023644 | 0.00021606 | 180190.6348 | 178764.8737 | 245485.3192 | 193558.2882 | 156125.2729 | 18809.0419 | -1.0362e-06 | -3.6018e-05 | 0.0025035 | 0.0046642
207 | 10_15_21_LAB_A |  | 06/17/2023 09:11:03 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_LAB_A.856 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.0002094 | 0 | 0.0019403 | 0.00024073 | 0.018273 | -9.4592e-05 | 0.00012058 | -7.7402e-05 | 0.0051518 | 0.0011236 | 0.0060111 | 0.00027025 | 0.04902 | -5.7243e-05 | 0.00013697 | 0.012331 | 0.004549 | 0.0019415 | -0.0016517 | 1.476e-05 | 3.1043e-05 | 0.00082432 | 0 | 1.8171e-06 | -5.7927e-06 | -3.8671e-05 | -3.7648e-06 | -0.00017126 | 0.00025267 | 179282.6427 | 179396.7392 | 235154.9706 | 192936.165 | 156928.5754 | 18933.2047 | -1.0284e-06 | -3.3792e-05 | 0.002647 | 0.0046453
208 | 10_15_21_LAB_B |  | 06/17/2023 09:17:03 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_LAB_B.857 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.00014283 | 3.6189e-05 | 0.0022411 | 0.0008566 | 0.013283 | 3.3601e-05 | 8.7679e-05 | -0.00018307 | 0.0023859 | 0.00044557 | 0.0029658 | 0.00012534 | 0.036705 | -5.5567e-05 | 0.00040398 | 0.012349 | 0.0069177 | 0.0044888 | -0.0016668 | 5.149e-05 | 9.7591e-06 | 0.00062423 | 0 | 5.6729e-06 | -5.5659e-06 | -0.00016097 | 5.5511e-08 | -0.00048397 | 3.429e-05 | 177619.3811 | 176058.2782 | 230011.567 | 190914.5945 | 153092.0135 | 18637.4837 | -3.7869e-06 | -5.0295e-05 | 0.0025535 | 0.0046407
209 | 10_15_21_LAB_B |  | 06/17/2023 09:23:20 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_LAB_B.858 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | -7.048e-05 | 0 | 0.0030003 | 0.0011654 | 0.0058905 | 0.00039181 | 8.5387e-05 | -7.5709e-05 | 0.0020275 | 0.00043839 | 0.0030157 | 9.6511e-05 | 0.030647 | -1.9914e-05 | 0.00011787 | 0.013801 | 0.004884 | 0.0044674 | -0.00055573 | -2.0341e-05 | 1.394e-05 | 0.0005322 | 3.7946e-06 | -1.8824e-06 | -1.6941e-05 | -0.00015373 | 6.7248e-08 | -0.00047336 | 3.0627e-05 | 177216.4264 | 174632.2434 | 232158.6703 | 192046.9696 | 152817.022 | 18630.809 | -9.889e-07 | -2.4627e-05 | 0.0026388 | 0.0047335
210 | 5_18_23_DB |  | 06/17/2023 09:29:20 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_DB.859 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.6781e-06 | 0 | 0.00090287 | 0.03428 | 0.011273 | -0.00014009 | 0.00012376 | -0.00018307 | 0.0004965 | 5.7488e-05 | 0.00072353 | 3.1455e-05 | 0.020408 | -1.9079e-05 | -0.00023904 | 0.010159 | 0.0041042 | -0.00042541 | -0.0018822 | -2.0341e-05 | 2.5204e-05 | 0.00034354 | 0 | -1.8824e-06 | -9.3667e-06 | -8.3117e-05 | -3.7648e-06 | -0.0005069 | 3.8593e-06 | 177957.0123 | 175151.9136 | 229255.5684 | 190128.9253 | 151968.1117 | 18532.0128 | -8.4731e-07 | -6.1861e-05 | 0.0027151 | 0.0050726
211 | 5_18_23_DB |  | 06/17/2023 09:35:39 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_DB.860 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | -3.4161e-05 | 0 | 0.00085609 | 0.045567 | 0.0075441 | 0.00030784 | 0.00012585 | -7.4089e-05 | 0.00079894 | 5.9989e-05 | 0.00086874 | 1.6568e-05 | 0.011626 | -5.403e-05 | -5.1723e-05 | 0.011277 | 0.0040508 | -0.00046794 | -0.00065249 | 1.564e-05 | 2.9921e-05 | 0.0004691 | 0 | 1.933e-06 | -1.3126e-05 | -9.9202e-05 | 3.7898e-06 | -0.00048678 | -3.5823e-06 | 175624.477 | 174104.6837 | 230217.568 | 192173.9529 | 153037.4124 | 18353.1159 | 2.0099e-06 | -3.5896e-05 | 0.0027389 | 0.0051156
212 | LBLANK |  | 06/17/2023 09:41:39 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.861 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | -0.00010549 | 0 | -0.0013628 | 0.010126 | 0.010611 | -3.4871e-05 | 5.3112e-05 | -0.00010996 | 0.00034494 | 4.1342e-05 | -0.00030996 | -1.3893e-05 | 0.017461 | -5.1762e-05 | -0.00040965 | 0.00024853 | -1.2492e-05 | -0.0022687 | -0.0015433 | 1.6648e-05 | 4.5117e-05 | 0.00020399 | 0 | 5.8576e-06 | -9.2489e-06 | 1.8871e-05 | 9.6859e-08 | -0.00043808 | 3.8801e-05 | 166693.3693 | 173101.7956 | 226292.9182 | 187967.1943 | 150968.8673 | 18109.4714 | -7.7227e-07 | -4.9957e-05 | -9.4926e-05 | -6.9286e-05
213 | LBLANK |  | 06/17/2023 09:47:59 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.862 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | -0.00010682 | 0 | -0.0012206 | 0.00018071 | 0.0052183 | 0.00055796 | 1.5946e-05 | -0.00018307 | 0.00057772 | 7.2887e-05 | -0.00025266 | -3.3075e-05 | 0.014105 | -9.0063e-05 | -0.00019233 | -0.00012567 | 3.1421e-06 | -0.0022444 | -0.0012012 | -2.0341e-05 | 1.0417e-05 | 0.00028155 | 3.8416e-06 | -1.8824e-06 | -1.6941e-05 | -0.00022448 | -3.7648e-06 | -0.00047122 | 4.249e-05 | 181900.4921 | 173570.7439 | 227585.8613 | 188469.7772 | 150483.1032 | 18361.7964 | -3.7458e-06 | -3.8322e-05 | -5.7841e-05 | -5.5743e-05
214 | CCV 1ppb TRACE |  | 06/17/2023 09:53:59 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.863 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.11458 | 0.053066 | 0.0083096 | 0.0085463 | 0.03673 | 0.075405 | 0.039048 | 0.081298 | 0.10371 | 0.023998 | 0.02154 | 0.045223 | 0.15475 | 0.18407 | 0.046714 | 0.11594 | 0.047041 | 0.019512 | 0.016628 | 0.0066721 | 0.020345 | 0.0032822 | 0.00046436 | 0.001559 | 0.0030202 | 0.0017973 | 0.00052611 | 0.041021 | 0.04381 | 172047.1632 | 170954.214 | 223223.9629 | 185156.9016 | 149846.4885 | 18022.6984 | 0.020096 | 0.048644 | 0.12033 | 0.22888
215 | CCV 1ppb TRACE |  | 06/17/2023 10:00:19 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.864 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.11837 | 0.052851 | 0.0090686 | 0.0079633 | 0.050156 | 0.077277 | 0.041057 | 0.084143 | 0.10701 | 0.024235 | 0.022331 | 0.045477 | 0.15523 | 0.18496 | 0.047724 | 0.11251 | 0.047419 | 0.019193 | 0.018114 | 0.0075087 | 0.021088 | 0.0030608 | 0.0004864 | 0.0013668 | 0.0029004 | 0.0018846 | 0.00060295 | 0.042085 | 0.043987 | 173312.9954 | 171369.0982 | 220720.6849 | 186393.4711 | 149362.7978 | 18163.5404 | 0.020681 | 0.049247 | 0.11939 | 0.22671
216 | LFB |  | 06/17/2023 10:06:18 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.865 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.105 | 0.48319 | 0.05338 | 0.11171 | 0.33216 | 0.70766 | 0.3696 | 0.76396 | 0.96716 | 0.21417 | 0.14603 | 0.41506 | 0.92514 | 1.7073 | 0.43404 | 1.036 | 0.41518 | 0.19357 | 0.20334 | 0.069396 | 0.20208 | 0.023329 | 0.0043546 | 0.013812 | 0.030896 | 0.019009 | 0.0056245 | 0.39489 | 0.41966 | 167212.9349 | 159085.5608 | 211967.1015 | 175396.1828 | 143686.9967 | 17333.1771 | 0.19267 | 0.46277 | 1.049 | 2.0824
217 | LFB |  | 06/17/2023 10:12:36 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.866 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.09 | 0.48406 | 0.051217 | 0.11631 | 0.3256 | 0.69072 | 0.36146 | 0.7436 | 0.93652 | 0.22241 | 0.14961 | 0.42251 | 0.9269 | 1.6949 | 0.43484 | 1.0383 | 0.42483 | 0.19736 | 0.20546 | 0.074502 | 0.20461 | 0.02378 | 0.0044078 | 0.014024 | 0.030599 | 0.019871 | 0.0057493 | 0.4044 | 0.42537 | 168531.8239 | 161963.7092 | 211751.6132 | 178718.3251 | 144880.3177 | 17822.444 | 0.19593 | 0.47318 | 1.0488 | 2.058
218 | CCV 10 ppb MAJORS |  | 06/17/2023 10:18:36 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.867 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.1676 | 0.51653 | 0.054695 | 0.094719 | 0.35119 | 0.74263 | 0.38913 | 0.80091 | 1.022 | 0.23066 | 0.15856 | 0.44336 | 0.99834 | 1.8149 | 0.44979 | 1.0907 | 0.44094 | 0.20684 | 0.2122 | 0.070422 | 0.20861 | 0.024153 | 0.0046606 | 0.014819 | 0.031047 | 0.019922 | 0.0056279 | 0.41684 | 0.44241 | 173663.9072 | 165362.9325 | 218880.2782 | 183155.9974 | 148057.3072 | 17829.122 | 0.19707 | 0.47545 | 1.0793 | 2.1335
219 | CCV 10 ppb MAJORS |  | 06/17/2023 10:24:55 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.868 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.1738 | 0.5103 | 0.052685 | 0.092301 | 0.35055 | 0.73238 | 0.38949 | 0.7997 | 1.0014 | 0.2334 | 0.157 | 0.44232 | 0.99781 | 1.8124 | 0.45785 | 1.0822 | 0.44361 | 0.20981 | 0.21258 | 0.071898 | 0.20695 | 0.023952 | 0.0046555 | 0.01436 | 0.030587 | 0.020726 | 0.0055434 | 0.42051 | 0.44387 | 175004.7591 | 167546.7906 | 219496.7395 | 187891.5718 | 151398.7252 | 18228.2921 | 0.1986 | 0.48073 | 1.0658 | 2.0819
220 | 5_18_23_DB |  | 06/17/2023 10:30:55 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_DB.869 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.00016918 | 0 | 0.0021498 | 0.036493 | 0.013921 | -9.4662e-06 | 0.00012039 | -7.8707e-05 | 0.0011863 | 0.00023524 | 0.024722 | 0.0003332 | 0.18151 | 0.00018606 | -0.00018681 | 0.010554 | 0.0037629 | -0.0003842 | -0.0014342 | 1.5078e-05 | 2.6708e-05 | 0.00066087 | 0 | 1.2865e-05 | -2.2357e-06 | -2.459e-06 | 7.255e-06 | 3.6192e-05 | 0.00010273 | 180432.4421 | 181122.9688 | 240309.8697 | 203881.6095 | 158121.6578 | 19149.4946 | -6.6811e-06 | -2.3589e-05 | 0.0027004 | 0.0047805
221 | 5_18_23_DB |  | 06/17/2023 10:37:14 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_DB.870 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.00010155 | 3.5116e-05 | 0.0015521 | 0.048221 | 0.011718 | 0.00020419 | 4.8973e-05 | -0.00014795 | 0.0012551 | 0.00018562 | 0.024562 | 0.00034862 | 0.17683 | -5.7381e-05 | -0.00022083 | 0.010573 | 0.0039861 | -0.00046085 | -0.0019534 | 4.8973e-05 | 4.1587e-05 | 0.00055542 | 3.6585e-06 | 1.8235e-06 | -1.3258e-05 | -0.0001197 | 3.6248e-06 | -0.0001384 | 2.5531e-05 | 183166.8723 | 181032.5228 | 237361.1006 | 196856.3968 | 156323.9866 | 19152.8313 | -3.8947e-06 | -5.1033e-05 | 0.0025242 | 0.0050347
222 | 5_18_23_DB |  | 06/17/2023 10:43:14 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_DB.871 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.2417e-05 | 0 | 0.0017787 | 0.050151 | 0.015277 | 0.00052743 | 0.00026047 | -0.0001134 | 0.001896 | 0.00037839 | 0.0027606 | 0.00015315 | 0.043614 | 0.01383 | 0.00030762 | 0.011442 | 0.0043124 | 0.00053868 | -0.0032711 | 4.9789e-05 | 4.3028e-05 | 0.0013683 | 0 | 5.619e-06 | -5.7307e-06 | 3.3568e-05 | -2.0557e-08 | -0.00027883 | 4.1024e-05 | 162678.0653 | 177862.465 | 234938.1038 | 193957.8415 | 155689.7924 | 19054.6997 | -1.0068e-06 | -3.9433e-05 | 0.0034926 | 0.0055268
223 | 5_18_23_DB |  | 06/17/2023 10:49:33 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_DB.872 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.00010602 | 0 | 0.0011187 | 0.048166 | 0.01578 | -0.00019254 | 8.6171e-05 | -0.00011195 | 0.0016116 | 0.00028727 | 0.0031247 | 0.00013627 | 0.043448 | 0.012981 | -0.00024469 | 0.01102 | 0.0044295 | 0.00035752 | -0.0010319 | -2.0341e-05 | 3.6368e-05 | 0.00078958 | 1.5077e-05 | 1.8654e-06 | -1.8432e-06 | -3.7131e-05 | -3.7648e-06 | -0.00033656 | 4.9292e-05 | 173898.7665 | 176212.8421 | 233088.8592 | 215892.0064 | 153733.5011 | 18778.3345 | 1.8796e-06 | -5.9164e-05 | 0.0030073 | 0.0045877
224 | 5_18_23_DB |  | 06/17/2023 10:55:33 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_DB.873 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.00021704 | 0 | 0.0013644 | 0.046615 | 0.01715 | 3.5155e-05 | 8.6575e-05 | -0.00014728 | 0.0010321 | 0.00012401 | 0.0015393 | 0.00024838 | 0.043616 | -1.9451e-05 | 4.8041e-05 | 0.0087339 | 0.003111 | 0.00055117 | -0.0011054 | -2.0341e-05 | 2.8424e-06 | 0.00044822 | 0 | -1.8824e-06 | 2.4121e-06 | -6.1972e-05 | -3.7648e-06 | -0.00032696 | 1.9581e-05 | 177857.0979 | 172906.0988 | 229777.2186 | 190276.1405 | 151666.2296 | 18581.4104 | 4.8948e-06 | -3.5774e-05 | 0.0019112 | 0.0035378
225 | 5_18_23_DB |  | 06/17/2023 11:01:51 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_DB.874 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.00017609 | 0 | 0.00091451 | 0.045052 | 0.014093 | 6.3052e-05 | 0.00015842 | -0.00018307 | 0.00073033 | 9.0346e-05 | 0.0014928 | 0.00019678 | 0.034712 | -5.3438e-05 | 0.00025576 | 0.0078984 | 0.0031746 | 0.00049108 | -0.0019252 | 8.7905e-05 | -2.0353e-05 | 0.00030303 | 0 | 1.7266e-05 | 2.2074e-06 | -6.2823e-05 | -3.7648e-06 | -0.00036302 | 1.5534e-05 | 177178.6385 | 173933.1065 | 227774.3647 | 189868.7608 | 149604.6124 | 18706.2423 | -3.7288e-06 | -3.2593e-05 | 0.001755 | 0.0038191
226 | 5_18_23_MFB |  | 06/17/2023 11:07:51 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_MFB.875 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.57491 | 0.24984 | 0.026613 | 0.10017 | 0.18392 | 0.36251 | 0.18851 | 0.39331 | 0.51422 | 0.11292 | 0.080395 | 0.22168 | 0.54334 | 0.90082 | 0.22478 | 0.55826 | 0.22106 | 0.097331 | 0.10109 | 0.036223 | 0.10407 | 0.012392 | 0.002461 | 0.0074186 | 0.015525 | 0.0096807 | 0.002708 | 0.19922 | 0.21765 | 169606.937 | 161533.7043 | 216413.0629 | 180530.2078 | 145646.8429 | 17403.261 | 0.095133 | 0.23028 | 0.5607 | 1.0876
227 | 5_18_23_MFB |  | 06/17/2023 11:14:09 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_MFB.876 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.56981 | 0.25723 | 0.02862 | 0.1 | 0.17221 | 0.35657 | 0.19076 | 0.38853 | 0.49395 | 0.11337 | 0.080701 | 0.21675 | 0.52518 | 0.88993 | 0.21917 | 0.54857 | 0.21988 | 0.097795 | 0.099792 | 0.035584 | 0.10107 | 0.012192 | 0.0024467 | 0.0069384 | 0.015251 | 0.010049 | 0.0026376 | 0.19968 | 0.21033 | 173987.1948 | 167143.5182 | 219026.2928 | 183715.8288 | 148100.3887 | 17964.6224 | 0.097009 | 0.23228 | 0.55066 | 1.0838
228 | 5_18_23_MFB |  | 06/17/2023 11:20:08 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_MFB.877 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.56534 | 0.24193 | 0.028314 | 0.10261 | 0.18139 | 0.35338 | 0.18497 | 0.38251 | 0.49576 | 0.10964 | 0.075341 | 0.21093 | 0.49233 | 0.87466 | 0.2215 | 0.54383 | 0.21618 | 0.094733 | 0.099702 | 0.032626 | 0.098134 | 0.012623 | 0.0023956 | 0.0071363 | 0.014726 | 0.0097047 | 0.0026612 | 0.19291 | 0.20476 | 172329.8817 | 170626.9733 | 223347.0644 | 168623.2992 | 152571.0843 | 18334.425 | 0.095122 | 0.22766 | 0.55301 | 1.1902
229 | 5_18_23_MFB |  | 06/17/2023 11:26:26 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_MFB.878 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.56006 | 0.25382 | 0.027226 | 0.10339 | 0.17768 | 0.35449 | 0.1854 | 0.38511 | 0.50365 | 0.11324 | 0.076606 | 0.2128 | 0.50168 | 0.88191 | 0.22793 | 0.55003 | 0.22083 | 0.096531 | 0.092186 | 0.035456 | 0.10064 | 0.01117 | 0.002244 | 0.0069928 | 0.014668 | 0.0088309 | 0.0025854 | 0.19996 | 0.20901 | 185179.2427 | 167839.5484 | 220168.9896 | 183578.0273 | 150692.6801 | 18120.1515 | 0.095591 | 0.22885 | 0.50347 | 1.0622
230 | LBLANK |  | 06/17/2023 11:32:26 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.879 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | -3.484e-05 | 3.5834e-05 | -0.0011511 | 0.011726 | 0.013539 | 0.00057301 | 1.5286e-05 | -3.9435e-05 | 0.00067144 | 5.916e-05 | -0.00018498 | -1.8732e-05 | 0.038704 | -9.1268e-05 | 0.00029783 | 0.00029941 | 1.0839e-05 | -0.0023357 | -0.0014661 | -2.0341e-05 | 4.7127e-05 | 0.00049833 | 0 | 5.677e-06 | 1.898e-06 | 1.0834e-05 | 1.0484e-08 | -1.9464e-05 | 6.7812e-05 | 183096.8862 | 176878.9632 | 232224.4978 | 191770.5901 | 153207.1672 | 18596.7635 | -3.7698e-06 | -2.755e-05 | -8.8235e-05 | -8.0571e-05
231 | LBLANK |  | 06/17/2023 11:38:51 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.880 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | -6.1197e-07 | 0 | -0.0011568 | 0.0011578 | 0.013203 | 0.00015904 | 1.5263e-05 | -0.00011186 | 0.0006878 | 4.4759e-05 | -0.00041452 | -2.6142e-05 | 0.024979 | -9.2162e-05 | -3.4952e-05 | 4.3689e-05 | -7.756e-05 | -0.0022615 | -0.0010686 | -2.0341e-05 | 9.6554e-06 | 0.00061688 | 0 | -1.8824e-06 | -1.7954e-06 | -2.6167e-05 | -3.7648e-06 | -0.00017716 | 3.7919e-05 | 181653.9068 | 176141.264 | 231901.7594 | 192738.8369 | 154062.961 | 18837.7455 | -6.6811e-06 | -4.7623e-05 | -8.7614e-05 | -8.8634e-05
232 | CCV 1ppb TRACE |  | 06/17/2023 11:44:51 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.881 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.11796 | 0.05486 | 0.011036 | 0.0098365 | 0.047018 | 0.072821 | 0.040526 | 0.081021 | 0.10671 | 0.024153 | 0.021967 | 0.045666 | 0.16856 | 0.18191 | 0.047978 | 0.11489 | 0.045195 | 0.019217 | 0.018164 | 0.0075068 | 0.021073 | 0.0030407 | 0.0004094 | 0.0014538 | 0.0032651 | 0.0021743 | 0.00055999 | 0.041292 | 0.043572 | 174742.9395 | 172655.0787 | 227959.8665 | 189577.6165 | 150964.8472 | 18596.7634 | 0.019781 | 0.047997 | 0.12061 | 0.22395
233 | CCV 1ppb TRACE |  | 06/17/2023 11:51:11 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.882 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.11845 | 0.055257 | 0.010392 | 0.0097034 | 0.05224 | 0.076641 | 0.038794 | 0.082986 | 0.10333 | 0.023332 | 0.021594 | 0.046701 | 0.16426 | 0.18589 | 0.049422 | 0.11654 | 0.046009 | 0.018829 | 0.018312 | 0.0079843 | 0.020436 | 0.0036569 | 0.00045498 | 0.0013904 | 0.0030894 | 0.0020361 | 0.00056759 | 0.040961 | 0.044205 | 173668.5873 | 171473.6434 | 226133.1703 | 187341.7412 | 151458.6845 | 18484.618 | 0.019595 | 0.048203 | 0.11962 | 0.22531
234 | CCV 10 ppb MAJORS |  | 06/17/2023 11:57:11 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.883 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.1608 | 0.50738 | 0.05156 | 0.092073 | 0.34004 | 0.74 | 0.38879 | 0.79608 | 1.0015 | 0.23245 | 0.15918 | 0.44526 | 0.98398 | 1.7843 | 0.45271 | 1.0802 | 0.44026 | 0.20793 | 0.20697 | 0.075532 | 0.20867 | 0.023944 | 0.0046679 | 0.014336 | 0.031524 | 0.020133 | 0.0055071 | 0.42094 | 0.44205 | 171142.452 | 160408.2368 | 212534.0175 | 177545.2355 | 143608.8938 | 17579.4768 | 0.19769 | 0.47637 | 1.0633 | 2.1425
235 | CAL BLANK |  | 06/17/2023 12:25:05 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CAL BLANK.884 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | [29 empty] | 0 | 0 | 0 | 0 | 0 | 0

## Net Intensities RSDs
| Sample Id | R | Acquisition Time | QC Status | Dataset File | Method File | Na KED 23/23 Helium KED_LowLow | Mg KED 24/24 Helium KED_LowLow | Al KED 27/27 Helium KED_LowLow | P Oshift 31/47 Oxygen DRC | S (34) -1 34/34 Helium KED_LowLow | K KED 39/39 Helium KED_LowLow | Ca KED 43/43 Helium KED_LowLow | V KED 51/51 Helium KED_LowLow | Cr (52) 52/52 Helium KED_LowLow | Cr (53) NH3 53/53 Ammonia DRC | Fe (54) NH3 54/54 Ammonia DRC | Mn Oshift 55/71 Oxygen DRC | Fe (56) KED 56/56 Helium KED_LowLow | Co KED 59/59 Helium KED_LowLow | Ni KED 60/60 Helium KED_LowLow | Cu (63) KED 63/63 Helium KED_LowLow | Cu (65) NH3 65/65 Ammonia DRC | Zn NH3 66/66 Ammonia DRC | As 75/75 | As KED 75/75 Helium KED_LowLow | As Oshift 75/91 Oxygen DRC | Se (77) 77/77 | Se (77) Oshift 77/93 Oxygen DRC | Se (78) Oshift 78/94 Oxygen DRC | Se (80) Oshift 80/96 Oxygen DRC | Se (82) 82/82 | Se (82) Oshift 82/98 Oxygen DRC | Mo (95) NH3 95/95 Ammonia DRC | Mo (95) OShift 95/127 Oxygen DRC | Rh (IS1) 103/103 (IS) | Rh (IS2) 103/103 (IS) Oxygen DRC | Rh (IS4) 103/103 (IS) Ammonia DRC | Rh (IS5) 103/103 (IS) Ammonia DRC | Rh (IS6) 103/103 (IS) Oxygen DRC | Rh (IS7) 103/103 (IS) Helium KED_LowLow | Cd (111) NH3 111/111 Ammonia DRC | Cd (114) NH3 114/114 Ammonia DRC | Pb 208/208 | Pb NH3 208/208 Ammonia DRC
1 | sac blank |  | 6/16/2023 11:58:56 AM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\sac blank .650 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 21.28% |  | 15.9% | 5.455% | 4.014% | 16.26% | 173.2% | 173.2% | 30.56% | 11.6% | 4.163% | 14.28% | 2.28% | 101.1% | 7.612% | 26.16% | 7.698% | 12.55% | 3.108% | 86.6% | 66.54% | 3.44% |  | 43.05% | 46.96% | 14.36% |  | 6.803% | 17.57% | 0.8211% | 1.123% | 1.118% | 1.066% | 2.031% | 1.658% | 86.61% | 26.03% | 9.653% | 13.69%
2 | sac blank cal |  | 6/16/2023 12:04:57 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\sac blank cal.651 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 23.5% | 173.2% | 25.86% | 3.136% | 2.118% | 33.93% |  | 92.3% | 14.27% | 28.8% | 4.109% | 42.6% | 2.122% | 64.37% | 19.3% | 42.09% | 17.9% | 4.567% | 1.407% | 99.53% | 49.84% | 2.871% | 86.61% | 49.32% | 125% | 54.96% | 0.7992% | 9.516% | 57.08% | 0.9443% | 0.8027% | 1.274% | 0.301% | 2.06% | 1.339% | 173.2% | 21.36% | 17.33% | 12.94%
3 | sac blank cal |  | 6/16/2023 12:10:57 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\sac blank cal.652 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 64.38% |  | 12.54% | 3.126% | 4.512% | 7.101% |  | 132% | 14.55% | 51.79% | 5.355% | 33.1% | 3.238% | 32.72% | 14.59% | 30.07% | 9.549% | 10.99% | 2.581% |  | 46.8% | 4.499% |  | 173.2% | 77.66% | 23.55% |  | 1.939% | 42.73% | 1.708% | 0.8716% | 0.9613% | 0.6692% | 0.5924% | 1.444% | 86.6% | 50.46% | 28.21% | 22.04%
4 | CAL BLANK |  | 6/16/2023 12:16:57 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CAL BLANK.653 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 50.02% |  | 13.37% | 6.303% | 4.852% | 17.27% | 173.2% | 66.12% | 30.31% | 10.02% | 9.663% | 18.23% | 1.655% | 93.59% | 16.81% | 26.47% | 15.7% | 4.727% | 4.106% | 173.2% | 42.95% | 4.192% |  | 173.2% | 66.56% | 13.58% | 86.6% | 14.84% | 99.9% | 0.6452% | 0.1482% | 0.7484% | 0.4378% | 0.8227% | 1.556% | 115.1% | 34.44% | 15.14% | 20.31%
5 | LOW 0.01 |  | 6/16/2023 12:22:58 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LOW 0.01.654 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 5.025% | 43.56% | 8.74% | 4.686% | 28.81% | 6.386% | 27.88% | 19.78% | 8.138% | 16.82% | 7.415% | 7.404% | 50.86% | 10.6% | 15.34% | 10.06% | 10.48% | 15.36% | 46.72% | 64.56% | 12.32% | 268.1% | 173.2% | 62.08% | 22.81% | 470.1% | 58.73% | 7.622% | 4.887% | 1.137% | 0.8423% | 0.9212% | 0.873% | 0.507% | 1.977% | 19.43% | 14.02% | 1.533% | 4.093%
6 | LOW 0.05 |  | 6/16/2023 12:28:58 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LOW 0.05.655 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 5.394% | 7.296% | 6.828% | 9.819% | 53.6% | 32.73% | 18.02% | 10.92% | 8.799% | 3.813% | 3.869% | 6.028% | 22.93% | 4.843% | 18.87% | 12.75% | 4.692% | 11.91% | 59.09% | 18.67% | 7.755% | 54.47% | 43.41% | 29.69% | 12.13% | 46.63% | 53.11% | 1.515% | 3.934% | 1.348% | 0.5856% | 1.247% | 0.3364% | 0.4161% | 0.9849% | 12.52% | 4.937% | 4.111% | 0.5748%
7 | LOW 0.1 |  | 6/16/2023 12:34:59 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LOW 0.1.656 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 6.411% | 6.176% | 707.6% | 3.359% | 121.4% | 16.45% | 2.986% | 7.133% | 14.73% | 2.891% | 0.4672% | 0.4631% | 9.451% | 13.85% | 12.46% | 10.31% | 4.091% | 3.182% | 46.94% | 24.38% | 5.045% | 16.45% | 65.92% | 9.52% | 12.15% | 9.369% | 56.28% | 1.375% | 3.092% | 0.7358% | 0.9433% | 1.578% | 0.8387% | 1.27% | 0.9627% | 1.814% | 4.791% | 5.015% | 0.9395%
8 | LOW 0.5 |  | 6/16/2023 12:40:59 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LOW 0.5.657 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 8.5% | 7.79% | 7.616% | 4.234% | 28.71% | 7.652% | 8.608% | 2.419% | 5.027% | 3.809% | 2.243% | 2.136% | 2.973% | 1.196% | 2.632% | 3.048% | 2.674% | 4.453% | 4.435% | 12.09% | 2.069% | 12.25% | 9.641% | 7.181% | 6.709% | 36.81% | 10.96% | 1.14% | 2.256% | 0.7953% | 0.8605% | 0.2624% | 1.017% | 2.533% | 0.759% | 2.451% | 1.076% | 0.7845% | 0.8135%
9 | STD 1 |  | 6/16/2023 12:46:58 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 1.658 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.613% | 2.614% | 5.776% | 1.269% | 8.832% | 2.69% | 7.465% | 5.349% | 1.467% | 1.761% | 1.148% | 1.218% | 1.017% | 1.35% | 3.935% | 3.144% | 3.477% | 3.669% | 7.707% | 16.95% | 2.534% | 12.77% | 17.24% | 6.087% | 2.336% | 9.205% | 13.28% | 2.799% | 2.532% | 0.3183% | 1.265% | 1.929% | 0.8632% | 1.834% | 0.8067% | 1.119% | 2.532% | 0.7859% | 1.881%
10 | STD 5 |  | 6/16/2023 12:52:58 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 5.659 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.784% | 1.728% | 4.627% | 0.5952% | 3.681% | 2.856% | 3.022% | 2.248% | 1.497% | 1.338% | 0.8646% | 1.117% | 0.9935% | 0.7165% | 2.254% | 1.252% | 2.021% | 1.106% | 0.9885% | 2.357% | 0.7316% | 3.725% | 7.26% | 3.535% | 0.1248% | 4.406% | 0.8257% | 1.31% | 1.547% | 0.536% | 1.333% | 1.775% | 1.024% | 1.027% | 1.792% | 2.81% | 1.445% | 1.655% | 0.3124%
11 | STD 10 |  | 6/16/2023 12:58:58 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 10.660 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.055% | 2.135% | 3.947% | 1.111% | 2.951% | 1.913% | 0.2306% | 0.5586% | 0.9864% | 0.2118% | 0.734% | 0.8017% | 1.746% | 0.9688% | 2.385% | 0.7701% | 0.4167% | 0.3109% | 0.7891% | 1.344% | 0.7174% | 2.749% | 3.723% | 2.31% | 1.872% | 1.639% | 4.194% | 0.5499% | 0.9027% | 1.287% | 0.3345% | 0.3715% | 0.8411% | 0.951% | 0.05589% | 1.029% | 0.1656% | 0.8505% | 1.657%
12 | STD 25 |  | 6/16/2023 1:04:58 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 25.661 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.7985% | 0.4825% | 1.912% | 15.29% | 1.857% | 1.405% | 0.09784% | 0.2777% | 1.453% | 0.9425% | 0.6778% | 1.584% | 0.8416% | 0.8121% | 0.4859% | 0.4295% | 1.483% | 0.9272% | 1.217% | 1.869% | 1.036% | 0.7324% | 2.785% | 0.7979% | 0.2029% | 0.6071% | 1.103% | 0.7384% | 0.618% | 0.4345% | 0.873% | 1.241% | 0.8751% | 12.49% | 0.2745% | 1.019% | 1.001% | 1.982% | 0.7443%
13 | STD 50 |  | 6/16/2023 1:10:58 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 50.662 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.28% | 0.7515% | 1.268% | 0.2055% | 1.985% | 1.425% | 0.2304% | 1.258% | 1.224% | 0.2961% | 0.1892% | 0.331% | 2.048% | 1.7% | 1.853% | 2.111% | 0.8869% | 0.8523% | 0.368% | 3.074% | 0.9183% | 0.7909% | 3.29% | 2.066% | 1.474% | 0.2983% | 2.865% | 0.5889% | 0.5153% | 0.3426% | 1.221% | 0.7681% | 0.7598% | 0.9539% | 1.873% | 0.7613% | 1.383% | 2.64% | 1.808%
14 | STD 100 |  | 6/16/2023 1:16:57 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 100.663 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.701% | 1.307% | 1.332% | 0.9243% | 0.4312% | 2.473% | 2.012% | 1.035% | 0.9789% | 1.074% | 0.3773% | 1.41% | 2.08% | 1.808% | 0.9255% | 0.4701% | 0.858% | 0.7931% | 0.7475% | 1.916% | 0.1657% | 0.4867% | 2.393% | 0.1669% | 0.5189% | 0.7125% | 0.8781% | 0.5236% | 0.4601% | 1.263% | 0.6764% | 0.9956% | 0.7512% | 0.532% | 0.3539% | 0.271% | 0.8062% | 2.834% | 0.7029%
16 | LBLANK |  | 6/16/2023 1:22:58 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.664 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 41.13% | 67.66% | 7.246% | 5.759% | 44.28% | 43.19% | 60.96% | 921.4% | 32.06% | 44.43% | 16.27% | 34.22% | 31.45% | 66.8% | 16.89% | 57.36% | 22.98% | 10.72% | 57.36% | 0% | 95.16% | 71.08% | 0.9771% | 31.51% | 4116% | 908.4% | 2.265e-06% | 1.567% | 20.59% | 0.9482% | 0.9773% | 0.6451% | 0.2886% | 1.639% | 1.666% | 58.64% | 110.4% | 77.48% | 749.6%
17 | LBLANK |  | 6/16/2023 1:41:50 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.665 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 27.31% | 173.2% | 577.2% | 19.85% | 111.8% | 71.67% | 3243% | 4761% | 17.75% | 44.33% | 94.16% | 151.4% | 715.3% | 260% | 93.49% | 196.5% | 275% | 620.1% | 264.9% | 0% | 17.9% | 874.2% | 86.6% | 71.42% | 533.1% | 280.8% | 2.265e-06% | 4.189% | 6.456% | 0.5292% | 0.7508% | 0.6203% | 0.286% | 0.6941% | 7.71% | 104.2% | 579.1% | 15.42% | 38.63%
18 | ICV PE |  | 6/16/2023 1:47:50 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\ICV PE.666 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.549% | 0.4419% | 2.199% | 1.045% | 1.039% | 1.398% | 1.262% | 2.059% | 0.5711% | 1.793% | 1.407% | 0.481% | 1.749% | 0.7222% | 1.62% | 0.3406% | 0.8917% | 0.6151% | 1.091% | 3.687% | 1.317% | 3.358% | 4.599% | 0.628% | 1.939% | 1.242% | 3.814% | 0.6344% | 1.031% | 0.2963% | 0.3461% | 1.092% | 0.656% | 0.4973% | 0.4802% | 1.364% | 1.29% | 0.3702% | 1.501%
19 | CCV 1ppb TRACE |  | 6/16/2023 1:53:50 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.667 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 25.46% | 173.2% | 8.774% | 2.846% | 1.219% | 61.57% | 1.903% | 57.87% | 78.63% | 77.32% | 10.31% | 12.31% | 6.59% | 151.5% | 53.34% | 197.4% | 41.49% | 8.29% | 2.88% | 1417% | 42.77% | 2.543% |  | 63.9% | 87.36% | 44.11% | 2.265e-06% | 4.824% | 608.2% | 2.49% | 2.482% | 2.464% | 1.203% | 7.071% | 1.19% | 304.4% | 37.21% | 18.2% | 14.63%
20 | CCV 1ppb TRACE |  | 6/16/2023 2:00:07 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.668 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.534% | 5.784% | 5.758% | 11.14% | 6.663% | 1.649% | 4.587% | 2.879% | 0.8274% | 0.06961% | 3.072% | 1.448% | 1.952% | 0.394% | 2.617% | 2.64% | 0.8567% | 2.66% | 6.565% | 2.559% | 1.809% | 11.31% | 7.716% | 5.365% | 4.02% | 4.012% | 4.908% | 1.126% | 3.055% | 1.609% | 2.152% | 0.7165% | 0.008299% | 6.066% | 1.123% | 2.269% | 2.106% | 2.012% | 0.9401%
21 | LFB |  | 6/16/2023 2:06:07 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.669 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9.343% | 9.796% | 10.87% | 0.4905% | 12.16% | 9.702% | 8.68% | 10.42% | 11.25% | 4.289% | 4.773% | 0.9218% | 12.24% | 10.29% | 10.32% | 10.62% | 3.286% | 4.289% | 1.47% | 14.66% | 1.092% | 2.552% | 1.711% | 2.934% | 2.56% | 2.869% | 3.761% | 5.366% | 1.34% | 0.4745% | 1.575% | 2.537% | 0.2438% | 1.024% | 7.511% | 4.806% | 4.648% | 0.5711% | 1.751%
22 | LLOQ 0.01 |  | 6/16/2023 2:12:05 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LLOQ 0.01.670 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 6.121% | 5.489% | 59.49% | 1.179% | 30.6% | 37.25% | 53.13% | 53.24% | 18.72% | 9.748% | 6.434% | 11.29% | 2.045% | 6.534% | 34.02% | 16.78% | 6.343% | 8.319% | 1653% | 84.65% | 9.01% | 48.16% | 100.5% | 29.11% | 1.313% | 176.4% | 188.4% | 2.681% | 9.671% | 0.9267% | 0.7248% | 0.4548% | 0.4669% | 0.7163% | 0.5086% | 15.06% | 14.34% | 8.613% | 2.804%
23 | CCV 10 ppb MAJORS |  | 6/16/2023 2:18:05 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.671 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.705% | 2.042% | 1.458% | 1.42% | 2.509% | 0.4968% | 1.979% | 0.7572% | 0.4382% | 1.003% | 0.9506% | 1.13% | 1.743% | 1.491% | 1.427% | 1.354% | 0.3777% | 0.8132% | 1.882% | 5.733% | 1.145% | 2.286% | 6.128% | 2.951% | 0.8824% | 3.275% | 2.879% | 1.265% | 1.507% | 0.5598% | 1.158% | 0.2376% | 0.9476% | 0.9511% | 0.4046% | 0.2624% | 0.3826% | 0.3777% | 1.141%
24 | LLOQ 0.1 |  | 6/16/2023 2:24:05 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LLOQ 0.1.672 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9.135% | 4.803% | 175.9% | 21.43% | 44.29% | 5.801% | 11.87% | 10.09% | 4.185% | 4.912% | 5.245% | 5.862% | 12.27% | 8.85% | 7.049% | 7.735% | 2.412% | 13.14% | 56.4% | 4.312% | 4.623% | 44.8% | 43.74% | 14.63% | 20.22% | 42.72% | 15.87% | 5.274% | 2.763% | 0.03196% | 0.558% | 1.709% | 0.535% | 0.9518% | 0.9947% | 4.882% | 3.717% | 2.503% | 1.079%
25 | 9_13_21_F_11_DUP |  | 6/16/2023 2:30:04 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_F_11_DUP.673 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.87% | 4.329% | 3.488% | 1.225% | 2.34% | 3.911% | 2.287% | 2.956% | 7.638% | 4.778% | 4.197% | 1.126% | 1.894% | 5.494% | 5.133% | 3.174% | 4.245% | 1.767% | 1.234% | 1.36% | 1.275% | 26.44% | 63.93% | 25.57% | 16.54% | 9.16% | 9.413% | 1.153% | 1.742% | 0.9009% | 1.324% | 2.934% | 0.765% | 0.9075% | 1.778% | 10.28% | 7.669% | 0.7524% | 2.405%
26 | LBLANK |  | 6/16/2023 2:47:08 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.675 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 88.17% | 173.2% | 105.7% | 7.384% | 26.43% | 101.5% | 139.8% | 1323% | 24.34% | 52.76% | 46.04% | 115.3% | 22.41% | 98.55% | 977.7% | 163.9% | 565% | 27.67% | 204.7% | 242% | 91.82% | 230.1% | 173.2% | 144.3% | 1733% | 414.7% | 264.8% | 7.65% | 40.75% | 0.3744% | 0.8687% | 1.274% | 1.54% | 2.42% | 1.249% | 71.19% | 225.7% | 12.99% | 34.26%
27 | 0.5 STD |  | 6/16/2023 2:53:09 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\0.5 STD.676 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 5.003% | 0.9287% | 17.54% | 2.139% | 15.57% | 3.977% | 3.118% | 2.635% | 2.502% | 2.9% | 1.759% | 2.973% | 4.252% | 3.879% | 1.602% | 3.306% | 3.445% | 1.983% | 4.368% | 14.97% | 4.806% | 20.15% | 26.93% | 13.36% | 11.67% | 5.698% | 10.84% | 1.733% | 1.259% | 0.8426% | 0.909% | 0.3572% | 0.3128% | 2.212% | 1.801% | 1.526% | 2.377% | 1.313% | 2.354%
28 | 0.1 STD |  | 6/16/2023 2:59:12 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\0.1 STD.677 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.015% | 2.891% | 77.67% | 6.717% | 24.31% | 8.172% | 12.46% | 3.368% | 4.681% | 4.117% | 5.217% | 2.436% | 26.06% | 3.406% | 16.32% | 16.73% | 2.575% | 5.727% | 16.03% | 16.77% | 4.499% | 87.47% | 19.44% | 37.75% | 19.71% | 68.57% | 48.01% | 1.535% | 6.522% | 2.147% | 0.2205% | 0.8267% | 0.4524% | 0.9052% | 0.3296% | 0.4024% | 2.982% | 4.307% | 3.005%
29 | CCV 10 ppb MAJORS |  | 6/16/2023 3:05:12 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.678 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.4366% | 2.685% | 9.209% | 0.6274% | 2.328% | 1.365% | 1.595% | 0.5378% | 0.7932% | 1.013% | 1.169% | 0.8408% | 1.013% | 1.478% | 1.743% | 1.756% | 1.333% | 2.026% | 1.995% | 1.769% | 0.9791% | 5.339% | 3.021% | 1.871% | 0.3716% | 3.439% | 0.965% | 0.3918% | 0.6338% | 1.111% | 0.4492% | 1.105% | 1.027% | 0.6158% | 0.921% | 0.9161% | 1.446% | 0.8593% | 1.517%
30 | 9_13_21_F_11_DUP |  | 6/16/2023 3:11:13 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_F_11_DUP.679 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.092% | 1.475% | 6.864% | 3.007% | 1.032% | 1.453% | 0.7652% | 3.143% | 7.055% | 4.867% | 1.298% | 1.763% | 0.4364% | 9.29% | 5.149% | 1.822% | 1.4% | 0.3487% | 0.6054% | 1.301% | 1.877% | 39.44% | 44.21% | 26.31% | 3.103% | 6.586% | 29.17% | 1.806% | 0.8245% | 0.5127% | 1.097% | 0.2767% | 0.5001% | 0.9798% | 0.2759% | 17.71% | 6.52% | 2.711% | 2.008%
31 | 9_13_21_F_11_DUP_LSP |  | 6/16/2023 3:17:13 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_F_11_DUP_LSP.680 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.064% | 2.296% | 4.284% | 0.8815% | 1.94% | 3.844% | 2.041% | 2.551% | 2.387% | 0.4907% | 0.9551% | 0.2349% | 2.447% | 0.1818% | 0.9012% | 1.015% | 0.9037% | 0.6647% | 2.064% | 6.485% | 0.3311% | 0.5873% | 4.784% | 0.3163% | 0.7201% | 2.308% | 4.202% | 1.605% | 0.6595% | 1.517% | 0.4821% | 0.4831% | 1.866% | 1.838% | 1.934% | 1.613% | 0.6047% | 0.4698% | 1.512%
32 | 9_13_21_UF_1 |  | 6/16/2023 3:23:13 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_UF_1.681 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.226% | 0.7033% | 4.157% | 0.3991% | 2.508% | 1.447% | 0.8772% | 3.158% | 16.94% | 6.445% | 1.786% | 0.6782% | 7.324% | 8.527% | 4.002% | 0.9735% | 1.452% | 4.559% | 2.358% | 9.115% | 1.378% | 30.87% | 8.507% | 7.527% | 11.45% | 3.386% | 29.89% | 0.7246% | 0.3864% | 0.281% | 0.7427% | 0.2023% | 1.003% | 2.241% | 0.5457% | 23.19% | 16.97% | 4.46% | 0.31%
33 | 9_13_21_UF_6 |  | 6/16/2023 3:29:13 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_UF_6.682 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.4631% | 2.541% | 11.69% | 4.039% | 1.831% | 1.279% | 1.548% | 2.28% | 11.23% | 6.158% | 1.905% | 1.201% | 2.373% | 4.965% | 8.828% | 2.213% | 1.725% | 1.589% | 2.695% | 1.964% | 0.5393% | 19.29% | 22.26% | 25.88% | 22.89% | 13.98% | 26.58% | 0.6887% | 1.121% | 1.183% | 0.7426% | 0.9056% | 2.44% | 0.7638% | 0.704% | 13.08% | 7.012% | 5.242% | 3.901%
34 | 9_13_21_UF_6_DUP |  | 6/16/2023 3:35:13 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_UF_6_DUP.683 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.458% | 2.573% | 15.27% | 0.9923% | 1.892% | 2.492% | 1.271% | 2.372% | 3.234% | 2.487% | 3.397% | 1.127% | 2.365% | 6.823% | 12.51% | 1.421% | 1.633% | 0.6518% | 0.2338% | 6.708% | 0.2503% | 5.585% | 35.65% | 20.07% | 0.6089% | 7.222% | 31.37% | 0.9976% | 0.9993% | 1.278% | 0.5713% | 0.4614% | 0.5448% | 1.12% | 1.845% | 24.44% | 9.259% | 2.667% | 0.8234%
35 | 9_13_21_UF_8 |  | 6/16/2023 3:41:13 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_UF_8.684 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.668% | 3.495% | 9.241% | 0.7875% | 2.547% | 2.038% | 1.419% | 3.879% | 9.637% | 2.429% | 3.825% | 0.941% | 1.941% | 4.378% | 15.92% | 1.687% | 1.039% | 0.09692% | 1.797% | 7.028% | 1.141% | 24.79% | 58.29% | 16.62% | 19.11% | 6.942% | 56% | 1.607% | 0.3971% | 1.747% | 0.2641% | 1.159% | 0.9784% | 1.742% | 0.5843% | 28.57% | 19.8% | 5.011% | 1.806%
36 | 9_13_21_UF_9 |  | 6/16/2023 3:47:13 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_UF_9.685 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.867% | 1.867% | 20.06% | 2.46% | 3.386% | 1.552% | 2.129% | 5.44% | 3.949% | 10.91% | 0.9448% | 0.2041% | 7.724% | 4.79% | 5.251% | 3.431% | 2.014% | 2.856% | 3.177% | 5.574% | 2.222% | 26.35% | 35.9% | 9.801% | 36.92% | 11.13% | 37.98% | 1.553% | 1.91% | 0.7681% | 1.131% | 1.681% | 1.3% | 1.13% | 1.956% | 33.88% | 20.03% | 3.908% | 2.64%
37 | LBLANK |  | 6/16/2023 3:53:13 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.686 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 97.92% |  | 21.16% | 6.619% | 119.3% | 72.32% | 146.2% | 68.19% | 334.3% | 57.5% | 133.6% | 300.9% | 28.24% | 38.31% | 354.4% | 126.6% | 69.17% | 4.046% | 81.49% | 33.85% | 7.592% | 27.99% |  | 40.76% | 2651% | 2284% | 2.265e-06% | 18.53% | 16.41% | 1.583% | 0.9188% | 0.8614% | 1.071% | 0.3249% | 1.125% | 75.83% | 21.63% | 65.05% | 34.23%
38 | CCV 1ppb TRACE |  | 6/16/2023 3:59:13 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.687 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9.69% | 12.12% | 4.809% | 1.043% | 25.29% | 6.787% | 8.881% | 11.05% | 13.3% | 1.383% | 2.976% | 1.547% | 20.39% | 13.66% | 14.99% | 9.645% | 0.7888% | 2.276% | 4.227% | 20% | 1.181% | 10.2% | 16.18% | 4.296% | 3.441% | 4.159% | 6.451% | 2.136% | 3.189% | 0.2891% | 1.007% | 1.056% | 1.099% | 1.01% | 7.83% | 1.537% | 1.674% | 1.469% | 0.346%
39 | CCV 1ppb TRACE |  | 6/16/2023 4:05:32 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.688 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.009% | 4.591% | 5.507% | 1.448% | 5.927% | 3.47% | 2.239% | 2.99% | 0.9036% | 1.681% | 4.494% | 0.1876% | 5.998% | 1.981% | 1.883% | 4.737% | 1.37% | 1.167% | 3.586% | 10.98% | 1.266% | 9.073% | 13.65% | 2.939% | 0.361% | 4.722% | 7.359% | 1.146% | 2.409% | 2.002% | 0.2311% | 1.067% | 0.8012% | 1.069% | 2.357% | 3.67% | 0.9021% | 1.439% | 0.9951%
40 | CCV 10 ppb MAJORS |  | 6/16/2023 4:11:33 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.689 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.544% | 2.383% | 3.048% | 1.697% | 4.874% | 4.327% | 1.933% | 2.621% | 1.892% | 0.6636% | 1.489% | 1.647% | 1.52% | 2.013% | 2.75% | 1.967% | 0.3061% | 0.6731% | 1.132% | 4.041% | 0.757% | 2.36% | 5.469% | 3.089% | 1.736% | 4.881% | 2.823% | 0.6093% | 0.6697% | 0.9888% | 0.5536% | 0.3194% | 0.3864% | 0.835% | 1.568% | 1.235% | 0.6601% | 0.9522% | 0.9133%
41 | CCV 10 ppb MAJORS |  | 6/16/2023 4:17:52 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.690 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.118% | 1.095% | 4.913% | 2.113% | 4.494% | 1.899% | 0.9169% | 2.681% | 1.102% | 1.538% | 1.172% | 0.9717% | 1.435% | 1.659% | 1.621% | 1.296% | 1.351% | 0.5049% | 1.308% | 2.169% | 1.621% | 0.2717% | 7.475% | 1.715% | 2.407% | 2.508% | 3.99% | 1.616% | 1.114% | 1.178% | 0.76% | 0.7053% | 1.42% | 0.6062% | 0.9535% | 1.158% | 0.8828% | 1.155% | 2.215%
42 | 9_13_21_UF_10 |  | 6/16/2023 4:23:53 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_UF_10.691 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.797% | 0.7524% | 9.813% | 1.333% | 2.264% | 4% | 0.9871% | 3.946% | 12.07% | 6.814% | 1.553% | 0.6607% | 4.934% | 13.08% | 7.213% | 1.511% | 2.171% | 4.014% | 1.803% | 4.645% | 0.4586% | 15.94% | 43.53% | 23.94% | 18.4% | 10.6% | 41.41% | 1.22% | 1.084% | 1.462% | 1.066% | 0.7566% | 0.5826% | 0.2239% | 1.46% | 18.96% | 21.04% | 2.479% | 3.012%
43 | 9_13_21_UF_11 |  | 6/16/2023 4:29:52 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_UF_11.692 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.6387% | 1.598% | 18.23% | 0.5235% | 0.6612% | 0.8846% | 0.5289% | 3.01% | 6.25% | 1.44% | 4.017% | 1.167% | 6.518% | 24.58% | 9.283% | 1.717% | 1.669% | 1.447% | 0.5325% | 2.796% | 1.707% | 55.1% | 24.24% | 28.52% | 13.91% | 13.06% | 16.44% | 0.9314% | 2.923% | 1.227% | 1.29% | 0.6789% | 3.544% | 0.4046% | 0.4105% | 13.1% | 14.26% | 3.604% | 4.479%
44 | 9_13_21_LAB_A |  | 6/16/2023 4:35:53 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_LAB_A.693 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 18.18% |  | 40.4% | 4.933% | 3.409% | 109% | 83.93% | 113.8% | 25.19% | 11.82% | 2.884% | 2.398% | 22.5% | 189.6% | 20.5% | 1.866% | 1.42% | 3.818% | 124.7% | 56.39% | 12.77% | 124.5% | 100.5% | 233.2% | 170.3% | 38.35% | 2.265e-06% | 18.4% | 67.27% | 1.869% | 0.8749% | 0.654% | 1.593% | 0.6699% | 0.573% | 77.34% | 88.76% | 3.574% | 2.51%
45 | 9_13_21_BB |  | 6/16/2023 4:41:52 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_BB.694 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 86.78% |  | 21.3% | 8.063% | 104.4% | 85.65% | 67.84% | 1209% | 19.82% | 26.57% | 2.566% | 16.33% | 21.11% | 138.9% | 27.2% | 10.05% | 4.742% | 14.27% | 1058% | 149.2% | 3.453% | 925.2% |  | 141.3% | 597.3% | 16.51% | 323% | 10.65% | 31.34% | 0.2672% | 0.4443% | 0.6803% | 1.461% | 0.6236% | 1.267% | 1540% | 360.5% | 5.364% | 6.941%
46 | 9_13_21_FB |  | 6/16/2023 4:47:52 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_FB.695 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 6.255% | 15.79% | 3.478% | 0.7358% | 21.39% | 23.3% | 5.52% | 31.93% | 11.72% | 8.799% | 3.276% | 1.592% | 14.73% | 17.01% | 6.614% | 6.571% | 2.438% | 0.5216% | 57.94% | 51.31% | 7.643% | 26.55% | 86.61% | 41.41% | 47.53% | 442.6% | 366.3% | 4.564% | 3.641% | 0.5812% | 0.8864% | 0.3722% | 0.5201% | 2.576% | 4.93% | 18.28% | 19.99% | 1.458% | 2.25%
47 | 5_18_23_MFB |  | 6/16/2023 4:53:52 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_MFB.696 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.312% | 2.93% | 3.196% | 1.203% | 3.638% | 1.233% | 6.703% | 0.5873% | 2.325% | 1.456% | 1.861% | 1.25% | 0.9987% | 0.6841% | 1.631% | 1.76% | 0.5904% | 0.7702% | 12.11% | 2.384% | 1.114% | 19.01% | 4.21% | 4.583% | 1.84% | 14.32% | 4.212% | 0.5202% | 0.7298% | 10.33% | 0.4279% | 0.3683% | 0.6024% | 0.7247% | 0.5646% | 0.9069% | 0.4858% | 13.42% | 1%
48 | 5_18_23_DB |  | 6/16/2023 4:59:52 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_DB.697 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4.246% | 34.01% | 7.665% | 1.033% | 17.67% | 16.83% | 24.36% | 42.95% | 19.09% | 7.782% | 7.079% | 21.43% | 4.399% | 5.589% | 68.31% | 6.47% | 3.424% | 114.2% | 33.18% | 101.2% | 14.18% | 70.26% | 119.5% | 38.41% | 13.79% | 607.4% | 180.8% | 51.16% | 1.633% | 1.234% | 1.118% | 1.078% | 1.644% | 0.7032% | 0.7146% | 10.31% | 16.27% | 3.789% | 5.079%
49 | 5_18_23_T_239 |  | 6/16/2023 5:05:52 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_T_239.698 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.132% | 3.192% | 3.569% | 1.294% | 18.62% | 4.459% | 3.346% | 3.957% | 7.592% | 2.354% | 1.686% | 0.849% | 3.214% | 3.945% | 1.368% | 4.043% | 2.607% | 0.8893% | 9.863% | 9.017% | 1.649% | 11.57% | 9.068% | 5.021% | 1.803% | 9.928% | 6.244% | 3% | 0.8838% | 6.404% | 0.5446% | 1.433% | 1.146% | 0.595% | 2.293% | 0.9839% | 2.157% | 8.31% | 2.473%
50 | 5_18_23_T_239 |  | 6/16/2023 5:11:52 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_T_239.699 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4.975% | 4.422% | 5.994% | 1.027% | 10.89% | 5.296% | 5.271% | 7.898% | 3.921% | 2.443% | 0.7334% | 1.084% | 6.075% | 7.702% | 5.33% | 6.87% | 1.707% | 0.6432% | 2.151% | 10.67% | 0.9165% | 1.81% | 11.91% | 6.084% | 4.108% | 5.765% | 14.47% | 1.906% | 3.27% | 1.134% | 1.015% | 0.3858% | 0.5237% | 0.1556% | 4.989% | 0.5031% | 0.6949% | 0.3759% | 0.6936%
51 | 5_18_23_M_224 |  | 6/16/2023 5:17:52 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_M_224.700 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.341% | 7.995% | 0.7817% | 1.488% | 5.484% | 4.492% | 3.222% | 2.619% | 5.811% | 9.377% | 3.299% | 1.373% | 15.09% | 17.19% | 29.06% | 8.852% | 5.95% | 0.8489% | 19.91% | 74.06% | 11.56% | 33.03% | 15.66% | 102.3% | 22.64% | 5.901% | 28.65% | 0.9433% | 2.111% | 0.8471% | 0.2931% | 0.3809% | 2.088% | 0.4269% | 1.043% | 15.55% | 47.14% | 12.67% | 4.691%
52 | LBLANK |  | 6/16/2023 5:23:52 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.701 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 62.36% | 173.2% | 13.17% | 4.054% | 190.3% | 4.171% | 136.9% | 250.7% | 90.29% | 78.49% | 55.36% | 39.29% | 60.84% | 326% | 50.19% | 90.99% | 1012% | 3.693% | 92.15% | 817.2% | 19.34% | 36.56% |  | 139.1% | 189% | 72.77% | 2.265e-06% | 12.8% | 41.33% | 7.484% | 1.298% | 3.372% | 1.101% | 0.909% | 1.297% | 0% | 28.29% | 31.88% | 10.45%
53 | CCV 1ppb TRACE |  | 6/16/2023 5:29:53 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.702 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.318% | 1.847% | 10.6% | 5.11% | 8.724% | 1.181% | 5.413% | 5.71% | 2.595% | 1.222% | 3.55% | 2.453% | 4.184% | 0.7635% | 10.5% | 2.28% | 1.98% | 1.786% | 0.6855% | 3.784% | 1.087% | 4.531% | 16.64% | 6.579% | 2.702% | 7.838% | 12.31% | 1.025% | 0.484% | 0.6882% | 0.8524% | 0.7412% | 1.292% | 1.373% | 0.2669% | 1.605% | 1.886% | 1.637% | 1.238%
54 | CCV 1ppb TRACE |  | 6/16/2023 5:36:12 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.703 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.149% | 1.163% | 20.61% | 2.416% | 11.94% | 1.726% | 7.527% | 3.56% | 5.343% | 3.101% | 1.694% | 0.9123% | 3.638% | 2.651% | 4.245% | 1.089% | 2.036% | 2.587% | 0.2953% | 16.67% | 1.453% | 17.96% | 4.674% | 5.798% | 7.38% | 7.319% | 14.19% | 2.302% | 2.984% | 1.197% | 0.7841% | 0.5756% | 0.9823% | 0.9679% | 2.215% | 1.762% | 1.764% | 1.629% | 0.3635%
55 | LFB |  | 6/16/2023 5:42:12 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.704 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.377% | 2.521% | 2.214% | 0.7193% | 3.204% | 1.419% | 1.507% | 1.374% | 0.5737% | 0.5818% | 0.6636% | 1.102% | 1.456% | 0.5441% | 0.9103% | 0.6597% | 0.7681% | 0.6412% | 1.761% | 3.017% | 0.962% | 4.535% | 3.839% | 3.355% | 0.369% | 4.764% | 3.39% | 0.8525% | 0.4637% | 0.8486% | 0.7544% | 0.7418% | 0.4742% | 1.207% | 1.048% | 1.629% | 1.533% | 0.7315% | 1.588%
56 | CCV 10 ppb MAJORS |  | 6/16/2023 5:48:11 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.705 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.9729% | 0.7618% | 2.925% | 2.152% | 0.5413% | 1.977% | 2.256% | 1.596% | 1.957% | 0.8776% | 1.398% | 0.1265% | 0.8207% | 1.807% | 1.36% | 1.065% | 0.6885% | 0.1166% | 1.277% | 1.527% | 0.61% | 2.035% | 2.426% | 1.878% | 1.027% | 2.457% | 2.364% | 1.008% | 0.314% | 1.081% | 0.4757% | 0.6096% | 0.7384% | 0.7501% | 1.019% | 0.3958% | 0.6549% | 1.586% | 2.123%
57 | CCV 10 ppb MAJORS |  | 6/16/2023 5:54:30 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.706 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.155% | 2.367% | 4.176% | 0.7853% | 1.117% | 1.122% | 2.779% | 0.6756% | 1.609% | 1.134% | 1.375% | 2.294% | 1.583% | 1.116% | 2.696% | 0.9457% | 0.4454% | 0.6214% | 2.964% | 0.2648% | 1.583% | 3.674% | 5.64% | 3.164% | 1.79% | 1.494% | 3.227% | 0.8932% | 1.623% | 1.002% | 1.19% | 0.2564% | 1.13% | 1.691% | 1.553% | 1.033% | 0.4199% | 2.245% | 1.691%
58 | 5_18_23_MFB |  | 6/16/2023 6:00:31 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_MFB.707 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.9324% | 2.942% | 2.206% | 0.5925% | 1.939% | 2.446% | 0.5602% | 1.913% | 3.347% | 13.22% | 13.08% | 0.3013% | 2.177% | 2.53% | 3.843% | 1.815% | 12.95% | 13.49% | 1.076% | 10.91% | 1.904% | 1.462% | 3.193% | 2.358% | 2.011% | 3.988% | 8.126% | 12.74% | 0.3504% | 1.004% | 0.5323% | 10.78% | 1.046% | 1.368% | 1.773% | 13.1% | 13.48% | 0.5489% | 0.3543%
59 | 10_1_21_W_1 |  | 6/16/2023 6:06:31 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_W_1.708 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4.736% | 2.695% | 5.73% | 2.161% | 0.6903% | 6.846% | 1.204% | 17.91% | 13.4% | 7.133% | 1.099% | 1.201% | 4.204% | 19.23% | 4.349% | 1.797% | 2.33% | 5.104% | 3.375% | 3.751% | 1.523% | 44.63% | 92.23% | 85.42% | 147.9% | 32.48% | 47.98% | 0.3417% | 1.241% | 0.8712% | 1.276% | 0.6219% | 0.2237% | 0.9858% | 2.059% | 40.47% | 75.45% | 3.61% | 7.649%
60 | 10_1_21_W_6 |  | 6/16/2023 6:12:31 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_W_6.709 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.358% | 1.093% | 4.875% | 1.235% | 2.309% | 4.565% | 1.661% | 5.632% | 18.77% | 14.94% | 3.244% | 0.5759% | 2.795% | 8.659% | 18.2% | 4.427% | 4.567% | 7.132% | 5.472% | 17.39% | 1.783% | 63.01% | 98.72% | 24.07% | 38.26% | 51.69% | 135.9% | 3.73% | 0.5656% | 0.9056% | 0.8356% | 2.591% | 1.434% | 1.621% | 0.7154% | 7.937% | 20.59% | 3.613% | 5.367%
61 | 10_1_21_W_8 |  | 6/16/2023 6:18:32 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_W_8.710 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.402% | 4.254% | 4.235% | 1.314% | 6.759% | 3.06% | 0.8418% | 15.62% | 26.53% | 3.835% | 1.038% | 0.3948% | 0.764% | 17.7% | 33.5% | 2.704% | 1.336% | 1.975% | 8.848% | 8.102% | 1.001% | 108.1% | 68.87% | 55.32% | 40.21% | 99.32% | 190.3% | 0.5932% | 1.472% | 1.643% | 0.409% | 0.09719% | 0.5937% | 0.6828% | 1.014% | 34.58% | 49.81% | 9.719% | 4.128%
62 | 10_1_21_W_9 |  | 6/16/2023 6:24:32 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_W_9.711 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.321% | 1.63% | 4.91% | 2.156% | 1.899% | 2.514% | 0.5194% | 15.13% | 10.52% | 4.998% | 2.215% | 0.8102% | 4.45% | 17.69% | 12.8% | 6.4% | 0.7462% | 4.266% | 11.57% | 9.118% | 2.377% | 1.732% | 86.6% | 62.84% | 20.61% | 57.85% | 128.7% | 1.659% | 1.302% | 0.8161% | 0.1883% | 1.016% | 0.9502% | 1.183% | 1.215% | 56.23% | 42.01% | 4.627% | 4.509%
63 | 10_1_21_W_10 |  | 6/16/2023 6:30:32 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_W_10.712 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.135% | 0.409% | 3.32% | 0.2588% | 3.944% | 14.3% | 0.4464% | 7.335% | 2.284% | 10.41% | 1.402% | 1.02% | 1.154% | 12.38% | 17.66% | 6.649% | 0.4781% | 3.704% | 5.814% | 4.28% | 1.928% | 11.38% | 86.62% | 21.94% | 68.11% | 118.4% | 188.3% | 2.776% | 1.843% | 1.056% | 1.327% | 1.253% | 0.1956% | 1.093% | 0.7761% | 37.28% | 13.82% | 2.876% | 0.8604%
64 | 10_1_21_W_11 |  | 6/16/2023 6:36:31 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_W_11.713 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.45% | 2.91% | 3.361% | 0.6973% | 5.677% | 4.471% | 4.139% | 7.096% | 8.821% | 0.6074% | 0.378% | 0.4668% | 0.9701% | 7.875% | 18.91% | 5.67% | 1.392% | 3.856% | 7.302% | 13.64% | 1.543% | 9.847% | 114.6% | 46.67% | 44.15% | 858.7% | 61.67% | 1.621% | 1.826% | 0.8303% | 0.9175% | 0.5402% | 0.6778% | 0.5215% | 1.719% | 30.55% | 23.71% | 2.27% | 1.453%
65 | 10_1_21_W_11_LD |  | 6/16/2023 6:42:31 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_W_11_LD.714 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 146.4% |  | 6.126% | 10.36% | 1.447% | 4.899% | 2853% | 20.76% | 58.63% | 155.1% | 14.11% | 34.28% | 6.398% | 61.92% | 21.18% | 72.69% | 50.34% | 6.951% | 2.397% | 329.7% | 113.7% | 2.082% |  | 2.265e-06% | 57.54% | 92.63% | 2.265e-06% | 8.3% | 515% | 3.855% | 2.406% | 1.592% | 7.767% | 3.211% | 5.794% | 63.47% | 22.2% | 8.759% | 7.134%
66 | 10_1_21_W_11_LD |  | 6/16/2023 6:48:48 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_W_11_LD.715 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.285% | 2.827% | 2.021% | 2.643% | 2.737% | 2.291% | 5.972% | 13.1% | 14.3% | 2.514% | 1.841% | 0.8981% | 1.546% | 36.24% | 25.5% | 2.188% | 2.631% | 1.106% | 2.281% | 10.39% | 2.757% | 24.33% | 91.72% | 52.92% | 22.47% | 7.29% | 291.1% | 4.136% | 1.131% | 1.642% | 0.2073% | 0.7535% | 0.251% | 1.196% | 0.2052% | 41.33% | 22.74% | 3.793% | 1.784%
67 | 10_1_21_W_11_DUP |  | 6/16/2023 6:54:48 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_W_11_DUP.716 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.992% | 4.337% | 3.212% | 3.396% | 4.791% | 6.591% | 2.484% | 7.327% | 20.96% | 3.994% | 0.6646% | 1.389% | 1.121% | 19.78% | 10.13% | 6.375% | 2.466% | 4.09% | 6.576% | 7.972% | 0.658% | 39.84% | 91.81% | 52.74% | 37.5% | 11.14% | 66.38% | 2.73% | 2.142% | 0.7506% | 0.3305% | 0.0735% | 0.6337% | 1.047% | 1.247% | 20.44% | 13.56% | 0.4327% | 0.6862%
68 | LBLANK |  | 6/16/2023 7:00:48 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.717 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 54.23% | 173.2% | 17.37% | 7.325% | 30.96% | 748% | 0% | 408.7% | 18.3% | 76.86% | 24.98% | 18.85% | 17.7% | 34.32% | 19.97% | 102.5% | 1341% | 2.481% | 170.7% | 0% | 561.9% | 131.3% |  | 588.3% | 32.37% | 405.3% | 134.2% | 9.67% | 62.23% | 1.322% | 1.149% | 1.42% | 0.7713% | 0.2646% | 1.73% | 84.93% | 23.27% | 12.01% | 21.27%
69 | CCV 1ppb TRACE |  | 6/16/2023 7:06:48 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.718 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 6.284% | 4.716% | 8.143% | 3.328% | 7.815% | 1.938% | 6.14% | 3.174% | 2.025% | 4.358% | 1.014% | 0.3137% | 4.333% | 4.579% | 2.015% | 1.211% | 2.111% | 1.344% | 5.627% | 13.25% | 0.86% | 9.974% | 18.04% | 2.059% | 6.548% | 10.48% | 23.69% | 1.858% | 1.621% | 0.2421% | 0.6822% | 0.3343% | 0.3424% | 1.82% | 0.8832% | 2.242% | 0.8252% | 1.249% | 1.281%
70 | CCV 1ppb TRACE |  | 6/16/2023 7:13:07 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.719 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.453% | 3.328% | 1.018% | 5.776% | 7.151% | 4.411% | 2.596% | 2.453% | 3.244% | 3.36% | 1.675% | 1.67% | 3.846% | 1.559% | 4.556% | 1.361% | 0.3688% | 1.803% | 2.598% | 12.14% | 1.409% | 8.074% | 19.22% | 6.754% | 8.348% | 11.8% | 14.42% | 2.297% | 0.9423% | 0.07654% | 0.387% | 1.265% | 1.57% | 1.173% | 0.4386% | 1.741% | 0.7201% | 0.8444% | 2.093%
71 | CCV 10 ppb MAJORS |  | 6/16/2023 7:19:08 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.720 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.2689% | 1.216% | 2.737% | 1.133% | 2.995% | 1.187% | 1.708% | 0.8203% | 1.175% | 0.3949% | 0.4681% | 0.7408% | 2.126% | 1.372% | 1.642% | 1.685% | 0.3235% | 0.8711% | 1.704% | 3.012% | 0.5141% | 5.743% | 4.27% | 2.688% | 2.158% | 4.036% | 3.412% | 1.514% | 0.1612% | 0.9196% | 0.458% | 1.032% | 0.6228% | 1.176% | 0.2518% | 1.223% | 1.082% | 2.323% | 0.853%
72 | CCV 10 ppb MAJORS |  | 6/16/2023 7:25:27 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.721 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.979% | 3.026% | 6.403% | 1.677% | 1.914% | 2.193% | 2.547% | 0.6023% | 2.091% | 1.827% | 1.579% | 0.7612% | 2.536% | 1.308% | 1.734% | 0.953% | 0.3514% | 0.7695% | 1.736% | 1.136% | 0.9204% | 1.076% | 2.75% | 2.122% | 1.329% | 2.722% | 4.197% | 0.5112% | 1.202% | 0.5563% | 0.8365% | 0.739% | 1.099% | 0.7696% | 1.348% | 1.634% | 0.3776% | 2.441% | 0.9254%
73 | 10_1_21_W_11_DUP_LSP |  | 6/16/2023 7:31:27 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_W_11_DUP_LSP.722 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.402% | 1.279% | 2.026% | 1.288% | 0.8883% | 1.24% | 0.9688% | 1.695% | 0.3673% | 1.604% | 0.6488% | 0.8618% | 0.3873% | 0.7274% | 0.9391% | 1.27% | 0.9034% | 0.6331% | 1.855% | 2.104% | 0.4513% | 1.072% | 3.616% | 3.033% | 2.28% | 2.79% | 0.4528% | 0.5079% | 0.9732% | 0.5045% | 0.874% | 0.3329% | 0.8878% | 0.4816% | 0.6117% | 0.2999% | 0.7454% | 1.212% | 1.903%
74 | 10_1_21_F_1 |  | 6/16/2023 7:37:27 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_F_1.723 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.591% | 0.9222% | 1.167% | 2.499% | 0.332% | 3.419% | 1.935% | 3.104% | 8.739% | 4.681% | 0.523% | 0.9396% | 1.788% | 2.015% | 5.739% | 5.861% | 1.119% | 0.7167% | 3.648% | 9.132% | 0.6829% | 20.86% | 78.61% | 47.18% | 44.69% | 1.574% | 28.57% | 1.515% | 0.9954% | 2.205% | 0.382% | 1.085% | 0.3892% | 0.6383% | 0.8666% | 24.7% | 7.964% | 2.679% | 3.732%
75 | 10_1_21_F_1_DUP |  | 6/16/2023 7:43:27 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_F_1_DUP.724 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.964% | 0.9929% | 4.715% | 0.9713% | 1.749% | 4.344% | 1.864% | 3.962% | 7.869% | 4.641% | 1.95% | 1.085% | 3.333% | 14.17% | 7.199% | 2.486% | 1.368% | 0.5957% | 1.077% | 7.545% | 1.184% | 28.05% | 78.97% | 15.32% | 14.44% | 19.39% | 24.14% | 0.9101% | 1.375% | 1.078% | 0.9386% | 0.4764% | 0.7746% | 0.9433% | 0.8129% | 17.57% | 26.94% | 2.187% | 1.19%
76 | 10_1_21_F_6 |  | 6/16/2023 7:49:26 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_F_6.725 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.595% | 1.663% | 0.7421% | 2.354% | 0.7971% | 1.218% | 1.084% | 2.982% | 13.08% | 5.375% | 1.32% | 0.8357% | 0.3085% | 16.19% | 9.734% | 2.947% | 0.6037% | 0.9525% | 0.6924% | 7.233% | 1.119% | 43.36% | 43.65% | 11.32% | 32.46% | 6.742% | 56.81% | 1.806% | 1.802% | 0.4811% | 0.5873% | 1.195% | 0.7977% | 1.239% | 0.628% | 14.14% | 11.54% | 6.682% | 1.606%
77 | 10_1_21_F_6_DUP |  | 6/16/2023 7:55:26 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_F_6_DUP.726 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.175% | 3.065% | 11.42% | 0.993% | 1.811% | 2.359% | 1.33% | 7.154% | 9.542% | 4.127% | 0.604% | 0.5288% | 0.5006% | 17.82% | 9.222% | 2.965% | 1.159% | 0.3584% | 1.586% | 4.396% | 0.8247% | 30.69% | 101.3% | 26.45% | 31.01% | 8.173% | 58.23% | 0.2191% | 0.6637% | 1.425% | 0.2286% | 0.498% | 0.5496% | 0.5948% | 0.9555% | 19.33% | 9.034% | 6.338% | 3.307%
78 | 10_1_21_F_8 |  | 6/16/2023 8:01:27 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_F_8.727 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.331% | 2.942% | 5.476% | 3.081% | 3.116% | 3.857% | 2.102% | 3.357% | 8.209% | 7.728% | 0.945% | 1.039% | 2.646% | 11.98% | 8.251% | 3.124% | 1.645% | 0.6631% | 1.974% | 3.031% | 2.198% | 9.459% | 54.21% | 16.71% | 12.4% | 16.83% | 23.71% | 1.898% | 0.6347% | 0.8553% | 0.4319% | 0.2814% | 0.9658% | 0.8528% | 2.239% | 46.87% | 3.8% | 3.052% | 1.322%
79 | 10_1_21_F_9 |  | 6/16/2023 8:07:27 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_F_9.728 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.092% | 0.9938% | 8.255% | 1.144% | 1.42% | 0.2774% | 2.545% | 5.805% | 8.872% | 4.269% | 0.4636% | 0.4159% | 1.396% | 13.83% | 11.76% | 2.268% | 0.224% | 1.446% | 2.537% | 7.299% | 2.517% | 5.957% | 32.77% | 6.558% | 21.76% | 5.949% | 40.98% | 0.5247% | 0.5453% | 1.202% | 0.5019% | 0.6072% | 0.4501% | 1.3% | 1.393% | 14.73% | 8.524% | 1.007% | 1.711%
80 | 10_1_21_F_9_DUP |  | 6/16/2023 8:13:27 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_F_9_DUP.729 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.698% | 1.338% | 4.112% | 3.603% | 1.373% | 4.842% | 1.094% | 3.474% | 10.25% | 3.173% | 1.16% | 0.07186% | 1.361% | 6.858% | 4.844% | 2.748% | 1.744% | 0.4749% | 1.019% | 7.697% | 0.1188% | 13.52% | 57.22% | 36.96% | 24.06% | 10.12% | 78.45% | 1.671% | 1.055% | 1.132% | 0.1082% | 0.6324% | 0.2692% | 1.257% | 0.7% | 17.75% | 3.987% | 3.431% | 3.191%
81 | 10_1_21_F_10 |  | 6/16/2023 8:19:27 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_F_10.730 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.047% | 3.074% | 1.697% | 2.785% | 0.8963% | 2.065% | 1.333% | 4.22% | 5.576% | 8.909% | 0.5237% | 0.8441% | 4.308% | 13.89% | 6.013% | 4.479% | 1.334% | 0.7484% | 0.6387% | 9.656% | 0.4012% | 8.771% | 41.11% | 31.21% | 25.68% | 10.6% | 59.4% | 0.7153% | 0.9913% | 0.8002% | 0.7045% | 0.2221% | 0.7774% | 0.7656% | 1.287% | 5.811% | 15.01% | 1.242% | 0.5703%
82 | 10_1_21_F_11 |  | 6/16/2023 8:25:27 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_F_11.731 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.9821% | 1.223% | 4.427% | 2.087% | 1.009% | 1.261% | 0.6213% | 3.487% | 6.563% | 1.624% | 0.5487% | 0.6693% | 1.165% | 10.87% | 9.995% | 1.598% | 1.68% | 1.31% | 2.717% | 2.56% | 1.223% | 25.33% | 32.69% | 19.47% | 26.2% | 14.87% | 30.08% | 0.4966% | 0.09071% | 1.547% | 0.4645% | 1.041% | 0.7159% | 0.54% | 1.854% | 8.935% | 11.12% | 0.7817% | 3.251%
83 | LBLANK |  | 6/16/2023 8:31:27 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.732 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 163.5% | 173.2% | 12.49% | 10.42% | 29.16% | 118.2% | 715.2% | 29.76% | 29.24% | 45.65% | 8.004% | 59.05% | 33.4% | 0% | 163.4% | 176.1% | 226.6% | 6.056% | 92.58% | 0% | 86.69% | 104.3% |  | 132.6% | 0% | 474.9% | 404.4% | 10.93% | 7.998% | 1.747% | 1.257% | 0.8401% | 0.8553% | 0.4416% | 1.881% | 0% | 16.54% | 66.56% | 24.44%
84 | CCV 1ppb TRACE |  | 6/16/2023 8:37:27 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.733 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 6.483% | 3.166% | 9.738% | 5.266% | 7.151% | 3.826% | 5.466% | 2.326% | 2.682% | 3.716% | 2.258% | 0.1425% | 5.124% | 0.8712% | 2.9% | 1.653% | 0.866% | 3.999% | 10.3% | 6.979% | 1.518% | 7.626% | 3.252% | 5.335% | 1.776% | 11.64% | 4.984% | 1.083% | 0.6756% | 1.783% | 0.6754% | 0.5846% | 0.6097% | 1.136% | 0.8885% | 1.336% | 2.063% | 1.705% | 1.263%
85 | CCV 1ppb TRACE |  | 6/16/2023 8:43:46 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.734 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.401% | 2.189% | 7.147% | 3.876% | 4.824% | 2.93% | 8.158% | 5.012% | 2.239% | 2.532% | 0.4264% | 0.8634% | 0.5535% | 0.8658% | 0.4691% | 5.033% | 0.8757% | 1.547% | 8.415% | 5.934% | 2.192% | 18.13% | 6.781% | 10.88% | 2.679% | 16.76% | 10.29% | 1.119% | 1.896% | 2.358% | 0.458% | 0.7415% | 1.106% | 0.4806% | 1.06% | 0.5633% | 0.1752% | 3.338% | 0.7485%
86 | LFB |  | 6/16/2023 8:49:46 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.735 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.16% | 1.854% | 5.655% | 0.787% | 2.586% | 3.288% | 1.124% | 0.3248% | 0.5298% | 0.8765% | 2.57% | 1.788% | 0.9173% | 0.8482% | 0.1846% | 1.195% | 1.337% | 1.241% | 4.468% | 1.39% | 1.755% | 4.598% | 4.447% | 3.118% | 2.125% | 5.943% | 3.355% | 1.476% | 1.294% | 2.448% | 0.8675% | 0.9842% | 0.7748% | 0.3062% | 0.6025% | 1.642% | 1.409% | 1.868% | 0.3544%
87 | CCV 10 ppb MAJORS |  | 6/16/2023 8:55:46 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.736 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.997% | 2.077% | 5.423% | 0.2403% | 2.218% | 1.183% | 1.589% | 0.6888% | 0.9238% | 1.414% | 0.7096% | 0.7985% | 0.8242% | 0.5786% | 0.9823% | 0.5906% | 0.9204% | 0.3708% | 1.534% | 1.629% | 0.5154% | 5.375% | 2.691% | 2.557% | 1.252% | 3.558% | 3.108% | 1.465% | 0.8864% | 0.9861% | 0.4187% | 0.4929% | 1.466% | 0.1154% | 0.2436% | 0.4793% | 0.3063% | 0.6554% | 2.008%
88 | 10_1_21_UF_1 |  | 6/16/2023 9:01:46 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_UF_1.737 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.719% | 2.527% | 15.86% | 1.677% | 0.8884% | 5.321% | 0.6359% | 10.23% | 12.57% | 12.38% | 3.755% | 1.953% | 5.106% | 9.175% | 4.392% | 5.235% | 1.916% | 0.9068% | 2.87% | 5.49% | 1.965% | 14.77% | 88.5% | 13.1% | 27.85% | 2.614% | 65.13% | 0.3539% | 1.205% | 0.6437% | 0.8076% | 0.3625% | 0.9473% | 1.633% | 1.092% | 5.693% | 17.08% | 5.447% | 3.078%
89 | 10_1_21_UF_6 |  | 6/16/2023 9:07:46 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_UF_6.738 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.106% | 2.261% | 7.614% | 0.9927% | 3.966% | 2.396% | 1.335% | 3.141% | 8.782% | 8.634% | 1.805% | 0.692% | 2.935% | 4.197% | 2.584% | 2.211% | 0.387% | 2.471% | 5.218% | 0.7776% | 2.221% | 44.54% | 34.29% | 24.14% | 16.13% | 5.474% | 60% | 0.7368% | 1.617% | 3.668% | 0.9077% | 0.2122% | 0.3631% | 0.7168% | 1.548% | 22.57% | 11.3% | 4.535% | 4.892%
90 | 10_1_21_UF_8 |  | 6/16/2023 9:13:46 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_UF_8.739 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.938% | 2.327% | 16.05% | 0.7867% | 1.878% | 4.486% | 0.7628% | 1.367% | 10.59% | 2.749% | 0.8693% | 0.9919% | 1.209% | 11.98% | 4.425% | 3.261% | 0.3484% | 1.267% | 0.4341% | 5.748% | 0.8099% | 5.507% | 114.4% | 20.18% | 13.53% | 25.35% | 29.75% | 0.1035% | 1.045% | 1.618% | 0.6212% | 0.3365% | 1.295% | 0.3257% | 1.388% | 19.28% | 14.48% | 9.288% | 2.251%
91 | 10_1_21_UF_8_LD |  | 6/16/2023 9:19:45 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_UF_8_LD.740 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.5251% | 1.959% | 2.391% | 0.5423% | 0.9722% | 1.535% | 1.918% | 6.95% | 3.828% | 1.594% | 3.089% | 0.9425% | 1.738% | 24.91% | 8.48% | 4.766% | 1.939% | 0.7625% | 2.272% | 5.619% | 1.229% | 11.7% | 66.57% | 27.11% | 22.38% | 14.58% | 33.17% | 0.8879% | 0.2118% | 0.1703% | 0.3394% | 0.2906% | 2.277% | 0.7436% | 0.7012% | 23.03% | 12.36% | 5.158% | 4.644%
92 | 10_1_21_UF_8_LD |  | 6/16/2023 9:26:03 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_UF_8_LD.741 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.201% | 2.301% | 13.47% | 1.96% | 1.271% | 2.356% | 0.4951% | 3.025% | 3.745% | 2.101% | 5.086% | 1.099% | 0.6072% | 6.596% | 3.56% | 2.229% | 2.252% | 2.148% | 1.554% | 6.595% | 2.176% | 12.33% | 77.97% | 29.12% | 21.42% | 8.75% | 63.85% | 2.674% | 0.8064% | 0.7377% | 0.7907% | 2.171% | 2.481% | 1.714% | 0.2614% | 16.35% | 11.98% | 3.778% | 4.521%
93 | 10_1_21_UF_8_DUP |  | 6/16/2023 9:32:02 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_UF_8_DUP.742 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.991% | 2.278% | 9.639% | 2.494% | 0.933% | 3.697% | 1.209% | 5.817% | 1.451% | 1.692% | 1.982% | 0.9763% | 4.178% | 21.28% | 10.25% | 0.4969% | 1.51% | 1.611% | 1.864% | 4.623% | 0.7094% | 34.83% | 39.69% | 24.89% | 17.6% | 6.037% | 42.15% | 1.862% | 0.3897% | 0.5614% | 0.3978% | 0.9743% | 1.165% | 1.111% | 0.7187% | 16.27% | 10.9% | 5.808% | 2.76%
94 | 10_1_21_UF_8_DUP_LSP |  | 6/16/2023 9:38:02 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_UF_8_DUP_LSP.743 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.17% | 1.184% | 4.87% | 1.435% | 0.8929% | 0.9759% | 2.76% | 0.8851% | 1.585% | 1.898% | 1.525% | 0.4342% | 0.8219% | 1.962% | 1.444% | 1.757% | 0.9359% | 1.214% | 0.8798% | 3.932% | 0.842% | 1.823% | 6.419% | 3.741% | 0.9892% | 1.749% | 2.089% | 0.5316% | 0.7296% | 0.5751% | 0.3661% | 0.8009% | 0.1717% | 1.2% | 0.8372% | 1.542% | 1.122% | 1.319% | 1.141%
95 | 10_1_21_UF_9 |  | 6/16/2023 9:44:02 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_UF_9.744 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.6274% | 2.372% | 12.42% | 1.216% | 1.433% | 2.751% | 1.609% | 1.895% | 6.457% | 12.83% | 2.386% | 2.162% | 2.258% | 7.255% | 2.904% | 4.531% | 1.129% | 0.9827% | 1.583% | 3.699% | 2.93% | 26.47% | 30.87% | 25.47% | 32.69% | 6.51% | 19.86% | 0.4942% | 1.631% | 1.32% | 0.8741% | 0.6536% | 0.8234% | 0.3847% | 1.231% | 8.643% | 18.74% | 7.189% | 7.76%
96 | 10_1_21_UF_10 |  | 6/16/2023 9:50:03 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_UF_10.745 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.52% | 0.9931% | 13.77% | 1.648% | 2.032% | 3.355% | 1.156% | 5.737% | 6.18% | 4.029% | 1.84% | 2.168% | 1.423% | 17.2% | 21.78% | 5.107% | 0.8545% | 1.148% | 6.048% | 5.592% | 1.504% | 32.15% | 12.79% | 34.61% | 29.7% | 10.42% | 27.28% | 0.9222% | 1.824% | 3.721% | 1.277% | 0.9107% | 1.123% | 0.3988% | 0.4589% | 42.93% | 16.96% | 10.49% | 1.72%
97 | 10_1_21_UF_11 |  | 6/16/2023 9:56:03 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_UF_11.746 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.284% | 0.3752% | 16.07% | 1.86% | 1.966% | 3.988% | 2.295% | 3.155% | 13.01% | 13.79% | 1.003% | 0.6564% | 3.068% | 6.378% | 5.305% | 2.479% | 0.7799% | 0.4675% | 6.415% | 7.982% | 1.193% | 21.02% | 44.79% | 7.283% | 12.62% | 24.83% | 17.83% | 0.3982% | 1.031% | 4.674% | 0.5039% | 0.4325% | 1.185% | 0.9599% | 1.006% | 23.61% | 10.91% | 0.225% | 2.419%
98 | LBLANK |  | 6/16/2023 10:02:04 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.747 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 157.8% | 86.61% | 54.19% | 12.03% | 30.87% | 193.3% | 212.1% | 37.73% | 29.64% | 23.35% | 76.41% | 26.49% | 7.724% | 89.91% | 74.12% | 72.84% | 296.7% | 3.005% | 253.9% | 270.9% | 9.251% | 58.24% | 173.2% | 40.64% | 148.9% | 73.2% | 2.265e-06% | 9.108% | 60.04% | 2.235% | 0.6251% | 0.5635% | 0.4907% | 0.4222% | 0.8836% | 4039% | 84.71% | 3.385% | 6.34%
99 | CCV 1ppb TRACE |  | 6/16/2023 10:08:04 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.748 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.823% | 4.302% | 9.179% | 4.464% | 4.02% | 1.578% | 1.556% | 5.873% | 3.335% | 1.547% | 3.2% | 1.876% | 8.436% | 3.492% | 4.981% | 1.366% | 0.8953% | 0.6437% | 6.177% | 7.721% | 0.4782% | 8.231% | 22.93% | 1.934% | 5.589% | 9.454% | 21.29% | 1.798% | 0.8211% | 1.193% | 0.5393% | 0.7116% | 0.9686% | 1.285% | 1.066% | 1.57% | 0.7921% | 1.75% | 0.592%
100 | CCV 1ppb TRACE |  | 6/16/2023 10:14:23 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.749 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.741% | 2.482% | 11.17% | 3.45% | 5.301% | 1.243% | 7.64% | 5.017% | 1.518% | 0.4888% | 2.992% | 2.408% | 3.887% | 0.2889% | 5.677% | 0.7032% | 1.294% | 0.9328% | 8.185% | 11.81% | 1.915% | 6.002% | 8.463% | 2.649% | 3.957% | 7.312% | 11.41% | 0.7836% | 0.1722% | 0.8387% | 0.3328% | 0.6684% | 3.154% | 0.9213% | 1.786% | 1.951% | 0.1712% | 0.7027% | 3.304%
101 | CCV 10 ppb MAJORS |  | 6/16/2023 10:20:23 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.750 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.443% | 3.628% | 10.24% | 0.2845% | 1.249% | 2.844% | 2.149% | 3.612% | 1.084% | 0.1596% | 2.034% | 1.286% | 3.404% | 1.998% | 0.9779% | 2.287% | 0.3343% | 1.068% | 0.8665% | 5.025% | 0.6997% | 0.6212% | 2.069% | 2.901% | 2.522% | 1.704% | 4.464% | 0.6721% | 1.543% | 0.3927% | 0.8377% | 0.3682% | 0.5189% | 0.7274% | 1.204% | 0.8821% | 0.3755% | 1.592% | 0.8669%
102 | 10_1_21_BB |  | 6/16/2023 10:26:24 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_BB.751 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 191.9% | 173.2% | 28.31% | 15.05% | 23.38% | 149.7% | 55.87% | 79.56% | 23.39% | 9.075% | 1.581% | 18.28% | 6.938% | 148.4% | 131.7% | 2.091% | 7.355% | 7.373% | 69.5% | 79.72% | 11.29% | 83.31% | 173.2% | 138.4% | 166.5% | 827.2% | 492.3% | 29.24% | 13.13% | 1.103% | 0.5277% | 1.434% | 1.109% | 0.3852% | 1.471% | 322.5% | 47.99% | 2.168% | 3.207%
103 | 10_1_21_FB |  | 6/16/2023 10:32:24 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_FB.752 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 7.973% | 86.62% | 2.336% | 1.225% | 22.98% | 20.84% | 80.09% | 206.6% | 19.48% | 0.3922% | 1.74% | 3.66% | 5.489% | 83.01% | 15.2% | 7.496% | 5.632% | 2.178% | 29.36% | 105.5% | 11.36% | 98.53% | 173.2% | 128.8% | 454.1% | 398.9% | 2.265e-06% | 57.15% | 22.86% | 1.537% | 0.631% | 0.6416% | 0.3558% | 0.3348% | 1.151% | 2603% | 612.4% | 6.367% | 6.428%
104 | 10_1_21_LAB_A |  | 6/16/2023 10:38:24 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_LAB_A.753 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 56.55% |  | 28.56% | 22.23% | 27.29% | 310.2% | 61.87% | 38.49% | 5.086% | 17.02% | 6.541% | 12.42% | 7.287% | 2116% | 63.76% | 12.6% | 2.474% | 1.98% | 741.7% | 131.1% | 21.19% | 27.68% | 173.2% | 112.2% | 235.1% | 232.7% | 533.8% | 3.34% | 13.57% | 0.9147% | 0.2723% | 0.6468% | 0.4915% | 2.048% | 0.8414% | 0% | 75.87% | 5.757% | 1.097%
105 | 10_1_21_LAB_B |  | 6/16/2023 10:44:24 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_LAB_B.754 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 85.51% | 173.2% | 14.39% | 19.43% | 26.71% | 143.1% | 50.13% | 39.22% | 16.25% | 2.907% | 1.142% | 17.27% | 9.289% | 1484% | 93.16% | 8.158% | 1.598% | 3.759% | 3969% | 606.4% | 42.28% | 381.2% |  | 508.9% | 382.9% | 82.69% | 557.3% | 5.234% | 38.25% | 1.494% | 0.3412% | 0.3601% | 0.1852% | 1.299% | 2.03% | 91.42% | 115.4% | 2.823% | 1.262%
106 | 5_18_23_DB |  | 6/16/2023 10:50:24 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_DB.755 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.8154% | 0.7885% | 1.028% | 1.08% | 16.84% | 1.391% | 0.7187% | 2.036% | 6.081% | 0.9481% | 0.8211% | 1.584% | 0.7206% | 1.2% | 1.794% | 6.144% | 1.416% | 0.693% | 2.596% | 10.01% | 1.034% | 5.084% | 8.481% | 6.662% | 5.712% | 12.94% | 3.851% | 1.537% | 0.6426% | 1.37% | 0.3016% | 0.7725% | 0.9521% | 1.032% | 0.8361% | 2.014% | 1.659% | 2.534% | 0.2394%
107 | 5_18_23_DB |  | 6/16/2023 10:56:42 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_DB.756 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.389% | 0.6112% | 1.084% | 1.201% | 11.74% | 3.012% | 2.119% | 1.201% | 2.88% | 0.5643% | 1.115% | 1.617% | 1.717% | 1.125% | 2.026% | 2.083% | 1.783% | 0.5802% | 5.692% | 4.073% | 1.39% | 2.258% | 16.6% | 8.358% | 1.622% | 4.028% | 0.8743% | 0.5982% | 1.602% | 0.2917% | 0.6116% | 0.6478% | 0.4366% | 0.3877% | 2.18% | 1.604% | 0.492% | 1.892% | 0.7187%
108 | 5_18_23_T_239 |  | 6/16/2023 11:02:42 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_T_239.757 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.705% | 1.907% | 1.179% | 1.288% | 15.43% | 0.3441% | 0.7533% | 2.132% | 1.168% | 1.058% | 0.9936% | 0.8387% | 1.55% | 2.327% | 0.9691% | 2.001% | 2.772% | 1.67% | 1.232% | 5.951% | 0.0314% | 1.361% | 3.764% | 1.618% | 4.651% | 6.869% | 16.11% | 0.8785% | 0.9314% | 0.3745% | 0.618% | 0.6688% | 0.9669% | 0.4168% | 0.5022% | 0.8261% | 1.933% | 1.448% | 1.796%
109 | 5_18_23_T_239 |  | 6/16/2023 11:09:00 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_T_239.758 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.4136% | 1.535% | 0.2847% | 0.5935% | 9.903% | 3.414% | 1.464% | 2.406% | 2.219% | 1.677% | 0.1356% | 0.5156% | 0.7631% | 2.182% | 1.639% | 1.215% | 2.436% | 0.9572% | 3.168% | 11.58% | 1.017% | 3.838% | 10.3% | 8.253% | 5.811% | 1.678% | 2.166% | 2.476% | 0.6743% | 0.8346% | 1.095% | 0.3458% | 1.506% | 0.8793% | 0.7705% | 0.4859% | 0.6726% | 0.603% | 2.031%
110 | 5_18_23_T_239 |  | 6/16/2023 11:15:00 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_T_239.759 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4.585% | 4.354% | 0.9746% | 1.768% | 4.297% | 4.308% | 5.743% | 2.542% | 16.82% | 8.016% | 3.074% | 0.4812% | 24.51% | 2.535% | 25.22% | 15.57% | 5.092% | 3.294% | 28.26% | 56.38% | 15.88% | 9.104% | 86.6% | 36.17% | 76.27% | 3.201% | 130.6% | 1.055% | 3.306% | 0.7488% | 0.4336% | 0.58% | 4.1% | 1.307% | 3.672% | 60.34% | 45.43% | 8.094% | 6.912%
111 | 5_18_23_M_224 |  | 6/16/2023 11:20:59 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_M_224.760 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.7644% | 3.332% | 7.859% | 1.667% | 1.394% | 0.8847% | 0.9669% | 1.444% | 1.036% | 0.5976% | 1.397% | 1.095% | 2.955% | 2.023% | 1.408% | 2.025% | 0.6704% | 1.204% | 8.83% | 0.909% | 1.534% | 5.584% | 7.933% | 2.946% | 1.541% | 12.32% | 5.566% | 0.8529% | 0.734% | 6.87% | 0.5318% | 0.1772% | 0.921% | 0.99% | 0.4839% | 1.6% | 0.2287% | 7.678% | 2.162%
112 | LBLANK |  | 6/16/2023 11:27:00 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.761 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 519.5% | 173.2% | 24.53% | 1.601% | 11.51% | 590.7% | 267.4% | 1201% | 84.02% | 100.8% | 6.58% | 1536% | 10.6% | 86.52% | 47.84% | 10623% | 230.2% | 3.41% | 70.27% | 0% | 20.84% | 167.5% | 173.2% | 127.9% | 35.29% | 2676% | 504.1% | 9.862% | 16.7% | 0.3356% | 0.8375% | 0.8353% | 0.9035% | 0.728% | 1.007% | 90.6% | 30.94% | 34.58% | 2.497%
113 | LBLANK |  | 6/16/2023 11:33:18 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.762 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 197.1% | 173.2% | 3.866% | 11.6% | 24.06% | 688.9% | 134% | 190.6% | 89.18% | 24.64% | 24.83% | 309.5% | 9.178% | 204.1% | 49.47% | 83.6% | 22.51% | 10.67% | 1006% | 579.3% | 178.9% | 214.8% |  | 255.7% | 36.01% | 46.25% | 2.265e-06% | 50.27% | 41.71% | 1.577% | 0.4858% | 2.63% | 0.5922% | 1.335% | 0.9195% | 336.8% | 36.33% | 57.96% | 14.43%
114 | CCV 1ppb TRACE |  | 6/16/2023 11:39:19 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.763 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.692% | 2.505% | 12.06% | 1.196% | 8.247% | 6.108% | 2.591% | 1.704% | 3.215% | 4.792% | 2.078% | 0.3748% | 3.146% | 1.221% | 5.967% | 1.755% | 2.556% | 1.586% | 3.03% | 11.61% | 4.606% | 5.331% | 9.893% | 4.262% | 5.937% | 11.93% | 3.884% | 1.87% | 1.288% | 1.072% | 1.182% | 1.366% | 0.5358% | 0.6714% | 0.9509% | 1.364% | 1.444% | 2.612% | 0.9951%
115 | CCV 1ppb TRACE |  | 6/16/2023 11:45:38 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.764 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.904% | 5.461% | 10.17% | 2.856% | 8.76% | 5.463% | 12.05% | 2.837% | 3.165% | 0.835% | 1.985% | 1.7% | 2.391% | 2.497% | 1.288% | 3.619% | 0.7561% | 1.853% | 3.202% | 19.42% | 2.567% | 7.447% | 32.39% | 4.63% | 5.148% | 4.927% | 21.37% | 1.674% | 0.9617% | 5.296% | 0.9594% | 0.7673% | 0.5262% | 0.5698% | 1.376% | 4.825% | 0.892% | 6.024% | 0.8635%
116 | LFB |  | 6/16/2023 11:51:38 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.765 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.604% | 0.7547% | 10.25% | 1.623% | 1.512% | 1.791% | 2.155% | 2.311% | 2.209% | 1.168% | 2.57% | 1.044% | 1.712% | 1.171% | 1.045% | 0.9954% | 0.291% | 0.866% | 2.885% | 2.994% | 1.595% | 3.173% | 4.439% | 2.106% | 1.53% | 4.889% | 2.987% | 0.397% | 0.2004% | 2.44% | 0.4442% | 0.8114% | 0.965% | 1.074% | 1.004% | 1.158% | 0.354% | 3.089% | 0.753%
117 | LFB |  | 6/16/2023 11:57:55 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.766 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.374% | 2.48% | 4.737% | 0.4439% | 1.554% | 2.136% | 1.587% | 2.39% | 1.7% | 0.5854% | 0.7738% | 0.3211% | 1.411% | 0.5273% | 3.956% | 1.206% | 0.522% | 0.2923% | 0.2423% | 4.092% | 0.6042% | 2.797% | 2.642% | 1.459% | 2.422% | 4.215% | 0.5878% | 1.45% | 0.6605% | 0.6494% | 0.9099% | 0.7759% | 0.3577% | 0.5826% | 1.404% | 0.8615% | 0.7794% | 1.528% | 1.333%
118 | CCV 10 ppb MAJORS |  | 6/17/2023 12:03:55 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.767 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.9584% | 0.9423% | 6.356% | 0.7835% | 3.474% | 1.306% | 1.492% | 0.9654% | 1.202% | 0.3607% | 1.288% | 0.3994% | 0.8147% | 1.66% | 1.681% | 2.273% | 0.1867% | 1.119% | 2.473% | 3.526% | 0.8933% | 2.777% | 3.107% | 0.5264% | 2.642% | 3.971% | 3.931% | 0.2169% | 0.9716% | 1.176% | 0.3713% | 0.6955% | 0.888% | 0.2153% | 0.2875% | 0.3733% | 0.7943% | 0.7546% | 1.26%
119 | 5_18_23_MFB |  | 6/17/2023 12:09:55 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_MFB.768 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 19.46% |  | 78.86% | 1.619% | 38.95% | 165.1% | 60.48% | 488% | 13.49% | 14.08% | 4.152% | 21.8% | 10.59% | 80.94% | 332.3% | 9.383% | 5.955% | 2.288% | 1152% | 648% | 114.6% | 12.76% | 173.2% | 105.9% | 845.6% | 137.1% | 2.265e-06% | 19.4% | 20.47% | 0.5741% | 0.343% | 0.7403% | 0.3774% | 0.4238% | 0.3288% | 458.3% | 102% | 5.731% | 5.033%
120 | 10_15_21_W_1 |  | 6/17/2023 12:15:55 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_W_1.769 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.159% | 3.327% | 7.252% | 1.333% | 2.416% | 6.353% | 1.574% | 9.671% | 20.92% | 1.638% | 1.17% | 0.6361% | 1.167% | 23.59% | 17.06% | 2.441% | 0.8656% | 4.304% | 2.409% | 8.988% | 3.253% | 41.59% | 86.61% | 61.41% | 54.9% | 58.77% | 224.4% | 1.641% | 1.449% | 0.902% | 0.4892% | 0.5454% | 0.07775% | 0.8708% | 1.325% | 24.26% | 40.04% | 5.285% | 2.126%
121 | 10_15_21_W_6 |  | 6/17/2023 12:21:56 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_W_6.770 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.537% | 0.9228% | 3.507% | 2.572% | 5.183% | 0.2241% | 2.033% | 3.152% | 26.67% | 53.16% | 1.301% | 1.653% | 2.122% | 19.05% | 7.99% | 5.949% | 2.602% | 2.22% | 7.924% | 14.5% | 1.319% | 54.82% | 35.02% | 82.57% | 12.37% | 91.68% | 905.8% | 0.7712% | 1.546% | 1.057% | 1.391% | 1.007% | 1.223% | 0.1285% | 1.332% | 29.13% | 46.52% | 2.157% | 1.174%
122 | 10_15_21_W_8 |  | 6/17/2023 12:27:56 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_W_8.771 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.18% | 8.427% | 2.567% | 3.34% | 1.605% | 1.384% | 0.8417% | 7.387% | 5.311% | 3.955% | 2.644% | 1.137% | 0.853% | 5.967% | 6.335% | 6.102% | 1.803% | 1.598% | 4.167% | 16.41% | 2.115% | 46.22% | 43.19% | 107.4% | 71.79% | 19.58% | 0.7066% | 3.306% | 1.457% | 0.5462% | 0.5638% | 0.9321% | 0.6423% | 0.1904% | 0.6107% | 51.99% | 41.63% | 3.947% | 2.398%
123 | 10_15_21_W_9 |  | 6/17/2023 12:33:56 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_W_9.772 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.633% | 1.214% | 1.005% | 3.226% | 3.96% | 3.609% | 2.794% | 19.59% | 11.9% | 6.459% | 0.6239% | 2% | 2.647% | 36.46% | 11.53% | 4.52% | 2.955% | 5.1% | 8.747% | 16.02% | 2.303% | 50% | 86.61% | 0.8137% | 61.82% | 99.05% | 120.8% | 1.981% | 2.147% | 15% | 0.7616% | 0.54% | 0.269% | 0.9237% | 1.057% | 20.26% | 52.83% | 14.38% | 1.535%
124 | 10_15_21_W_9 |  | 6/17/2023 12:40:15 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_W_9.773 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.229% | 3.263% | 5.511% | 1.344% | 8.496% | 8.131% | 3.339% | 17.21% | 15.47% | 5.56% | 0.5437% | 2.051% | 1.037% | 12.96% | 10.31% | 1.136% | 0.856% | 0.8098% | 15% | 28.38% | 6.732% | 19.39% | 173.2% | 55.13% | 145.5% | 77.85% | 90.14% | 0.8056% | 1.831% | 0.2768% | 1.192% | 1.315% | 1.345% | 0.6338% | 1.182% | 46.46% | 48.29% | 1.136% | 2.308%
125 | 10_15_21_W_10 |  | 6/17/2023 12:46:15 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_W_10.774 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.238% | 2.868% | 0.9185% | 2.563% | 3.288% | 3.187% | 3.784% | 4.015% | 8.581% | 21.63% | 1.112% | 1.511% | 0.7182% | 11.73% | 15.18% | 3.752% | 2.364% | 1.122% | 2.998% | 16.33% | 4.928% | 32.75% | 43.6% | 43.92% | 51.53% | 61.93% | 946.5% | 2.461% | 0.6256% | 1.148% | 0.2914% | 0.4235% | 0.5708% | 0.6537% | 0.5292% | 49.5% | 14.17% | 2.018% | 2.175%
126 | 10_15_21_W_10 |  | 6/17/2023 12:52:33 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_W_10.775 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.698% | 2.844% | 1.736% | 2.575% | 4.286% | 8.09% | 2.798% | 16.07% | 16.28% | 13.09% | 0.2579% | 0.7195% | 0.7119% | 9.115% | 10.68% | 1.466% | 1.072% | 2.008% | 10.86% | 3.737% | 3.821% | 33.54% | 86.6% | 61.93% | 33.92% | 38.96% | 165.9% | 0.7952% | 2.681% | 0.7711% | 0.5148% | 0.3204% | 0.5286% | 0.2867% | 1.01% | 26.35% | 9.403% | 2.066% | 2.279%
127 | 10_15_21_W_10_DUP |  | 6/17/2023 12:58:33 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_W_10_DUP.776 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.003% | 1.47% | 0.5446% | 1.349% | 7.918% | 9.782% | 1.137% | 6.306% | 1.614% | 6.328% | 3.278% | 0.4858% | 2.736% | 10.56% | 28.5% | 1.312% | 2.445% | 1.273% | 14% | 18.84% | 1.55% | 19.24% |  | 79.32% | 112.1% | 18% | 287.5% | 1.786% | 2.275% | 0.5889% | 0.9428% | 1.56% | 0.8012% | 0.5468% | 1.441% | 4.822% | 18.29% | 0.6556% | 0.6765%
128 | 10_15_21_W_10_DUP |  | 6/17/2023 1:04:52 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_W_10_DUP.777 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4.765% | 2.152% | 4.218% | 1.095% | 2.697% | 4.068% | 1.845% | 8.441% | 11.52% | 5.76% | 2.439% | 1.456% | 1.502% | 25.43% | 29.66% | 1.474% | 1.27% | 0.4662% | 7.578% | 15.85% | 3.594% | 27.37% |  | 35.36% | 17.24% | 18.58% | 422.3% | 1.384% | 1.03% | 0.5219% | 0.3854% | 0.5241% | 1.382% | 0.4935% | 1.045% | 9.672% | 31.15% | 2.571% | 4.393%
129 | LBLANK |  | 6/17/2023 1:10:52 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.778 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 176.1% | 173.2% | 8.396% | 4.913% | 23.1% | 142.6% | 533.4% | 237.7% | 287.3% | 70.03% | 6.515% | 96.78% | 81.33% | 40.46% | 2052% | 51.19% | 21.41% | 3.144% | 97.53% | 0% | 536.2% | 69.93% |  | 248.1% | 39.05% | 191.4% | 800.9% | 35.99% | 11.15% | 0.3605% | 0.9115% | 0.5611% | 0.5531% | 0.133% | 1.395% | 0% | 47.14% | 32.83% | 31.55%
130 | LBLANK |  | 6/17/2023 1:17:12 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.779 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 452.1% |  | 14.2% | 57.08% | 23.03% | 41.59% | 519.7% | 34.95% | 109.1% | 39.1% | 25.94% | 130.1% | 34.08% | 130.4% | 304.3% | 201.4% | 84.24% | 2.765% | 394.9% | 0% | 131.4% | 262.7% |  | 438.7% | 518.1% | 4932% | 2.265e-06% | 9.916% | 10.82% | 2.965% | 0.5999% | 0.3541% | 0.7981% | 0.9525% | 1.186% | 102.3% | 11.02% | 3.838% | 10.25%
131 | CCV 1ppb TRACE |  | 6/17/2023 1:23:12 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.780 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.265% | 9.595% | 8.317% | 4.2% | 13.16% | 6.401% | 6.751% | 2.455% | 3.802% | 2.875% | 0.4987% | 1.999% | 0.9098% | 0.9675% | 0.6687% | 4.071% | 1.924% | 4.322% | 3.269% | 23.9% | 2.709% | 12.94% | 14.07% | 5.25% | 6.942% | 10.74% | 11.92% | 2.403% | 1.092% | 0.5928% | 0.4277% | 0.8454% | 0.6498% | 0.3416% | 0.6929% | 0.82% | 2.559% | 0.6266% | 1.412%
132 | CCV 1ppb TRACE |  | 6/17/2023 1:29:31 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.781 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.688% | 4.615% | 17.26% | 5.27% | 14.14% | 4.259% | 6.295% | 0.8652% | 2.455% | 1.581% | 1.354% | 1.234% | 4.847% | 3.201% | 7.945% | 5.281% | 1.026% | 0.8549% | 6.168% | 13.43% | 1.587% | 9.641% | 15.55% | 5.379% | 2.576% | 20.49% | 6.509% | 0.7208% | 1.085% | 0.7887% | 1.005% | 0.8591% | 0.5559% | 0.7531% | 1.07% | 1.166% | 0.5027% | 1.583% | 1.343%
133 | CCV 10 ppb MAJORS |  | 6/17/2023 1:35:32 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.782 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.014% | 2.445% | 7.093% | 1.105% | 2.318% | 1.074% | 2.278% | 1.838% | 1.381% | 0.4251% | 1.699% | 1.522% | 1.584% | 0.4675% | 2.686% | 2.345% | 0.4022% | 1.317% | 2.156% | 1.952% | 0.2037% | 2.131% | 4.887% | 4.565% | 1.114% | 0.4636% | 2.862% | 0.3259% | 0.4194% | 1.054% | 0.3696% | 1.044% | 0.2332% | 0.4327% | 0.5378% | 0.6583% | 0.7843% | 2.855% | 0.5706%
134 | CCV 10 ppb MAJORS |  | 6/17/2023 1:41:51 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.783 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.878% | 0.4532% | 4.244% | 0.816% | 0.7356% | 0.269% | 0.9189% | 0.3192% | 1.008% | 3.601% | 3.846% | 0.7115% | 0.6324% | 0.8256% | 2.204% | 1.515% | 3.465% | 4.347% | 1.719% | 5.408% | 0.8444% | 2.902% | 4.206% | 3.207% | 1.787% | 3.752% | 5.228% | 3.819% | 1.194% | 0.593% | 0.9523% | 2.859% | 0.9469% | 0.4514% | 0.6567% | 5.009% | 4.018% | 0.7733% | 1.249%
135 | 10_15_21_W_11 |  | 6/17/2023 1:47:51 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_W_11.784 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.567% | 3.089% | 2.36% | 0.3885% | 5.469% | 8.929% | 0.9165% | 3.512% | 24.86% | 4.381% | 1.063% | 2.364% | 1.143% | 6.933% | 30.01% | 5.329% | 1.22% | 1.097% | 6.091% | 26.69% | 4.054% | 67.76% | 173.2% | 124.5% | 52.72% | 180.7% | 166.1% | 3.139% | 0.8738% | 2.839% | 0.7473% | 0.2272% | 0.6943% | 1.043% | 0.6431% | 31.79% | 7.53% | 5.656% | 0.7969%
136 | 10_15_21_W_11 |  | 6/17/2023 1:54:10 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_W_11.785 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.274% | 2.723% | 2.283% | 1.656% | 2.842% | 5.047% | 0.504% | 3.469% | 5.009% | 3.83% | 1.197% | 1.086% | 0.2496% | 15.23% | 49.09% | 3.968% | 0.3686% | 1.101% | 5.54% | 13.81% | 0.969% | 35.8% | 86.6% | 59.9% | 59.06% | 63.92% | 163.7% | 1.864% | 0.4367% | 0.6382% | 0.3568% | 0.8426% | 1.27% | 1.042% | 0.5285% | 17.02% | 19.37% | 2.771% | 0.708%
137 | 10_15_21_F_1 |  | 6/17/2023 2:00:09 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_1.786 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.357% | 2.966% | 7.624% | 0.51% | 1.516% | 1.16% | 0.06207% | 10.39% | 19.76% | 7.285% | 0.2269% | 0.4545% | 1.816% | 24.48% | 2.511% | 3.901% | 0.7691% | 0.5746% | 1.249% | 1.704% | 1.972% | 5.816% | 25.6% | 9.767% | 5.691% | 17.94% | 76.41% | 0.9988% | 0.5172% | 0.5382% | 0.6462% | 0.6477% | 3.372% | 1.247% | 0.7562% | 45.61% | 13.87% | 0.9543% | 4.765%
138 | 10_15_21_F_1 |  | 6/17/2023 2:06:28 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_1.787 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.712% | 1.065% | 9.028% | 2.442% | 0.4809% | 2.588% | 2.417% | 3.018% | 14.48% | 2.259% | 0.5512% | 0.1365% | 0.8894% | 16.4% | 13.03% | 1.162% | 1.576% | 0.7913% | 2.228% | 1.293% | 0.6856% | 9.769% | 41.78% | 29.35% | 11.54% | 6.794% | 118.5% | 0.2832% | 0.8434% | 1.082% | 0.1814% | 0.2782% | 0.1876% | 1.107% | 0.8423% | 21.21% | 26.74% | 1.036% | 0.6192%
139 | 10_15_21_F_1_DUP |  | 6/17/2023 2:12:28 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_1_DUP.788 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.9868% | 1.758% | 13.86% | 1.174% | 0.8851% | 3.564% | 1.102% | 10.36% | 16.5% | 9.391% | 0.2197% | 1.154% | 1.408% | 6.854% | 1.609% | 4.144% | 1.926% | 1.821% | 1.515% | 7.785% | 0.3797% | 16.45% | 56.17% | 29.36% | 14.9% | 10.9% | 50.95% | 1.104% | 0.6436% | 0.8147% | 0.7142% | 0.5063% | 0.2845% | 0.3679% | 0.535% | 3.978% | 18.75% | 2.604% | 1.728%
140 | 10_15_21_F_1_DUP |  | 6/17/2023 2:18:46 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_1_DUP.789 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.858% | 2.069% | 14.04% | 3.134% | 2.222% | 1.792% | 2.244% | 4.022% | 14.16% | 8.478% | 2.156% | 0.4631% | 3.545% | 13.72% | 8.381% | 2.948% | 1.197% | 0.3024% | 3.252% | 5.895% | 0.9851% | 24.26% | 48.98% | 14.55% | 15.56% | 21.63% | 18.57% | 0.7921% | 1.511% | 1.767% | 1.071% | 0.5785% | 0.5509% | 0.4484% | 1.734% | 45.66% | 16.46% | 3.91% | 1.624%
141 | 10_15_21_F_6 |  | 6/17/2023 2:24:45 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_6.790 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.5278% | 2.162% | 5.554% | 2.284% | 3.804% | 2.458% | 1.126% | 10.78% | 9.173% | 4.728% | 1.15% | 0.9176% | 2.196% | 9.531% | 4.617% | 2.966% | 2.242% | 0.911% | 1.277% | 2.708% | 1.548% | 37.06% | 39.85% | 32.68% | 19.26% | 7.359% | 38.13% | 1.103% | 1.002% | 0.6614% | 0.647% | 1.256% | 4.023% | 0.812% | 1.467% | 24.57% | 5.029% | 2.58% | 9.436%
142 | 10_15_21_F_6 |  | 6/17/2023 2:31:03 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_6.791 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.4672% | 1.797% | 0.6418% | 1.235% | 0.8111% | 1.355% | 1.163% | 3.037% | 5.906% | 6.214% | 0.5036% | 1.107% | 0.4424% | 8.836% | 8.851% | 2.041% | 1.004% | 2.065% | 1.229% | 8.774% | 0.8952% | 12.51% | 19.69% | 43.11% | 24% | 5.899% | 24.43% | 0.3201% | 0.6179% | 0.7932% | 0.2577% | 0.2251% | 0.2746% | 0.8431% | 0.5364% | 14.96% | 7.413% | 3.116% | 2.614%
143 | 10_15_21_F_6_LD |  | 6/17/2023 2:37:02 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_6_LD.792 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.875% | 2.66% | 1.189% | 2.955% | 1.218% | 3.048% | 1.843% | 0.5972% | 7.13% | 6.376% | 0.7485% | 1.48% | 2.123% | 18.03% | 1.434% | 0.8665% | 0.9495% | 1.476% | 1.127% | 7.061% | 2.884% | 7.839% | 65.96% | 9.37% | 6.896% | 16.94% | 16.61% | 0.7403% | 1.539% | 0.7848% | 1.222% | 0.9164% | 0.7052% | 0.2116% | 1.097% | 19.22% | 2.095% | 3.837% | 1.849%
144 | 10_15_21_F_6_LD |  | 6/17/2023 2:43:20 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_6_LD.793 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.167% | 2.972% | 4.415% | 2.162% | 2.079% | 1.242% | 0.5963% | 1.764% | 13.03% | 3.472% | 5.008% | 0.7059% | 2.949% | 10.41% | 12.65% | 1.244% | 3.695% | 4.904% | 6.914% | 12.95% | 0.8491% | 47.35% | 43.08% | 13.2% | 3.696% | 4.575% | 46.59% | 4.463% | 0.6121% | 5.774% | 0.3839% | 3.872% | 0.1897% | 1.387% | 1.006% | 6.147% | 11.19% | 5.042% | 0.3324%
145 | LBLANK |  | 6/17/2023 2:49:20 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.794 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 131.1% |  | 16.02% | 3.092% | 30.55% | 56.45% | 2.462% | 87.27% | 16.08% | 22.75% | 28.76% | 121.9% | 10.03% | 118% | 142.8% | 127.3% | 101.1% | 5.149% | 2134% | 130.7% | 316.2% | 123.4% |  | 126% | 48.18% | 67.18% | 721.4% | 5.496% | 23.71% | 5.908% | 0.6351% | 0.924% | 1.533% | 0.5081% | 1.933% | 0% | 24.91% | 14.12% | 12.01%
146 | CCV 1ppb TRACE |  | 6/17/2023 2:55:21 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.795 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.746% | 1.901% | 5.872% | 4.177% | 8.723% | 4.28% | 5.069% | 6.305% | 3.021% | 0.2865% | 2.457% | 1.544% | 5.069% | 2.544% | 3.2% | 0.7601% | 0.3716% | 2.019% | 3.752% | 9.26% | 2.134% | 8.647% | 10.97% | 13.42% | 2.557% | 16.96% | 2.458% | 1.665% | 1.077% | 1.788% | 0.1805% | 0.9617% | 0.7834% | 0.5332% | 0.4923% | 2.109% | 2.609% | 0.7533% | 1.31%
147 | CCV 1ppb TRACE |  | 6/17/2023 3:01:39 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.796 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.944% | 5.262% | 7.958% | 5.487% | 12.28% | 2.574% | 2.772% | 2.501% | 1.845% | 1.124% | 2.043% | 1.057% | 4.899% | 3.135% | 6.265% | 5.331% | 1.46% | 3.41% | 2.938% | 22.06% | 2.796% | 21.35% | 32.66% | 3.324% | 3.57% | 11.55% | 0.8195% | 1.671% | 1.099% | 7.55% | 0.8115% | 0.5285% | 0.6279% | 0.6499% | 0.6774% | 0.6277% | 1.178% | 9.038% | 0.7631%
148 | LFB |  | 6/17/2023 3:07:39 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.797 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.177% | 1.353% | 9.23% | 0.6962% | 3.525% | 1.367% | 3.19% | 1.218% | 3.286% | 1.361% | 0.2212% | 0.3317% | 1.794% | 2.098% | 1.477% | 1.675% | 1.567% | 0.5318% | 1.636% | 1.98% | 1.311% | 3.887% | 6.433% | 1.691% | 3.429% | 4.105% | 3.149% | 0.7474% | 0.7684% | 1.999% | 1.025% | 0.5075% | 3.364% | 0.3256% | 1.357% | 1.39% | 0.6682% | 1.189% | 2.758%
149 | LFB |  | 6/17/2023 3:13:57 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.798 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.2348% | 2.946% | 2.964% | 0.6771% | 2.505% | 1.17% | 2.835% | 2.843% | 1.314% | 1.48% | 0.7824% | 1.551% | 1.159% | 1.583% | 2.099% | 1.824% | 0.6074% | 0.8179% | 1.718% | 3.983% | 0.3362% | 1.355% | 6.94% | 0.447% | 0.3963% | 5.396% | 4.832% | 1.274% | 0.671% | 0.6623% | 0.3639% | 0.5003% | 0.7924% | 0.8585% | 0.9356% | 1.862% | 1.167% | 1.192% | 1.673%
150 | CCV 10 ppb MAJORS |  | 6/17/2023 3:19:56 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.799 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.353% | 1.876% | 3.581% | 1.426% | 5.49% | 1.094% | 4.928% | 1.963% | 2.917% | 0.7546% | 0.8856% | 0.0879% | 1.595% | 3.264% | 1.874% | 2.247% | 0.5498% | 0.2185% | 1.471% | 8.059% | 0.3988% | 1.728% | 9.077% | 2.095% | 2.062% | 2.148% | 1.079% | 0.3529% | 0.7393% | 0.2922% | 0.5046% | 0.4959% | 0.7286% | 0.1791% | 2.148% | 0.2161% | 0.9851% | 0.8039% | 1.445%
151 | CCV 10 ppb MAJORS |  | 6/17/2023 3:26:14 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.800 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.127% | 4.693% | 4.989% | 2.312% | 1.154% | 1.819% | 2.979% | 1.002% | 1.219% | 0.3387% | 1.642% | 1.1% | 1.652% | 3.283% | 2.689% | 2.029% | 0.2038% | 1.469% | 3.771% | 1.683% | 1.079% | 6.584% | 5.424% | 2.696% | 3.14% | 6.176% | 0.6838% | 0.393% | 0.6423% | 3.224% | 0.144% | 0.6114% | 0.323% | 0.3116% | 1.771% | 0.6004% | 1.088% | 4.752% | 0.67%
152 | 10_15_21_F_6_DUP |  | 6/17/2023 3:32:14 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_6_DUP.801 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.44% | 1.751% | 3.15% | 0.945% | 1.347% | 5.328% | 1.175% | 7.537% | 6.758% | 9.653% | 0.763% | 0.4112% | 0.421% | 15.02% | 10.54% | 2.741% | 0.3542% | 0.7167% | 2.142% | 2.856% | 1.148% | 26.93% | 12.53% | 49.7% | 53.37% | 16.85% | 73.24% | 0.3911% | 0.8336% | 1.558% | 0.292% | 0.5471% | 0.4767% | 0.8061% | 0.7152% | 1.611% | 9.685% | 1.792% | 1.512%
153 | 10_15_21_F_6_DUP |  | 6/17/2023 3:38:32 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_6_DUP.802 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.247% | 0.7876% | 6.417% | 1.894% | 2.021% | 0.5495% | 1.517% | 9.726% | 17.9% | 7.509% | 1.386% | 1.302% | 1.473% | 11.57% | 11.2% | 2.233% | 1.239% | 0.7864% | 3.094% | 8.386% | 0.6538% | 14.64% | 89.15% | 9.989% | 20.34% | 27.01% | 57.59% | 0.4462% | 2.105% | 2.836% | 1.254% | 0.4233% | 2.772% | 0.4706% | 0.6008% | 7.679% | 4.495% | 3.015% | 3.993%
154 | 10_15_21_F_6_DUP_LSP |  | 6/17/2023 3:44:32 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_6_DUP_LSP.803 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.102% | 2.515% | 0.9933% | 1.284% | 0.2239% | 2.139% | 1.779% | 1.915% | 0.3925% | 0.5379% | 0.3655% | 0.3351% | 1.441% | 2.235% | 2.143% | 0.7121% | 1.088% | 0.6616% | 0.9231% | 4.195% | 0.8581% | 2.483% | 0.1632% | 0.3215% | 2.499% | 2.799% | 7.313% | 0.6912% | 0.3114% | 0.8807% | 0.1632% | 0.4564% | 0.9774% | 1.092% | 1.435% | 0.3568% | 0.4267% | 1.301% | 1.236%
155 | 10_15_21_F_6_DUP_LSP |  | 6/17/2023 3:50:52 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_6_DUP_LSP.804 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4.37% | 3.896% | 4.707% | 0.7159% | 3.69% | 2.009% | 3.466% | 3.468% | 3.77% | 0.9822% | 1.346% | 0.7914% | 3.469% | 3.349% | 3.884% | 3.383% | 0.7578% | 1.706% | 1.178% | 6.838% | 1.194% | 3.249% | 5.933% | 3.003% | 1.914% | 1.602% | 5.33% | 0.1656% | 0.52% | 0.6434% | 0.7323% | 0.6684% | 0.6743% | 0.5631% | 3.299% | 0.4971% | 1.167% | 1.336% | 1.423%
156 | 10_15_21_F_9 |  | 6/17/2023 3:56:52 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_9.805 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.44% | 1.434% | 4.387% | 0.7991% | 2.783% | 2.133% | 2.236% | 4.217% | 6.745% | 5.591% | 0.6864% | 0.924% | 0.9017% | 6.858% | 9.24% | 2.688% | 1.168% | 1.636% | 4.57% | 10.27% | 1.344% | 36.47% | 57.86% | 38.86% | 19.16% | 24.42% | 47.93% | 1.781% | 0.9165% | 1.257% | 0.7266% | 0.6754% | 0.232% | 0.7402% | 0.3964% | 14.89% | 15.8% | 4.434% | 1.649%
157 | 10_15_21_F_9 |  | 6/17/2023 4:03:12 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_9.806 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.8639% | 2.562% | 5.447% | 1.67% | 1.912% | 5.534% | 0.8325% | 3.636% | 5.812% | 6.768% | 1.023% | 0.3914% | 1.925% | 11.98% | 13.95% | 3.243% | 1.562% | 1.103% | 1.297% | 8.328% | 0.8488% | 3.535% | 33.76% | 20.1% | 24.56% | 14.84% | 12.12% | 1.017% | 1.707% | 0.6497% | 0.4503% | 0.2038% | 0.6271% | 0.5051% | 1.061% | 12.73% | 16.22% | 2.94% | 1.554%
158 | 10_15_21_F_10 |  | 6/17/2023 4:09:13 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_10.807 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.31% | 0.06311% | 4.209% | 0.8246% | 2.738% | 1.446% | 1.075% | 4.545% | 10.26% | 4.93% | 0.239% | 5.352% | 2.604% | 17.47% | 1.136% | 2.347% | 1.533% | 1.674% | 1.715% | 7.343% | 7.023% | 17.86% | 34.02% | 29.66% | 24.57% | 5.731% | 105.3% | 0.979% | 6.915% | 2.2% | 4.762% | 0.5503% | 0.9162% | 0.5534% | 0.5655% | 38.27% | 25.51% | 3.991% | 2.132%
159 | 10_15_21_F_10 |  | 6/17/2023 4:15:32 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_10.808 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.035% | 1.784% | 4.662% | 0.7119% | 1.374% | 5.553% | 2.389% | 3.801% | 11.14% | 14.57% | 17.51% | 1.388% | 3.021% | 33.82% | 8.614% | 4.501% | 19.18% | 16.66% | 2.091% | 2.348% | 2.392% | 17.8% | 48.18% | 18.38% | 15.2% | 3.902% | 56.77% | 18.09% | 1.682% | 0.1382% | 0.7617% | 13.17% | 0.7009% | 0.663% | 2.296% | 39.84% | 29.48% | 1.9% | 1.611%
160 | 10_15_21_F_10_DUP |  | 6/17/2023 4:21:32 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_10_DUP.809 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.5488% | 2.954% | 6.632% | 0.371% | 3.7% | 0.8359% | 1.134% | 6.366% | 14.2% | 4.57% | 0.7253% | 2.147% | 3.186% | 11.3% | 2.77% | 3.902% | 0.8227% | 1.422% | 1.436% | 10.2% | 2.491% | 19.42% | 85.76% | 43.94% | 29.19% | 13.59% | 132% | 0.5748% | 1.521% | 1.174% | 1.045% | 1.064% | 0.9918% | 0.909% | 0.6035% | 38.93% | 23.32% | 3.181% | 1.367%
161 | 10_15_21_F_10_DUP |  | 6/17/2023 4:27:51 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_10_DUP.810 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.19% | 1.729% | 9.254% | 1.013% | 0.168% | 2.03% | 2.604% | 3.962% | 23.01% | 5.401% | 1.382% | 0.7548% | 2.851% | 34.38% | 6.926% | 2.227% | 1.688% | 1.912% | 1.708% | 4.866% | 0.8761% | 21.02% | 86.6% | 27.65% | 31.63% | 13.97% | 69.78% | 1.793% | 0.9037% | 2.058% | 0.6494% | 0.7846% | 0.9728% | 1.412% | 0.5801% | 28.43% | 26.68% | 3.371% | 0.7971%
162 | LBLANK |  | 6/17/2023 4:33:51 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.811 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 205.6% | 173.2% | 8.041% | 3.002% | 6.427% | 44.35% | 104.1% | 0% | 65.15% | 16.66% | 63.57% | 1243% | 29.59% | 41.14% | 107.1% | 159.1% | 160.1% | 4.014% | 119.3% | 0% | 166.6% | 189.7% | 173.2% | 133.4% | 147.3% | 122.2% | 2.265e-06% | 8.697% | 51.24% | 1.042% | 0.6218% | 0.2723% | 1.28% | 0.4278% | 2.298% | 108.4% | 44.01% | 35.76% | 31.19%
163 | LBLANK |  | 6/17/2023 4:40:11 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.812 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 69.78% | 173.2% | 13.85% | 11.24% | 15.31% | 619.8% | 449.1% | 205.5% | 80.26% | 12.32% | 38.47% | 24.34% | 45.18% | 44.25% | 181.8% | 117% | 157.3% | 1.332% | 22.15% | 494.8% | 149.6% | 78.4% |  | 2.265e-06% | 0% | 277.2% | 2.265e-06% | 14.18% | 67.14% | 3.627% | 1.124% | 0.9879% | 0.4514% | 0.4492% | 2.821% | 0% | 15.81% | 3.96% | 7.205%
164 | CCV 1ppb TRACE |  | 6/17/2023 4:46:11 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.813 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4.192% | 2.781% | 8.378% | 4.069% | 10% | 7.493% | 4.271% | 2.06% | 1.315% | 1.625% | 5.489% | 1.6% | 7.659% | 4.338% | 8.936% | 3.578% | 1.614% | 1.435% | 6.371% | 10.41% | 2.168% | 6.441% | 25.88% | 6.204% | 7.984% | 18.03% | 15.5% | 0.6867% | 1.216% | 1.316% | 0.5856% | 0.763% | 0.7265% | 0.7338% | 1.232% | 2.112% | 2.304% | 1.933% | 1.563%
165 | CCV 1ppb TRACE |  | 6/17/2023 4:52:31 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.814 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4.866% | 3.342% | 9.975% | 3.624% | 10.23% | 3.996% | 4.171% | 5.512% | 5.38% | 0.7703% | 2.252% | 1.057% | 8.538% | 3.578% | 6.472% | 5.01% | 1.476% | 1.426% | 3.203% | 19.93% | 1.215% | 24.23% | 14.41% | 9.818% | 5.021% | 29.31% | 4.969% | 1.029% | 1.575% | 7.182% | 0.6419% | 0.9373% | 0.7019% | 0.02218% | 2.09% | 2.429% | 1.258% | 10.71% | 1.476%
166 | CCV 10 ppb MAJORS |  | 6/17/2023 4:58:31 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.815 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.291% | 1.348% | 7.849% | 1.429% | 3.53% | 1.817% | 4.024% | 2.343% | 2.548% | 1.316% | 1.205% | 0.4426% | 1.143% | 0.2008% | 2.314% | 0.5505% | 1.105% | 0.8735% | 0.6905% | 3.935% | 0.3461% | 1.493% | 3.729% | 4.339% | 3.418% | 2.596% | 6.43% | 0.3831% | 0.8138% | 0.9511% | 0.3233% | 0.7959% | 0.6926% | 0.3525% | 0.6979% | 0.6314% | 0.7513% | 1.768% | 0.9957%
167 | CCV 10 ppb MAJORS |  | 6/17/2023 5:04:50 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.816 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.844% | 3.413% | 1.819% | 1.837% | 2.713% | 2.573% | 1.57% | 2.563% | 1.529% | 1.134% | 0.8562% | 0.7042% | 2.651% | 3.877% | 5.022% | 2.237% | 0.9798% | 1.249% | 1.447% | 4.414% | 1.043% | 4.111% | 8.026% | 0.6408% | 1.411% | 1.655% | 4.405% | 0.9121% | 0.711% | 2.251% | 0.5143% | 0.8808% | 0.516% | 0.8302% | 2.247% | 0.941% | 0.7781% | 3.304% | 0.4811%
168 | 10_15_21_F_11 |  | 6/17/2023 5:10:50 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_11.817 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.286% | 2.241% | 7.427% | 0.7476% | 3.856% | 2.238% | 1.938% | 4.232% | 19.7% | 9.996% | 2.114% | 0.5314% | 3.205% | 15.74% | 19.66% | 3.537% | 0.8457% | 1.137% | 2.662% | 9.57% | 1.667% | 19.53% | 48.3% | 38.15% | 19.43% | 18.77% | 58% | 1.616% | 0.3514% | 0.07403% | 0.4769% | 0.5942% | 0.1247% | 0.2647% | 1.043% | 16.41% | 7.468% | 2.748% | 1.079%
169 | 10_15_21_F_11 |  | 6/17/2023 5:17:09 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_11.818 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.377% | 0.2476% | 3.317% | 3.311% | 2.294% | 2.696% | 1.871% | 2.384% | 11.23% | 4.966% | 0.809% | 1.6% | 2.051% | 20.34% | 12.93% | 3.294% | 1.836% | 1.021% | 3.359% | 6.822% | 0.9324% | 21.24% | 23.81% | 44.65% | 33.33% | 8.796% | 18.3% | 1.548% | 0.8391% | 0.4129% | 0.9654% | 0.5893% | 3.357% | 0.7831% | 0.9304% | 19.49% | 3.736% | 2.745% | 5.82%
170 | 10_15_21_UF_1 |  | 6/17/2023 5:23:09 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_1.819 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.476% | 2.191% | 20.84% | 1.446% | 0.8617% | 1.869% | 0.7702% | 6.856% | 6.072% | 5.013% | 2.779% | 0.5548% | 2.703% | 15.37% | 10.34% | 1.15% | 2.63% | 1.228% | 1.968% | 10.16% | 0.9385% | 19.76% | 86.6% | 49.9% | 31.96% | 12.06% | 116.7% | 1.781% | 0.7251% | 0.9167% | 0.4293% | 0.508% | 1.344% | 0.3541% | 0.7725% | 21.44% | 18.78% | 3.65% | 8.344%
171 | 10_15_21_UF_1 |  | 6/17/2023 5:29:28 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_1.820 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.72% | 0.1893% | 11.69% | 1.709% | 0.2564% | 1.582% | 0.8396% | 6.165% | 4.268% | 3.683% | 1.901% | 1.597% | 1.985% | 7.095% | 3.341% | 4.349% | 2.024% | 1.962% | 2.293% | 7.971% | 2.272% | 11.21% | 91.13% | 13.65% | 12.86% | 8.905% | 108.4% | 1.19% | 1.11% | 0.1832% | 1.218% | 0.2944% | 0.4764% | 0.9472% | 0.7133% | 12.81% | 23.49% | 5.75% | 7.149%
172 | 10_15_21_UF_6 |  | 6/17/2023 5:35:27 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_6.821 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.187% | 3.343% | 12.66% | 2.137% | 3.844% | 2.336% | 2.566% | 0.4413% | 6.466% | 12.9% | 3.541% | 0.4486% | 4.971% | 23.95% | 15.1% | 1.689% | 2.434% | 1.588% | 2.527% | 4.819% | 2.892% | 7.945% | 26.04% | 35.37% | 22.99% | 18.45% | 26.68% | 0.9583% | 0.8413% | 1.196% | 1.086% | 0.1733% | 0.6111% | 0.8935% | 1.225% | 13.04% | 5.917% | 5.017% | 3.01%
173 | 10_15_21_UF_6 |  | 6/17/2023 5:41:46 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_6.822 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.9705% | 0.5068% | 14.28% | 4.175% | 2.148% | 4.69% | 0.156% | 5.639% | 11.52% | 6.611% | 2.414% | 0.1559% | 2.681% | 21.3% | 4.315% | 2.808% | 0.9258% | 0.6723% | 4.818% | 1.652% | 0.7622% | 28.12% | 0.2844% | 43.57% | 13.73% | 11.24% | 99.7% | 1.126% | 0.2293% | 2.371% | 0.2845% | 0.7375% | 0.683% | 0.8663% | 0.3766% | 7.883% | 8.374% | 1.761% | 1.419%
174 | 10_15_21_UF_8 |  | 6/17/2023 5:47:46 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_8.823 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.717% | 1.284% | 9.51% | 1.573% | 0.6269% | 0.389% | 0.9051% | 1.537% | 3.061% | 10.4% | 0.9429% | 0.3297% | 4.351% | 19.8% | 10.29% | 3.817% | 1.107% | 0.4324% | 2.278% | 17.25% | 1.083% | 22.93% | 34.3% | 36.97% | 28.6% | 12.13% | 53.64% | 0.9745% | 0.4214% | 0.1926% | 0.3072% | 0.5153% | 0.5438% | 0.6148% | 0.7879% | 13.51% | 5.592% | 1.7% | 4.106%
175 | 10_15_21_UF_8 |  | 6/17/2023 5:54:04 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_8.824 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.976% | 2.433% | 3.699% | 2.158% | 1.552% | 3.379% | 2.205% | 2.825% | 10.5% | 2.097% | 1.658% | 1.33% | 5.292% | 13.9% | 11.14% | 4.283% | 1.139% | 1.571% | 3.68% | 5.991% | 0.3012% | 26.93% | 41.93% | 48.99% | 14.49% | 5.165% | 18.59% | 0.3213% | 1.39% | 1.203% | 0.6943% | 0.8069% | 2.57% | 0.7964% | 1.393% | 18.03% | 24.07% | 4.632% | 2.57%
176 | 10_15_21_UF_8_LD |  | 6/17/2023 6:00:03 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_8_LD.825 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.296% | 2.606% | 6.224% | 2.214% | 2.292% | 4.446% | 1.666% | 3.311% | 9.211% | 3.326% | 1.797% | 7.151% | 4.235% | 3.806% | 5.26% | 3.336% | 1.955% | 0.225% | 0.6612% | 4.847% | 5.216% | 9.305% | 38.54% | 48.82% | 14.85% | 11.78% | 19.31% | 0.849% | 7.332% | 0.5062% | 5.726% | 0.41% | 0.1099% | 0.5577% | 1.264% | 14.95% | 14.27% | 1.057% | 2.593%
177 | 10_15_21_UF_8_LD |  | 6/17/2023 6:06:21 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_8_LD.826 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.5627% | 2.228% | 1.447% | 0.253% | 2.003% | 1.909% | 2.046% | 1.328% | 4.486% | 10.18% | 2.238% | 0.04406% | 3.754% | 20.34% | 12.63% | 1.463% | 2.965% | 1.393% | 2.8% | 9.771% | 0.4222% | 52.63% | 21.93% | 30.21% | 18.63% | 3.185% | 26.94% | 1.782% | 0.1594% | 0.6671% | 0.3272% | 1.419% | 0.8862% | 0.7562% | 0.8365% | 20.59% | 9.707% | 2.817% | 3.413%
178 | LBLANK |  | 6/17/2023 6:12:21 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.827 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1066% | 86.61% | 7.757% | 8.829% | 31.65% | 319.1% | 125.4% | 115.7% | 85.19% | 37.18% | 58.94% | 33.27% | 25.71% | 0% | 102.2% | 114.1% | 26.37% | 10.63% | 246.6% | 102.4% | 22.5% | 84.04% |  | 391.7% | 312.3% | 202% | 1499% | 11.21% | 20.6% | 4.406% | 0.3825% | 0.1172% | 0.5691% | 0.8644% | 0.7997% | 667.8% | 27.21% | 22.61% | 33.51%
179 | LBLANK |  | 6/17/2023 6:18:40 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.828 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 176.7% | 173.2% | 27.87% | 25.35% | 48.59% | 877.8% | 125.6% | 200.5% | 76.43% | 22.17% | 34.43% | 209.4% | 14.65% | 381.6% | 1065% | 200.6% | 49.51% | 1.595% | 28.73% | 439.4% | 34.05% | 20.41% |  | 389.6% | 153.1% | 293.5% | 2.265e-06% | 8.935% | 29.11% | 1.018% | 0.2777% | 0.4676% | 2.53% | 0.6979% | 1.457% | 113.4% | 15.87% | 33.6% | 63.36%
180 | CCV 1ppb TRACE |  | 6/17/2023 6:24:40 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.829 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4.763% | 7.165% | 2.938% | 2.32% | 10.41% | 2.877% | 4.798% | 3.44% | 2.651% | 1.971% | 3.864% | 0.9751% | 11.24% | 3.139% | 2.51% | 3.529% | 1.988% | 2.341% | 7.759% | 5.863% | 1.487% | 9.284% | 4.926% | 6.959% | 6.793% | 13.04% | 19.22% | 0.6042% | 1.952% | 1.844% | 0.3839% | 0.2344% | 0.6877% | 0.6185% | 0.492% | 1.71% | 1.027% | 2.032% | 1.103%
181 | CCV 1ppb TRACE |  | 6/17/2023 6:31:00 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.830 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.635% | 6.113% | 10.35% | 2.667% | 3.266% | 3.526% | 7.278% | 1.791% | 4.868% | 1.284% | 2.786% | 2.709% | 1.159% | 1.726% | 1.398% | 2.199% | 0.7486% | 3.737% | 4.877% | 5.824% | 1.737% | 11.68% | 9.237% | 12.76% | 5.92% | 9.361% | 16.67% | 1.768% | 0.7222% | 1.818% | 0.9932% | 0.4423% | 0.5217% | 1.183% | 1.402% | 1.146% | 2.022% | 3.796% | 0.8782%
182 | LFB |  | 6/17/2023 6:37:00 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.831 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.3749% | 1.303% | 7.924% | 1.944% | 2.1% | 1.652% | 0.8032% | 1.808% | 1.105% | 2.624% | 1.413% | 0.7119% | 2.721% | 0.5264% | 1.769% | 2.131% | 1.662% | 3.737% | 1.955% | 3.046% | 1.05% | 1.16% | 3.362% | 1.31% | 0.7819% | 2.27% | 2.497% | 1.138% | 0.9901% | 0.2936% | 0.09303% | 1.796% | 0.6008% | 0.2231% | 1.394% | 2.028% | 1.939% | 1.143% | 0.7656%
183 | LFB |  | 6/17/2023 6:43:18 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.832 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.973% | 1.206% | 1.236% | 1.056% | 1.755% | 3.021% | 2.011% | 3.285% | 1.442% | 0.1805% | 0.5378% | 0.5927% | 2.06% | 1.289% | 1.443% | 2.787% | 0.5666% | 0.445% | 1.067% | 4.661% | 1.291% | 2.599% | 5.215% | 3.664% | 0.5587% | 1.543% | 3.422% | 0.4209% | 0.6971% | 0.2805% | 0.7076% | 0.7562% | 3.879% | 1.099% | 1.651% | 0.6314% | 0.4677% | 0.3099% | 5.153%
184 | CCV 10 ppb MAJORS |  | 6/17/2023 6:49:17 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.833 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.254% | 1.262% | 4.954% | 0.6241% | 1.987% | 1.496% | 2.707% | 1.148% | 0.4314% | 1.366% | 8.574% | 1.02% | 3.273% | 0.9114% | 1.451% | 1.107% | 0.6878% | 0.1319% | 3.532% | 1.164% | 0.5965% | 3.763% | 5.938% | 2.004% | 0.933% | 4.079% | 3.306% | 0.347% | 0.3778% | 2.554% | 0.4529% | 0.4295% | 0.9437% | 0.3084% | 0.8975% | 1.672% | 0.5652% | 2.363% | 1.48%
185 | CCV 10 ppb MAJORS |  | 6/17/2023 6:55:37 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.834 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.621% | 5.241% | 7.026% | 2.375% | 3.404% | 1.26% | 2.374% | 2.421% | 2.159% | 0.1906% | 1.962% | 1.025% | 3.734% | 1.826% | 1.642% | 1.252% | 0.4349% | 1.182% | 2.39% | 8.073% | 0.9846% | 1.855% | 4.697% | 2.268% | 0.378% | 1.493% | 5.256% | 0.5069% | 0.3829% | 0.6864% | 0.1965% | 1.142% | 0.3485% | 0.7907% | 1.439% | 0.747% | 1.158% | 1.089% | 1.055%
186 | 10_15_21_UF_8_DUP |  | 6/17/2023 7:01:37 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_8_DUP.835 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.865% | 3.993% | 16.54% | 2.881% | 3.098% | 2.38% | 3.685% | 3.846% | 13.81% | 1.821% | 0.4532% | 0.5931% | 5.223% | 19.47% | 9.596% | 2.498% | 1.298% | 0.9585% | 3.32% | 7.574% | 2.571% | 26.48% | 66.74% | 31.2% | 13.81% | 17.37% | 14.58% | 0.5161% | 0.8114% | 0.1749% | 0.2182% | 1.236% | 1.021% | 0.4684% | 2.809% | 37.3% | 6.132% | 1.762% | 2.786%
187 | 10_15_21_UF_8_DUP |  | 6/17/2023 7:07:56 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_8_DUP.836 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.719% | 2.83% | 2.884% | 3.039% | 1.356% | 3.237% | 1.65% | 5.969% | 9.042% | 4.109% | 2.896% | 1.992% | 1.955% | 12.86% | 10.93% | 3.9% | 1.151% | 3.003% | 0.7604% | 5.619% | 1.755% | 25.49% | 0.514% | 40.67% | 21.05% | 27.46% | 52.32% | 0.4943% | 1.553% | 0.5484% | 0.5125% | 1.34% | 0.9152% | 0.5244% | 1.625% | 11.07% | 24.4% | 3.546% | 3.261%
188 | 10_15_21_UF_8_DUP_LSP |  | 6/17/2023 7:13:55 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_8_DUP_LSP.837 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.777% | 0.2693% | 5.703% | 5.413% | 0.1913% | 1.56% | 1.742% | 0.5517% | 1.671% | 0.7577% | 0.2585% | 0.9826% | 1.383% | 1.884% | 1.689% | 0.7931% | 0.6083% | 0.4918% | 0.5871% | 0.8738% | 0.4298% | 1.235% | 7.387% | 1.784% | 2.237% | 2.232% | 2.218% | 0.8136% | 0.4757% | 0.4237% | 0.3315% | 0.4178% | 1.069% | 0.8426% | 1.13% | 0.3112% | 0.8816% | 1.51% | 0.3766%
189 | 10_15_21_UF_8_DUP_LSP |  | 6/17/2023 7:20:13 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_8_DUP_LSP.838 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.857% | 2.039% | 3.162% | 1.204% | 1.213% | 1.753% | 1.722% | 1.906% | 1.928% | 1.483% | 0.6229% | 0.6074% | 0.5072% | 1.57% | 2.005% | 1.205% | 0.5694% | 1.155% | 3.132% | 5.385% | 1.163% | 5.301% | 3.18% | 3.569% | 0.6459% | 4.096% | 3.329% | 1.151% | 0.05412% | 2.649% | 0.4814% | 0.8073% | 0.515% | 0.4411% | 1.224% | 0.3916% | 0.5217% | 3.731% | 1.137%
190 | 10_15_21_UF_9 |  | 6/17/2023 7:26:13 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_9.839 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.033% | 2.016% | 3.528% | 0.9727% | 1.764% | 2.696% | 0.3651% | 2.715% | 4.873% | 1.173% | 1.414% | 0.6749% | 1.961% | 25.83% | 3.395% | 0.739% | 2.268% | 0.3228% | 0.6776% | 14.84% | 1.226% | 18.96% | 22.15% | 30.42% | 28.18% | 22.71% | 61.61% | 1.288% | 1.045% | 0.2182% | 0.6003% | 0.7352% | 0.3603% | 1.181% | 0.8849% | 31.27% | 19.68% | 1.537% | 3.877%
191 | 10_15_21_UF_9 |  | 6/17/2023 7:32:31 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_9.840 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.411% | 2.225% | 1.34% | 2.197% | 3.006% | 2.5% | 2.315% | 7.075% | 16.15% | 3.298% | 2.264% | 1.768% | 3.321% | 19.84% | 19.21% | 1.688% | 1.591% | 1.745% | 3.701% | 7.065% | 1.946% | 12.52% | 68.86% | 21.56% | 25.39% | 29.67% | 70.37% | 0.3902% | 1.215% | 0.9742% | 0.5266% | 1.202% | 0.3178% | 0.5594% | 1.763% | 12.61% | 7% | 8.983% | 3.821%
192 | 10_15_21_UF_10 |  | 6/17/2023 7:38:31 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_10.841 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.006% | 0.8705% | 18.96% | 1.884% | 1.609% | 0.2636% | 1.189% | 4.573% | 8.512% | 1.536% | 2.759% | 2.113% | 5.75% | 23.2% | 8.147% | 3.874% | 2.06% | 0.3378% | 2.261% | 6.186% | 1.633% | 18.18% | 56.12% | 36.33% | 5.918% | 1.781% | 29.32% | 2.526% | 2.254% | 0.1584% | 1.186% | 0.9719% | 14.82% | 0.1394% | 0.3658% | 19.31% | 6.224% | 2.484% | 22.87%
193 | 10_15_21_UF_10 |  | 6/17/2023 7:44:50 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_10.842 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.9737% | 1.606% | 27.55% | 2.741% | 2.633% | 3.233% | 2.519% | 4.684% | 9.141% | 6.886% | 1.444% | 1.321% | 4.927% | 9.436% | 9.179% | 1.515% | 0.5605% | 1.049% | 1.238% | 5.01% | 3.127% | 10.57% | 22.03% | 19.41% | 35.93% | 44.36% | 57.92% | 1.16% | 0.7293% | 0.7467% | 0.8777% | 0.2577% | 2.528% | 0.6488% | 1.17% | 7.655% | 26.99% | 1.992% | 2.025%
194 | 10_15_21_UF_11 |  | 6/17/2023 7:50:50 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_11.843 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.844% | 1.84% | 4.517% | 1.779% | 3.603% | 4.25% | 1.25% | 2.828% | 0.2851% | 5.771% | 4.159% | 1.186% | 3.38% | 20.17% | 10.02% | 0.4753% | 0.3712% | 0.2367% | 0.9736% | 2.673% | 1.42% | 20.02% | 125.1% | 27.67% | 11.84% | 17.95% | 75.89% | 0.9605% | 1.415% | 0.505% | 0.2873% | 0.1994% | 3.137% | 0.8055% | 0.9374% | 55.62% | 12.5% | 4.614% | 4.893%
195 | 10_15_21_UF_11 |  | 6/17/2023 7:57:09 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_11.844 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.611% | 3.456% | 4.698% | 1.477% | 2.121% | 2.078% | 2.095% | 8.549% | 3.753% | 16.98% | 0.7066% | 0.3526% | 3.318% | 12.3% | 21.04% | 3.006% | 1.845% | 1.351% | 2.054% | 7.918% | 1.05% | 9.619% | 77.75% | 27.89% | 20.36% | 30% | 79.85% | 1.033% | 1.216% | 0.1615% | 0.3486% | 0.1205% | 0.1712% | 1.083% | 1.883% | 19.84% | 14.37% | 2.121% | 1.709%
196 | LBLANK |  | 6/17/2023 8:03:10 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.845 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 184.6% | 173.2% | 13.61% | 14.34% | 20.69% | 125% | 0% | 40.96% | 48.63% | 165.3% | 75.86% | 269.7% | 13.97% | 0% | 258.3% | 1.055% | 54.94% | 5.117% | 104% | 413.3% | 28.6% | 53.66% |  | 101.9% | 495.8% | 373% | 182.3% | 8.637% | 54.38% | 3.274% | 0.9434% | 4.597% | 1.051% | 0.982% | 0.5099% | 0% | 59.44% | 11.36% | 40.75%
197 | LBLANK |  | 6/17/2023 8:09:28 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.846 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 55.22% |  | 22.09% | 7.618% | 15.79% | 88.86% | 51.55% | 39.97% | 102.8% | 39.03% | 69.12% | 335.8% | 36.37% | 472% | 98.1% | 495.3% | 176% | 1.379% | 132.8% | 123% | 14.64% | 111.1% | 173.2% | 120.8% | 1113% | 115.6% | 3945% | 3.977% | 12.11% | 11.6% | 0.3508% | 3.078% | 0.775% | 1.091% | 0.5998% | 118.5% | 26.98% | 90.41% | 34.87%
198 | CCV 1ppb TRACE |  | 6/17/2023 8:15:29 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.847 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.562% | 4.662% | 1.543% | 9.42% | 11.49% | 2.848% | 2.607% | 1.023% | 3.35% | 0.6648% | 2.958% | 2.9% | 4.82% | 2.09% | 8.39% | 2.582% | 3.578% | 0.3717% | 5.691% | 19.32% | 2.9% | 7.151% | 15.44% | 8.323% | 1.675% | 20.96% | 14.72% | 2.618% | 0.7795% | 0.4053% | 0.1566% | 0.143% | 0.3961% | 0.2567% | 0.4616% | 1.369% | 1.671% | 0.5638% | 1.317%
199 | CCV 1ppb TRACE |  | 6/17/2023 8:21:48 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.848 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.574% | 3.154% | 16.57% | 5.968% | 13.28% | 5.56% | 2.271% | 4.368% | 3.85% | 1.659% | 2.349% | 1.149% | 2.432% | 0.9318% | 1.069% | 3.433% | 1.909% | 2.996% | 3.183% | 5.07% | 1.188% | 11.34% | 25.32% | 6.287% | 9.421% | 7.562% | 19.75% | 0.7087% | 1.538% | 0.7409% | 0.5037% | 0.6126% | 0.645% | 0.4516% | 0.822% | 0.3409% | 2.281% | 0.3926% | 1.556%
200 | CCV 10 ppb MAJORS |  | 6/17/2023 8:27:48 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.849 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.2697% | 1.064% | 9.36% | 1.554% | 1.787% | 1.612% | 3.507% | 0.8178% | 1.587% | 0.4366% | 1.048% | 0.3435% | 1.576% | 0.5768% | 2.167% | 0.6739% | 0.3921% | 0.1115% | 1.794% | 6.135% | 1.169% | 0.4463% | 5.332% | 1.069% | 2.03% | 1.758% | 0.5246% | 0.316% | 1.731% | 0.09716% | 0.8362% | 0.7998% | 0.1872% | 1.181% | 0.7117% | 0.517% | 0.3037% | 1.527% | 0.8679%
201 | CCV 10 ppb MAJORS |  | 6/17/2023 8:34:07 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.850 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.347% | 1.933% | 4.073% | 2.605% | 0.9413% | 1.11% | 1.38% | 1.505% | 1.338% | 1.003% | 2.155% | 1.669% | 0.9055% | 2.215% | 3.986% | 0.7508% | 0.9425% | 1.345% | 2.739% | 3.076% | 1.196% | 4.506% | 3.114% | 5.702% | 1.851% | 3.69% | 3.074% | 0.7198% | 1.975% | 2.453% | 0.9137% | 0.7203% | 0.4736% | 0.2292% | 0.7842% | 0.3928% | 0.8561% | 4.914% | 1.362%
202 | 10_15_21_FB |  | 6/17/2023 8:40:08 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_FB.851 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10.95% | 32.94% | 7.339% | 4.448% | 13.41% | 63.81% | 54.82% | 126.9% | 5.795% | 5.753% | 0.9622% | 11.94% | 1.646% | 70.06% | 47.28% | 9.537% | 6.721% | 0.6885% | 57.34% | 429% | 24.01% | 43.01% |  | 114.4% | 1159% | 505.8% | 2.265e-06% | 41.97% | 5.453% | 0.1586% | 11.15% | 0.5913% | 0.3092% | 0.5303% | 1.478% | 175.2% | 90.95% | 4.061% | 4.804%
203 | 10_15_21_FB |  | 6/17/2023 8:46:26 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_FB.852 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4.805% | 34.64% | 1.108% | 2.364% | 22.21% | 52.01% | 78.13% | 125.3% | 9.468% | 7.139% | 1.539% | 6.97% | 4.421% | 206.5% | 59.11% | 17.62% | 6.645% | 0.9449% | 42.52% | 410.1% | 10.14% | 31.56% | 173.2% | 232.4% | 191.2% | 646.9% | 2.265e-06% | 19.95% | 23.77% | 0.9108% | 0.6744% | 0.6422% | 0.3913% | 0.1658% | 0.6935% | 511.6% | 3257% | 7.11% | 2.662%
204 | 10_15_21_BB |  | 6/17/2023 8:52:26 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_BB.853 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 118.6% | 86.65% | 14.09% | 21.1% | 9.088% | 258.7% | 53.67% | 112.4% | 29.01% | 12.8% | 13.2% | 68% | 29.87% | 69.8% | 46.65% | 14.55% | 3.185% | 6.163% | 83.52% | 203% | 15.7% | 66.68% | 173.2% | 116.1% | 573.9% | 361.9% | 2.265e-06% | 5.137% | 89.78% | 0.1579% | 0.6144% | 1.012% | 0.5925% | 1.232% | 1.993% | 130.2% | 48.19% | 6.365% | 6.976%
205 | 10_15_21_BB |  | 6/17/2023 8:58:45 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_BB.854 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 59.17% |  | 19.55% | 4.199% | 6.044% | 390.9% | 84.33% | 0% | 26.29% | 11.35% | 0.9335% | 40.21% | 19.52% | 48.32% | 92.03% | 6.216% | 1.688% | 23.07% | 130.3% | 0% | 15.16% | 50.45% |  | 121% | 829.6% | 465.3% | 49640% | 7.693% | 74.98% | 0.4912% | 1.114% | 1.255% | 1.123% | 0.5711% | 0.7145% | 0% | 14.25% | 4.387% | 3.648%
206 | 10_15_21_LAB_A |  | 6/17/2023 9:04:45 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_LAB_A.855 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 64.08% |  | 22.85% | 50.04% | 27.04% | 123.9% | 122.8% | 0% | 20.52% | 6.443% | 4.745% | 15.98% | 25.07% | 47.78% | 1045% | 4.266% | 4.598% | 29.65% | 38.6% | 0% | 55.13% | 37.38% |  | 2.265e-06% | 49.05% | 148.3% | 2.265e-06% | 17.05% | 43.78% | 0.5217% | 0.2406% | 4.992% | 2.44% | 1.429% | 1.157% | 943.6% | 52.33% | 7.397% | 6.067%
207 | 10_15_21_LAB_A |  | 6/17/2023 9:11:03 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_LAB_A.856 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 104.7% |  | 19.86% | 68.44% | 23.52% | 583.2% | 50.84% | 136.8% | 8.188% | 1.375% | 0.9304% | 11.97% | 11.57% | 184% | 231.4% | 11.44% | 6.755% | 11.62% | 50.03% | 411.9% | 54.22% | 13.69% |  | 352.6% | 1.069% | 87.31% | 2.265e-06% | 25.55% | 13.34% | 0.4016% | 0.5544% | 0.3027% | 0.7753% | 0.5973% | 0.2759% | 952% | 95.3% | 8.321% | 4.713%
208 | 10_15_21_LAB_B |  | 6/17/2023 9:17:03 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_LAB_B.857 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 113.6% | 173.2% | 30.28% | 12.37% | 8.991% | 1472% | 123.8% | 0% | 9.336% | 14.85% | 6.687% | 45.28% | 27.68% | 191.6% | 83.24% | 7.78% | 52.58% | 5.253% | 39.27% | 120.8% | 244.6% | 78.42% |  | 230.7% | 203.6% | 62.48% | 11920% | 6.451% | 89% | 0.09203% | 0.7893% | 0.1261% | 0.6512% | 1.254% | 1.026% | 132.4% | 26.45% | 2.45% | 4.714%
209 | 10_15_21_LAB_B |  | 6/17/2023 9:23:20 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_LAB_B.858 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 176.7% |  | 16.95% | 24.1% | 59.03% | 37.33% | 214.5% | 1.916% | 11.1% | 11.04% | 7.33% | 25.45% | 14.59% | 621.2% | 195% | 20.89% | 7.246% | 6.034% | 133.7% | 0% | 49.3% | 5.509% | 173.2% | 2.265e-06% | 0% | 18.22% | 9870% | 0.8837% | 93.47% | 0.238% | 0.5297% | 1.005% | 0.773% | 0.5101% | 1.359% | 997% | 54.19% | 6.522% | 2.13%
210 | 5_18_23_DB |  | 6/17/2023 9:29:20 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_DB.859 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 14870% |  | 67.49% | 3.68% | 2.788% | 158.1% | 51.07% | 0% | 44.4% | 17.57% | 12.27% | 193.8% | 13.93% | 321% | 102.5% | 5.47% | 0.6581% | 2.453% | 46.9% | 0% | 69.47% | 131.4% |  | 2.265e-06% | 70.03% | 118.2% | 2.265e-06% | 3.572% | 450% | 1.367% | 0.8639% | 0.2967% | 0.9445% | 1.284% | 0.6405% | 1193% | 8.138% | 8.079% | 2.17%
211 | 5_18_23_DB |  | 6/17/2023 9:35:39 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_DB.860 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 316% |  | 101.1% | 1.71% | 36.75% | 16.24% | 134% | 1.543% | 34.08% | 34.11% | 5.769% | 42.58% | 49.25% | 199.8% | 414% | 9.792% | 6.242% | 27.24% | 60.66% | 398.5% | 179.1% | 25.13% |  | 341.9% | 50.35% | 13.12% | 345.3% | 7.354% | 186.5% | 1.684% | 1.583% | 0.691% | 0.3408% | 1.09% | 1.047% | 430.8% | 36.39% | 5.192% | 1.126%
212 | LBLANK |  | 6/17/2023 9:41:39 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.861 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 60.58% |  | 16.74% | 3.45% | 50.44% | 1942% | 119.8% | 115.2% | 94.06% | 111% | 50.83% | 216.6% | 15.9% | 371.3% | 55.35% | 93.68% | 413.9% | 7.081% | 70.75% | 384.8% | 52.37% | 227.7% |  | 228.9% | 72.03% | 766.7% | 6905% | 6.238% | 45.96% | 6.235% | 0.6584% | 0.5743% | 0.1421% | 0.631% | 0.6248% | 662.6% | 20.14% | 19.18% | 87.42%
213 | LBLANK |  | 6/17/2023 9:47:59 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.862 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 57.67% |  | 8.365% | 64.56% | 88.69% | 84.96% | 394.1% | 0% | 61.52% | 41.51% | 22.86% | 72.79% | 80.16% | 69.92% | 171.3% | 178.3% | 2108% | 4.025% | 82.94% | 0% | 355.7% | 124.9% | 173.2% | 2.265e-06% | 0% | 35.72% | 2.265e-06% | 5.69% | 15.97% | 3.775% | 0.2978% | 0.3872% | 0.5834% | 0.4001% | 2.111% | 135.7% | 23.2% | 86.21% | 58.38%
214 | CCV 1ppb TRACE |  | 6/17/2023 9:53:59 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.863 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.763% | 2.611% | 18.99% | 8.224% | 10.92% | 4.987% | 7.593% | 1.056% | 5.492% | 1.408% | 2.264% | 0.9288% | 7.772% | 2.004% | 7.331% | 2.865% | 2.235% | 1.135% | 8.332% | 15.18% | 1.523% | 4.681% | 17.02% | 14.26% | 4.196% | 6.594% | 23.85% | 1.714% | 2.667% | 0.5838% | 1.076% | 1.098% | 0.8227% | 0.8006% | 2.185% | 2.474% | 2.289% | 1.298% | 1.636%
215 | CCV 1ppb TRACE |  | 6/17/2023 10:00:19 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.864 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.982% | 1.496% | 14.03% | 6.096% | 9.317% | 6.765% | 8.01% | 2.499% | 1.224% | 7.436% | 10.34% | 0.9958% | 5.605% | 0.3654% | 5.358% | 2.658% | 8.154% | 10.31% | 2.258% | 12.55% | 2.485% | 13.09% | 14.93% | 19.07% | 2.991% | 19.44% | 13.28% | 11.56% | 2.433% | 0.6072% | 0.7237% | 7.333% | 0.6219% | 0.8248% | 1.009% | 9.973% | 9.674% | 1.845% | 0.3796%
216 | LFB |  | 6/17/2023 10:06:18 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.865 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.562% | 2.687% | 7.865% | 0.3101% | 1.654% | 2.522% | 4.856% | 1.761% | 3.707% | 0.9374% | 1.075% | 0.6715% | 2.67% | 1.541% | 1.777% | 2.43% | 0.3662% | 0.7081% | 1.672% | 4.352% | 0.6176% | 5.311% | 9.392% | 2.174% | 2.602% | 1.502% | 6.995% | 1.206% | 0.6867% | 0.5999% | 0.4423% | 1.127% | 0.5146% | 0.2978% | 1.436% | 0.6129% | 0.1443% | 0.8389% | 0.5001%
217 | LFB |  | 6/17/2023 10:12:36 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.866 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.54% | 0.5538% | 9.565% | 1.837% | 5.6% | 1.141% | 1.607% | 2.04% | 1.387% | 3.165% | 4.473% | 1.192% | 1.518% | 0.7671% | 1.404% | 1.387% | 3.709% | 3.217% | 0.5059% | 3.487% | 0.4909% | 4.714% | 9.765% | 4.339% | 3.259% | 3.501% | 6.001% | 3.566% | 0.4086% | 1.821% | 1.218% | 2.141% | 0.6679% | 0.7912% | 0.5973% | 4.553% | 4.468% | 2.096% | 0.5973%
218 | CCV 10 ppb MAJORS |  | 6/17/2023 10:18:36 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.867 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.703% | 0.8758% | 0.6375% | 1.112% | 2.763% | 1.497% | 2.122% | 2.2% | 2.156% | 1.167% | 1.196% | 1.536% | 2.698% | 1.103% | 1.042% | 2.544% | 0.642% | 1.141% | 0.5778% | 4.624% | 1.095% | 0.8349% | 2.879% | 1.346% | 2.3% | 1.846% | 2.353% | 0.6817% | 1.139% | 0.8282% | 0.7479% | 1.438% | 0.6097% | 0.7542% | 2.072% | 1.437% | 1.729% | 1.442% | 1.189%
219 | CCV 10 ppb MAJORS |  | 6/17/2023 10:24:55 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.868 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.461% | 2.865% | 4.144% | 1.385% | 3.856% | 1.982% | 3.188% | 2.503% | 1.559% | 1.41% | 1.133% | 0.3845% | 2.442% | 3.211% | 2.027% | 2.122% | 0.8121% | 0.891% | 1.083% | 1.679% | 0.9736% | 1.795% | 5.337% | 3.203% | 1.156% | 0.8082% | 6.213% | 1.003% | 0.4736% | 0.5097% | 0.1366% | 0.1719% | 0.6183% | 0.7821% | 2.089% | 0.9214% | 0.3083% | 1.425% | 1.045%
220 | 5_18_23_DB |  | 6/17/2023 10:30:55 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_DB.869 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 105.1% |  | 27.53% | 1.671% | 48.58% | 4635% | 135.1% | 131.4% | 19.27% | 23.52% | 2.491% | 25.43% | 8.688% | 35.49% | 56.1% | 14.88% | 3.957% | 53.8% | 47.25% | 406.9% | 71.36% | 76.13% |  | 131.3% | 569.6% | 1610% | 151.6% | 154.5% | 34.13% | 3.284% | 0.3298% | 1.188% | 2.003% | 0.9762% | 1.486% | 0% | 60.51% | 10.12% | 3.497%
221 | 5_18_23_DB |  | 6/17/2023 10:37:14 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_DB.870 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 238.1% | 173.2% | 4.065% | 1.827% | 4.563% | 267.3% | 122.6% | 41.11% | 9.418% | 19.71% | 2.168% | 31.86% | 12.71% | 318% | 74.68% | 14.01% | 5.782% | 40.72% | 82.5% | 122.6% | 81.72% | 37.7% | 173.2% | 352% | 48.12% | 32.32% | 176.6% | 31.28% | 49.95% | 1.134% | 0.6438% | 0.9004% | 0.3227% | 0.3216% | 0.9927% | 123.9% | 18.98% | 8.921% | 7.061%
222 | 5_18_23_DB |  | 6/17/2023 10:43:14 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_DB.871 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 372.6% |  | 25.02% | 2.2% | 13.88% | 54.67% | 102% | 106.4% | 19.7% | 9.467% | 11.66% | 23.99% | 36.61% | 8.874% | 155.9% | 13.16% | 5.258% | 17.47% | 37.6% | 122% | 41.72% | 29.85% |  | 115.6% | 194.2% | 540% | 31547% | 10.57% | 55.58% | 15.75% | 1.103% | 0.1436% | 0.6102% | 0.5082% | 0.5605% | 488.1% | 21.71% | 19.74% | 2.296%
223 | 5_18_23_DB |  | 6/17/2023 10:49:33 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_DB.872 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 115.8% |  | 34.89% | 0.2935% | 36.67% | 104.9% | 1.158% | 110% | 42% | 17.61% | 2.522% | 44.96% | 10.55% | 9.548% | 110.3% | 11.78% | 5.701% | 35.68% | 59.04% | 0% | 111.2% | 26.61% | 86.61% | 348% | 347.1% | 419.4% | 2.265e-06% | 12.66% | 35.64% | 1.045% | 0.8654% | 0.355% | 11.84% | 1.151% | 0.9395% | 455.2% | 22.23% | 0.7375% | 17.82%
224 | 5_18_23_DB |  | 6/17/2023 10:55:33 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_DB.873 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 103.5% |  | 57% | 1.559% | 44.3% | 781.6% | 213.9% | 42.09% | 36.75% | 7.858% | 21.33% | 34.09% | 21.91% | 314.1% | 238.7% | 3.509% | 11.03% | 24.38% | 74.31% | 0% | 841.2% | 6.204% |  | 2.265e-06% | 1004% | 375% | 2.265e-06% | 16.42% | 90.74% | 1.29% | 0.4664% | 0.6213% | 0.4251% | 1.018% | 0.8267% | 204.8% | 50.77% | 8.438% | 10.11%
225 | 5_18_23_DB |  | 6/17/2023 11:01:51 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_DB.874 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 102.9% |  | 87.9% | 3.302% | 31.63% | 515.5% | 40.16% | 0% | 43.7% | 34.96% | 22.71% | 12.22% | 13.35% | 354.2% | 189.9% | 12.99% | 4.387% | 67.07% | 28.06% | 124.3% | 65.42% | 93.93% |  | 138.5% | 795% | 188.8% | 2.265e-06% | 10.92% | 113% | 0.4389% | 0.1387% | 1.361% | 0.1499% | 0.3278% | 1.918% | 137.1% | 61.71% | 3.383% | 6.19%
226 | 5_18_23_MFB |  | 6/17/2023 11:07:51 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_MFB.875 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.432% | 3.422% | 6.348% | 1.198% | 2.863% | 1.935% | 1.252% | 2.341% | 2.497% | 0.8721% | 0.7998% | 1.97% | 0.7885% | 1.306% | 2.165% | 1.982% | 0.7451% | 0.3855% | 2.823% | 7.606% | 0.8316% | 3.608% | 4.506% | 2.063% | 4.964% | 1.309% | 2.681% | 0.5386% | 2.683% | 0.8757% | 2.025% | 1.277% | 0.4159% | 1.774% | 0.576% | 1.23% | 1.477% | 0.8246% | 1.73%
227 | 5_18_23_MFB |  | 6/17/2023 11:14:09 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_MFB.876 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.986% | 1.504% | 6.124% | 0.994% | 5.283% | 2.374% | 0.885% | 2.284% | 1.997% | 1.298% | 0.9174% | 1.635% | 0.5375% | 2.621% | 2.944% | 4.28% | 0.5587% | 0.6122% | 1.912% | 6.723% | 0.7605% | 5.249% | 14.86% | 0.3806% | 2.305% | 5.893% | 5.157% | 0.1383% | 0.4085% | 1.223% | 1.105% | 0.5077% | 0.9751% | 0.09416% | 0.9342% | 1.245% | 0.3788% | 2.589% | 1.301%
228 | 5_18_23_MFB |  | 6/17/2023 11:20:08 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_MFB.877 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.3973% | 1.626% | 8.258% | 0.6226% | 4.271% | 3.382% | 3.489% | 3.998% | 1.875% | 1.666% | 1.18% | 0.9603% | 2.615% | 0.6166% | 0.2325% | 2.047% | 0.9698% | 0.8973% | 1.334% | 6.075% | 2.907% | 6.704% | 5.779% | 1.181% | 2.964% | 6.058% | 7.92% | 0.7479% | 0.4238% | 0.6761% | 0.8451% | 0.1682% | 5.476% | 0.5712% | 0.8766% | 1.382% | 1.207% | 0.2192% | 5.48%
229 | 5_18_23_MFB |  | 6/17/2023 11:26:26 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_MFB.878 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.279% | 4.854% | 5.068% | 0.8736% | 2.159% | 2.576% | 1.508% | 2.371% | 1.331% | 1.378% | 1.419% | 0.2227% | 0.3136% | 1.657% | 3.9% | 0.5079% | 0.4493% | 0.2843% | 3.802% | 6.487% | 2.471% | 9.316% | 3.468% | 5.632% | 7.665% | 4.314% | 7.335% | 0.7457% | 1.045% | 3.141% | 0.5436% | 0.4914% | 0.5853% | 1.073% | 0.5426% | 0.6245% | 0.404% | 3.079% | 1.229%
230 | LBLANK |  | 6/17/2023 11:32:26 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.879 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.982% | 173.2% | 11.02% | 4.01% | 30.47% | 2.169% | 403.7% | 160.1% | 15.32% | 46.6% | 77.15% | 174.8% | 26.9% | 67.81% | 116.2% | 113.8% | 775.1% | 7.728% | 77.72% | 0% | 60.25% | 82.72% |  | 115.3% | 342.9% | 819.8% | 62370% | 290.5% | 28.54% | 2.088% | 0.5003% | 1.362% | 0.3981% | 0.5041% | 0.6416% | 133.8% | 36.29% | 28.47% | 83.44%
231 | LBLANK |  | 6/17/2023 11:38:51 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.880 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10168% |  | 5.139% | 17.91% | 43.55% | 78.4% | 404% | 110.3% | 32.08% | 0.2848% | 9.583% | 42.68% | 30.26% | 66.31% | 353.5% | 651.2% | 29.11% | 5.201% | 55.33% | 0% | 134.9% | 25.84% |  | 2.265e-06% | 968.9% | 506.8% | 2.265e-06% | 41.25% | 17.59% | 1.885% | 0.4801% | 0.1343% | 0.5602% | 0.4477% | 0.5455% | 0% | 18.06% | 16.91% | 43.25%
232 | CCV 1ppb TRACE |  | 6/17/2023 11:44:51 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.881 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.511% | 5.958% | 9.425% | 2.16% | 7.023% | 4.07% | 1.192% | 1.765% | 8.044% | 2.825% | 4.464% | 0.171% | 4.699% | 1.214% | 2.453% | 2.451% | 2.129% | 2.93% | 8.96% | 13.61% | 5.188% | 11.67% | 21.63% | 5.998% | 3.074% | 2.158% | 6.685% | 1.827% | 1.401% | 0.7664% | 0.2276% | 1.049% | 0.4844% | 0.9052% | 0.5254% | 1.024% | 0.8526% | 0.3766% | 1.378%
233 | CCV 1ppb TRACE |  | 6/17/2023 11:51:11 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.882 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.6303% | 2.68% | 2.338% | 2.02% | 7.818% | 4.142% | 6.076% | 1.115% | 2.984% | 3.499% | 2.553% | 0.9863% | 6.998% | 2.023% | 4.429% | 2.84% | 1.753% | 3.081% | 6.869% | 8.966% | 1.448% | 1.589% | 13.81% | 18.28% | 4.774% | 2.398% | 19.4% | 0.729% | 1.407% | 0.2555% | 0.2396% | 1.031% | 0.4663% | 0.7229% | 0.5741% | 3.375% | 1.046% | 1.809% | 1.283%
234 | CCV 10 ppb MAJORS |  | 6/17/2023 11:57:11 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.883 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.283% | 1.424% | 4.869% | 0.5222% | 4.939% | 1.296% | 0.5212% | 1.958% | 0.1686% | 1.364% | 2.691% | 0.7222% | 0.2226% | 0.2676% | 3.043% | 0.8966% | 1.099% | 0.5196% | 2.612% | 4.67% | 1.636% | 2.347% | 7.065% | 3.876% | 2.808% | 4.307% | 6.95% | 0.3074% | 0.6465% | 1.104% | 0.5806% | 0.6938% | 1.186% | 0.5447% | 0.7936% | 1.944% | 0.8455% | 1.896% | 1.627%
235 | CAL BLANK |  | 6/17/2023 12:25:05 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CAL BLANK.884 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth

## Concentrations
| Sample Id | R | Acquisition Time | QC Status | Dataset File | Method File | Na KED 23/23 Helium KED_LowLow (ppb) | Mg KED 24/24 Helium KED_LowLow (ppb) | Al KED 27/27 Helium KED_LowLow (ppb) | P Oshift 31/47 Oxygen DRC (ppb) | S (34) -1 34/34 Helium KED_LowLow (ppb) | K KED 39/39 Helium KED_LowLow (ppb) | Ca KED 43/43 Helium KED_LowLow (ppb) | V KED 51/51 Helium KED_LowLow (ppb) | Cr (52) 52/52 Helium KED_LowLow (ppb) | Cr (53) NH3 53/53 Ammonia DRC (ppb) | Fe (54) NH3 54/54 Ammonia DRC (ppb) | Mn Oshift 55/71 Oxygen DRC (ppb) | Fe (56) KED 56/56 Helium KED_LowLow (ppb) | Co KED 59/59 Helium KED_LowLow (ppb) | Ni KED 60/60 Helium KED_LowLow (ppb) | Cu (63) KED 63/63 Helium KED_LowLow (ppb) | Cu (65) NH3 65/65 Ammonia DRC (ppb) | Zn NH3 66/66 Ammonia DRC (ppb) | As 75/75 (ppb) | As KED 75/75 Helium KED_LowLow (ppb) | As Oshift 75/91 Oxygen DRC (ppb) | Se (77) 77/77 (ppb) | Se (77) Oshift 77/93 Oxygen DRC (ppb) | Se (78) Oshift 78/94 Oxygen DRC (ppb) | Se (80) Oshift 80/96 Oxygen DRC (ppb) | Se (82) 82/82 (ppb) | Se (82) Oshift 82/98 Oxygen DRC (ppb) | Mo (95) NH3 95/95 Ammonia DRC (ppb) | Mo (95) OShift 95/127 Oxygen DRC (ppb) | Rh (IS1) 103/103 (IS) | Rh (IS2) 103/103 (IS) Oxygen DRC | Rh (IS4) 103/103 (IS) Ammonia DRC | Rh (IS5) 103/103 (IS) Ammonia DRC | Rh (IS6) 103/103 (IS) Oxygen DRC | Rh (IS7) 103/103 (IS) Helium KED_LowLow | Cd (111) NH3 111/111 Ammonia DRC (ppb) | Cd (114) NH3 114/114 Ammonia DRC (ppb) | Pb 208/208 (ppb) | Pb NH3 208/208 Ammonia DRC (ppb)
1 | sac blank |  | 06/16/2023 11:58:56 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\sac blank .650 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth
2 | sac blank cal |  | 06/16/2023 12:04:57 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\sac blank cal.651 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth
3 | sac blank cal |  | 06/16/2023 12:10:57 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\sac blank cal.652 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth
4 | CAL BLANK |  | 06/16/2023 12:16:57 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CAL BLANK.653 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | [29 empty] | 100% | 100% | 100% | 100% | 100% | 100%
5 | LOW 0.01 |  | 06/16/2023 12:22:58 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LOW 0.01.654 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | [29 empty] | 100.4% | 100.2% | 100.2% | 99.9% | 98.19% | 101.8%
6 | LOW 0.05 |  | 06/16/2023 12:28:58 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LOW 0.05.655 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | [7 empty] | 0.05 | 0.05 | 0.05 |  |  |  | 0.05 | 0.05 | 0.05 | 0.05 |  |  | 0.05 | 0.05 | [8 empty] | 99.33% | 98.34% | 98.99% | 98.96% | 95.43% | 99.25% | 0.05 | 0.05 | 0.05 | 0.05
7 | LOW 0.1 |  | 06/16/2023 12:34:59 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LOW 0.1.656 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | [7 empty] | 0.091491 | 0.10172 | 0.097655 |  | 0.1 |  | 0.099402 | 0.10001 | 0.10345 | 0.10293 |  | 0.1 | 0.11294 | 0.096974 |  | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 98.4% | 98.02% | 100.3% | 99.99% | 96.31% | 99.11% | 0.098881 | 0.096188 | 0.10299 | 0.10182
8 | LOW 0.5 |  | 06/16/2023 12:40:59 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LOW 0.5.657 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth |  |  |  | 0.5 |  |  |  | 0.47899 | 0.54495 | 0.51826 |  | 0.49806 |  | 0.51852 | 0.49897 | 0.56111 | 0.56962 | 0.5 | 0.62329 | 0.56584 | 0.49801 | 0.5 | 0.55552 | 0.52402 | 0.51467 | 0.55061 | 0.55207 | 0.51936 | 0.50551 | 96.44% | 95.8% | 96.57% | 97.76% | 93.95% | 98.62% | 0.50456 | 0.49928 | 0.53645 | 0.5312
9 | STD 1 |  | 06/16/2023 12:46:58 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 1.658 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1000 | 1000 | 1 | 0.523 |  | 1000 | 1000 | 0.96784 | 1.0723 | 1.0004 | 1 | 0.99303 | 1 | 1.0119 | 0.97409 | 1.0383 | 1.0389 | 0.89245 | 1.0612 | 1.0017 | 0.98309 | 1.0627 | 0.99131 | 1.0406 | 1.0155 | 1.0001 | 1.0004 | 1.0047 | 1.0047 | 95.19% | 95.52% | 95.89% | 96.38% | 94.22% | 98.28% | 0.97751 | 0.97957 | 1.0241 | 1.0424
10 | STD 5 |  | 06/16/2023 12:52:58 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 5.659 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 5021.2304 | 5049.1253 | 5.6267 | 2.9923 | 5000 | 4966.6842 | 5131.8876 | 5.0136 | 5.3783 | 4.9234 | 4.6709 | 5.0395 | 4.7022 | 5.1448 | 4.8947 | 5.1802 | 5.0303 | 4.5954 | 5.5479 | 5.1629 | 4.9865 | 5.0238 | 5.1474 | 5.2097 | 4.9795 | 5.3795 | 5.329 | 5.03 | 4.9866 | 93.85% | 90.94% | 92.39% | 92.66% | 91.44% | 92.35% | 4.9537 | 4.9369 | 4.9653 | 5.1031
11 | STD 10 |  | 06/16/2023 12:58:58 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 10.660 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9999.5136 | 10083.0832 | 10.3725 | 5.7932 | 9888.1205 | 9980.4144 | 10192.9569 | 10.1267 | 10.5701 | 9.8979 | 9.8012 | 10.2265 | 9.6613 | 10.203 | 9.6861 | 10.1208 | 9.9206 | 9.4306 | 10.8636 | 10.0086 | 9.9641 | 10.6374 | 9.9419 | 10.1848 | 10.0889 | 10.3578 | 10.3331 | 10.0956 | 10.0999 | 93.11% | 89.26% | 90.58% | 91.62% | 89.43% | 91.1% | 9.8008 | 9.801 | 9.4087 | 9.9173
12 | STD 25 |  | 06/16/2023 13:04:58 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 25.661 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 25949.6012 | 25382.7433 | 26.4382 | 15.7537 | 25263.2589 | 25792.5995 | 25454.0498 | 25.9986 | 26.7776 | 26.1222 | 25.926 | 26.1442 | 25.0156 | 26.2376 | 24.8838 | 25.7058 | 25.6116 | 24.4674 | 27.7521 | 26.3788 | 26.3983 | 25.8426 | 26.8878 | 26.6488 | 26.8519 | 26.6799 | 27.4171 | 26.391 | 26.2628 | 92.8% | 84.29% | 84.33% | 86.04% | 93.17% | 87.43% | 25.3755 | 25.3304 | 21.5059 | 24.7359
13 | STD 50 |  | 06/16/2023 13:10:58 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 50.662 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 51920.6916 | 49981.3479 | 52.4153 | 35.8084 | 50235.6864 | 52623.4836 | 51975.4887 | 53.2049 | 54.0967 | 52.3163 | 50.8468 | 51.6675 | 52.3849 | 52.2692 | 49.0791 | 50.1304 | 49.6667 | 48.1366 | 54.8391 | 51.1131 | 52.4779 | 51.2563 | 53.1132 | 53.2205 | 52.8092 | 52.6819 | 53.9795 | 52.4782 | 52.5072 | 88.05% | 79.22% | 79.29% | 81.35% | 80.66% | 82.78% | 49.3176 | 48.9369 | 35.3187 | 46.7403
14 | STD 100 |  | 06/16/2023 13:16:57 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 100.663 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 108940.0898 | 104423.2102 | 109.062 | 76.9405 | 102894.999 | 111261.013 | 108386.8988 | 116.1759 | 116.2272 | 112.535 | 102.9693 | 106.2473 | 103.1562 | 109.9238 | 102.8191 | 102.6854 | 100.5096 | 96.8635 | 117.6454 | 106.6672 | 105.7244 | 107.4687 | 107.6878 | 109.1336 | 108.317 | 114.3442 | 111.5977 | 110.7929 | 108.8356 | 78.14% | 72.24% | 69.65% | 71.72% | 73.39% | 74.82% | 98.5794 | 97.2169 | 57.984 | 102.7341
15 | Calibration Curves |  | 06/16/2023 13:16:57
16 | LBLANK |  | 06/16/2023 13:22:58 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.664 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 60.4509 | 4.5053 | -0.89613 | -0.23606 | 324.0337 | 71.0745 | 42.4779 | 0.0012092 | -0.0036826 | 0.00097579 | -0.59692 | 0.00078078 | -0.28103 | -0.0045444 | 0.034784 | -0.015084 | -0.014802 | 0.45922 | 0.032511 | -0.011484 | -0.0019304 | -0.079577 | 0.018001 | 0.02332 | -0.006465 | 0.0037711 | -0.021474 | 0.20341 | 0.012281 | 98.14% | 97.31% | 96% | 97.6% | 95.78% | 95.62% | 0.0012708 | -0.0008358 | -0.011131 | -0.007241
17 | LBLANK |  | 06/16/2023 13:41:50 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.665 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 56.0325 | 0.92206 | -0.78861 | -0.82037 | -463.6058 | 66.3065 | 39.6685 | 0.0014031 | -0.0072756 | -0.00021147 | -0.57632 | -0.00038413 | -0.59158 | -0.0069095 | 0.006604 | -0.012116 | -0.0099493 | -0.10807 | 0.046967 | -0.01135 | -0.0022639 | -0.16067 | 0.014389 | 0.016408 | -0.0074087 | 0.010201 | -0.021474 | 0.05543 | 0.0072977 | 94.29% | 90.98% | 91.27% | 94.82% | 89.08% | 99.79% | 0.00068284 | -0.0015896 | -0.010754 | -0.0043189
18 | ICV PE |  | 06/16/2023 13:47:50 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\ICV PE.666 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10056.5045 | 10399.9312 | 9.8698 | 18.7224 | 9864.5338 | 9803.376 | 9913.8204 | 10.0101 | 10.5126 | 10.1321 | 10.071 | 10.2031 | 9.9179 | 10.0963 | 9.7904 | 10.0569 | 10.0276 | 9.713 | 9.7284 | 9.0755 | 10.0024 | 6.9848 | 9.9015 | 10.1402 | 10.1402 | 9.8368 | 9.9585 | 9.5255 | 9.5507 | 90.84% | 84.9% | 86.62% | 88.49% | 84.45% | 87.2% | 9.5729 | 9.6206 | 9.2728 | 9.5202
19 | CCV 1ppb TRACE |  | 06/16/2023 13:53:50 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.667 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 55.6556 | 0.91518 | -1.0076 | -0.99296 | -2402.6525 | 50.0019 | 40.5457 | 7.111e-06 | -0.013837 | -0.0030719 | -0.6073 | -0.0012092 | -2.2335 | -0.00661 | -0.0051573 | -0.01323 | -0.01183 | -0.21218 | 0.71528 | -0.0089217 | -0.004222 | -1.9308 | 0.0055366 | 0.0096123 | -0.0090549 | -0.090629 | -0.021474 | 0.035588 | 0.0014145 | 102.1% | 101% | 101.5% | 103.3% | 100.1% | 111.2% | 0.00074945 | -0.0022094 | -0.010929 | -0.0047576
20 | CCV 1ppb TRACE |  | 06/16/2023 14:00:07 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.668 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1051.1896 | 1076.671 | 1.2775 | 0.14651 | 1270.6449 | 1104.0053 | 989.4019 | 1.0496 | 1.09 | 1.0467 | 0.8539 | 1.0223 | 1.167 | 1.0494 | 1.0433 | 1.0774 | 1.0596 | 0.82662 | 0.96458 | 0.93718 | 0.9793 | 1.1859 | 0.9275 | 0.99008 | 1.0313 | 0.96014 | 0.9736 | 1.0233 | 0.98635 | 92.39% | 90.3% | 91.8% | 93.04% | 90.48% | 91.64% | 0.97624 | 0.97117 | 1.0298 | 0.98783
21 | LFB |  | 06/16/2023 14:06:07 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.669 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10631.0355 | 10937.2198 | 10.7055 | 11.5311 | 10903.3394 | 10615.1967 | 10398.4198 | 10.7744 | 11.3367 | 9.9786 | 9.5522 | 9.9937 | 10.6854 | 10.9986 | 10.5865 | 10.8235 | 9.914 | 10.0895 | 9.9048 | 11.207 | 9.9616 | 10.3388 | 10.1419 | 10.1303 | 9.9546 | 10.4514 | 10.5175 | 9.7079 | 9.7943 | 89.22% | 83.36% | 85.25% | 86.67% | 82.9% | 80.88% | 9.669 | 9.5549 | 9.1443 | 9.426
22 | LLOQ 0.01 |  | 06/16/2023 14:12:05 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LLOQ 0.01.670 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 68.7627 | 11.998 | 0.31844 | 1.2213 | 404.9005 | 74.1205 | 49.5998 | 0.011602 | 0.0069612 | 0.013291 | -0.064712 | 0.029726 | 0.26583 | 0.0096876 | 0.02016 | 0.01962 | 0.02022 | -0.0037104 | 0.05627 | 0.0034775 | 0.0077831 | 0.028379 | 0.019141 | 0.023973 | 9.0899e-05 | 0.031049 | 0.001602 | 0.050454 | 0.021506 | 93.13% | 89.42% | 92.11% | 92.22% | 87.86% | 91.82% | 0.010088 | 0.0088365 | 0.0077222 | 0.013587
23 | CCV 10 ppb MAJORS |  | 06/16/2023 14:18:05 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.671 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10122.8889 | 10275.35 | 9.7477 | 9.1705 | 10255.1656 | 9972.5124 | 9698.068 | 10.1342 | 10.5512 | 10.4501 | 10.0215 | 10.4576 | 9.83 | 10.4451 | 10.2808 | 10.3056 | 10.327 | 10.4275 | 9.9744 | 10.2396 | 10.2403 | 10.5057 | 9.9803 | 9.9928 | 10.1894 | 10.2047 | 10.2849 | 10.1979 | 10.2561 | 88.31% | 81.2% | 83.3% | 85.97% | 82.26% | 85.11% | 9.8422 | 9.8537 | 9.3505 | 9.6711
24 | LLOQ 0.1 |  | 06/16/2023 14:24:05 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LLOQ 0.1.672 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 157.5781 | 102.3591 | -0.73809 | -0.51602 | 462.677 | 164.6774 | 129.5481 | 0.10451 | 0.088081 | 0.10195 | -0.2596 | 0.10732 | 0.13125 | 0.09552 | 0.10538 | 0.10931 | 0.10692 | 0.014436 | 0.13764 | 0.094082 | 0.094888 | 0.13829 | 0.096396 | 0.10776 | 0.091667 | 0.056754 | 0.13188 | 0.13578 | 0.10093 | 93.22% | 89.09% | 90.83% | 92.57% | 86.35% | 90.45% | 0.097918 | 0.09334 | 0.089389 | 0.093548
25 | 9_13_21_F_11_DUP |  | 06/16/2023 14:30:04 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_F_11_DUP.673 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 6749.2116 | 8683.4952 | 8.7958 | 2.7535 | 14706.9096 | 1924.6586 | 33015.485 | 0.76408 | 0.11728 | 0.11583 | 4.4686 | 5.0477 | 5.0885 | 0.016785 | 0.24271 | 2.2292 | 2.293 | 3.555 | 4.1806 | 4.0028 | 4.1759 | 0.29654 | 0.10333 | 0.11553 | 0.093087 | 0.43196 | 0.065493 | 2.3338 | 2.2977 | 82.33% | 78.59% | 79.73% | 82.14% | 79.72% | 82.75% | 0.012929 | 0.011039 | 0.11831 | 0.11785
26 | LBLANK |  | 06/16/2023 14:47:08 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.675 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 55.7227 | 1.5596 | -0.86348 | -0.86945 | 243.5678 | 64.2909 | 40.9968 | 0.0012129 | -0.0026043 | -0.00061988 | -0.57775 | -0.00058098 | -0.27505 | -0.0062026 | 0.00013072 | -0.011968 | -0.010147 | -0.1891 | 0.029022 | -0.001768 | -0.0025415 | -0.10492 | 0.010392 | 0.0056793 | -0.0070266 | -0.017754 | -0.017342 | 0.020267 | 0.0038164 | 86.29% | 83.37% | 84.58% | 86.4% | 80.13% | 84.33% | 0.0005995 | -0.0016388 | -0.01057 | -0.0043334
27 | 0.5 STD |  | 06/16/2023 14:53:09 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\0.5 STD.676 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 567.1011 | 543.0869 | 0.48963 | 1.4303 | 822.519 | 587.9992 | 554.3056 | 0.5089 | 0.53802 | 0.52638 | 0.57417 | 0.51448 | 0.74484 | 0.51448 | 0.51989 | 0.56375 | 0.5828 | 0.55951 | 0.5425 | 0.47652 | 0.50708 | 0.42628 | 0.51372 | 0.53534 | 0.50474 | 0.45577 | 0.56183 | 0.51778 | 0.50407 | 85.29% | 82.06% | 83.44% | 85.04% | 80.46% | 83.16% | 0.4982 | 0.50077 | 0.5187 | 0.51064
28 | 0.1 STD |  | 06/16/2023 14:59:12 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\0.1 STD.677 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 155.8817 | 102.4158 | -0.81077 | -0.012616 | 408.199 | 159.4332 | 148.202 | 0.088929 | 0.097049 | 0.099312 | -0.25647 | 0.10293 | -0.028854 | 0.09906 | 0.10391 | 0.097358 | 0.10278 | -0.032258 | 0.17035 | 0.12607 | 0.099857 | -0.085497 | 0.13845 | 0.10755 | 0.096673 | 0.068767 | 0.045051 | 0.11071 | 0.10245 | 86% | 82.14% | 84.5% | 85.76% | 80.04% | 85.3% | 0.10497 | 0.10023 | 0.096042 | 0.094023
29 | CCV 10 ppb MAJORS |  | 06/16/2023 15:05:12 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.678 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9998.191 | 10263.8003 | 8.857 | 9.0615 | 10259.6438 | 10055.9356 | 9931.5792 | 10.1704 | 10.489 | 10.3977 | 10.0399 | 10.3779 | 9.8225 | 10.2719 | 10.0946 | 10.1712 | 10.319 | 10.349 | 10.0115 | 10.2913 | 10.2486 | 10.1398 | 10.4099 | 10.2028 | 10.292 | 10.3317 | 10.3477 | 10.1689 | 10.3237 | 84.03% | 74.88% | 78.34% | 80.25% | 76.75% | 80.04% | 9.8926 | 9.9207 | 9.3714 | 9.774
30 | 9_13_21_F_11_DUP |  | 06/16/2023 15:11:13 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_F_11_DUP.679 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 6661.3322 | 8884.6197 | 9.2147 | 2.6296 | 14635.657 | 1901.3119 | 32311.5268 | 0.7198 | 0.11733 | 0.11484 | 4.4854 | 5.0046 | 5.0335 | 0.016368 | 0.24765 | 2.1914 | 2.2814 | 3.5103 | 4.0506 | 3.7434 | 4.1249 | 0.20292 | 0.1016 | 0.094639 | 0.076293 | 0.37117 | 0.036995 | 2.316 | 2.3 | 80.86% | 75.84% | 76.82% | 78.73% | 76.93% | 81% | 0.010934 | 0.0095419 | 0.11362 | 0.11771
31 | 9_13_21_F_11_DUP_LSP |  | 06/16/2023 15:17:13 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_F_11_DUP_LSP.680 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 16372.2944 | 18220.7526 | 18.303 | 13.5298 | 23548.6851 | 11957.1968 | 40450.9585 | 10.8758 | 10.5394 | 10.4667 | 14.8681 | 14.7905 | 14.809 | 10.2056 | 9.9521 | 11.8482 | 11.8968 | 13.1843 | 13.7685 | 13.8019 | 13.7087 | 10.1245 | 9.7167 | 9.8906 | 9.5655 | 10.7426 | 9.7061 | 12.1925 | 12.2044 | 79.68% | 73.4% | 73.97% | 75.96% | 73.32% | 76.61% | 9.76 | 9.7339 | 9.3837 | 9.6345
32 | 9_13_21_UF_1 |  | 06/16/2023 15:23:13 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_UF_1.681 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 8600.8984 | 6172.4502 | 1.1442 | 7.7774 | 21031.1995 | 2115.0093 | 33630.2459 | 0.31841 | 0.070932 | 0.056883 | 0.39907 | 13.1481 | 1.3661 | 0.026365 | 0.49956 | 0.59207 | 0.6441 | 1.1866 | 3.9346 | 3.6791 | 3.9304 | 0.32996 | 0.10727 | 0.036967 | 0.039942 | 0.3644 | 0.055555 | 15.4627 | 15.6921 | 80.7% | 75.48% | 76.67% | 79.63% | 76.7% | 78.5% | 0.0076473 | 0.005176 | 0.041612 | 0.04545
33 | 9_13_21_UF_6 |  | 06/16/2023 15:29:13 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_UF_6.682 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9119.8134 | 7168.0192 | 0.040541 | 2.4342 | 15318.6539 | 1909.3491 | 33678.1192 | 0.64699 | 0.083507 | 0.061443 | 0.12175 | 12.0435 | 1.1733 | 0.016883 | 0.25807 | 1.8047 | 1.868 | 2.2027 | 4.0122 | 3.8794 | 4.1069 | 0.39017 | 0.11304 | 0.095389 | 0.07704 | 0.47657 | 0.074028 | 5.9911 | 6.0122 | 80.64% | 75.14% | 74.99% | 79.33% | 76.74% | 78.88% | 0.015542 | 0.013462 | 0.058878 | 0.062798
34 | 9_13_21_UF_6_DUP |  | 06/16/2023 15:35:13 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_UF_6_DUP.683 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9054.7298 | 7120.8688 | 0.026518 | 2.4373 | 15410.9294 | 1870.7874 | 33131.715 | 0.65816 | 0.10579 | 0.091806 | 0.2807 | 11.9175 | 1.1758 | 0.01591 | 0.26493 | 1.8118 | 1.8602 | 1.7857 | 4.1471 | 3.9179 | 4.1674 | 0.3714 | 0.077033 | 0.065904 | 0.074097 | 0.462 | 0.062133 | 5.9135 | 5.9981 | 78.33% | 73.56% | 75.16% | 76.43% | 74.06% | 77.46% | 0.016588 | 0.011743 | 0.062784 | 0.063414
35 | 9_13_21_UF_8 |  | 06/16/2023 15:41:13 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_UF_8.684 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9413.63 | 6987.0859 | -0.0090554 | 2.914 | 13899.2824 | 1983.0529 | 32356.6161 | 0.84496 | 0.074651 | 0.057237 | 0.028408 | 10.6395 | 0.88943 | 0.011977 | 0.25932 | 1.6028 | 1.6104 | 0.98706 | 4.5497 | 3.9851 | 4.5172 | 0.20632 | 0.077473 | 0.080379 | 0.072963 | 0.35976 | 0.058022 | 4.9521 | 5.0425 | 77.83% | 73.13% | 74.93% | 76.85% | 74.09% | 76.26% | 0.011199 | 0.0070041 | 0.049838 | 0.051821
36 | 9_13_21_UF_9 |  | 06/16/2023 15:47:13 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_UF_9.685 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 8672.9482 | 7152.9978 | 0.2647 | 2.6899 | 12879.5084 | 1879.952 | 29543.5492 | 0.73641 | 0.10046 | 0.077492 | 0.19619 | 3.8521 | 1.1171 | 0.010035 | 0.20771 | 1.4859 | 1.479 | 0.73751 | 3.7713 | 3.6218 | 3.875 | 0.41215 | 0.088717 | 0.071676 | 0.063054 | 0.38646 | 0.08164 | 3.7576 | 3.8449 | 78.72% | 73.1% | 74.63% | 76.01% | 73.74% | 76.56% | 0.006981 | 0.0034572 | 0.050297 | 0.051291
37 | LBLANK |  | 06/16/2023 15:53:13 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.686 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 54.4446 | 0.23633 | -0.96263 | -0.87372 | 236.4523 | 59.288 | 40.4154 | 2.2133e-05 | -0.011409 | -0.0013358 | -0.57434 | 0.00021345 | -0.40485 | -0.0076191 | 0.00016619 | -0.010467 | -0.011086 | -0.22461 | 0.09778 | 0.035163 | 0.038651 | -0.021279 | 0.0055366 | 0.0076658 | -0.0066793 | -0.0046485 | -0.021474 | -0.0033911 | 0.0030236 | 84% | 78.59% | 81.52% | 83.52% | 77.53% | 81.84% | 0.00060382 | -0.0025156 | -0.010272 | -0.0044217
38 | CCV 1ppb TRACE |  | 06/16/2023 15:59:13 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.687 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1190.5987 | 1169.6923 | 1.2344 | 0.06955 | 1723.162 | 1147.5669 | 1120.4383 | 1.1169 | 1.1415 | 1.0426 | 0.85164 | 1.0106 | 1.3864 | 1.1093 | 1.1316 | 1.1304 | 1.0666 | 0.84399 | 0.96898 | 1.1141 | 1.0042 | 1.0799 | 0.98937 | 1.0329 | 0.95564 | 1.0227 | 0.99694 | 0.97943 | 1.0019 | 83.14% | 78.55% | 79.84% | 81.62% | 78% | 77.21% | 1.0047 | 0.9778 | 1.0507 | 1.0158
39 | CCV 1ppb TRACE |  | 06/16/2023 16:05:32 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.688 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1057.7035 | 1073.3082 | 1.2211 | 0.097238 | 1427.3953 | 1096.0919 | 1040.0608 | 1.0547 | 1.0688 | 1.0731 | 0.92835 | 1.0313 | 1.2401 | 1.0406 | 1.1001 | 1.044 | 1.068 | 0.83757 | 0.92826 | 1.0796 | 1.0157 | 1.1541 | 0.94276 | 1.0118 | 0.97498 | 0.95912 | 1.1056 | 0.99301 | 1.0103 | 83.25% | 78.52% | 80.81% | 82.55% | 78.15% | 82.08% | 0.99726 | 0.98164 | 1.0647 | 1.0064
40 | CCV 10 ppb MAJORS |  | 06/16/2023 16:11:33 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.689 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9961.7153 | 10510.6335 | 9.7507 | 8.8572 | 10582.1499 | 10164.7328 | 9939.21 | 10.1587 | 10.4579 | 10.3163 | 10.2467 | 10.2249 | 9.8969 | 10.2916 | 9.9266 | 10.1153 | 10.131 | 10.3464 | 9.7164 | 10.2678 | 9.9965 | 10.2444 | 9.977 | 10.1643 | 10.0885 | 10.0647 | 10.1743 | 9.9111 | 10.0907 | 82.64% | 73.78% | 76.88% | 78.33% | 74.42% | 78.4% | 9.8342 | 9.8947 | 9.6594 | 9.8891
41 | CCV 10 ppb MAJORS |  | 06/16/2023 16:17:52 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.690 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10132.4338 | 10727.9424 | 9.4693 | 8.9241 | 10627.4135 | 10161.8632 | 9907.9138 | 10.2231 | 10.6301 | 10.4494 | 10.0884 | 10.3631 | 10.0249 | 10.3471 | 9.9207 | 10.238 | 10.3375 | 10.5247 | 9.926 | 9.9607 | 10.1212 | 10.1338 | 9.9371 | 10.052 | 9.8966 | 10.2714 | 10.1326 | 10.0575 | 10.1465 | 83.34% | 74.42% | 77.21% | 79.09% | 75.88% | 78.81% | 9.7552 | 9.8221 | 9.5135 | 9.8035
42 | 9_13_21_UF_10 |  | 06/16/2023 16:23:53 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_UF_10.691 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 7990.4275 | 8816.9371 | 0.52768 | 3.6403 | 14848.7104 | 2008.2639 | 33636.0982 | 0.77593 | 0.075592 | 0.056051 | -0.10492 | 4.3968 | 1.0208 | 0.0087647 | 0.19679 | 1.18 | 1.2209 | 0.85804 | 4.2053 | 3.9632 | 4.2088 | 0.14297 | 0.071487 | 0.07471 | 0.072311 | 0.4533 | 0.057613 | 2.4596 | 2.4743 | 78.82% | 73.64% | 74.04% | 77.44% | 74.68% | 77.82% | 0.0097692 | 0.0042645 | 0.05325 | 0.057264
43 | 9_13_21_UF_11 |  | 06/16/2023 16:29:52 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_UF_11.692 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 7994.5118 | 9055.3802 | 0.10884 | 3.5752 | 14604.9877 | 2018.1785 | 33979.6864 | 0.76714 | 0.16724 | 0.14546 | 0.82035 | 4.8666 | 1.7196 | 0.0074467 | 0.19257 | 1.4241 | 1.4436 | 1.1555 | 4.1652 | 3.7122 | 4.1654 | 0.2023 | 0.073047 | 0.094714 | 0.073834 | 0.40225 | 0.078638 | 2.2978 | 2.3288 | 77.24% | 71.76% | 73.9% | 74.61% | 72.72% | 76.55% | 0.0080234 | 0.0069188 | 0.064531 | 0.066184
44 | 9_13_21_LAB_A |  | 06/16/2023 16:35:53 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_LAB_A.693 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 63.8156 | 0.23633 | -0.12208 | -0.62349 | 285.1991 | 59.7555 | 42.3034 | -2.0077e-06 | 0.010343 | 0.014273 | -0.34594 | 0.038506 | -0.25411 | -0.0065759 | 0.027704 | 0.58511 | 0.59532 | 0.12296 | 0.093665 | 0.011442 | 0.023397 | -0.061019 | 0.021136 | 0.0044888 | -0.0058471 | 0.032777 | -0.021474 | -0.0044156 | 0.0019025 | 83.93% | 77.69% | 80.4% | 82.41% | 76.32% | 83.78% | 0.00060519 | -0.0019527 | 0.018159 | 0.02355
45 | 9_13_21_BB |  | 06/16/2023 16:41:52 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_BB.694 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 56.6196 | 0.23633 | 0.09474 | -0.63072 | 229.6033 | 62.8683 | 45.752 | 0.0012963 | 0.0071925 | 0.016015 | -0.36435 | 0.0026256 | -0.21291 | -0.0073475 | 0.0066767 | 0.12189 | 0.12284 | -0.04338 | 0.056527 | 0.0054457 | 0.013177 | -0.1341 | 0.0055366 | 0.0028146 | -0.0058124 | -0.03456 | -0.017025 | -0.0055398 | 0.0031128 | 84.28% | 77.19% | 80.23% | 82% | 75.47% | 80.58% | 0.00081119 | -0.001438 | 0.019392 | 0.024504
46 | 9_13_21_FB |  | 06/16/2023 16:47:52 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_FB.695 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 236.6815 | 43.7084 | 14.4759 | 0.6 | 637.7343 | 126.7295 | 270.1861 | 0.015376 | 0.075294 | 0.080747 | 0.76598 | 0.30634 | 1.1792 | 0.0046164 | 0.048366 | 0.2146 | 0.25594 | 18.6562 | 0.017427 | 0.05652 | 0.04947 | 0.074959 | 0.016188 | 0.0079667 | 0.0070413 | 0.0023331 | -0.012523 | 0.096043 | 0.10539 | 81.45% | 75.99% | 78.95% | 80.53% | 75.59% | 75.89% | 0.0097213 | 0.0067231 | 0.089345 | 0.092484
47 | 5_18_23_MFB |  | 06/16/2023 16:53:52 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_MFB.696 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4836.2346 | 5082.0378 | 4.6388 | 8.5802 | 5184.0438 | 4837.1276 | 4882.1192 | 4.9302 | 5.2641 | 5.2 | 4.6128 | 5.0652 | 4.6849 | 5.037 | 5.0001 | 5.1198 | 5.1607 | 4.8593 | 4.3888 | 5.0084 | 4.8807 | 4.4037 | 5.1275 | 4.8968 | 4.9452 | 4.5838 | 5.1712 | 4.8166 | 4.8727 | 88.19% | 74.05% | 77.27% | 79.41% | 73.98% | 78.45% | 4.8335 | 4.8479 | 4.658 | 5.0361
48 | 5_18_23_DB |  | 06/16/2023 16:59:52 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_DB.697 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 69.2273 | 12.5222 | -0.27238 | 4.1911 | 484.1611 | 72.9149 | 59.3009 | 0.013734 | 0.010734 | 0.018163 | -0.31381 | 0.01657 | -0.025073 | 0.0073802 | 0.010834 | 0.073513 | 0.074039 | -0.11093 | 0.014335 | 0.012004 | 0.015188 | 0.069414 | 0.052083 | 0.016129 | 0.012074 | 0.0042037 | 0.00037767 | 0.0061381 | 0.014442 | 85.33% | 77.5% | 81.83% | 82.51% | 76.68% | 81.98% | 0.014689 | 0.011044 | 0.024244 | 0.030415
49 | 5_18_23_T_239 |  | 06/16/2023 17:05:52 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_T_239.698 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 46409.2332 | 13062.6293 | 115.9872 | 12.2818 | 1173.2444 | 3943.0834 | 49867.5428 | 1.9186 | 1.4536 | 1.4751 | 43.2952 | 3.69 | 41.8006 | 1.4258 | 1.7129 | 0.77367 | 0.85401 | 6.0938 | 2.5039 | 2.2333 | 2.4468 | 5.416 | 2.0766 | 1.97 | 1.9029 | 2.3068 | 1.9684 | 1.0927 | 1.1055 | 72.24% | 65.79% | 69.66% | 72.21% | 65.36% | 71.35% | 3.3754 | 3.3183 | 1.4041 | 1.2647
50 | 5_18_23_T_239 |  | 06/16/2023 17:11:52 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_T_239.699 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 48071.2603 | 13505.3814 | 121.1505 | 12.8964 | 1362.1759 | 4062.0503 | 53018.5233 | 2.0132 | 1.602 | 1.4479 | 42.6029 | 3.6219 | 44.1082 | 1.5019 | 1.8124 | 0.85556 | 0.87471 | 6.0507 | 2.434 | 2.4273 | 2.3709 | 4.9176 | 1.9015 | 1.7989 | 1.8994 | 2.1095 | 1.938 | 1.1136 | 1.0972 | 75.63% | 65.42% | 69.33% | 71.95% | 65.46% | 68.77% | 3.3567 | 3.2663 | 1.2908 | 1.2357
51 | 5_18_23_M_224 |  | 06/16/2023 17:17:52 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_M_224.700 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4052.0246 | 642.9194 | 7.3955 | 17.0722 | 916.4976 | 519.1015 | 2963.3336 | 1.0225 | 0.031492 | 0.031926 | -0.31224 | 5.5129 | -0.091135 | 0.0065512 | 0.051456 | 0.064557 | 0.062569 | 0.69528 | 0.23479 | 0.040533 | 0.059438 | 0.23278 | 0.0679 | 0.019298 | 0.016721 | 0.76179 | 0.036066 | 0.2415 | 0.20725 | 81.37% | 71.3% | 76.78% | 79.24% | 72.12% | 79.09% | 0.0031649 | 0.00078166 | 0.0089421 | 0.012469
52 | LBLANK |  | 06/16/2023 17:23:52 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.701 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 55.8376 | 0.70899 | -1.0141 | -0.18597 | 157.8543 | 64.6363 | 41.096 | 0.0010043 | -0.0086365 | -0.0021681 | -0.58445 | -0.00097667 | -0.38029 | -0.007202 | 0.0047496 | -0.0094054 | -0.010432 | -0.22261 | -0.020689 | -0.0078919 | -0.00011531 | 0.21802 | 0.0055366 | 0.0028999 | -0.0080639 | -0.014407 | -0.021474 | 0.017648 | 0.0020502 | 74.67% | 75.09% | 82.71% | 82.32% | 74.77% | 79.98% | 0.00050427 | -0.0021097 | -0.010416 | -0.004357
53 | CCV 1ppb TRACE |  | 06/16/2023 17:29:53 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.702 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 987.8331 | 1010.777 | 1.2439 | 0.23021 | 1094.2251 | 1046.5265 | 971.0796 | 1.0204 | 1.0526 | 1.0616 | 0.90075 | 1.0129 | 1.035 | 1.0071 | 0.96643 | 1.0699 | 1.0575 | 0.8381 | 0.96922 | 0.9777 | 0.97344 | 1.046 | 1.1939 | 1.0451 | 0.97312 | 0.98967 | 0.95556 | 1.0014 | 1.0082 | 82.72% | 75.25% | 78.92% | 79.98% | 73.49% | 79.36% | 0.9601 | 0.97936 | 1.0383 | 1.0069
54 | CCV 1ppb TRACE |  | 06/16/2023 17:36:12 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.703 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1009.6197 | 1032.865 | 1.0845 | 0.13398 | 1218.9479 | 1035.6432 | 996.5243 | 0.99993 | 1.0281 | 1.0567 | 0.85323 | 1.0119 | 1.2074 | 1.005 | 1.0396 | 1.0711 | 1.0542 | 0.83458 | 0.90204 | 1.0035 | 0.99199 | 1.1499 | 0.89965 | 0.94656 | 0.99771 | 1.0148 | 0.96448 | 0.99645 | 1.0146 | 83.29% | 75.52% | 79.03% | 81.67% | 74.72% | 81.02% | 0.97104 | 0.98178 | 1.0613 | 1.0075
55 | LFB |  | 06/16/2023 17:42:12 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.704 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9318.2332 | 9688.4626 | 9.7361 | 11.4387 | 10212.8116 | 9450.1073 | 9627.478 | 9.6244 | 10.0298 | 9.8195 | 9.6886 | 9.8477 | 9.3763 | 9.8275 | 9.545 | 9.8513 | 9.8229 | 9.7185 | 9.5215 | 10.0351 | 9.8384 | 10.3697 | 10.479 | 9.8586 | 9.9861 | 10.1133 | 9.8555 | 9.5679 | 9.8013 | 81.51% | 70.96% | 75.83% | 77.13% | 71.52% | 76.66% | 9.4964 | 9.4689 | 9.4318 | 9.6216
56 | CCV 10 ppb MAJORS |  | 06/16/2023 17:48:11 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.705 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9642.9197 | 10168.3327 | 9.0177 | 8.9859 | 10240.0218 | 9932.3643 | 9801.3678 | 9.9058 | 10.2795 | 10.4564 | 10.2546 | 10.287 | 9.766 | 10.216 | 9.7392 | 10.0267 | 10.2615 | 10.4433 | 9.5325 | 9.8829 | 9.9566 | 10.3203 | 9.9852 | 9.9118 | 10.1161 | 9.7431 | 9.7936 | 10.0474 | 10.2082 | 82.99% | 71.63% | 76.33% | 77.5% | 72.42% | 78.59% | 9.7785 | 9.8097 | 9.5589 | 9.9094
57 | CCV 10 ppb MAJORS |  | 06/16/2023 17:54:30 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.706 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9842.1164 | 10306.0446 | 9.6058 | 9.0394 | 10768.3658 | 9919.952 | 10003.6006 | 10.227 | 10.4489 | 10.5586 | 10.1865 | 10.2453 | 9.9996 | 10.4423 | 10.1498 | 10.3751 | 10.3184 | 10.385 | 9.8983 | 9.9591 | 9.8456 | 10.486 | 9.6585 | 10.0605 | 9.9058 | 10.3007 | 10.396 | 10.0906 | 10.1703 | 83.53% | 73.2% | 76.77% | 77.85% | 73.62% | 78.01% | 9.6967 | 9.7855 | 9.6044 | 9.8314
58 | 5_18_23_MFB |  | 06/16/2023 18:00:31 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_MFB.707 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4976.5158 | 5147.4868 | 4.9909 | 8.543 | 5454.665 | 4914.5063 | 5048.8721 | 5.0602 | 5.2755 | 5.6289 | 5.309 | 5.1463 | 4.9703 | 5.2693 | 5.1257 | 5.25 | 5.5752 | 5.242 | 4.7392 | 5.1705 | 4.9592 | 4.9333 | 4.7759 | 4.9131 | 4.8191 | 5.0688 | 5.0131 | 5.2372 | 4.9371 | 84.17% | 74.32% | 75.02% | 80.68% | 74.65% | 80.22% | 5.1679 | 5.1922 | 4.9494 | 4.8554
59 | 10_1_21_W_1 |  | 06/16/2023 18:06:31 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_W_1.708 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1809.7496 | 1366.4408 | 5.0964 | 3.1972 | 5506.4698 | 502.6889 | 7806.1746 | 0.053229 | 0.0098102 | 0.016988 | 7.5556 | 5.789 | 7.7203 | 0.0014734 | 0.079968 | 0.46544 | 0.46547 | 0.16074 | 0.69481 | 0.66979 | 0.67812 | 0.12079 | 0.032968 | 0.022083 | -0.0022877 | 0.077872 | 0.0015327 | 3.5298 | 3.5381 | 82.38% | 74.06% | 76.88% | 79.67% | 73.56% | 79.66% | 0.0026279 | 6.599e-05 | 0.022528 | 0.026239
60 | 10_1_21_W_6 |  | 06/16/2023 18:12:31 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_W_6.709 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1829.2879 | 1606.9129 | 8.9897 | 1.7914 | 4261.9224 | 455.8822 | 7633.578 | 0.14421 | 0.013414 | 0.017521 | 9.3953 | 3.2925 | 9.5866 | 0.00098277 | 0.075184 | 0.84881 | 0.83956 | 0.64956 | 0.7535 | 0.65615 | 0.78522 | 0.040281 | 0.044644 | 0.01374 | 0.013881 | 0.056784 | -0.0026399 | 1.5877 | 1.6668 | 81.59% | 72.24% | 77.83% | 71.77% | 72.5% | 79.23% | 0.0057574 | 0.0015786 | 0.067478 | 0.07909
61 | 10_1_21_W_8 |  | 06/16/2023 18:18:32 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_W_8.710 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1861.2963 | 1519.0654 | 3.4101 | 1.9807 | 3849.6145 | 458.502 | 7373.0043 | 0.16039 | 0.019854 | 0.020336 | 5.1233 | 2.4486 | 5.2766 | -0.0034375 | 0.045822 | 0.82341 | 0.82485 | 0.51531 | 0.80668 | 0.69827 | 0.85654 | 0.10225 | 0.033808 | 0.026485 | 0.006546 | 0.061177 | -0.0070859 | 1.4505 | 1.4482 | 80.76% | 71.38% | 74.63% | 78.22% | 71.42% | 76.42% | 0.0030228 | 6.2098e-05 | 0.026713 | 0.0328
62 | 10_1_21_W_9 |  | 06/16/2023 18:24:32 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_W_9.711 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1568.3974 | 1380.7452 | 5.1934 | 1.7039 | 3104.5776 | 418.9389 | 6409.642 | 0.13384 | 0.043758 | 0.045602 | 5.3406 | 1.8014 | 5.4711 | -0.0028335 | 0.033099 | 0.56246 | 0.56041 | 0.1864 | 0.64656 | 0.67178 | 0.66101 | 0.10495 | 0.016749 | 0.013764 | 0.0055303 | 0.052236 | 0.0022145 | 0.99206 | 1.0037 | 81.66% | 72.06% | 76.22% | 78.25% | 72.58% | 78.2% | 0.0043542 | 0.00058309 | 0.025359 | 0.031942
63 | 10_1_21_W_10 |  | 06/16/2023 18:30:32 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_W_10.712 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1389.6206 | 1711.7712 | 30.1756 | 3.435 | 3384.5393 | 420.0594 | 6934.8936 | 0.14515 | 0.027355 | 0.03211 | 14.9574 | 3.7925 | 14.4594 | 0.0030418 | 0.051527 | 0.61913 | 0.63275 | 0.67065 | 0.73793 | 0.62318 | 0.75786 | 0.18985 | 0.0169 | 0.030373 | 0.0032508 | 0.085612 | -0.00689 | 0.58556 | 0.61175 | 81.98% | 70.86% | 75.66% | 78.61% | 71.22% | 77.31% | 0.0040781 | 0.002379 | 0.10042 | 0.10142
64 | 10_1_21_W_11 |  | 06/16/2023 18:36:31 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_W_11.713 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1334.4543 | 1653.3531 | 21.1115 | 3.12 | 3350.5535 | 419.7848 | 6718.1624 | 0.16741 | 0.047703 | 0.040418 | 17.5419 | 3.8911 | 16.7815 | 0.0035638 | 0.055717 | 0.78855 | 0.81081 | 0.84412 | 0.77576 | 0.71771 | 0.7898 | 0.17462 | 0.027946 | 0.024552 | 0.0030542 | -0.00010942 | 0.011738 | 0.55225 | 0.56414 | 81.93% | 71.81% | 75.79% | 78.32% | 72.08% | 76.75% | 0.0047143 | 0.002778 | 0.21187 | 0.20855
65 | 10_1_21_W_11_LD |  | 06/16/2023 18:42:31 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_W_11_LD.714 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 54.4392 | 0.23633 | -1.0383 | -0.92404 | -2381.0441 | 46.3568 | 39.6077 | -0.00066774 | -0.016067 | -0.0033762 | -0.61044 | -0.0009563 | -2.3984 | -0.0076572 | -0.0085046 | -0.013579 | -0.01327 | -0.23067 | 0.73465 | -0.0056089 | -0.0041382 | -1.8771 | 0.0055366 | -0.00050913 | -0.010059 | -0.062944 | -0.021474 | -0.0012166 | 0.0013969 | 96.01% | 90.46% | 91.58% | 96.28% | 87.67% | 101.6% | 0.00059194 | -0.0027606 | -0.011182 | -0.0047675
66 | 10_1_21_W_11_LD |  | 06/16/2023 18:48:48 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_W_11_LD.715 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1339.1526 | 1688.0961 | 22.0867 | 2.483 | 3244.2172 | 424.2943 | 6864.056 | 0.15839 | 0.047825 | 0.052847 | 17.8769 | 4.057 | 16.8064 | 0.0058586 | 0.050277 | 0.76652 | 0.81322 | 0.8674 | 0.80021 | 0.87521 | 0.84892 | 0.16215 | 0.034303 | 0.021451 | -2.8715e-05 | 0.031092 | -0.011745 | 0.57455 | 0.58916 | 79.18% | 70.23% | 73.83% | 75.82% | 69.5% | 77.46% | 0.0059181 | 0.0033105 | 0.21835 | 0.21327
67 | 10_1_21_W_11_DUP |  | 06/16/2023 18:54:48 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_W_11_DUP.716 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1391.7828 | 1697.8068 | 23.0783 | 2.4225 | 3279.9744 | 401.5499 | 6786.8219 | 0.15423 | 0.026191 | 0.027036 | 18.466 | 4.2314 | 17.8155 | 0.0023811 | 0.045374 | 0.80496 | 0.81933 | 0.87544 | 0.78012 | 0.74155 | 0.81046 | 0.0074674 | 0.034973 | 0.031283 | 0.014287 | 0.11711 | -0.0015615 | 0.55047 | 0.57562 | 79.96% | 68.7% | 73.72% | 76.3% | 69.15% | 74.9% | 0.0049392 | 0.0041534 | 0.13131 | 0.12775
68 | LBLANK |  | 06/16/2023 19:00:48 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.717 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 54.1008 | 0.74518 | -0.99077 | -0.58503 | 373.8525 | 57.4128 | 39.1036 | 0.0011378 | -0.0072258 | -0.0012028 | -0.58948 | -0.00056814 | -0.25972 | -0.0077432 | 0.0035129 | -0.00485 | -0.010419 | -0.22553 | 0.077648 | -0.01135 | -0.0034741 | -0.026405 | 0.0055366 | 0.0013136 | -0.011205 | 0.0074769 | -0.0022024 | -0.0050221 | 0.0024088 | 81.81% | 71.35% | 76.1% | 77.27% | 70.63% | 74.68% | 0.00061183 | -0.0020735 | -0.010642 | -0.0044551
69 | CCV 1ppb TRACE |  | 06/16/2023 19:06:48 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.718 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1002.6883 | 1004.4991 | 1.3914 | 0.16075 | 1337.7963 | 1006.692 | 987.1159 | 1.0297 | 1.0642 | 1.0889 | 0.92501 | 0.99538 | 1.0695 | 1.0281 | 1.0015 | 1.0313 | 1.0631 | 0.85139 | 0.91214 | 1.0631 | 0.9722 | 1.0056 | 1.0267 | 0.98058 | 0.97723 | 0.95232 | 1.0954 | 1.0007 | 1.0144 | 79.09% | 70.91% | 74.53% | 76.32% | 69.14% | 75.39% | 0.983 | 0.98388 | 1.087 | 1.0288
70 | CCV 1ppb TRACE |  | 06/16/2023 19:13:07 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.719 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1013.9286 | 1074.6779 | 1.2329 | 0.11892 | 1416.6104 | 1024.4187 | 1009.3335 | 1.0007 | 1.0725 | 1.0779 | 0.94758 | 1.0146 | 1.0703 | 1.009 | 0.98826 | 1.0691 | 1.0609 | 0.84908 | 0.96223 | 0.93815 | 0.98826 | 1.0093 | 1.0598 | 0.96533 | 0.97458 | 1.0173 | 1.0478 | 1.0036 | 1.0065 | 80.14% | 70.96% | 75% | 76.9% | 70.67% | 76.66% | 0.97373 | 1.0071 | 1.0945 | 1.0464
71 | CCV 10 ppb MAJORS |  | 06/16/2023 19:19:08 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.720 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9635.2676 | 10151.3372 | 9.4319 | 8.8801 | 10594.1633 | 9744.4391 | 9907.7371 | 9.8795 | 10.1958 | 10.4225 | 10.0754 | 10.2562 | 9.6368 | 10.0051 | 9.6712 | 10.0261 | 10.1498 | 10.3049 | 9.8162 | 9.9649 | 9.9857 | 10.6706 | 9.934 | 10.0132 | 10.0792 | 10.4874 | 9.8281 | 9.9381 | 10.214 | 78.21% | 66.14% | 71.46% | 73.35% | 66.94% | 73% | 9.6553 | 9.7762 | 10.0501 | 10.1056
72 | CCV 10 ppb MAJORS |  | 06/16/2023 19:25:27 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.721 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9689.1906 | 9953.3571 | 9.403 | 8.9625 | 10606.1927 | 9782.6699 | 9988.6639 | 10.0685 | 10.5402 | 10.6579 | 10.2485 | 10.4787 | 9.9744 | 10.2245 | 9.8293 | 10.151 | 10.2338 | 10.3994 | 9.7787 | 10.1096 | 10.0639 | 10.5556 | 10.3708 | 10.2393 | 10.1633 | 10.5873 | 10.6943 | 10.0407 | 10.2687 | 80.29% | 67.91% | 72.92% | 75.02% | 68.57% | 74.83% | 9.7836 | 9.7555 | 9.7335 | 9.9283
73 | 10_1_21_W_11_DUP_LSP |  | 06/16/2023 19:31:27 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_W_11_DUP_LSP.722 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 21099.4095 | 21681.7785 | 50.5648 | 23.655 | 23812.2807 | 20403.4455 | 27095.5336 | 20.7957 | 21.2822 | 20.8528 | 40.2941 | 25.1748 | 39.462 | 20.7043 | 19.8801 | 20.7994 | 20.4969 | 20.9631 | 20.7963 | 20.5839 | 21.1385 | 20.3535 | 20.5938 | 20.0012 | 20.2836 | 21.0746 | 20.9631 | 20.5648 | 21.1591 | 79.54% | 66.48% | 70.87% | 72.6% | 66.31% | 71.74% | 19.0252 | 19.1029 | 17.47 | 19.1188
74 | 10_1_21_F_1 |  | 06/16/2023 19:37:27 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_F_1.723 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 8742.7159 | 6405.1866 | 3.167 | 6.0707 | 25855.2236 | 2306.763 | 37692.4243 | 0.23152 | 0.073985 | 0.073037 | 11.4893 | 18.1682 | 12.0924 | 0.022389 | 0.42714 | 0.84493 | 0.90022 | 14.784 | 2.9138 | 2.704 | 2.9412 | 0.44004 | 0.071566 | 0.043453 | 0.02738 | 0.47343 | 0.039489 | 17.0741 | 17.6609 | 76% | 67.31% | 70.03% | 72.32% | 68.35% | 72.87% | 0.010785 | 0.0063999 | 0.05286 | 0.056564
75 | 10_1_21_F_1_DUP |  | 06/16/2023 19:43:27 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_F_1_DUP.724 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9263.3257 | 6778.5892 | 2.9322 | 5.5897 | 26615.3499 | 2410.8315 | 39017.5819 | 0.22279 | 0.062003 | 0.072469 | 8.2742 | 18.699 | 8.941 | 0.028896 | 0.48295 | 0.79325 | 0.88599 | 14.4302 | 2.9967 | 2.7682 | 3.0565 | 0.46945 | 0.054116 | 0.061047 | 0.040309 | 0.41476 | 0.060321 | 17.7716 | 18.1895 | 75.66% | 66.9% | 69.1% | 71.4% | 68.35% | 72.19% | 0.010208 | 0.0058715 | 0.051917 | 0.054907
76 | 10_1_21_F_6 |  | 06/16/2023 19:49:26 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_F_6.725 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9298.8645 | 7693.4769 | 2.9702 | 2.1459 | 19871.8072 | 2124.291 | 39223.3973 | 0.61224 | 0.068168 | 0.073008 | 7.3023 | 11.2861 | 8.1134 | 0.014791 | 0.28575 | 2.1475 | 2.1909 | 5.2438 | 3.4724 | 3.1372 | 3.5601 | 0.32161 | 0.077464 | 0.090957 | 0.058505 | 0.41388 | 0.079956 | 8.0802 | 8.2927 | 76.28% | 67.51% | 69.56% | 71.78% | 68.01% | 72.08% | 0.014942 | 0.012828 | 0.093 | 0.096764
77 | 10_1_21_F_6_DUP |  | 06/16/2023 19:55:26 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_F_6_DUP.726 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9073.1547 | 7643.3788 | 3.1145 | 2.0632 | 19827.4014 | 2064.2721 | 37819.6465 | 0.6252 | 0.0774 | 0.071969 | 7.3302 | 11.3088 | 8.1017 | 0.015033 | 0.24745 | 2.1239 | 2.2132 | 5.949 | 3.5128 | 3.1918 | 3.5796 | 0.38132 | 0.090731 | 0.067165 | 0.074158 | 0.51767 | 0.055653 | 7.9611 | 8.3101 | 75.61% | 66.49% | 69.15% | 70.86% | 68.02% | 72.25% | 0.013986 | 0.0092485 | 0.10072 | 0.10339
78 | 10_1_21_F_8 |  | 06/16/2023 20:01:27 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_F_8.727 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9802.4587 | 7873.6674 | 2.9216 | 2.1788 | 18423.7304 | 2146.8071 | 38184.1905 | 0.80314 | 0.059471 | 0.057574 | 8.5179 | 10.7656 | 9.2095 | 0.012868 | 0.21698 | 2.3688 | 2.4325 | 5.689 | 4.0224 | 3.7941 | 4.0989 | 0.53463 | 0.090242 | 0.087958 | 0.083625 | 0.45329 | 0.060293 | 7.1881 | 7.4649 | 75.86% | 66.88% | 69.44% | 72.22% | 68.14% | 71.31% | 0.0090979 | 0.0088977 | 0.20797 | 0.2049
79 | 10_1_21_F_9 |  | 06/16/2023 20:07:27 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_F_9.728 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 8271.2277 | 7374.2568 | 3.0834 | 2.8137 | 15606.1718 | 1905.2783 | 33829.3697 | 0.64553 | 0.11516 | 0.12303 | 5.8875 | 5.2989 | 6.6274 | 0.011546 | 0.2179 | 2.2008 | 2.2179 | 6.9589 | 3.1205 | 2.9042 | 3.1696 | 0.57604 | 0.089734 | 0.098979 | 0.068709 | 0.4393 | 0.10568 | 4.9655 | 5.1406 | 77.04% | 67.22% | 69.96% | 72.49% | 68.47% | 72.23% | 0.01079 | 0.009053 | 0.56867 | 0.54358
80 | 10_1_21_F_9_DUP |  | 06/16/2023 20:13:27 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_F_9_DUP.729 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 8050.5994 | 7277.0856 | 3.9511 | 2.8481 | 15276.1024 | 1878.9129 | 33176.8647 | 0.66988 | 0.096024 | 0.10924 | 5.5729 | 5.2829 | 6.3201 | 0.026412 | 0.24651 | 2.2178 | 2.1847 | 6.9416 | 3.0468 | 2.9396 | 3.1595 | 0.492 | 0.053814 | 0.05712 | 0.071754 | 0.39993 | 0.070376 | 4.9542 | 5.0875 | 75.94% | 66.97% | 69.98% | 71.82% | 68.02% | 71.67% | 0.010424 | 0.008658 | 0.090634 | 0.092873
81 | 10_1_21_F_10 |  | 06/16/2023 20:19:27 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_F_10.730 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 7215.9152 | 8769.5475 | 7.2734 | 3.7305 | 15496.8067 | 1938.2461 | 35354.5845 | 0.71378 | 0.053591 | 0.058758 | 5.449 | 5.7516 | 6.1681 | 0.013086 | 0.18023 | 1.642 | 1.6867 | 7.6751 | 3.6392 | 3.2996 | 3.6945 | 0.62804 | 0.072647 | 0.061787 | 0.092025 | 0.43305 | 0.045708 | 3.0536 | 3.1571 | 80.36% | 66.14% | 68.64% | 71.4% | 67.79% | 72.24% | 0.0081257 | 0.0063848 | 18.7896 | 20.6667
82 | 10_1_21_F_11 |  | 06/16/2023 20:25:27 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_F_11.731 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 7110.057 | 8822.8049 | 5.9823 | 2.7911 | 15956.1457 | 1911.5726 | 35337.4308 | 0.73616 | 0.090888 | 0.099734 | 6.8677 | 5.0239 | 7.4942 | 0.013871 | 0.19325 | 1.5776 | 1.6761 | 7.2731 | 3.6826 | 3.3321 | 3.8473 | 0.48399 | 0.090918 | 0.08483 | 0.078173 | 0.3447 | 0.045618 | 2.918 | 2.9617 | 75.28% | 66.27% | 68.33% | 70.69% | 66.89% | 71.1% | 0.0087624 | 0.0053262 | 0.10438 | 0.10478
83 | LBLANK |  | 06/16/2023 20:31:27 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.732 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 55.3543 | 0.72782 | -1.0117 | -0.91666 | 447.6718 | 60.514 | 39.8091 | -0.00057673 | -0.0069438 | -0.0012405 | -0.57586 | -0.00052007 | -0.26616 | -0.0078885 | 0.0018702 | -0.011305 | -0.0096492 | -0.22954 | -0.016738 | -0.01135 | -0.0037458 | 0.0040065 | 0.0055366 | 0.0048521 | -0.01206 | -0.013452 | -0.016689 | -0.006875 | 0.0028274 | 80.56% | 72.06% | 75.43% | 76.14% | 70.8% | 74.53% | 0.00050427 | -0.0023379 | -0.010438 | -0.004459
84 | CCV 1ppb TRACE |  | 06/16/2023 20:37:27 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.733 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1036.9693 | 1078.0255 | 1.2511 | 0.12445 | 1274.8891 | 1063.1539 | 952.9247 | 1.058 | 1.0708 | 1.104 | 0.95774 | 1.0155 | 1.1228 | 1.0253 | 0.97374 | 1.0594 | 1.0654 | 0.85663 | 0.93263 | 1.0259 | 0.98818 | 1.0653 | 0.97813 | 0.96768 | 0.95858 | 0.99066 | 0.90136 | 0.97415 | 0.99372 | 79.15% | 71.08% | 73.5% | 75.55% | 70.54% | 75.67% | 0.96191 | 0.99069 | 1.0722 | 1.0157
85 | CCV 1ppb TRACE |  | 06/16/2023 20:43:46 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.734 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1040.4587 | 1103.9276 | 1.193 | 0.10895 | 1444.0459 | 1062.3671 | 1072.8531 | 1.0397 | 1.0729 | 1.0873 | 0.96254 | 1.0395 | 1.2168 | 1.029 | 1.0613 | 1.0792 | 1.0569 | 0.84771 | 0.9166 | 0.94519 | 0.9981 | 1.0994 | 1.105 | 0.94652 | 0.94985 | 0.97181 | 0.9856 | 0.98237 | 1.0205 | 80.81% | 70.6% | 74.1% | 76.12% | 69.51% | 75.24% | 0.97908 | 0.97651 | 1.0438 | 1.0203
86 | LFB |  | 06/16/2023 20:49:46 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.735 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9466.6541 | 9904.8327 | 9.4772 | 11.9438 | 10401.5438 | 9535.8912 | 9640.5655 | 9.5913 | 10.0779 | 10.0633 | 9.7932 | 9.9504 | 9.2658 | 9.741 | 9.4367 | 9.7016 | 9.8291 | 9.7731 | 9.8013 | 10.0316 | 9.739 | 10.624 | 9.8981 | 9.9175 | 9.9831 | 10.3206 | 9.9436 | 9.5256 | 9.8087 | 76.21% | 66.1% | 70.78% | 72.1% | 66.42% | 72.54% | 9.4004 | 9.4604 | 9.7719 | 9.6345
87 | CCV 10 ppb MAJORS |  | 06/16/2023 20:55:46 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.736 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9907.5254 | 10350.9786 | 9.9565 | 9.377 | 10682.0945 | 9945.1814 | 10117.4077 | 10.2223 | 10.5432 | 10.7419 | 10.3071 | 10.4417 | 9.7082 | 10.1628 | 9.7941 | 10.0052 | 10.3036 | 10.5344 | 9.7441 | 10.2478 | 9.9968 | 10.3172 | 10.0712 | 10.1348 | 10.1377 | 10.582 | 10.2023 | 9.9813 | 10.2245 | 79.38% | 66.35% | 70.25% | 72.9% | 67.44% | 72.79% | 9.7546 | 9.7844 | 9.7097 | 9.9626
88 | 10_1_21_UF_1 |  | 06/16/2023 21:01:46 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_UF_1.737 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9565.9098 | 6356.3857 | 0.34999 | 4.9861 | 25052.5414 | 2310.0364 | 37921.4774 | 0.23106 | 0.081269 | 0.064605 | 0.37956 | 17.867 | 1.4258 | 0.021832 | 0.38227 | 0.57217 | 0.62899 | 12.3857 | 2.7295 | 2.6363 | 2.8453 | 0.43792 | 0.060909 | 0.046443 | 0.037455 | 0.42122 | 0.046161 | 17.0459 | 17.3986 | 76.65% | 65.75% | 68.42% | 70.48% | 66.52% | 72.35% | 0.011255 | 0.0054113 | 0.022377 | 0.029689
89 | 10_1_21_UF_6 |  | 06/16/2023 21:07:46 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_UF_6.738 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10567.788 | 7987.2439 | 0.52271 | 3.1185 | 20518.5617 | 2195.1584 | 40248.9315 | 0.62622 | 0.095584 | 0.081501 | 0.28123 | 11.4151 | 1.4518 | 0.014599 | 0.26024 | 1.4437 | 1.5485 | 4.4138 | 3.3061 | 3.3731 | 3.5376 | 0.42696 | 0.067389 | 0.076407 | 0.080424 | 0.42684 | 0.10384 | 8.1382 | 8.2675 | 78.64% | 65.29% | 68.02% | 70.93% | 66.31% | 69.81% | 0.012277 | 0.0088445 | 0.046355 | 0.054753
90 | 10_1_21_UF_8 |  | 06/16/2023 21:13:46 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_UF_8.739 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10287.4416 | 7656.4409 | 0.52817 | 4.573 | 18040.6633 | 2133.4152 | 37895.7799 | 0.8091 | 0.16492 | 0.16073 | 0.88635 | 10.322 | 2.0103 | 0.017304 | 0.21873 | 1.2939 | 1.3682 | 4.395 | 3.8363 | 3.8934 | 3.9829 | 0.52551 | 0.080604 | 0.077414 | 0.074223 | 0.46841 | 0.068935 | 6.9984 | 7.159 | 75.12% | 64.33% | 67.96% | 69.74% | 66.44% | 70.18% | 0.01133 | 0.0056188 | 0.046361 | 0.051983
91 | 10_1_21_UF_8_LD |  | 06/16/2023 21:19:45 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_UF_8_LD.740 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10423.8202 | 7773.2355 | 1.4202 | 7.7731 | 17452.2012 | 2117.3708 | 36943.8326 | 0.79313 | 0.16307 | 0.14829 | 0.88894 | 10.1712 | 1.8357 | 0.011939 | 0.23289 | 1.2906 | 1.3532 | 4.3519 | 3.865 | 3.6391 | 3.9593 | 0.70811 | 0.061687 | 0.090947 | 0.071889 | 0.45243 | 0.084089 | 7.0082 | 7.1914 | 75.06% | 64.73% | 67.7% | 68.27% | 65.04% | 70.38% | 0.010528 | 0.0060094 | 0.044366 | 0.052771
92 | 10_1_21_UF_8_LD |  | 06/16/2023 21:26:03 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_UF_8_LD.741 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10273.4334 | 7508.2639 | 1.4141 | 8.1758 | 17632.898 | 2164.6772 | 36960.6201 | 0.82172 | 0.17465 | 0.15192 | 0.86339 | 10.2554 | 1.758 | 0.012183 | 0.24061 | 1.2665 | 1.36 | 4.3385 | 3.8132 | 3.5223 | 3.9473 | 0.52993 | 0.05594 | 0.087908 | 0.075636 | 0.48323 | 0.047787 | 6.9729 | 7.2246 | 74.74% | 64.06% | 67.6% | 70.05% | 64.94% | 70.56% | 0.011505 | 0.0077663 | 0.045513 | 0.050741
93 | 10_1_21_UF_8_DUP |  | 06/16/2023 21:32:02 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_UF_8_DUP.742 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10395.9339 | 7712.0563 | 0.49512 | 5.3695 | 17918.3544 | 2159.0861 | 37599.3298 | 0.79388 | 0.28663 | 0.2676 | 1.699 | 10.2653 | 2.6554 | 0.015724 | 0.23428 | 1.3991 | 1.406 | 4.8656 | 3.9155 | 3.4748 | 3.9794 | 0.42577 | 0.10131 | 0.084821 | 0.089093 | 0.51932 | 0.081039 | 7.0821 | 7.2285 | 73.66% | 63.27% | 66.21% | 68.96% | 64.11% | 68.33% | 0.0096322 | 0.0088035 | 0.053256 | 0.055655
94 | 10_1_21_UF_8_DUP_LSP |  | 06/16/2023 21:38:02 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_UF_8_DUP_LSP.743 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 19395.3071 | 16899.0275 | 11.1808 | 16.4541 | 26417.594 | 12033.9032 | 43587.2902 | 11.0948 | 10.6658 | 10.7603 | 12.718 | 19.5298 | 12.6553 | 10.179 | 9.692 | 10.9061 | 11.1057 | 14.1191 | 13.1521 | 13.3973 | 13.3559 | 10.7136 | 9.8977 | 10.2027 | 9.9796 | 10.9439 | 9.791 | 16.3919 | 16.5672 | 74.52% | 62.62% | 65.2% | 66.82% | 63.18% | 67.69% | 9.5021 | 9.5257 | 9.4286 | 9.751
95 | 10_1_21_UF_9 |  | 06/16/2023 21:44:02 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_UF_9.744 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9181.1223 | 7265.0967 | 0.090679 | 5.4856 | 15397.3048 | 1878.5572 | 34058.2355 | 0.65891 | 0.07963 | 0.05955 | 0.56798 | 5.1421 | 1.5824 | 0.011856 | 0.20379 | 1.3465 | 1.3943 | 6.0107 | 3.016 | 3.0197 | 3.158 | 0.43566 | 0.07489 | 0.10003 | 0.073556 | 0.38209 | 0.079943 | 4.9891 | 5.1125 | 76.64% | 64.05% | 67.62% | 68.8% | 65.96% | 70.59% | 0.0099357 | 0.0089202 | 0.040606 | 0.047614
96 | 10_1_21_UF_10 |  | 06/16/2023 21:50:03 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_UF_10.745 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 8285.4207 | 8770.3211 | 0.012551 | 4.5596 | 15678.3237 | 1938.0638 | 35826.3709 | 0.73358 | 0.29517 | 0.28494 | 1.6836 | 5.4361 | 2.5709 | 0.010188 | 0.19433 | 1.0301 | 1.0762 | 5.0537 | 3.276 | 3.2359 | 3.6643 | 0.3434 | 0.094532 | 0.094293 | 0.074442 | 0.38877 | 0.080798 | 3.038 | 3.1485 | 81.01% | 63.62% | 67.48% | 69.79% | 64.73% | 69.23% | 0.0069081 | 0.0042631 | 0.057783 | 0.063994
97 | 10_1_21_UF_11 |  | 06/16/2023 21:56:03 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_UF_11.746 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 7796.2775 | 8747.7439 | -0.045743 | 3.8452 | 15584.5364 | 1913.1878 | 35413.076 | 0.71627 | 0.067322 | 0.060274 | 1.6133 | 4.6069 | 2.4692 | 0.010103 | 0.16813 | 0.94025 | 1.0007 | 5.2103 | 3.3255 | 3.3699 | 3.7391 | 0.38255 | 0.11298 | 0.091824 | 0.085954 | 0.31002 | 0.074532 | 2.8191 | 2.873 | 79.93% | 64.06% | 67.96% | 70.32% | 64.94% | 70.16% | 0.0073545 | 0.0044562 | 0.045802 | 0.054742
98 | LBLANK |  | 06/16/2023 22:02:04 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.747 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 55.3996 | 2.2875 | -0.90892 | -0.86876 | 710.7681 | 58.875 | 42.0219 | -0.00021605 | -0.006893 | -0.0011939 | -0.57746 | -0.00077284 | -0.32539 | -0.0075884 | 0.0085943 | -0.012158 | -0.010057 | -0.22863 | 0.037665 | -0.0037923 | 0.01697 | 0.0048278 | 0.011408 | 0.0089106 | -0.0076485 | -0.040154 | -0.021474 | -0.0075587 | 0.003048 | 80% | 68.37% | 73.37% | 74.96% | 68.27% | 72.86% | 0.00083952 | -0.0021333 | -0.010903 | -0.0044755
99 | CCV 1ppb TRACE |  | 06/16/2023 22:08:04 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.748 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1025.6458 | 1070.4949 | 1.142 | 0.089874 | 1758.7854 | 1038.7734 | 1043.2635 | 1.0473 | 1.0654 | 1.1002 | 0.95648 | 1.0125 | 0.99979 | 1.0335 | 1.0321 | 1.0676 | 1.0825 | 0.83832 | 0.94036 | 1.0034 | 0.98286 | 1.1848 | 1.1063 | 0.996 | 0.96327 | 0.93271 | 0.98896 | 0.97105 | 1.0138 | 78.75% | 68.28% | 71.13% | 73.65% | 66.5% | 72.75% | 0.98231 | 0.96731 | 1.1049 | 1.0381
100 | CCV 1ppb TRACE |  | 06/16/2023 22:14:23 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.749 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1042.9753 | 1077.1368 | 1.0946 | 0.055224 | 1588.5975 | 1046.2009 | 1073.8464 | 1.0625 | 1.0754 | 1.08 | 0.90942 | 1.0248 | 1.1486 | 1.0834 | 1.0595 | 1.065 | 1.0643 | 0.82146 | 0.89691 | 1.0153 | 0.99907 | 1.2416 | 1.0274 | 0.96342 | 0.97571 | 0.95118 | 1.1049 | 0.97055 | 1.017 | 78.85% | 67.64% | 72.11% | 79.46% | 66.12% | 72.15% | 0.97399 | 0.98445 | 1.0887 | 0.93946
101 | CCV 10 ppb MAJORS |  | 06/16/2023 22:20:23 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.750 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9730.104 | 10148.7696 | 9.5046 | 9.0768 | 10884.6744 | 9980.2188 | 9967.1371 | 10.1639 | 10.6508 | 10.6538 | 10.304 | 10.5466 | 9.7354 | 10.2861 | 9.7647 | 10.1525 | 10.228 | 10.429 | 9.6983 | 9.8855 | 10.1663 | 10.4224 | 10.3729 | 10.0838 | 9.9256 | 10.2341 | 10.3914 | 9.995 | 10.3779 | 76.72% | 61.69% | 67.72% | 68.64% | 62.94% | 68.68% | 9.7695 | 9.8099 | 9.8755 | 10.1874
102 | 10_1_21_BB |  | 06/16/2023 22:26:24 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_BB.751 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 55.1546 | 1.2239 | -0.25314 | -0.61745 | 692.0038 | 58.1697 | 43.4113 | 0.00010834 | 0.012387 | 0.024695 | 0.080283 | 0.0045032 | 0.32342 | -0.0062728 | 0.0030991 | 0.079735 | 0.081493 | -0.031276 | 0.023456 | 0.0068852 | 0.0058004 | -0.075634 | 0.011455 | 0.0070818 | -0.009398 | -0.0091788 | -0.016418 | -0.0046728 | 0.0035826 | 80.84% | 67.97% | 73.66% | 74.56% | 66.73% | 74.27% | 0.00095124 | -0.0017145 | 0.020115 | 0.02603
103 | 10_1_21_FB |  | 06/16/2023 22:32:24 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_FB.752 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 248.2807 | 1.2699 | 4.3483 | 0.32434 | 485.9773 | 63.7857 | 52.5276 | 0.0022868 | 0.047236 | 0.056925 | 0.47676 | 0.040819 | 0.65784 | -0.0052927 | 0.052968 | 0.22894 | 0.23379 | 4.1399 | -0.031358 | 0.0038511 | 0.0057717 | 0.090935 | 0.011596 | 0.0053161 | -0.0057177 | -0.015839 | -0.021474 | -0.00084081 | 0.0082393 | 79.01% | 66.48% | 71.82% | 73.41% | 65.65% | 72.11% | 0.00084424 | -0.0016373 | 0.019352 | 0.024384
104 | 10_1_21_LAB_A |  | 06/16/2023 22:38:24 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_LAB_A.753 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 56.5068 | 0.23633 | -0.21045 | -0.67577 | 481.8848 | 59.104 | 45.7502 | -0.00020634 | 0.0084091 | 0.01495 | -0.11956 | 0.0076566 | -0.073696 | -0.0069724 | 0.0065624 | 0.083533 | 0.078229 | 0.12564 | 0.061907 | -6.1539e-05 | 0.0012466 | 0.0036966 | 0.011675 | 0.011181 | -0.0084087 | 0.0031808 | -0.016313 | -0.0075698 | 0.0030503 | 78.87% | 66.06% | 72.49% | 73.44% | 64.68% | 72.2% | 0.00050427 | -0.0021193 | 0.018135 | 0.02193
105 | 10_1_21_LAB_B |  | 06/16/2023 22:44:24 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_LAB_B.754 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 56.7511 | 0.75978 | -0.33986 | -0.60508 | 356.0908 | 54.3773 | 42.0522 | -0.00019717 | 0.043227 | 0.046668 | -0.073799 | 0.010293 | 0.035383 | -0.0069807 | 0.0084083 | 0.092568 | 0.091436 | 0.089311 | 0.054174 | -0.0075353 | 0.000731 | -0.18223 | 0.0055366 | 0.0014405 | -0.0074628 | -0.035078 | -0.01626 | -0.0082757 | 0.0021832 | 79.39% | 65.72% | 72.29% | 73.68% | 64.35% | 72.14% | 0.00061721 | -0.0018803 | 0.040749 | 0.042802
106 | 5_18_23_DB |  | 06/16/2023 22:50:24 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_DB.755 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 43769.9619 | 12318.5837 | 113.1533 | 10.8613 | 1207.6708 | 3716.8541 | 49307.7328 | 1.9543 | 1.4534 | 1.4731 | 42.6719 | 3.5835 | 39.5921 | 1.3396 | 1.6309 | 0.77295 | 0.86221 | 5.8381 | 2.2272 | 2.1675 | 2.3597 | 4.8171 | 2.0437 | 1.8249 | 1.8442 | 1.9192 | 2.0587 | 1.0445 | 1.08 | 70.23% | 54.54% | 61.13% | 64.17% | 54.65% | 61.01% | 3.267 | 3.2022 | 1.3019 | 1.2691
107 | 5_18_23_DB |  | 06/16/2023 22:56:42 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_DB.756 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 43362.4959 | 12410.459 | 108.812 | 11.8107 | 961.5035 | 3682.17 | 49526.3751 | 1.8454 | 1.4003 | 1.4197 | 41.8412 | 3.573 | 38.854 | 1.357 | 1.5545 | 0.7449 | 0.85922 | 5.795 | 2.2864 | 2.1086 | 2.3171 | 5.0834 | 1.9291 | 1.8769 | 1.853 | 2.1861 | 1.9397 | 1.0555 | 1.0916 | 68.53% | 54.89% | 62.38% | 63.77% | 54.65% | 62.32% | 3.2344 | 3.1803 | 1.3582 | 1.2892
108 | 5_18_23_T_239 |  | 06/16/2023 23:02:42 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_T_239.757 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 43818.6967 | 12401.9544 | 112.5601 | 12.1886 | 1049.0914 | 3786.4273 | 50300.4608 | 1.8823 | 1.4043 | 1.4782 | 42.5428 | 3.6191 | 39.7527 | 1.3633 | 1.5978 | 0.75129 | 0.84652 | 5.9218 | 2.2436 | 2.1529 | 2.3504 | 5.183 | 1.8762 | 1.911 | 1.8645 | 2.1089 | 1.7647 | 1.1012 | 1.0772 | 69.01% | 54.9% | 61.65% | 64.91% | 54.28% | 61.45% | 3.2441 | 3.2062 | 1.3426 | 1.2786
109 | 5_18_23_T_239 |  | 06/16/2023 23:09:00 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_T_239.758 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 43638.5231 | 12318.0966 | 114.4458 | 11.7129 | 1252.4876 | 3804.2649 | 50684.1641 | 1.886 | 1.4278 | 1.4217 | 41.8146 | 3.451 | 39.7999 | 1.3533 | 1.7091 | 0.74796 | 0.83342 | 5.7601 | 2.2777 | 2.1467 | 2.2225 | 5.0414 | 1.6799 | 1.8186 | 1.7585 | 2.1142 | 1.806 | 1.0965 | 1.0362 | 68.16% | 56.3% | 61.92% | 63.18% | 54.43% | 60.83% | 3.2232 | 3.183 | 1.3481 | 1.3047
110 | 5_18_23_T_239 |  | 06/16/2023 23:15:00 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_T_239.759 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4013.3513 | 652.6814 | 7.8734 | 16.8415 | 1351.5605 | 524.8745 | 3247.7153 | 1.0284 | 0.029052 | 0.037332 | -0.29472 | 5.436 | -0.050935 | 0.0093901 | 0.049561 | 0.052033 | 0.070904 | 0.59276 | 0.20613 | 0.033469 | 0.053173 | 0.26488 | 0.018852 | 0.018528 | 0.012771 | 0.82724 | 0.017952 | 0.23726 | 0.19843 | 75.54% | 60.86% | 67.78% | 73.71% | 60.62% | 66.04% | 0.0044874 | -0.00018506 | 0.0082028 | 0.013524
111 | 5_18_23_M_224 |  | 06/16/2023 23:20:59 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_M_224.760 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4417.0847 | 4583.8643 | 4.3343 | 9.4264 | 5249.3873 | 4611.0071 | 4862.5717 | 4.8126 | 4.9756 | 5.1286 | 4.7419 | 5.0416 | 4.4366 | 4.8558 | 4.6759 | 4.9378 | 5.063 | 4.9123 | 4.9229 | 4.7641 | 4.8185 | 5.9733 | 5.0512 | 4.5848 | 4.7783 | 5.345 | 5.0127 | 4.8095 | 4.942 | 69.37% | 61.15% | 68.63% | 70.21% | 60.93% | 69.05% | 4.7627 | 4.7858 | 5.9992 | 5.2479
112 | LBLANK |  | 06/16/2023 23:27:00 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.761 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 54.9966 | 0.765 | -0.98765 | 0.27807 | 453.1504 | 57.124 | 40.6215 | 0.0015915 | -0.0054195 | -0.0026971 | -0.57853 | -0.00011699 | -0.28595 | -0.00588 | -0.0025338 | -0.012831 | -0.010844 | -0.22593 | 0.014198 | -0.01135 | -0.0014082 | -0.099117 | 0.018063 | 0.0054514 | -0.011141 | -0.0027456 | -0.011061 | 0.015293 | 0.0032846 | 79.88% | 65.1% | 72.29% | 73.47% | 63.15% | 71.26% | 0.00061655 | -0.0017854 | -0.01059 | -0.004417
113 | LBLANK |  | 06/16/2023 23:33:18 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.762 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 55.6865 | 0.75871 | -0.99821 | -0.70648 | 531.5513 | 57.2547 | 40.6163 | 0.00052196 | -0.0056343 | -0.0020015 | -0.57838 | -0.00023853 | -0.28665 | -0.0072719 | -0.0024492 | -0.013827 | -0.011843 | -0.22239 | 0.054473 | -0.0074625 | -0.0027578 | -0.084283 | 0.0055366 | 0.0034757 | -0.011125 | -0.045514 | -0.021474 | 0.0079127 | 0.0024118 | 78.97% | 64.45% | 76.36% | 73.61% | 62.69% | 70.84% | 0.00072017 | -0.0025324 | -0.010044 | -0.0041253
114 | CCV 1ppb TRACE |  | 06/16/2023 23:39:19 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.763 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 973.5156 | 996.1043 | 1.0589 | 0.14864 | 1343.6848 | 993.3618 | 1012.2498 | 1.0111 | 1.0245 | 1.0526 | 0.85519 | 1.0001 | 1.0524 | 0.97807 | 0.94431 | 1.0144 | 1.0229 | 0.78053 | 0.92841 | 0.99216 | 0.95211 | 1.2711 | 0.88487 | 0.95893 | 0.97227 | 0.95235 | 0.9513 | 0.97658 | 0.99264 | 78.33% | 63.95% | 71.94% | 73.62% | 62.34% | 70.62% | 0.95932 | 0.96684 | 1.1404 | 1.0837
115 | CCV 1ppb TRACE |  | 06/16/2023 23:45:38 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.764 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 988.5811 | 943.7088 | 1.1674 | 0.089159 | 1565.2944 | 990.3717 | 1077.1318 | 0.99564 | 1.0509 | 1.0646 | 0.93036 | 0.99354 | 1.0632 | 0.99528 | 0.99444 | 1.0052 | 1.0578 | 0.82988 | 0.99061 | 0.90047 | 0.97965 | 1.6046 | 1.0735 | 1.0641 | 0.97289 | 1.148 | 0.93666 | 1.0003 | 1.0057 | 68.17% | 63.54% | 70.08% | 72.26% | 62.14% | 69.83% | 0.95725 | 0.97728 | 1.3077 | 1.0615
116 | LFB |  | 06/16/2023 23:51:38 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.765 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 8856.5512 | 9233.4125 | 9.0278 | 11.5031 | 10225.7146 | 8971.6175 | 9523.1364 | 9.4157 | 9.8454 | 10.0143 | 9.6841 | 9.7627 | 8.711 | 9.4044 | 9.2996 | 9.5611 | 9.7698 | 9.8099 | 9.0941 | 9.6085 | 9.6807 | 10.0371 | 9.4857 | 9.6289 | 9.9364 | 9.6183 | 9.8764 | 9.6503 | 9.8539 | 76.6% | 59.15% | 65.61% | 67.66% | 58.8% | 66.09% | 9.5391 | 9.5837 | 9.4655 | 10.0337
117 | LFB |  | 06/16/2023 23:57:55 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.766 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 8924.9738 | 9503.9294 | 9.1653 | 11.5659 | 10370.3317 | 9199.315 | 9709.2367 | 9.5653 | 10.0675 | 10.037 | 9.7676 | 9.8653 | 8.7551 | 9.6751 | 9.3953 | 9.6844 | 9.9154 | 9.7938 | 9.2828 | 9.5281 | 9.8623 | 10.2711 | 9.9673 | 10.1907 | 10.0648 | 9.8316 | 10.3033 | 9.6507 | 9.8937 | 75.47% | 59.81% | 66.63% | 68.64% | 60.52% | 66.87% | 9.5445 | 9.6102 | 9.6924 | 9.9435
118 | CCV 10 ppb MAJORS |  | 06/17/2023 00:03:55 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.767 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9146.4813 | 9732.3602 | 9.1596 | 9.2258 | 10290.665 | 9419.1562 | 9950.6236 | 9.838 | 10.2238 | 10.3956 | 10.111 | 10.221 | 9.3725 | 9.9119 | 9.504 | 9.8709 | 10.1029 | 10.2163 | 9.8929 | 9.2593 | 9.9075 | 10.9237 | 9.476 | 9.8004 | 10.0228 | 10.2829 | 10.5384 | 9.9149 | 10.2118 | 73.91% | 60.71% | 67.92% | 69.11% | 61.08% | 69.03% | 9.6348 | 9.7148 | 10.4492 | 10.3041
119 | 5_18_23_MFB |  | 06/17/2023 00:09:55 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_MFB.768 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 56.6753 | 0.23633 | -0.59115 | 2.7087 | 406.8146 | 61.5672 | 47.099 | 0.0011718 | 0.0025575 | 0.0076277 | -0.43906 | 0.0022547 | -0.074802 | -0.0066936 | 0.0010111 | 0.066099 | 0.065083 | 0.062457 | 0.05115 | -0.0076308 | -0.0021745 | 0.063854 | 0.011577 | 0.0090925 | -0.0075515 | -0.040273 | -0.021474 | 0.01862 | 0.0028276 | 82.84% | 66.85% | 74.76% | 75.53% | 66.3% | 73.48% | 0.00094041 | -0.0021917 | 0.0081761 | 0.013714
120 | 10_15_21_W_1 |  | 06/17/2023 00:15:55 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_W_1.769 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2005.5658 | 1413.4402 | 4.2652 | 3.7352 | 6577.8116 | 561.4068 | 8786.5903 | 0.047893 | 0.027796 | 0.028673 | 11.1255 | 5.9621 | 10.6707 | 0.0026552 | 0.079964 | 0.29786 | 0.32803 | 0.15963 | 0.61128 | 0.62592 | 0.67172 | 0.28938 | 0.031875 | 0.018321 | 0.004674 | 0.11832 | -0.010342 | 4.1547 | 4.2854 | 75.91% | 61.36% | 68.55% | 71.07% | 61.78% | 69.65% | 0.0039643 | 3.6953e-05 | 0.027694 | 0.029919
121 | 10_15_21_W_6 |  | 06/17/2023 00:21:56 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_W_6.770 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1820.2365 | 1468.9528 | 14.72 | 2.0336 | 4891.6762 | 491.3656 | 8091.1895 | 0.12697 | 0.024065 | 0.051147 | 17.161 | 3.8165 | 15.9952 | 0.0028552 | 0.056043 | 1.1194 | 1.1462 | 0.73276 | 0.64307 | 0.62888 | 0.70034 | 0.10601 | 0.03898 | 0.03133 | 0.0099013 | 0.09291 | -0.015736 | 2.3016 | 2.3506 | 75.43% | 60.52% | 67.62% | 69.47% | 61.33% | 68.51% | 0.0081055 | 0.0029772 | 0.12268 | 0.12325
122 | 10_15_21_W_8 |  | 06/17/2023 00:27:56 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_W_8.771 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1833.5767 | 1389.3738 | 5.9307 | 1.9088 | 4423.7611 | 464.205 | 7863.5214 | 0.13799 | 0.015746 | 0.017335 | 8.2087 | 2.5864 | 7.6748 | 0.00036112 | 0.053662 | 0.67381 | 0.72028 | 0.27556 | 0.72553 | 0.66762 | 0.72841 | 0.20391 | 0.032373 | 0.018776 | 0.0070127 | 0.06854 | 0.012562 | 2.0589 | 2.0931 | 75.42% | 60.24% | 66.51% | 69.72% | 60.63% | 67.85% | 0.0029558 | 0.00054757 | 0.052617 | 0.054571
123 | 10_15_21_W_9 |  | 06/17/2023 00:33:56 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_W_9.772 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1572.1129 | 1292.7546 | 9.5485 | 1.8094 | 3728.0764 | 403.3684 | 6875.8952 | 0.11606 | 0.030872 | 0.034331 | 9.0725 | 2.3189 | 8.4067 | -0.0014588 | 0.056117 | 0.58915 | 0.62044 | 0.25387 | 0.5951 | 0.54605 | 0.58819 | 0.78538 | 0.019037 | 0.018882 | 0.006106 | 0.16521 | 0.0071705 | 1.4202 | 1.4846 | 64.24% | 59.69% | 67.57% | 69.72% | 60.6% | 68.45% | 0.0044981 | 0.00048436 | 0.06522 | 0.054579
124 | 10_15_21_W_9 |  | 06/17/2023 00:40:15 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_W_9.773 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1590.7698 | 1274.7171 | 8.7663 | 1.8495 | 3698.9078 | 410.1781 | 6898.2629 | 0.1177 | 0.029519 | 0.037816 | 9.099 | 2.28 | 8.5242 | -0.0010056 | 0.036501 | 0.62082 | 0.63553 | 0.26782 | 0.57476 | 0.47825 | 0.59738 | 0.19708 | 0.012288 | 0.012467 | 0.0042995 | 0.11182 | 0.00718 | 1.4431 | 1.4681 | 74.86% | 59.47% | 66.53% | 68.83% | 60.02% | 66.94% | 0.0033199 | 0.00028658 | 0.051986 | 0.05653
125 | 10_15_21_W_10 |  | 06/17/2023 00:46:15 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_W_10.774 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1285.8078 | 1505.7855 | 30.7228 | 2.9233 | 3222.2775 | 403.2486 | 6651.2359 | 0.15196 | 0.033668 | 0.034283 | 25.7102 | 5.1996 | 23.6581 | 0.0072383 | 0.048663 | 0.86917 | 0.88971 | 1.0798 | 0.65321 | 0.6412 | 0.6658 | 0.16078 | 0.032907 | 0.023412 | 0.0083637 | 0.046454 | -0.015696 | 0.63991 | 0.66836 | 74.27% | 59.13% | 66.39% | 68.91% | 59.31% | 66.02% | 0.0067906 | 0.0042742 | 0.1793 | 0.16979
126 | 10_15_21_W_10 |  | 06/17/2023 00:52:33 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_W_10.775 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1268.8095 | 1473.1627 | 31.2012 | 2.7824 | 3205.0898 | 367.1392 | 6757.8141 | 0.15041 | 0.025625 | 0.031782 | 25.6163 | 5.2486 | 23.8756 | 0.0071848 | 0.040801 | 0.86183 | 0.88084 | 1.0925 | 0.60909 | 0.48603 | 0.68398 | 0.075173 | 0.032777 | 0.0168 | 0.0092596 | 0.060801 | -0.0043024 | 0.63699 | 0.66042 | 75.35% | 59.5% | 67.14% | 69.75% | 60.2% | 66.31% | 0.0064709 | 0.0032196 | 0.17381 | 0.17129
127 | 10_15_21_W_10_DUP |  | 06/17/2023 00:58:33 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_W_10_DUP.776 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1229.5166 | 1420.7281 | 28.7666 | 2.8389 | 3184.337 | 382.2393 | 6571.0237 | 0.1441 | 0.057515 | 0.063047 | 24.983 | 5.0238 | 22.7382 | 0.0059246 | 0.044269 | 0.85086 | 0.8821 | 0.96952 | 0.59848 | 0.54385 | 0.67152 | 0.24149 | 0.0055366 | 0.010246 | -0.00095537 | 0.071134 | -0.0042822 | 0.62367 | 0.63807 | 75.79% | 59.93% | 68.01% | 68.51% | 59.87% | 66.78% | 0.007231 | 0.0034866 | 0.16196 | 0.16484
128 | 10_15_21_W_10_DUP |  | 06/17/2023 01:04:52 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_W_10_DUP.777 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1247.1092 | 1458.0539 | 27.6383 | 2.85 | 3208.9612 | 387.2612 | 6409.5866 | 0.14541 | 0.055372 | 0.057049 | 25.2636 | 5.0287 | 22.9964 | 0.0059069 | 0.048256 | 0.86047 | 0.87087 | 0.95387 | 0.62072 | 0.60296 | 0.69031 | 0.16657 | 0.0055366 | 0.019184 | 0.013621 | 0.12264 | -0.0098875 | 0.61297 | 0.65128 | 73.61% | 58.74% | 66.43% | 68.29% | 58.67% | 65.15% | 0.0075176 | 0.0030841 | 0.16071 | 0.15901
129 | LBLANK |  | 06/17/2023 01:10:52 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.778 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 54.3812 | 0.80206 | -0.99466 | -0.53709 | 462.2207 | 59.9633 | 39.8954 | 0.00063728 | -0.0098968 | -0.0023999 | -0.58696 | -0.00071708 | -0.45669 | -0.0077222 | -0.00023988 | -0.010848 | -0.012549 | -0.23468 | 0.0044616 | -0.01135 | -0.0031847 | -0.035661 | 0.0055366 | 0.0037498 | -0.011061 | -0.02219 | -0.015854 | -0.0025302 | 0.0025004 | 75.47% | 60.31% | 68.86% | 69.07% | 59.2% | 66.46% | 0.00050427 | -0.0021725 | -0.010437 | -0.0044141
130 | LBLANK |  | 06/17/2023 01:17:12 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.779 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 54.8716 | 0.23633 | -0.95262 | -0.80843 | 506.4749 | 61.3376 | 39.9056 | -0.00052778 | -0.0079865 | -0.0020988 | -0.58614 | -0.00064534 | -0.34749 | -0.007392 | 0.0018527 | -0.011356 | -0.011045 | -0.22706 | 0.066471 | -0.01135 | -0.0036289 | -0.21459 | 0.0055366 | 0.0016159 | -0.0060409 | -0.0019992 | -0.021474 | -0.0033717 | 0.002501 | 78.66% | 60.28% | 67.57% | 68.82% | 58.69% | 65.99% | 0.00062566 | -0.0023092 | -0.01067 | -0.0043961
131 | CCV 1ppb TRACE |  | 06/17/2023 01:23:12 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.780 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 972.3873 | 1032.7939 | 1.2642 | 0.080488 | 1454.5502 | 1054.1315 | 1036.1287 | 1.0181 | 1.0324 | 1.0541 | 0.89712 | 0.99441 | 0.96018 | 1.0006 | 0.96316 | 1.0235 | 1.043 | 0.80002 | 0.89649 | 1.0668 | 0.95919 | 1.105 | 0.86769 | 1.0299 | 0.95419 | 0.95715 | 1.0801 | 0.98211 | 1.0186 | 73.11% | 59.55% | 66.73% | 67.87% | 58.29% | 65.02% | 0.97143 | 0.99687 | 1.1638 | 1.0874
132 | CCV 1ppb TRACE |  | 06/17/2023 01:29:31 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.781 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 965.4585 | 962.0715 | 1.0941 | 0.10977 | 1521.441 | 1017.0475 | 1037.8145 | 1.022 | 1.0321 | 1.0666 | 0.93879 | 1.0066 | 0.95728 | 0.98594 | 0.98441 | 1.0047 | 1.034 | 0.83014 | 0.91458 | 1.0204 | 0.97087 | 1.174 | 0.95901 | 1.029 | 0.98943 | 0.95493 | 1.0794 | 0.98612 | 1.0012 | 73.05% | 59.32% | 67.15% | 66.97% | 58.04% | 66.05% | 0.97472 | 0.97825 | 1.1556 | 1.1375
133 | CCV 10 ppb MAJORS |  | 06/17/2023 01:35:32 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.782 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9198.8992 | 9642.5842 | 8.8068 | 8.8042 | 10567.5571 | 9467.1743 | 10040.5209 | 9.9471 | 10.2946 | 10.3894 | 10.2326 | 10.2764 | 9.1999 | 9.9296 | 9.4652 | 9.813 | 9.9928 | 10.2096 | 9.5101 | 10.0143 | 9.8402 | 10.1058 | 9.9663 | 9.8738 | 9.9824 | 10.0215 | 9.9297 | 9.8692 | 10.3203 | 71.84% | 55.53% | 63.03% | 64.79% | 55.81% | 62.55% | 9.7451 | 9.8086 | 10.3905 | 10.5557
134 | CCV 10 ppb MAJORS |  | 06/17/2023 01:41:51 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.783 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9210.9318 | 9666.5118 | 9.5342 | 8.785 | 10531.8412 | 9347.6247 | 10086.3066 | 9.9852 | 10.1448 | 10.6126 | 10.2737 | 10.2175 | 9.2229 | 9.8789 | 9.5724 | 9.6853 | 10.1415 | 10.2698 | 9.4948 | 9.719 | 9.8399 | 10.1636 | 10.0204 | 9.9062 | 9.97 | 10.3088 | 10.1556 | 9.9226 | 10.2306 | 74.38% | 57.44% | 64.92% | 66.46% | 57.54% | 64.95% | 9.7404 | 9.771 | 10.2194 | 10.4249
135 | 10_15_21_W_11 |  | 06/17/2023 01:47:51 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_W_11.784 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1202.0649 | 1500.0398 | 41.9447 | 2.8231 | 3194.1838 | 398.3913 | 6621.6961 | 0.17238 | 0.043371 | 0.046943 | 33.9716 | 5.7171 | 31.3393 | 0.010605 | 0.057631 | 1.0295 | 1.0937 | 1.7004 | 0.62582 | 0.60114 | 0.70033 | -0.049296 | 0.019075 | 0.0059015 | 0.011138 | 0.014905 | -0.0042934 | 0.56267 | 0.55949 | 81.75% | 59.94% | 66.66% | 69.27% | 59.93% | 67.66% | 0.0061421 | 0.0065104 | 0.18482 | 0.20569
136 | 10_15_21_W_11 |  | 06/17/2023 01:54:10 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_W_11.785 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1185.5387 | 1477.3299 | 40.491 | 3.0542 | 2985.4201 | 371.5053 | 6594.0206 | 0.17399 | 0.040188 | 0.045111 | 34.2345 | 5.6832 | 30.7074 | 0.01021 | 0.055387 | 0.99603 | 1.0887 | 1.6934 | 0.736 | 0.74405 | 0.7189 | 0.099548 | 0.033358 | 0.023794 | 0.0014201 | 0.044907 | -0.0038902 | 0.55994 | 0.57723 | 73.51% | 58.22% | 65.48% | 67.98% | 58.67% | 66.76% | 0.0093589 | 0.0053224 | 0.21451 | 0.20386
137 | 10_15_21_F_1 |  | 06/17/2023 02:00:09 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_1.786 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10219.3731 | 6847.8694 | 1.9268 | 6.4141 | 32544.0982 | 2787.0104 | 45317.1903 | 0.19546 | 0.067471 | 0.075495 | 12.227 | 19.9237 | 12.1574 | 0.024947 | 0.39748 | 0.77675 | 0.83846 | 1.2518 | 2.7852 | 2.5831 | 2.8186 | 0.61552 | 0.092999 | 0.068957 | 0.046631 | 0.52158 | 0.046131 | 21.3393 | 21.8333 | 71.71% | 55.53% | 62.58% | 71.65% | 55.64% | 60.92% | 0.0086021 | 0.0075661 | 3.6005 | 3.1027
138 | 10_15_21_F_1 |  | 06/17/2023 02:06:28 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_1.787 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10386.9913 | 7032.3456 | 1.9451 | 6.3023 | 32560.42 | 2844.9986 | 45525.2282 | 0.20668 | 0.067385 | 0.074497 | 12.4622 | 20.0952 | 12.4585 | 0.024122 | 0.34814 | 0.76734 | 0.8441 | 1.2882 | 2.7229 | 2.5591 | 2.8237 | 0.6889 | 0.082766 | 0.070961 | 0.047669 | 0.45425 | 0.07357 | 21.2365 | 21.9279 | 73.09% | 57.6% | 62.86% | 64.85% | 58.08% | 63.5% | 0.010252 | 0.0079057 | 3.497 | 3.4967
139 | 10_15_21_F_1_DUP |  | 06/17/2023 02:12:28 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_1_DUP.788 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9901.0171 | 6587.3786 | 1.1049 | 5.1146 | 30527.5174 | 2650.462 | 42640.337 | 0.203 | 0.053164 | 0.050694 | 11.6102 | 19.3078 | 11.6988 | 0.023372 | 0.35659 | 0.70677 | 0.77966 | 1.2538 | 2.6899 | 2.4528 | 2.7299 | 0.69758 | 0.081364 | 0.056555 | 0.050794 | 0.5058 | 0.071988 | 20.3201 | 20.8434 | 71.5% | 58.55% | 63.61% | 64.86% | 58.9% | 65.68% | 0.010912 | 0.006162 | 0.05058 | 0.053585
140 | 10_15_21_F_1_DUP |  | 06/17/2023 02:18:46 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_1_DUP.789 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10138.9519 | 6771.2102 | 1.1569 | 5.1716 | 30874.0792 | 2702.4277 | 43277.5899 | 0.17748 | 0.04635 | 0.050235 | 11.6423 | 19.1878 | 11.8251 | 0.021165 | 0.35103 | 0.72759 | 0.77051 | 1.2365 | 2.6951 | 2.2757 | 2.7049 | 0.66134 | 0.053359 | 0.084417 | 0.039096 | 0.47696 | 0.042306 | 20.1097 | 20.7729 | 71.84% | 59% | 64% | 65.73% | 59.23% | 65.02% | 0.0092028 | 0.0066434 | 0.04779 | 0.049862
141 | 10_15_21_F_6 |  | 06/17/2023 02:24:45 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_6.790 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9629.9114 | 7669.1188 | 11.2546 | 2.3267 | 24141.9699 | 2245.8268 | 41752.8675 | 0.5675 | 0.080044 | 0.086269 | 18.7791 | 8.2352 | 18.7771 | 0.015592 | 0.28296 | 2.0075 | 2.1282 | 4.9782 | 2.8731 | 2.6881 | 2.8923 | 0.69087 | 0.13604 | 0.099793 | 0.099692 | 0.50131 | 0.030706 | 11.4307 | 11.707 | 71.79% | 58.9% | 63.23% | 72.67% | 59.7% | 64.51% | 0.015011 | 0.012246 | 0.091037 | 0.079212
142 | 10_15_21_F_6 |  | 06/17/2023 02:31:03 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_6.791 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9430.877 | 7611.3164 | 11.0525 | 2.3269 | 24142.7901 | 2294.784 | 41845.4663 | 0.54157 | 0.08211 | 0.089646 | 18.656 | 8.2356 | 18.2018 | 0.020376 | 0.31111 | 2.0248 | 2.083 | 5.0895 | 2.7877 | 2.5437 | 2.964 | 0.53509 | 0.1666 | 0.077532 | 0.068499 | 0.46481 | 0.096984 | 11.4216 | 11.8225 | 70.77% | 57.71% | 62.34% | 64.11% | 58.35% | 63.68% | 0.015578 | 0.014996 | 0.092431 | 0.09269
143 | 10_15_21_F_6_LD |  | 06/17/2023 02:37:02 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_6_LD.792 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9320.0319 | 7348.1645 | 11.938 | 4.6263 | 23009.2172 | 2170.4842 | 40625.0249 | 0.52264 | 0.08457 | 0.094454 | 15.5315 | 8.2293 | 15.1796 | 0.023452 | 0.30503 | 1.994 | 2.0703 | 4.9787 | 2.8201 | 2.5914 | 2.913 | 0.6383 | 0.089652 | 0.097468 | 0.064297 | 0.4884 | 0.084981 | 11.4579 | 11.7377 | 70.8% | 57.71% | 62.43% | 65.85% | 58.21% | 64.51% | 0.016067 | 0.014251 | 0.089412 | 0.090078
144 | 10_15_21_F_6_LD |  | 06/17/2023 02:43:20 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_6_LD.793 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9444.1291 | 7475.209 | 11.822 | 5.3002 | 23798.8203 | 2237.406 | 41932.4511 | 0.5491 | 0.083496 | 0.089356 | 15.8424 | 8.2016 | 15.5376 | 0.026349 | 0.30831 | 2.0273 | 2.1363 | 5.0521 | 2.6901 | 2.6837 | 2.9146 | 0.43153 | 0.089894 | 0.086774 | 0.090903 | 0.40693 | 0.073686 | 11.7513 | 11.7456 | 73.42% | 57.45% | 60.98% | 63.53% | 57.64% | 63.45% | 0.017802 | 0.013349 | 0.085059 | 0.091604
145 | LBLANK |  | 06/17/2023 02:49:20 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.794 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 54.1513 | 0.23633 | -0.98894 | -0.12799 | 465.4333 | 59.8055 | 41.4958 | -0.00014348 | -0.0075797 | -0.00058329 | -0.5883 | -0.00033068 | -0.23142 | -0.0072372 | 0.0046778 | -0.0084419 | -0.010987 | -0.22901 | 0.055995 | -0.0032668 | -0.0029648 | -0.040557 | 0.0055366 | 0.0056923 | -0.010118 | 0.0074721 | -0.015962 | -0.0057121 | 0.0044433 | 83.04% | 62.37% | 68.15% | 69.16% | 61.39% | 67.07% | 0.00050427 | -0.0025665 | -0.010739 | -0.0043062
146 | CCV 1ppb TRACE |  | 06/17/2023 02:55:21 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.795 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1005.7882 | 1048.4076 | 1.304 | 0.3476 | 1727.1377 | 1054.9772 | 1057.608 | 1.086 | 1.0813 | 1.1036 | 0.97469 | 1.0112 | 1.0803 | 0.99081 | 1.0656 | 1.041 | 1.0678 | 0.83716 | 0.86468 | 1.013 | 0.98192 | 1.2262 | 0.89459 | 0.96245 | 0.92098 | 0.96661 | 0.94546 | 0.98616 | 1.0248 | 73.82% | 61.85% | 67.65% | 70.31% | 60.25% | 67.06% | 0.99923 | 0.98593 | 1.1311 | 1.0721
147 | CCV 1ppb TRACE |  | 06/17/2023 03:01:39 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.796 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1039.1883 | 1046.4968 | 1.1749 | 0.1367 | 1691.0827 | 1014.6489 | 1058.595 | 1.0321 | 1.0505 | 1.1139 | 0.95184 | 1.0209 | 1.2062 | 1.0493 | 1.0383 | 1.0472 | 1.0801 | 0.82147 | 0.92262 | 0.97322 | 0.99683 | 1.0198 | 0.83139 | 1.0285 | 0.99269 | 0.82037 | 0.95962 | 0.99313 | 1.0076 | 77.68% | 60.61% | 66.43% | 68.2% | 60.4% | 66.76% | 0.95779 | 0.96766 | 1.0514 | 1.0706
148 | LFB |  | 06/17/2023 03:07:39 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.797 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9123.3373 | 9345.9849 | 9.3734 | 11.6738 | 10722.5977 | 9204.8392 | 9932.1496 | 9.6151 | 10.052 | 10.2338 | 9.729 | 9.9726 | 9.1071 | 9.643 | 9.3195 | 9.6476 | 9.8048 | 9.8163 | 9.5112 | 9.7902 | 9.8163 | 10.2543 | 10.7167 | 10.0953 | 10.0937 | 10.3666 | 10.3098 | 9.5825 | 9.8779 | 71.49% | 55.88% | 62.55% | 61.51% | 56.6% | 63.05% | 9.5213 | 9.5604 | 9.9174 | 10.4426
149 | LFB |  | 06/17/2023 03:13:57 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.798 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 8972.1227 | 9656.0419 | 9.6568 | 11.9956 | 10536.2556 | 9338.0058 | 9596.5639 | 9.7802 | 10.0446 | 10.2039 | 9.831 | 9.9603 | 9.1934 | 9.7123 | 9.1212 | 9.6313 | 9.8862 | 9.8124 | 9.3311 | 9.4447 | 9.7939 | 10.1977 | 9.7979 | 10.1906 | 10.1112 | 10.0706 | 10.0154 | 9.5099 | 9.9398 | 72.51% | 56.35% | 62.72% | 63.99% | 56.37% | 63.42% | 9.5094 | 9.6022 | 9.6689 | 10.0364
150 | CCV 10 ppb MAJORS |  | 06/17/2023 03:19:56 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.799 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9501.9728 | 10048.5273 | 9.3208 | 9.5016 | 10864.3158 | 9851.0134 | 10385.0432 | 10.3208 | 10.5674 | 10.6308 | 10.5023 | 10.3868 | 9.8696 | 10.1599 | 9.681 | 10.0892 | 10.2151 | 10.3869 | 9.4912 | 10.348 | 9.9173 | 10.5936 | 9.8965 | 10.0221 | 10.1009 | 10.2473 | 10.755 | 9.902 | 10.3375 | 73.59% | 56.86% | 63.74% | 65.15% | 57.57% | 63.49% | 9.7112 | 9.7188 | 10.031 | 10.3503
151 | CCV 10 ppb MAJORS |  | 06/17/2023 03:26:14 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.800 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9456.6663 | 9965.1394 | 9.5303 | 9.2573 | 10804.9655 | 9722.0743 | 10241.1771 | 10.3568 | 10.5133 | 10.7993 | 10.4783 | 10.5147 | 9.6045 | 10.1885 | 9.6873 | 9.9861 | 10.2508 | 10.472 | 9.3425 | 9.7767 | 9.9689 | 10.1017 | 9.7904 | 9.905 | 10.1544 | 9.6905 | 10.3999 | 10.0468 | 10.3589 | 75.33% | 57.22% | 63.82% | 65.71% | 58.15% | 65.02% | 9.7324 | 9.8083 | 9.7941 | 10.3318
152 | 10_15_21_F_6_DUP |  | 06/17/2023 03:32:14 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_6_DUP.801 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 8811.1153 | 6992.3724 | 6.8005 | 2.3788 | 22820.0661 | 2141.8205 | 39844.1552 | 0.50363 | 0.081727 | 0.093434 | 10.2046 | 7.7409 | 10.3714 | 0.024951 | 0.31427 | 1.9886 | 2.0722 | 4.8828 | 2.766 | 2.6448 | 2.8368 | 0.59003 | 0.10713 | 0.066391 | 0.072447 | 0.46184 | 0.052217 | 11.2417 | 11.4655 | 69.09% | 55.72% | 61.47% | 63.14% | 55.86% | 62.48% | 0.019252 | 0.014072 | 0.10432 | 0.10827
153 | 10_15_21_F_6_DUP |  | 06/17/2023 03:38:32 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_6_DUP.802 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 8910.3035 | 7051.8656 | 7.1826 | 2.3587 | 22781.2926 | 2139.8691 | 40020.173 | 0.50234 | 0.093405 | 0.097135 | 9.9386 | 7.8319 | 10.4932 | 0.026948 | 0.29225 | 1.9905 | 2.0379 | 4.8686 | 3.0144 | 2.65 | 2.864 | 0.68447 | 0.057486 | 0.079187 | 0.094594 | 0.53331 | 0.065473 | 10.9494 | 11.57 | 64.34% | 54.86% | 61.69% | 60.62% | 56.07% | 61.74% | 0.018388 | 0.015016 | 0.11671 | 0.11496
154 | 10_15_21_F_6_DUP_LSP |  | 06/17/2023 03:44:32 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_6_DUP_LSP.803 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 17563.3886 | 16273.8317 | 16.946 | 14.3103 | 31736.4955 | 11646.9161 | 46706.9872 | 10.6931 | 10.4294 | 10.5597 | 20.3091 | 17.5141 | 19.4848 | 9.9816 | 9.5785 | 11.2966 | 11.593 | 14.3953 | 12.1378 | 11.9845 | 12.4204 | 10.5632 | 10.0777 | 10.0092 | 10.0159 | 10.5568 | 9.4897 | 20.1646 | 20.9537 | 69.23% | 53.83% | 59.46% | 60.99% | 55% | 60.05% | 9.3567 | 9.4372 | 9.8518 | 10.1426
155 | 10_15_21_F_6_DUP_LSP |  | 06/17/2023 03:50:52 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_6_DUP_LSP.804 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 18749.7413 | 17093.0077 | 18.5224 | 15.1751 | 33563.1567 | 12318.1988 | 49989.2252 | 11.3525 | 11.2101 | 10.702 | 20.5879 | 17.8444 | 20.7693 | 10.6252 | 10.1331 | 11.9968 | 11.7464 | 14.5267 | 12.069 | 12.398 | 12.3779 | 10.5394 | 10.3708 | 10.065 | 9.8342 | 10.5624 | 10.2297 | 20.3677 | 20.9185 | 69.95% | 54.34% | 59.62% | 61.31% | 55.48% | 58.36% | 9.5494 | 9.5547 | 9.7951 | 10.1567
156 | 10_15_21_F_9 |  | 06/17/2023 03:56:52 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_9.805 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 8102.7062 | 6954.4507 | 4.8346 | 2.9816 | 18217.2736 | 1955.6885 | 36413.1928 | 0.57781 | 0.10968 | 0.12751 | 7.9751 | 4.7308 | 8.3261 | 0.012742 | 0.23535 | 2.363 | 2.4114 | 4.1252 | 2.6276 | 2.6289 | 2.7552 | 0.42187 | 0.070854 | 0.082572 | 0.065844 | 0.38652 | 0.052124 | 7.3815 | 7.5648 | 69.92% | 55.76% | 60.72% | 62.6% | 56.63% | 62.25% | 0.013028 | 0.0081197 | 0.086792 | 0.090247
157 | 10_15_21_F_9 |  | 06/17/2023 04:03:12 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_9.806 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 8150.1238 | 6942.8337 | 4.8298 | 2.4293 | 18057.9904 | 1921.0344 | 36289.5407 | 0.55985 | 0.10642 | 0.12468 | 7.9007 | 4.7679 | 8.0582 | 0.0095882 | 0.22446 | 2.3313 | 2.4272 | 4.0293 | 2.6614 | 2.5905 | 2.7507 | 0.62515 | 0.07126 | 0.071393 | 0.047871 | 0.43153 | 0.077226 | 7.333 | 7.547 | 69.62% | 55.42% | 60.65% | 62.32% | 56.18% | 61.59% | 0.015327 | 0.0085878 | 0.088836 | 0.092073
158 | 10_15_21_F_10 |  | 06/17/2023 04:09:13 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_10.807 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 6299.0039 | 7642.8526 | 5.376 | 3.9528 | 14518.378 | 1736.9165 | 34055.4244 | 0.63737 | 0.13138 | 0.14164 | 7.1129 | 5.3923 | 7.6122 | 0.011977 | 0.33376 | 1.3945 | 1.4571 | 3.226 | 2.9094 | 2.7114 | 3.1781 | 0.53466 | 0.14439 | 0.080377 | 0.066856 | 0.36566 | 0.017569 | 3.2005 | 3.5127 | 69.08% | 52.96% | 60.85% | 63.27% | 55.44% | 61.37% | 0.0084272 | 0.0047624 | 0.23764 | 0.22821
159 | 10_15_21_F_10 |  | 06/17/2023 04:15:32 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_10.808 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 6195.3918 | 7319.3991 | 5.6836 | 3.7716 | 14253.6366 | 1690.3526 | 33381.6621 | 0.60657 | 0.12683 | 0.16107 | 8.5265 | 5.0844 | 7.3288 | 0.0087652 | 0.32234 | 1.3962 | 1.7344 | 3.8172 | 2.9324 | 2.6849 | 2.9969 | 0.4144 | 0.10164 | 0.077081 | 0.074064 | 0.3276 | 0.072154 | 3.777 | 3.3264 | 68.79% | 54.69% | 54.14% | 62.08% | 55.7% | 61.54% | 0.0078382 | 0.0058589 | 0.23614 | 0.23336
160 | 10_15_21_F_10_DUP |  | 06/17/2023 04:21:32 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_10_DUP.809 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 6217.952 | 7506.678 | 6.1406 | 3.5061 | 14037.0184 | 1770.1345 | 34553.1907 | 0.64102 | 0.099026 | 0.095613 | 6.8738 | 4.9928 | 7.4375 | 0.012175 | 0.15514 | 1.3159 | 1.3214 | 2.582 | 2.4829 | 2.6549 | 2.6722 | 0.65078 | 0.050031 | 0.058871 | 0.044605 | 0.30788 | 0.029009 | 3.0638 | 3.2789 | 70.18% | 54.25% | 61.67% | 63.43% | 55.29% | 60.18% | 0.0085942 | 0.0035429 | 0.38288 | 0.35871
161 | 10_15_21_F_10_DUP |  | 06/17/2023 04:27:51 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_10_DUP.810 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 6073.8187 | 7368.8303 | 5.6486 | 3.4927 | 13772.809 | 1723.6082 | 33934.5645 | 0.61669 | 0.10096 | 0.097561 | 7.0339 | 4.9706 | 7.3089 | 0.0073058 | 0.16144 | 1.2764 | 1.3309 | 2.6238 | 2.421 | 2.3436 | 2.6484 | 0.3797 | 0.050117 | 0.058691 | 0.045758 | 0.38427 | 0.041455 | 3.1715 | 3.2682 | 71.4% | 54.25% | 60.34% | 62.5% | 54.7% | 60.65% | 0.0086233 | 0.0038039 | 0.36084 | 0.36895
162 | LBLANK |  | 06/17/2023 04:33:51 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.811 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 55.1512 | 0.82977 | -0.99851 | -0.9124 | 474.7171 | 66.3204 | 43.2566 | -0.00091385 | -0.0086804 | -0.00088664 | -0.58225 | -0.00014399 | -0.37868 | -0.0077199 | 0.011284 | -0.010375 | -0.0073362 | -0.23426 | 0.0091799 | -0.01135 | -0.0037193 | -0.079842 | 0.01249 | 0.0083662 | -0.007874 | 0.021953 | -0.021474 | -0.0057551 | 0.0029491 | 73.86% | 57.76% | 65.06% | 66.13% | 56.79% | 64.46% | 0.00063013 | -0.0023352 | -0.010447 | -0.0041987
163 | LBLANK |  | 06/17/2023 04:40:11 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.812 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 54.1682 | 0.84263 | -1.0169 | -0.93954 | 495.8656 | 57.3309 | 39.974 | 0.00033378 | -0.0077697 | -0.00062604 | -0.58195 | -0.0011218 | -0.40732 | -0.0077098 | 0.0020553 | -0.010273 | -0.011342 | -0.22944 | 0.032275 | -0.0071572 | -0.0024656 | -0.091799 | 0.0055366 | -0.00050913 | -0.01206 | -0.010603 | -0.021474 | -0.0068265 | 0.0022609 | 74.08% | 57.24% | 64.76% | 65.52% | 56.33% | 63.18% | 0.00050427 | -0.0025427 | -0.010484 | -0.0040756
164 | CCV 1ppb TRACE |  | 06/17/2023 04:46:11 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.813 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 983.7392 | 1072.1924 | 1.1539 | 0.055486 | 1537.7798 | 1029.8857 | 1079.6466 | 1.0459 | 1.0545 | 1.07 | 0.89737 | 1.0298 | 0.96944 | 0.97811 | 1.0569 | 1.0311 | 1.0457 | 0.82232 | 0.9227 | 0.99993 | 0.96635 | 1.3222 | 0.92116 | 0.996 | 0.98525 | 1.0635 | 0.96342 | 0.98122 | 1.0053 | 70.53% | 56.91% | 64.34% | 65.48% | 55.44% | 62.27% | 0.97088 | 0.96993 | 1.1709 | 1.1087
165 | CCV 1ppb TRACE |  | 06/17/2023 04:52:31 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.814 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 984.0166 | 996.0293 | 1.0865 | 0.068408 | 1722.9088 | 1051.2473 | 1043.3635 | 1.0329 | 1.0552 | 1.0901 | 0.96717 | 1.0441 | 1.0858 | 1.0268 | 0.98463 | 1.0304 | 1.0665 | 0.83325 | 0.85805 | 0.89092 | 0.98626 | 1.5568 | 0.94617 | 1.0005 | 0.99509 | 0.87096 | 1.0018 | 0.98554 | 1.0256 | 68.59% | 56.76% | 64.06% | 65.48% | 56.04% | 62.69% | 0.97611 | 0.98955 | 1.2189 | 1.1084
166 | CCV 10 ppb MAJORS |  | 06/17/2023 04:58:31 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.815 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9082.5039 | 9580.7748 | 8.9554 | 9.0381 | 10710.8149 | 9556.8428 | 10161.4654 | 9.8986 | 10.3285 | 10.4954 | 10.2133 | 10.198 | 8.9676 | 9.7084 | 9.3707 | 9.7398 | 10.0391 | 10.1715 | 9.4579 | 9.7517 | 9.8535 | 10.3911 | 9.875 | 9.7636 | 9.9566 | 10.2759 | 9.9948 | 9.8994 | 10.2608 | 68.97% | 52.07% | 59.53% | 60.56% | 52.49% | 59.64% | 9.734 | 9.7572 | 10.5532 | 10.797
167 | CCV 10 ppb MAJORS |  | 06/17/2023 05:04:50 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.816 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9381.6047 | 9735.7882 | 9.5173 | 8.783 | 10954.4274 | 9589.6856 | 10173.9004 | 10.0587 | 10.435 | 10.4507 | 10.1238 | 10.2891 | 9.4533 | 9.9703 | 9.6684 | 9.7898 | 10.0039 | 10.2206 | 9.5555 | 10.1011 | 9.9143 | 10.3912 | 10.362 | 9.79 | 9.9271 | 10.0439 | 10.2064 | 9.9061 | 10.2198 | 70.01% | 53.02% | 60.58% | 61.85% | 53.53% | 59.43% | 9.702 | 9.7842 | 10.4492 | 10.6178
168 | 10_15_21_F_11 |  | 06/17/2023 05:10:50 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_11.817 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 5592.9865 | 7242.6501 | 5.3809 | 2.8956 | 13472.9781 | 1637.2603 | 32214.5355 | 0.59902 | 0.045408 | 0.052248 | 6.4671 | 4.3891 | 6.6122 | 0.010852 | 0.16026 | 1.3174 | 1.3805 | 3.8596 | 2.9805 | 2.7283 | 3.0592 | 0.48999 | 0.10636 | 0.06849 | 0.075806 | 0.48096 | 0.083389 | 2.7031 | 2.8195 | 67.48% | 52.18% | 59.74% | 61.58% | 53.35% | 60.27% | 0.006384 | 0.0037648 | 0.074611 | 0.075565
169 | 10_15_21_F_11 |  | 06/17/2023 05:17:09 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_11.818 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 5709.2532 | 7276.4977 | 6.0489 | 2.9814 | 14036.4157 | 1622.1786 | 32411.0066 | 0.63879 | 0.047943 | 0.056146 | 6.4569 | 4.3605 | 6.5465 | 0.0062589 | 0.16589 | 1.3106 | 1.388 | 3.9062 | 2.9168 | 2.8841 | 3.0314 | 0.42432 | 0.060216 | 0.093855 | 0.066084 | 0.44742 | 0.044583 | 2.7347 | 2.8179 | 67.05% | 51.69% | 59.01% | 71.43% | 52.47% | 58.92% | 0.0081176 | 0.0047789 | 0.073105 | 0.059691
170 | 10_15_21_UF_1 |  | 06/17/2023 05:23:09 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_1.819 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10677.3627 | 6508.6069 | -0.091374 | 4.4228 | 30497.0149 | 2696.2837 | 43622.0162 | 0.17706 | 0.14795 | 0.13663 | 0.90798 | 18.9121 | 1.8741 | 0.020279 | 0.32587 | 0.51999 | 0.59483 | 1.7301 | 2.5021 | 2.4878 | 2.5515 | 0.66694 | 0.036776 | 0.046769 | 0.043952 | 0.54831 | 0.0049962 | 20.2131 | 20.74 | 67.15% | 51.76% | 58.55% | 60.23% | 52.18% | 58.42% | 0.0094326 | 0.0058988 | 0.027038 | 0.032148
171 | 10_15_21_UF_1 |  | 06/17/2023 05:29:28 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_1.820 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10912.8776 | 6553.5644 | 0.036679 | 4.675 | 31010.5811 | 2777.1165 | 44445.534 | 0.1867 | 0.15768 | 0.13581 | 0.88232 | 19.2145 | 2.0467 | 0.022778 | 0.3689 | 0.52478 | 0.60167 | 1.7154 | 2.416 | 2.3614 | 2.5725 | 0.66903 | 0.043898 | 0.050948 | 0.037463 | 0.48918 | 0.030998 | 20.2482 | 20.9722 | 69.17% | 52.44% | 58.93% | 60.53% | 53.04% | 58.71% | 0.0091044 | 0.0062453 | 0.026177 | 0.030681
172 | 10_15_21_UF_6 |  | 06/17/2023 05:35:27 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_6.821 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9778.1275 | 6907.1205 | 0.0462 | 2.8905 | 22982.8535 | 2177.2132 | 40607.0865 | 0.50769 | 0.069257 | 0.068044 | 1.7509 | 7.575 | 2.737 | 0.014942 | 0.2576 | 1.227 | 1.3191 | 3.8486 | 2.5823 | 2.4143 | 2.6986 | 0.63864 | 0.096684 | 0.081297 | 0.071793 | 0.44489 | 0.055334 | 10.8516 | 11.2565 | 67.7% | 53.32% | 59.25% | 61.12% | 53.65% | 59.68% | 0.013609 | 0.013152 | 0.044547 | 0.044508
173 | 10_15_21_UF_6 |  | 06/17/2023 05:41:46 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_6.822 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9509.7211 | 6878.5393 | 0.13504 | 2.8128 | 22388.3122 | 2150.7365 | 39413.0655 | 0.47552 | 0.071657 | 0.065158 | 1.7019 | 7.5941 | 2.5708 | 0.013491 | 0.23676 | 1.264 | 1.3155 | 3.8334 | 2.6357 | 2.7498 | 2.6374 | 0.62927 | 0.074189 | 0.099048 | 0.072203 | 0.51086 | 0.036481 | 10.7183 | 11.203 | 66.32% | 53% | 59.97% | 60.89% | 53.89% | 60.11% | 0.018228 | 0.012559 | 0.044285 | 0.046527
174 | 10_15_21_UF_8 |  | 06/17/2023 05:47:46 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_8.823 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10167.1898 | 7133.965 | 1.4165 | 2.6727 | 22209.4721 | 2276.7367 | 40130.2502 | 0.65756 | 0.10186 | 0.079606 | 0.76854 | 7.8594 | 1.8432 | 0.014035 | 0.256 | 1.1536 | 1.2398 | 3.4563 | 3.272 | 2.9149 | 3.2465 | 0.59386 | 0.12179 | 0.080902 | 0.06668 | 0.49962 | 0.070402 | 9.9962 | 10.4041 | 66.64% | 52.13% | 58.84% | 60.16% | 53.04% | 58.63% | 0.0099515 | 0.0074862 | 0.03899 | 0.04215
175 | 10_15_21_UF_8 |  | 06/17/2023 05:54:04 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_8.824 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10308.5276 | 7292.6146 | 1.4987 | 2.7549 | 22366.0507 | 2222.1842 | 40438.484 | 0.69611 | 0.10715 | 0.08496 | 0.74773 | 7.8134 | 1.8393 | 0.014904 | 0.24387 | 1.1724 | 1.2325 | 3.4805 | 3.1527 | 2.9835 | 3.2151 | 0.70476 | 0.091005 | 0.075842 | 0.084091 | 0.35646 | 0.044069 | 10.0139 | 10.3562 | 66.7% | 52.1% | 58.41% | 60.65% | 52.45% | 57.81% | 0.0093285 | 0.0067423 | 0.041126 | 0.042234
176 | 10_15_21_UF_8_LD |  | 06/17/2023 06:00:03 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_8_LD.825 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10279.8382 | 7278.2515 | 3.003 | 5.1627 | 22462.4541 | 2243.611 | 40395.2592 | 0.70522 | 0.1198 | 0.097023 | 1.1432 | 7.3959 | 2.2455 | 0.018584 | 0.26318 | 1.2075 | 1.3062 | 7.5269 | 3.1388 | 2.8911 | 3.1269 | 0.61915 | 0.043555 | 0.072032 | 0.060847 | 0.56655 | 0.06132 | 9.9496 | 9.8832 | 67.22% | 53.95% | 58.4% | 60% | 52.6% | 58.05% | 0.010438 | 0.007625 | 0.047093 | 0.04997
177 | 10_15_21_UF_8_LD |  | 06/17/2023 06:06:21 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_8_LD.826 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10118.7739 | 7023.569 | 3.1024 | 5.6801 | 21919.3357 | 2272.9515 | 39953.4492 | 0.69466 | 0.10874 | 0.099702 | 1.162 | 7.7619 | 2.0974 | 0.020192 | 0.26859 | 1.2615 | 1.3177 | 7.5997 | 3.0957 | 2.9964 | 3.2702 | 0.55472 | 0.1301 | 0.10105 | 0.076206 | 0.59828 | 0.070642 | 9.9627 | 10.3167 | 66.7% | 51.95% | 58.12% | 59.96% | 51.98% | 58.7% | 0.010205 | 0.0070167 | 0.044442 | 0.04924
178 | LBLANK |  | 06/17/2023 06:12:21 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.827 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 54.6956 | 1.442 | -0.99661 | -0.49387 | 622.9496 | 55.1933 | 41.6973 | 0.00031926 | -0.0091733 | -0.0024377 | -0.58138 | -0.00087292 | -0.30531 | -0.0078885 | 0.0037548 | -0.009191 | -0.011487 | -0.22872 | 0.022123 | 0.0062932 | 0.010213 | -0.063336 | 0.0055366 | 0.0017967 | -0.0056003 | 0.014232 | -0.015347 | -0.0066375 | 0.0044315 | 74.46% | 55.99% | 63.74% | 64.78% | 55.15% | 61.9% | 0.0008888 | -0.0025361 | -0.010858 | -0.0043132
179 | LBLANK |  | 06/17/2023 06:18:40 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.828 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 55.203 | 0.83572 | -0.90719 | -0.87994 | 515.7658 | 57.3635 | 40.8248 | 0.00031991 | -0.0091201 | -0.001297 | -0.57984 | -0.00024085 | -0.31988 | -0.0071727 | -0.00096481 | -0.013315 | -0.01122 | -0.22905 | -0.046507 | -0.0068514 | 0.0038727 | 0.1286 | 0.0055366 | 0.0018063 | -0.0033927 | -0.011958 | -0.021474 | -0.00815 | 0.0030957 | 70.72% | 55.68% | 62.84% | 64.97% | 54.29% | 61.61% | 0.00063363 | -0.0025293 | -0.010509 | -0.0042858
180 | CCV 1ppb TRACE |  | 06/17/2023 06:24:40 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.829 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 968.0848 | 959.1468 | 1.2952 | 0.059317 | 1514.0892 | 1011.1718 | 996.4636 | 1.0541 | 1.0725 | 1.0959 | 0.95095 | 1.0273 | 0.94525 | 0.9819 | 0.97449 | 0.98646 | 1.0628 | 0.84725 | 0.88584 | 1.0034 | 1.0054 | 1.1615 | 1.1205 | 1.0212 | 0.92404 | 0.90827 | 1.1525 | 1.0126 | 0.99527 | 70.87% | 54.75% | 61.81% | 63.28% | 53.75% | 61.49% | 0.98473 | 0.98321 | 1.1529 | 1.1166
181 | CCV 1ppb TRACE |  | 06/17/2023 06:31:00 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.830 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 989.569 | 1004.8621 | 1.1962 | 0.085023 | 1760.5622 | 1005.7186 | 1055.8563 | 1.0353 | 1.0936 | 1.0879 | 0.92351 | 1.0383 | 1.0326 | 1.0053 | 0.9927 | 1.0485 | 1.0576 | 0.83569 | 0.83491 | 1.0503 | 0.9721 | 1.1549 | 0.86982 | 1.0134 | 0.98219 | 0.97531 | 1.0308 | 0.99013 | 1.0449 | 72.14% | 54.29% | 61.78% | 63.44% | 53.57% | 61.09% | 0.96264 | 0.97011 | 1.1216 | 1.1282
182 | LFB |  | 06/17/2023 06:37:00 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.831 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 8970.3074 | 9183.3327 | 9.2986 | 11.5407 | 10881.8268 | 9176.0433 | 9770.5853 | 9.6765 | 10.1655 | 10.2248 | 9.854 | 9.8512 | 8.8593 | 9.5357 | 9.1662 | 9.5941 | 9.9232 | 10.0133 | 9.4709 | 9.8337 | 9.8297 | 9.9138 | 9.7686 | 9.9008 | 9.8871 | 10.1287 | 10.0335 | 9.6339 | 9.9816 | 67.11% | 50.06% | 57.27% | 59.5% | 51.03% | 56.9% | 9.6126 | 9.6824 | 10.1483 | 10.3675
183 | LFB |  | 06/17/2023 06:43:18 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.832 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 8867.5677 | 9124.608 | 9.0696 | 11.6353 | 10358.6309 | 9136.001 | 9792.1015 | 9.651 | 9.9827 | 9.9947 | 9.5324 | 9.8232 | 8.8429 | 9.5589 | 9.2102 | 9.5548 | 9.6558 | 9.4595 | 9.0273 | 9.728 | 9.669 | 9.5589 | 10.1294 | 9.8792 | 10.122 | 9.8282 | 9.9723 | 9.3787 | 9.762 | 69.45% | 50.89% | 59.35% | 56.14% | 51.01% | 57.72% | 9.3795 | 9.4313 | 9.8528 | 11.2168
184 | CCV 10 ppb MAJORS |  | 06/17/2023 06:49:17 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.833 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9306.831 | 9713.4271 | 9.1386 | 9.2248 | 10647.9564 | 9569.9926 | 10321.301 | 10.2197 | 10.4124 | 10.5272 | 10.534 | 10.2872 | 9.3711 | 9.8855 | 9.4574 | 9.9306 | 10.1239 | 10.3161 | 9.2763 | 9.8277 | 9.9325 | 10.1599 | 9.835 | 10.2077 | 9.7385 | 10.0393 | 10.3327 | 9.9262 | 10.3355 | 69.18% | 51.46% | 59.22% | 61.5% | 52.07% | 58.33% | 9.7562 | 9.8718 | 10.3118 | 10.5736
185 | CCV 10 ppb MAJORS |  | 06/17/2023 06:55:37 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.834 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9321.9399 | 9621.4959 | 9.6796 | 9.0552 | 10832.7867 | 9651.5489 | 10077.7649 | 10.1472 | 10.5289 | 10.5721 | 10.2148 | 10.5022 | 9.5828 | 9.9356 | 9.5294 | 9.8754 | 10.1215 | 10.1735 | 9.5486 | 10.0409 | 9.8948 | 10.1073 | 10.3434 | 10.1782 | 9.9266 | 10.2306 | 10.2963 | 9.9969 | 10.3937 | 69.38% | 51.94% | 60.3% | 61.29% | 52.48% | 59.72% | 9.6756 | 9.772 | 10.4837 | 10.7976
186 | 10_15_21_UF_8_DUP |  | 06/17/2023 07:01:37 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_8_DUP.835 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10384.274 | 7292.7311 | 0.011584 | 2.6336 | 22873.8333 | 2289.1992 | 41227.5619 | 0.69471 | 0.12987 | 0.10074 | 1.5959 | 7.6724 | 2.7848 | 0.013787 | 0.25072 | 1.2456 | 1.2584 | 2.6476 | 3.1381 | 2.9242 | 3.2768 | 0.61622 | 0.078082 | 0.13022 | 0.089055 | 0.50312 | 0.067117 | 10.07 | 10.3259 | 65.07% | 50.15% | 56.87% | 58.89% | 50.96% | 55.35% | 0.0097243 | 0.0070001 | 0.041976 | 0.046897
187 | 10_15_21_UF_8_DUP |  | 06/17/2023 07:07:56 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_8_DUP.836 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9850.2514 | 6887.3176 | -0.048696 | 2.8155 | 22047.339 | 2217.6119 | 39826.8979 | 0.68307 | 0.11885 | 0.10378 | 1.6106 | 7.6629 | 2.5062 | 0.011616 | 0.24294 | 1.1865 | 1.2277 | 2.6478 | 3.1327 | 2.9767 | 3.1707 | 0.55869 | 0.078707 | 0.079626 | 0.086299 | 0.39406 | 0.067957 | 9.8998 | 10.2526 | 64.78% | 49.73% | 56.66% | 58.3% | 50.21% | 56.33% | 0.012761 | 0.007404 | 0.043746 | 0.046002
188 | 10_15_21_UF_8_DUP_LSP |  | 06/17/2023 07:13:55 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_8_DUP_LSP.837 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 18258.2928 | 15957.0251 | 11.0449 | 15.7559 | 30341.579 | 11662.155 | 46464.476 | 10.7604 | 10.5319 | 10.3998 | 12.0851 | 16.998 | 11.9649 | 9.7962 | 9.4143 | 10.5352 | 10.681 | 11.525 | 12.0351 | 12.7318 | 12.5689 | 10.1763 | 10.1885 | 9.8168 | 9.7248 | 10.3647 | 9.7098 | 18.8764 | 19.5579 | 65.14% | 48.71% | 55.52% | 56.7% | 48.89% | 54.13% | 9.4169 | 9.424 | 10.2377 | 10.4079
189 | 10_15_21_UF_8_DUP_LSP |  | 06/17/2023 07:20:13 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_8_DUP_LSP.838 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 18290.0209 | 15456.0348 | 10.8726 | 15.6574 | 29759.8218 | 11437.4433 | 46012.4646 | 10.7682 | 10.3547 | 10.4974 | 11.9713 | 16.9751 | 11.7553 | 9.7419 | 9.3663 | 10.2005 | 10.663 | 11.5272 | 12.353 | 12.1812 | 12.4703 | 10.5643 | 10.1203 | 9.74 | 10.0045 | 10.7222 | 9.9553 | 18.9821 | 19.5368 | 64.76% | 49.64% | 56.32% | 57.22% | 50.47% | 56.17% | 9.4894 | 9.5252 | 10.3605 | 10.4457
190 | 10_15_21_UF_9 |  | 06/17/2023 07:26:13 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_9.839 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 8681.2319 | 6514.3072 | 12.7994 | 6.7919 | 17805.5469 | 1935.3941 | 35857.3928 | 0.55291 | 0.11357 | 0.11179 | 0.52013 | 4.3694 | 1.4831 | 0.0095453 | 0.19127 | 1.0147 | 1.0846 | 2.8383 | 2.5441 | 2.4415 | 2.596 | 0.69525 | 0.069559 | 0.080838 | 0.061884 | 0.42996 | 0.10037 | 7.0353 | 7.2486 | 65.56% | 50.56% | 57.25% | 58.82% | 51.24% | 57.38% | 0.011369 | 0.0087918 | 0.02821 | 0.028778
191 | 10_15_21_UF_9 |  | 06/17/2023 07:32:31 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_9.840 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 8598.2121 | 6573.2885 | 12.3818 | 6.0986 | 18024.4361 | 1882.2505 | 35627.2489 | 0.5333 | 0.11919 | 0.10858 | 0.48428 | 4.4017 | 1.4087 | 0.0092074 | 0.1751 | 0.99183 | 1.0864 | 2.8711 | 2.453 | 2.3753 | 2.632 | 0.50747 | 0.046004 | 0.08197 | 0.050878 | 0.36164 | 0.074505 | 7.0407 | 7.275 | 65.38% | 49.86% | 56.62% | 58.36% | 50.2% | 56.63% | 0.010457 | 0.0080746 | 0.026571 | 0.030637
192 | 10_15_21_UF_10 |  | 06/17/2023 07:38:31 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_10.841 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 6908.1782 | 7277.051 | -0.047034 | 4.3194 | 14283.6831 | 1647.0466 | 33688.9073 | 0.61551 | 0.10345 | 0.086187 | 0.28748 | 4.5774 | 1.2897 | 0.0072488 | 0.17595 | 0.80388 | 0.8599 | 2.1185 | 2.7805 | 2.8519 | 2.8633 | 0.54562 | 0.14189 | 0.076286 | 0.070962 | 0.36588 | 0.094407 | 3.1106 | 3.1919 | 65.22% | 50.19% | 56.81% | 60.84% | 50.58% | 55.87% | 0.0078488 | 0.0029637 | 0.028569 | 0.033127
193 | 10_15_21_UF_10 |  | 06/17/2023 07:44:50 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_10.842 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 6816.0919 | 7000.3872 | -0.18804 | 4.0854 | 14019.0841 | 1695.7911 | 33118.622 | 0.59768 | 0.098621 | 0.087848 | 0.26931 | 4.5917 | 1.1492 | 0.0065144 | 0.13791 | 0.79565 | 0.85683 | 2.1546 | 2.7302 | 2.7956 | 2.8861 | 0.61787 | 0.070913 | 0.098112 | 0.059739 | 0.40034 | 0.026922 | 3.0991 | 3.2061 | 65.38% | 49.5% | 56.53% | 65.53% | 49.67% | 56.39% | 0.0075895 | 0.0031747 | 0.030267 | 0.027758
194 | 10_15_21_UF_11 |  | 06/17/2023 07:50:50 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_11.843 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 6563.6549 | 7219.8262 | -0.040966 | 3.3337 | 13955.4355 | 1598.3116 | 32909.7825 | 0.58637 | 0.073131 | 0.056315 | 1.4423 | 4.0029 | 2.1636 | 0.0069465 | 0.12095 | 0.78428 | 0.85014 | 3.0595 | 2.9087 | 2.7622 | 2.9895 | 0.34193 | 0.046214 | 0.071946 | 0.053482 | 0.38851 | 0.095383 | 2.6352 | 2.7668 | 64.29% | 49.72% | 56.69% | 60.77% | 50.14% | 56.3% | 0.0054095 | 0.0046742 | 0.032175 | 0.032871
195 | 10_15_21_UF_11 |  | 06/17/2023 07:57:09 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_11.844 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 6694.7831 | 7168.185 | 0.095456 | 3.5003 | 14096.4545 | 1566.5985 | 32446.2584 | 0.58537 | 0.077728 | 0.063078 | 1.4567 | 4.0649 | 2.1564 | 0.0075975 | 0.16379 | 0.80875 | 0.86873 | 3.0912 | 2.8091 | 2.7826 | 3.0339 | 0.52947 | 0.071978 | 0.099996 | 0.067411 | 0.38781 | 0.069767 | 2.6288 | 2.7605 | 64.63% | 48.61% | 55.97% | 57.11% | 48.85% | 55.36% | 0.0066366 | 0.0043425 | 0.03311 | 0.036126
196 | LBLANK |  | 06/17/2023 08:03:10 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.845 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 55.035 | 0.87745 | -1.013 | -0.87717 | 568.7107 | 58.9971 | 39.1036 | -0.00047409 | -0.0055501 | -0.0021796 | -0.57735 | -0.00023199 | -0.31439 | -0.0078885 | 0.00083711 | -0.0086675 | -0.011112 | -0.23141 | 0.014612 | -0.0066593 | 0.0085284 | 0.022527 | 0.0055366 | 0.011718 | -0.0074358 | -0.0091548 | -0.0084538 | -0.0067453 | 0.0027929 | 70.13% | 52.47% | 61.14% | 61.84% | 51.43% | 57.98% | 0.00050427 | -0.0024531 | -0.010606 | -0.0041835
197 | LBLANK |  | 06/17/2023 08:09:28 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.846 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 53.943 | 0.23633 | -0.99374 | -0.94322 | 454.9155 | 62.6287 | 42.7364 | -0.00048271 | -0.007886 | -0.0021642 | -0.58084 | -0.00022223 | -0.37327 | -0.0071464 | 0.0083564 | -0.012219 | -0.011096 | -0.23312 | -0.067417 | -0.0021395 | 0.0089333 | 0.23586 | 0.013317 | 0.0068941 | -0.0062653 | -0.01877 | -0.014927 | -0.0084082 | 0.0025485 | 63.11% | 52.16% | 61.39% | 61.27% | 50.77% | 58.98% | 0.0010413 | -0.0022897 | -0.010089 | -0.0042789
198 | CCV 1ppb TRACE |  | 06/17/2023 08:15:29 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.847 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1002.3381 | 996.7207 | 1.3114 | 0.063046 | 1835.2714 | 1027.1575 | 1055.0865 | 1.0733 | 1.0884 | 1.0695 | 0.95485 | 1.0151 | 1.0531 | 1.0411 | 1.0045 | 1.0406 | 1.0572 | 0.81105 | 0.8506 | 1.0771 | 0.99088 | 1.287 | 0.90792 | 0.9629 | 0.98127 | 0.93088 | 1.0329 | 0.99225 | 1.0192 | 66.38% | 51.53% | 58.78% | 60.64% | 50.55% | 57.11% | 1.0105 | 0.99141 | 1.2018 | 1.1506
199 | CCV 1ppb TRACE |  | 06/17/2023 08:21:48 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.848 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 997.0932 | 995.0847 | 1.4091 | 0.051982 | 1719.7104 | 1061.1018 | 978.5833 | 1.0568 | 1.0815 | 1.0769 | 0.93359 | 1.0394 | 1.0595 | 1.0099 | 0.94677 | 1.0184 | 1.0648 | 0.81691 | 0.90047 | 0.92465 | 0.98077 | 1.2687 | 1.067 | 1.0128 | 1.0302 | 0.95362 | 0.97098 | 0.99058 | 1.0264 | 65.35% | 50.63% | 58.33% | 59.72% | 50.47% | 57.66% | 0.99939 | 0.99145 | 1.1946 | 1.146
200 | CCV 10 ppb MAJORS |  | 06/17/2023 08:27:48 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.849 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9196.0967 | 9579.6142 | 9.6821 | 8.8967 | 11025.4809 | 9642.3472 | 10209.7829 | 9.8544 | 10.2666 | 10.5153 | 9.984 | 10.2443 | 9.1805 | 9.7764 | 9.3451 | 9.8159 | 9.9135 | 10.1591 | 9.4528 | 9.8617 | 9.9315 | 10.4755 | 10.0874 | 10.0847 | 10.1008 | 10.2158 | 10.2704 | 10.0021 | 10.3519 | 64.18% | 47.35% | 55.14% | 56.66% | 47.89% | 54.53% | 9.8043 | 9.8706 | 10.8231 | 11.1097
201 | CCV 10 ppb MAJORS |  | 06/17/2023 08:34:07 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.850 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9343.0282 | 9610.9488 | 9.238 | 8.8704 | 10996.6849 | 9457.8814 | 10283.2552 | 10.1219 | 10.5213 | 10.4927 | 10.1373 | 10.3406 | 9.3419 | 9.8617 | 9.4333 | 9.8444 | 10.0404 | 10.1828 | 8.909 | 10.0913 | 9.973 | 9.5723 | 10.678 | 9.9592 | 10.05 | 10.1431 | 10.2223 | 9.9624 | 10.4288 | 69.06% | 48.3% | 56.36% | 57.59% | 48.85% | 55.02% | 9.759 | 9.8639 | 10.1266 | 11.0322
202 | 10_15_21_FB |  | 06/17/2023 08:40:08 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_FB.851 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 309.5635 | 5.9134 | 8.7846 | 0.64889 | 660.5355 | 73.1704 | 50.8432 | 0.00037281 | 0.081431 | 0.073486 | 1.4713 | 0.058921 | 1.571 | -0.0056663 | 0.0097588 | 0.10545 | 0.11742 | 7.5315 | -0.052131 | -0.00678 | 0.0090817 | 0.13273 | 0.0055366 | 0.0051056 | -0.0070843 | -0.021988 | -0.021474 | 0.00096373 | 0.012359 | 67.97% | 46.53% | 59.49% | 61.32% | 51.47% | 59.04% | 0.0011888 | -0.00094353 | 0.050446 | 0.053852
203 | 10_15_21_FB |  | 06/17/2023 08:46:26 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_FB.852 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 300.7921 | 6.6407 | 8.8997 | 0.585 | 577.2082 | 80.9715 | 51.9309 | 0.0026006 | 0.079389 | 0.076795 | 1.478 | 0.050685 | 1.4772 | -0.0061818 | 0.014396 | 0.09774 | 0.11869 | 7.5547 | -0.035538 | -0.006632 | 0.0082709 | 0.16972 | 0.013568 | 0.0045383 | -0.0084892 | 0.0132 | -0.021474 | -0.00030517 | 0.010018 | 67.01% | 50.73% | 58.38% | 60.81% | 50.53% | 58.16% | 0.00078386 | -0.0014934 | 0.047406 | 0.052525
204 | 10_15_21_BB |  | 06/17/2023 08:52:26 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_BB.853 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 56.1217 | 1.5306 | -0.36638 | -0.72755 | 443.6407 | 59.1756 | 44.6656 | -8.8222e-06 | 0.01434 | 0.019423 | -0.35299 | 0.002982 | -0.23378 | -0.0061632 | 0.011964 | 0.075239 | 0.085513 | -0.06074 | 0.0029923 | 0.0072197 | 0.00084329 | 0.10906 | 0.013477 | 0.0045487 | -0.0073402 | 0.0022803 | -0.021474 | -0.007672 | 0.0025852 | 67.81% | 51.03% | 58.8% | 60.92% | 49.91% | 57.45% | 0.00064456 | -0.0020845 | 0.013018 | 0.019339
205 | 10_15_21_BB |  | 06/17/2023 08:58:45 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_BB.854 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 56.1369 | 0.23633 | -0.19821 | -0.72812 | 487.6436 | 56.437 | 46.5609 | -0.00091385 | 0.012905 | 0.020062 | -0.3422 | 0.0043811 | -0.16623 | -0.0076969 | 0.0053235 | 0.08384 | 0.082649 | -0.061945 | 0.014109 | -0.01135 | -0.0010676 | 0.00016846 | 0.0055366 | 0.0071314 | -0.0061376 | -0.022243 | -0.014663 | -0.0064993 | 0.0025805 | 66.93% | 50.83% | 59.3% | 60.3% | 49.92% | 57.49% | 0.00050427 | -0.0022672 | 0.012778 | 0.019384
206 | 10_15_21_LAB_A |  | 06/17/2023 09:04:45 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_LAB_A.855 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 56.9618 | 0.23633 | -0.39347 | -0.7624 | 312.4661 | 55.5559 | 41.9002 | -0.00091385 | 0.038638 | 0.043247 | -0.18434 | 0.0070936 | -0.07752 | -0.0076986 | 0.00026748 | 0.088188 | 0.08877 | -0.039874 | -0.018785 | -0.01135 | -0.0012161 | 0.023675 | 0.0055366 | -0.00050913 | -0.010861 | 0.032175 | -0.021474 | -0.0030283 | 0.0064621 | 66.76% | 50.48% | 61.5% | 60.16% | 49.67% | 57.39% | 0.00078049 | -0.0022502 | 0.015723 | 0.020626
207 | 10_15_21_LAB_A |  | 06/17/2023 09:11:03 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_LAB_A.856 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 56.3986 | 0.23633 | -0.38434 | -0.75607 | 619.7993 | 55.4549 | 42.808 | 0.00041334 | 0.040328 | 0.046871 | -0.16125 | 0.0061968 | -0.053997 | -0.0073192 | 0.0023391 | 0.097948 | 0.093321 | -0.025752 | -0.019415 | -0.0066474 | -0.0017623 | 0.19994 | 0.0055366 | 0.0020185 | -0.0084854 | -0.023321 | -0.021474 | -0.001479 | 0.0073253 | 66.42% | 50.65% | 58.92% | 59.96% | 49.93% | 57.77% | 0.00078087 | -0.0022045 | 0.017173 | 0.020528
208 | 10_15_21_LAB_B |  | 06/17/2023 09:17:03 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_LAB_B.857 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 55.8835 | 0.89899 | -0.32662 | -0.69175 | 463.7233 | 57.0996 | 41.9431 | -0.00091385 | 0.012211 | 0.016121 | -0.36697 | 0.0028359 | -0.1743 | -0.0073101 | 0.007829 | 0.098107 | 0.14728 | 0.098667 | -0.020084 | -0.0017266 | -0.0027726 | 0.11412 | 0.0055366 | 0.0046528 | -0.0084127 | -0.085032 | -0.014538 | -0.0089123 | 0.0021771 | 65.81% | 49.71% | 57.63% | 59.33% | 48.71% | 56.87% | 0.00064589 | -0.0025438 | 0.016229 | 0.020504
209 | 10_15_21_LAB_B |  | 06/17/2023 09:23:20 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_LAB_B.858 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 54.2331 | 0.23633 | -0.18095 | -0.6595 | 232.4913 | 61.6954 | 41.8829 | 0.00043462 | 0.0085678 | 0.015795 | -0.3636 | 0.0021673 | -0.23348 | -0.0071177 | 0.0019464 | 0.11115 | 0.10095 | 0.097623 | 0.029191 | -0.01135 | -0.0025741 | 0.074651 | 0.013686 | -0.00050913 | -0.01206 | -0.081376 | -0.014517 | -0.0086601 | 0.0020907 | 65.66% | 49.31% | 58.17% | 59.69% | 48.62% | 56.85% | 0.00078281 | -0.002016 | 0.017091 | 0.020987
210 | 5_18_23_DB |  | 06/17/2023 09:29:20 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_DB.859 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 54.7914 | 0.23633 | -0.5834 | 2.7988 | 400.8324 | 54.8712 | 42.8917 | -0.00091385 | -0.0069955 | -0.0014792 | -0.51844 | 0.00065848 | -0.3335 | -0.0071132 | -0.005392 | 0.07844 | 0.083187 | -0.14136 | -0.02964 | -0.01135 | -0.0020394 | -0.0062654 | 0.0055366 | -0.00050913 | -0.0096315 | -0.045747 | -0.021474 | -0.0094574 | 0.0014597 | 65.93% | 49.46% | 57.44% | 59.09% | 48.35% | 56.54% | 0.00078973 | -0.0027816 | 0.017861 | 0.022752
211 | 5_18_23_DB |  | 06/17/2023 09:35:39 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_DB.860 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 54.5141 | 0.23633 | -0.59238 | 3.9776 | 284.2138 | 60.6181 | 42.9465 | 0.00045495 | -0.0039211 | -0.0013658 | -0.50863 | 0.00031322 | -0.41929 | -0.0073018 | -0.0015406 | 0.088481 | 0.08197 | -0.14344 | 0.024899 | -0.0065295 | -0.0018156 | 0.047586 | 0.0055366 | 0.0020976 | -0.010837 | -0.053863 | -0.007759 | -0.0089791 | 0.0012842 | 65.07% | 49.16% | 57.68% | 59.72% | 48.69% | 56% | 0.00092955 | -0.0022477 | 0.018101 | 0.022976
212 | LBLANK |  | 06/17/2023 09:41:39 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.861 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 53.9622 | 0.23633 | -1.0181 | 0.27629 | 380.1367 | 56.2211 | 41.0345 | 4.3949e-06 | -0.0085363 | -0.0022115 | -0.58825 | -0.00039326 | -0.3623 | -0.0072896 | -0.0088998 | -0.010558 | -0.010593 | -0.23139 | -0.014607 | -0.0063945 | -0.0010943 | -0.06612 | 0.0055366 | 0.004779 | -0.0095937 | 0.0057133 | -0.014463 | -0.0078214 | 0.0022834 | 61.76% | 48.88% | 56.7% | 58.42% | 48.03% | 55.26% | 0.00079341 | -0.0025369 | -0.010532 | -0.0040134
213 | LBLANK |  | 06/17/2023 09:47:59 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.862 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 53.9519 | 0.23633 | -0.99086 | -0.76234 | 211.4654 | 63.8271 | 40.0575 | -0.00091385 | -0.0061699 | -0.00078087 | -0.58438 | -0.00083815 | -0.39507 | -0.0074963 | -0.0044315 | -0.013919 | -0.010237 | -0.2302 | 0.00056601 | -0.01135 | -0.0027413 | -0.032854 | 0.013787 | -0.00050913 | -0.01206 | -0.11708 | -0.021474 | -0.0086093 | 0.0023704 | 67.39% | 49.01% | 57.02% | 58.57% | 47.88% | 56.03% | 0.00064791 | -0.0022976 | -0.010158 | -0.0039429
214 | CCV 1ppb TRACE |  | 06/17/2023 09:53:59 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.863 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 941.3352 | 971.9387 | 0.8378 | 0.11132 | 1197.1077 | 1024.1048 | 1066.1094 | 1.0225 | 1.0422 | 1.0843 | 0.88773 | 1.0488 | 0.97888 | 0.98638 | 0.96001 | 1.0284 | 1.0613 | 0.83248 | 0.79129 | 0.88526 | 0.96246 | 1.2541 | 1.0028 | 1.0659 | 0.96182 | 0.90306 | 0.94047 | 0.97769 | 1.0342 | 63.74% | 48.27% | 55.93% | 57.54% | 47.68% | 54.99% | 0.9842 | 0.99871 | 1.2063 | 1.1878
215 | CCV 1ppb TRACE |  | 06/17/2023 10:00:19 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.864 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 970.6048 | 967.9932 | 0.98345 | 0.050434 | 1617.0556 | 1048.122 | 1118.9109 | 1.0582 | 1.0757 | 1.095 | 0.94119 | 1.0547 | 0.98358 | 0.99119 | 0.98077 | 0.99756 | 1.0699 | 0.81686 | 0.85719 | 0.99734 | 0.99774 | 1.1592 | 1.0502 | 0.93461 | 0.92341 | 0.94712 | 1.08 | 1.003 | 1.0383 | 64.21% | 48.39% | 55.3% | 57.93% | 47.52% | 55.42% | 1.0128 | 1.0111 | 1.1968 | 1.1765
216 | LFB |  | 06/17/2023 10:06:18 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.865 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 8604.3312 | 8848.0206 | 9.4861 | 10.8847 | 10437.6528 | 9135.7801 | 9755.5838 | 9.5967 | 9.8197 | 9.7091 | 9.2974 | 9.6264 | 8.5047 | 9.2071 | 8.9238 | 9.2904 | 9.4477 | 9.3339 | 9.072 | 9.2887 | 9.5888 | 9.8521 | 9.3578 | 9.4377 | 9.9006 | 9.5877 | 10.1963 | 9.3895 | 9.8947 | 61.95% | 44.92% | 53.11% | 54.51% | 45.72% | 52.89% | 9.4288 | 9.514 | 10.5896 | 10.8361
217 | LFB |  | 06/17/2023 10:12:36 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.866 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 8488.4552 | 8863.8757 | 9.0709 | 11.3655 | 10232.5317 | 8918.4416 | 9541.5482 | 9.341 | 9.5082 | 10.0825 | 9.539 | 9.7991 | 8.522 | 9.1402 | 8.9403 | 9.3112 | 9.6675 | 9.5192 | 9.166 | 9.9727 | 9.709 | 10.0455 | 9.4721 | 9.5825 | 9.8053 | 10.0227 | 10.4228 | 9.6155 | 10.0292 | 62.44% | 45.73% | 53.05% | 55.54% | 46.1% | 54.38% | 9.588 | 9.728 | 10.5883 | 10.7092
218 | CCV 10 ppb MAJORS |  | 06/17/2023 10:18:36 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.867 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9088.8882 | 9458.5212 | 9.7384 | 9.1108 | 11032.9994 | 9584.4373 | 10268.8251 | 10.0607 | 10.3768 | 10.457 | 10.1439 | 10.2826 | 9.2199 | 9.7878 | 9.2477 | 9.7816 | 10.0346 | 9.9822 | 9.4649 | 9.426 | 9.8989 | 10.2056 | 10.015 | 10.1258 | 9.9487 | 10.0483 | 10.2024 | 9.9111 | 10.431 | 64.34% | 46.69% | 54.84% | 56.92% | 47.11% | 54.4% | 9.6437 | 9.7748 | 10.8962 | 11.1018
219 | CCV 10 ppb MAJORS |  | 06/17/2023 10:24:55 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.868 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9136.6389 | 9344.4324 | 9.3527 | 8.8583 | 11013.0283 | 9452.938 | 10278.224 | 10.0455 | 10.1677 | 10.581 | 10.0385 | 10.2585 | 9.2147 | 9.7741 | 9.4132 | 9.7052 | 10.0953 | 10.1275 | 9.4815 | 9.6238 | 9.8198 | 10.1195 | 10.0042 | 9.8117 | 9.8013 | 10.4542 | 10.049 | 9.9984 | 10.4653 | 64.84% | 47.31% | 54.99% | 58.39% | 48.17% | 55.62% | 9.7188 | 9.8832 | 10.7595 | 10.8331
220 | 5_18_23_DB |  | 06/17/2023 10:30:55 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_DB.869 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 56.0874 | 0.23633 | -0.34413 | 3.0299 | 483.6795 | 56.5471 | 42.803 | 0.00039696 | 1.6621e-05 | 0.0065821 | 1.1027 | 0.0076569 | 1.2403 | -0.0060061 | -0.0043181 | 0.081987 | 0.075413 | -0.13935 | -0.0097701 | -0.0066049 | -0.0019681 | 0.12984 | 0.0055366 | 0.0095666 | -0.0073449 | -0.0050493 | -0.0014681 | 0.0034523 | 0.0037905 | 66.85% | 51.14% | 60.21% | 63.36% | 50.31% | 58.43% | 0.00050427 | -0.0019947 | 0.017713 | 0.021232
221 | 5_18_23_DB |  | 06/17/2023 10:37:14 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_DB.870 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 55.5641 | 0.87935 | -0.45883 | 4.2547 | 414.7607 | 59.2883 | 40.9256 | -0.00047279 | 0.00071555 | 0.0043316 | 1.0919 | 0.0080144 | 1.1945 | -0.0073199 | -0.0050174 | 0.082157 | 0.080496 | -0.14309 | -0.032794 | -0.0020638 | -0.0012618 | 0.084607 | 0.013394 | 0.0020228 | -0.010879 | -0.064205 | -0.0080586 | -0.00069779 | 0.0019706 | 67.86% | 51.12% | 59.47% | 61.18% | 49.74% | 58.44% | 0.00064062 | -0.002559 | 0.015932 | 0.022555
222 | 5_18_23_DB |  | 06/17/2023 10:43:14 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_DB.871 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 55.0292 | 0.23633 | -0.41535 | 4.4563 | 526.0722 | 63.4354 | 46.4853 | -3.8842e-05 | 0.007231 | 0.013075 | -0.38083 | 0.0034809 | -0.10681 | 0.06763 | 0.0058478 | 0.089958 | 0.08793 | -0.09427 | -0.091237 | -0.0019545 | -0.0011934 | 0.43327 | 0.0055366 | 0.004616 | -0.0084655 | 0.013129 | -0.014676 | -0.004036 | 0.0023358 | 60.27% | 50.22% | 58.86% | 60.28% | 49.53% | 58.14% | 0.00078193 | -0.0023205 | 0.025718 | 0.025116
223 | 5_18_23_DB |  | 06/17/2023 10:49:33 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_DB.872 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 55.5987 | 0.23633 | -0.542 | 4.249 | 541.8147 | 54.1982 | 41.9035 | -2.0521e-05 | 0.0043401 | 0.0089419 | -0.35624 | 0.0030895 | -0.10843 | 0.063046 | -0.0055082 | 0.086176 | 0.090599 | -0.10312 | 0.0080738 | -0.01135 | -0.0015096 | 0.18504 | 0.037918 | 0.0020514 | -0.007219 | -0.022544 | -0.021474 | -0.0054082 | 0.0025307 | 64.43% | 49.76% | 58.4% | 67.1% | 48.91% | 57.3% | 0.00092317 | -0.0027262 | 0.020813 | 0.020228
224 | 5_18_23_DB |  | 06/17/2023 10:55:33 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_DB.873 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 56.4577 | 0.23633 | -0.49485 | 4.087 | 584.6872 | 57.1195 | 41.9141 | -0.00046435 | -0.0015507 | 0.0015378 | -0.46333 | 0.0056896 | -0.10679 | -0.0071152 | 0.00051067 | 0.065643 | 0.060561 | -0.09366 | 0.0048129 | -0.01135 | -0.0031009 | 0.03863 | 0.0055366 | -0.00050913 | -0.0058545 | -0.035078 | -0.021474 | -0.00518 | 0.0018303 | 65.9% | 48.82% | 57.57% | 59.14% | 48.25% | 56.7% | 0.0010707 | -0.0022452 | 0.0097389 | 0.014763
225 | 5_18_23_DB |  | 06/17/2023 11:01:51 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_DB.874 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 56.1408 | 0.23633 | -0.58117 | 3.9238 | 489.0645 | 57.4775 | 43.8027 | -0.00091385 | -0.0046186 | 1.0946e-05 | -0.46647 | 0.0044928 | -0.19377 | -0.0072986 | 0.0047816 | 0.05814 | 0.06201 | -0.096595 | -0.031544 | 0.003152 | -0.0042019 | -0.023642 | 0.0055366 | 0.012573 | -0.0059201 | -0.035507 | -0.021474 | -0.0060373 | 0.0017349 | 65.64% | 49.11% | 57.07% | 59.01% | 47.6% | 57.08% | 0.00064873 | -0.0021798 | 0.0081603 | 0.016227
226 | 5_18_23_MFB |  | 06/17/2023 11:07:51 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_MFB.875 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4502.9532 | 4575.0836 | 4.35 | 9.6796 | 5801.114 | 4707.6221 | 4994.9937 | 4.9413 | 5.2152 | 5.117 | 4.8635 | 5.1413 | 4.775 | 4.8546 | 4.6213 | 5.0005 | 5.0256 | 4.6334 | 4.5372 | 4.8443 | 4.9364 | 5.1613 | 5.2911 | 5.0694 | 4.9715 | 4.8808 | 4.9016 | 4.7381 | 5.1323 | 62.84% | 45.61% | 54.22% | 56.11% | 46.34% | 53.1% | 4.6559 | 4.7335 | 5.656 | 5.6578
227 | 5_18_23_MFB |  | 06/17/2023 11:14:09 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_MFB.876 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4463.4639 | 4710.4861 | 4.735 | 9.6626 | 5434.759 | 4631.4346 | 5054.1685 | 4.8814 | 5.0092 | 5.1374 | 4.8842 | 5.0269 | 4.5976 | 4.7959 | 4.5058 | 4.9135 | 4.9987 | 4.6561 | 4.4796 | 4.7587 | 4.7943 | 5.0754 | 5.2604 | 4.7412 | 4.8838 | 5.0667 | 4.7738 | 4.749 | 4.9597 | 64.46% | 47.19% | 54.88% | 57.1% | 47.12% | 54.81% | 4.7477 | 4.7746 | 5.5546 | 5.638
228 | 5_18_23_MFB |  | 06/17/2023 11:20:08 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_MFB.877 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4428.9243 | 4430.2363 | 4.6763 | 9.9353 | 5721.9346 | 4590.4168 | 4902.0196 | 4.8057 | 5.0277 | 4.9683 | 4.5221 | 4.8921 | 4.2767 | 4.7135 | 4.5537 | 4.871 | 4.9143 | 4.5065 | 4.4756 | 4.3624 | 4.6548 | 5.2603 | 5.1506 | 4.8765 | 4.7152 | 4.8929 | 4.8166 | 4.5883 | 4.8284 | 63.85% | 48.18% | 55.96% | 52.41% | 48.54% | 55.94% | 4.6554 | 4.6797 | 5.5783 | 6.1916
229 | 5_18_23_MFB |  | 06/17/2023 11:26:26 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_MFB.878 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4388.0558 | 4647.9886 | 4.4676 | 10.0162 | 5605.7713 | 4604.6739 | 4913.2522 | 4.8383 | 5.1078 | 5.1317 | 4.6075 | 4.9353 | 4.368 | 4.7526 | 4.6859 | 4.9266 | 5.0203 | 4.5944 | 4.1423 | 4.7415 | 4.7736 | 4.6372 | 4.8249 | 4.7784 | 4.6967 | 4.4521 | 4.679 | 4.7558 | 4.9286 | 68.61% | 47.39% | 55.16% | 57.05% | 47.94% | 55.29% | 4.6784 | 4.704 | 5.0777 | 5.5252
230 | LBLANK |  | 06/17/2023 11:32:26 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.879 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 54.5088 | 0.8925 | -0.97753 | 0.44344 | 471.7096 | 64.0202 | 40.0401 | 0.00089021 | -0.0052172 | -0.0014034 | -0.57981 | -0.0005055 | -0.15477 | -0.0075028 | 0.0056466 | -0.010101 | -0.010062 | -0.23466 | -0.011186 | -0.01135 | -0.00099884 | 0.060126 | 0.0055366 | 0.0046556 | -0.0060193 | 0.0016578 | -0.01462 | 0.0021293 | 0.0029673 | 67.84% | 49.94% | 58.18% | 59.6% | 48.74% | 56.74% | 0.00064673 | -0.0020761 | -0.010465 | -0.0040721
231 | LBLANK |  | 06/17/2023 11:38:51 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.880 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 54.7736 | 0.23633 | -0.97863 | -0.6603 | 461.2309 | 58.709 | 40.0395 | -1.947e-05 | -0.0050509 | -0.0020565 | -0.59531 | -0.00067734 | -0.28885 | -0.0075076 | -0.0011957 | -0.012398 | -0.012075 | -0.23104 | 0.0064449 | -0.01135 | -0.0027775 | 0.11097 | 0.0055366 | -0.00050913 | -0.0072037 | -0.017012 | -0.021474 | -0.0016192 | 0.0022626 | 67.3% | 49.73% | 58.1% | 59.9% | 49.02% | 57.48% | 0.00050427 | -0.0024889 | -0.010459 | -0.0041141
232 | CCV 1ppb TRACE |  | 06/17/2023 11:44:51 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.881 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 967.4726 | 1004.7918 | 1.361 | 0.24606 | 1518.8956 | 990.955 | 1104.9739 | 1.019 | 1.0727 | 1.0913 | 0.91663 | 1.0591 | 1.1138 | 0.97472 | 0.98598 | 1.0189 | 1.0193 | 0.81806 | 0.85939 | 0.99708 | 0.99701 | 1.1505 | 0.8848 | 0.99406 | 1.0403 | 1.0933 | 1.002 | 0.98413 | 1.0285 | 64.74% | 48.75% | 57.11% | 58.92% | 48.03% | 56.74% | 0.96878 | 0.98541 | 1.2091 | 1.1621
233 | CCV 1ppb TRACE |  | 06/17/2023 11:51:11 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.882 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 971.2691 | 1012.0527 | 1.2375 | 0.23216 | 1682.2554 | 1039.9667 | 1059.4263 | 1.0437 | 1.0384 | 1.0541 | 0.89143 | 1.083 | 1.0717 | 0.99625 | 1.0157 | 1.0338 | 1.0378 | 0.79909 | 0.86597 | 1.0611 | 0.96681 | 1.4148 | 0.9827 | 0.9507 | 0.98401 | 1.0236 | 1.0158 | 0.97625 | 1.0435 | 64.34% | 48.42% | 56.66% | 58.22% | 48.19% | 56.4% | 0.95965 | 0.98965 | 1.1992 | 1.1692
234 | CCV 10 ppb MAJORS |  | 06/17/2023 11:57:11 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.883 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9036.4133 | 9290.9252 | 9.1368 | 8.8344 | 10684.1627 | 9550.7685 | 10259.8601 | 10.0001 | 10.1686 | 10.5382 | 10.1855 | 10.3267 | 9.0796 | 9.6226 | 9.3076 | 9.688 | 10.0191 | 10.0356 | 9.2331 | 10.1107 | 9.9015 | 10.1159 | 10.0307 | 9.7954 | 10.1018 | 10.1547 | 9.9832 | 10.0086 | 10.4225 | 63.41% | 45.29% | 53.25% | 55.18% | 45.69% | 53.64% | 9.6742 | 9.7936 | 10.7344 | 11.149
235 | CAL BLANK |  | 06/17/2023 12:25:05 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CAL BLANK.884 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth

## Concentrations RSDs
| Sample Id | R | Acquisition Time | QC Status | Dataset File | Method File | Na KED 23/23 Helium KED_LowLow | Mg KED 24/24 Helium KED_LowLow | Al KED 27/27 Helium KED_LowLow | P Oshift 31/47 Oxygen DRC | S (34) -1 34/34 Helium KED_LowLow | K KED 39/39 Helium KED_LowLow | Ca KED 43/43 Helium KED_LowLow | V KED 51/51 Helium KED_LowLow | Cr (52) 52/52 Helium KED_LowLow | Cr (53) NH3 53/53 Ammonia DRC | Fe (54) NH3 54/54 Ammonia DRC | Mn Oshift 55/71 Oxygen DRC | Fe (56) KED 56/56 Helium KED_LowLow | Co KED 59/59 Helium KED_LowLow | Ni KED 60/60 Helium KED_LowLow | Cu (63) KED 63/63 Helium KED_LowLow | Cu (65) NH3 65/65 Ammonia DRC | Zn NH3 66/66 Ammonia DRC | As 75/75 | As KED 75/75 Helium KED_LowLow | As Oshift 75/91 Oxygen DRC | Se (77) 77/77 | Se (77) Oshift 77/93 Oxygen DRC | Se (78) Oshift 78/94 Oxygen DRC | Se (80) Oshift 80/96 Oxygen DRC | Se (82) 82/82 | Se (82) Oshift 82/98 Oxygen DRC | Mo (95) NH3 95/95 Ammonia DRC | Mo (95) OShift 95/127 Oxygen DRC | Rh (IS1) 103/103 (IS) | Rh (IS2) 103/103 (IS) | Rh (IS4) 103/103 (IS) | Rh (IS5) 103/103 (IS) | Rh (IS6) 103/103 (IS) | Rh (IS7) 103/103 (IS) | Cd (111) NH3 111/111 Ammonia DRC | Cd (114) NH3 114/114 Ammonia DRC | Pb 208/208 | Pb NH3 208/208 Ammonia DRC
1 | sac blank |  | 6/16/2023 11:58:56 AM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\sac blank .650 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth
2 | sac blank cal |  | 6/16/2023 12:04:57 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\sac blank cal.651 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth
3 | sac blank cal |  | 6/16/2023 12:10:57 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\sac blank cal.652 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth
4 | CAL BLANK |  | 6/16/2023 12:16:57 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CAL BLANK.653 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | [29 empty] | 0.6452% | 0.1482% | 0.7484% | 0.4378% | 0.8227% | 1.556%
5 | LOW 0.01 |  | 6/16/2023 12:22:58 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LOW 0.01.654 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | [29 empty] | 1.137% | 0.8423% | 0.9212% | 0.873% | 0.507% | 1.977%
6 | LOW 0.05 |  | 6/16/2023 12:28:58 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LOW 0.05.655 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | [7 empty] | 10.35% | 11.78% | 4.13% |  |  |  | 5.625% | 18.9% | 19.73% | 6.783% |  |  | 24.82% | 8.189% | [8 empty] | 1.348% | 0.5856% | 1.247% | 0.3364% | 0.4161% | 0.9849% | 12.25% | 4.998% | 5.011% | 0.6759%
7 | LOW 0.1 |  | 6/16/2023 12:34:59 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LOW 0.1.656 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | [7 empty] | 7.008% | 17.09% | 3.029% |  | 0.4664% |  | 14.99% | 12.47% | 12.85% | 4.921% |  | 327.1% | 26.84% | 5.222% |  | 99.07% | 11.74% | 14.62% | 15.47% | 107.5% | 1.537% | 3.205% | 0.7358% | 0.9433% | 1.578% | 0.8387% | 1.27% | 0.9627% | 1.797% | 4.853% | 5.502% | 1.016%
8 | LOW 0.5 |  | 6/16/2023 12:40:59 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LOW 0.5.657 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth |  |  |  | 4.782% |  |  |  | 2.412% | 5.16% | 3.841% |  | 2.144% |  | 1.214% | 2.632% | 3.169% | 2.757% | 4.289% | 4.461% | 12.29% | 2.084% | 17.01% | 9.586% | 7.261% | 6.841% | 37.46% | 11.4% | 1.139% | 2.259% | 0.7953% | 0.8605% | 0.2624% | 1.017% | 2.533% | 0.759% | 2.446% | 1.079% | 0.7974% | 0.8248%
9 | STD 1 |  | 6/16/2023 12:46:58 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 1.658 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.599% | 2.732% | 21.35% | 1.714% |  | 2.599% | 8.05% | 5.343% | 1.486% | 1.769% | 1.531% | 1.221% | 1.33% | 1.36% | 3.936% | 3.209% | 3.534% | 3.651% | 7.705% | 17.11% | 2.543% | 14.78% | 17.19% | 6.106% | 2.357% | 9.295% | 13.58% | 2.797% | 2.532% | 0.3183% | 1.265% | 1.929% | 0.8632% | 1.834% | 0.8067% | 1.118% | 2.535% | 0.7926% | 1.894%
10 | STD 5 |  | 6/16/2023 12:52:58 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 5.659 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.777% | 1.734% | 5.353% | 0.6488% | 3.649% | 2.846% | 3.023% | 2.247% | 1.501% | 1.34% | 0.9761% | 1.117% | 1.108% | 0.7175% | 2.254% | 1.257% | 2.028% | 1.107% | 0.9878% | 2.361% | 0.7321% | 3.848% | 7.255% | 3.536% | 0.125% | 4.411% | 0.8287% | 1.31% | 1.548% | 0.536% | 1.333% | 1.775% | 1.024% | 1.027% | 1.792% | 2.81% | 1.445% | 1.658% | 0.3128%
11 | STD 10 |  | 6/16/2023 12:58:58 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 10.660 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.053% | 2.138% | 4.263% | 1.176% | 2.943% | 1.909% | 0.2303% | 0.5585% | 0.9875% | 0.2119% | 0.7808% | 0.8018% | 1.851% | 0.9695% | 2.385% | 0.7717% | 0.4174% | 0.311% | 0.7886% | 1.345% | 0.7177% | 2.789% | 3.721% | 2.31% | 1.873% | 1.64% | 4.201% | 0.5498% | 0.9027% | 1.287% | 0.3345% | 0.3715% | 0.8411% | 0.951% | 0.05589% | 1.028% | 0.1656% | 0.8513% | 1.658%
12 | STD 25 |  | 6/16/2023 1:04:58 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 25.661 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.7975% | 0.4826% | 1.969% | 15.67% | 1.855% | 1.404% | 0.09777% | 0.2777% | 1.454% | 0.9426% | 0.6935% | 1.584% | 0.8611% | 0.8123% | 0.4859% | 0.4298% | 1.484% | 0.9274% | 1.217% | 1.87% | 1.037% | 0.7368% | 2.785% | 0.7979% | 0.203% | 0.6072% | 1.103% | 0.7384% | 0.618% | 0.4345% | 0.873% | 1.241% | 0.8751% | 12.49% | 0.2745% | 1.019% | 1.001% | 1.982% | 0.7445%
13 | STD 50 |  | 6/16/2023 1:10:58 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 50.662 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.279% | 0.7516% | 1.286% | 0.2079% | 1.984% | 1.424% | 0.2302% | 1.258% | 1.225% | 0.2961% | 0.1914% | 0.331% | 2.07% | 1.701% | 1.853% | 2.112% | 0.8872% | 0.8524% | 0.3679% | 3.075% | 0.9184% | 0.7932% | 3.289% | 2.066% | 1.475% | 0.2984% | 2.866% | 0.5889% | 0.5153% | 0.3426% | 1.221% | 0.7681% | 0.7598% | 0.9539% | 1.873% | 0.7613% | 1.383% | 2.64% | 1.808%
14 | STD 100 |  | 6/16/2023 1:16:57 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 100.663 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.7% | 1.307% | 1.341% | 0.9296% | 0.431% | 2.471% | 2.012% | 1.035% | 0.979% | 1.074% | 0.3794% | 1.41% | 2.091% | 1.808% | 0.9255% | 0.4702% | 0.8581% | 0.7932% | 0.7474% | 1.916% | 0.1657% | 0.4874% | 2.393% | 0.1669% | 0.519% | 0.7126% | 0.8783% | 0.5236% | 0.4601% | 1.263% | 0.6764% | 0.9956% | 0.7512% | 0.532% | 0.3539% | 0.271% | 0.8062% | 2.835% | 0.7029%
16 | LBLANK |  | 6/16/2023 1:22:58 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.664 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.86% | 64.11% | 1.372% | 5.183% | 37.68% | 8.754% | 4.075% | 134.4% | 72.77% | 230.5% | 0.3693% | 37.33% | 30% | 36.25% | 17.13% | 22.57% | 3.517% | 10.91% | 41.85% | 2.483e-06% | 66.26% | 66.13% | 0.6766% | 30.46% | 103.7% | 1826% | 0% | 1.547% | 18.29% | 0.9482% | 0.9773% | 0.6451% | 0.2886% | 1.639% | 1.666% | 20.29% | 88.98% | 3.631% | 5.325%
17 | LBLANK |  | 6/16/2023 1:41:50 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.665 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.6113% | 128.8% | 23.4% | 0.9477% | 123.4% | 10.42% | 2.466% | 59.85% | 11.63% | 812.3% | 1.471% | 123.4% | 70.99% | 3.791% | 100.2% | 10.93% | 9.931% | 71.78% | 38.75% | 1.777e-06% | 7.684% | 38.4% | 53.28% | 68.04% | 56.18% | 385.7% | 0% | 3.993% | 5.245% | 0.5292% | 0.7508% | 0.6203% | 0.286% | 0.6941% | 7.71% | 22.65% | 29.14% | 1.693% | 5.958%
18 | ICV PE |  | 6/16/2023 1:47:50 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\ICV PE.666 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.54% | 0.4419% | 2.368% | 1.088% | 1.033% | 1.39% | 1.257% | 2.059% | 0.5718% | 1.794% | 1.486% | 0.481% | 1.843% | 0.7227% | 1.62% | 0.3411% | 0.8926% | 0.6228% | 1.085% | 3.691% | 1.317% | 3.432% | 4.596% | 0.6279% | 1.94% | 1.243% | 3.82% | 0.6342% | 1.031% | 0.2963% | 0.3461% | 1.092% | 0.656% | 0.4973% | 0.4802% | 1.364% | 1.291% | 0.3706% | 1.502%
19 | CCV 1ppb TRACE |  | 6/16/2023 1:53:50 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.667 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.4014% | 128.5% | 2.185% | 0.6069% | 1.243% | 8.21% | 0.04259% | 11218% | 10.2% | 25.53% | 0.6791% | 11.59% | 5.018% | 9.173% | 48.41% | 6.57% | 5.336% | 3.579% | 2.663% | 47.15% | 9.991% | 2.341% | 0% | 58.73% | 23.41% | 42.25% | 0% | 4.473% | 19.67% | 2.49% | 2.482% | 2.464% | 1.203% | 7.071% | 1.19% | 33.2% | 11.78% | 2.257% | 3.397%
20 | CCV 1ppb TRACE |  | 6/16/2023 2:00:07 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.668 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.402% | 5.783% | 9.168% | 70.55% | 6.41% | 1.564% | 4.403% | 2.875% | 0.8365% | 0.06988% | 5.113% | 1.448% | 2.843% | 0.3966% | 2.618% | 2.672% | 0.8651% | 3.048% | 6.198% | 2.583% | 1.815% | 12.78% | 7.67% | 5.36% | 4.046% | 4.028% | 4.982% | 1.123% | 3.051% | 1.609% | 2.152% | 0.7165% | 0.008299% | 6.066% | 1.123% | 2.268% | 2.109% | 2.031% | 0.9436%
21 | LFB |  | 6/16/2023 2:06:07 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.669 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9.294% | 9.795% | 11.63% | 0.5237% | 12.11% | 9.65% | 8.647% | 10.41% | 11.27% | 4.291% | 5.056% | 0.9218% | 12.85% | 10.29% | 10.32% | 10.63% | 3.289% | 4.34% | 1.462% | 14.67% | 1.092% | 2.59% | 1.71% | 2.934% | 2.562% | 2.87% | 3.766% | 5.365% | 1.34% | 0.4745% | 1.575% | 2.537% | 0.2438% | 1.024% | 7.511% | 4.806% | 4.649% | 0.5717% | 1.751%
22 | LLOQ 0.01 |  | 6/16/2023 2:12:05 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LLOQ 0.01.670 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.245% | 5.38% | 200.8% | 1.933% | 26.96% | 8.771% | 10.67% | 46.88% | 51.1% | 12.75% | 49.97% | 11.32% | 6.145% | 11.26% | 34.83% | 27.72% | 9.577% | 262% | 71.46% | 294.6% | 12.76% | 308.8% | 71.42% | 28.17% | 97.05% | 198% | 1910% | 2.543% | 9.055% | 0.9267% | 0.7248% | 0.4548% | 0.4669% | 0.7163% | 0.5086% | 13.82% | 16.79% | 19.29% | 3.558%
23 | CCV 10 ppb MAJORS |  | 6/16/2023 2:18:05 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.671 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.695% | 2.042% | 1.571% | 1.54% | 2.497% | 0.494% | 1.971% | 0.7571% | 0.4387% | 1.003% | 1.004% | 1.13% | 1.838% | 1.492% | 1.427% | 1.356% | 0.3781% | 0.8226% | 1.872% | 5.737% | 1.145% | 2.319% | 6.124% | 2.95% | 0.883% | 3.276% | 2.883% | 1.265% | 1.507% | 0.5598% | 1.158% | 0.2376% | 0.9476% | 0.9511% | 0.4046% | 0.2624% | 0.3827% | 0.3781% | 1.141%
24 | LLOQ 0.1 |  | 6/16/2023 2:24:05 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LLOQ 0.1.672 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 5.96% | 4.792% | 4.423% | 11.01% | 39.67% | 3.805% | 8.24% | 9.959% | 4.757% | 5.109% | 6.217% | 5.866% | 62.07% | 9.499% | 7.081% | 8.64% | 2.645% | 122.9% | 34.34% | 4.707% | 4.781% | 94.57% | 41.23% | 14.52% | 21.68% | 45.59% | 17.63% | 5.174% | 2.726% | 0.03196% | 0.558% | 1.709% | 0.535% | 0.9518% | 0.9947% | 4.84% | 3.777% | 2.771% | 1.121%
25 | 9_13_21_F_11_DUP |  | 6/16/2023 2:30:04 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_F_11_DUP.673 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.855% | 4.329% | 3.788% | 1.573% | 2.332% | 3.796% | 2.284% | 2.951% | 8.423% | 4.946% | 4.73% | 1.126% | 2.093% | 7.789% | 5.143% | 3.192% | 4.264% | 1.827% | 1.218% | 1.363% | 1.276% | 40.14% | 60.51% | 25.4% | 17.72% | 9.241% | 11.52% | 1.152% | 1.741% | 0.9009% | 1.324% | 2.934% | 0.765% | 0.9075% | 1.778% | 9.622% | 8.717% | 0.8132% | 2.48%
26 | LBLANK |  | 6/16/2023 2:47:08 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.675 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.494% | 147% | 13.07% | 0.7494% | 21.2% | 12.03% | 4.634% | 188.2% | 88.22% | 295.1% | 0.8314% | 101.2% | 21.01% | 12.83% | 4546% | 11.25% | 8.977% | 10.03% | 175% | 938.7% | 25.09% | 106.8% | 80.92% | 124.6% | 98.34% | 325.8% | 41.27% | 6.671% | 26.13% | 0.3744% | 0.8687% | 1.274% | 1.54% | 2.42% | 1.249% | 27.51% | 17.79% | 1.225% | 5.382%
27 | 0.5 STD |  | 6/16/2023 2:53:09 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\0.5 STD.676 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4.52% | 0.9283% | 44.66% | 3.307% | 14.66% | 3.594% | 2.895% | 2.627% | 2.558% | 2.922% | 3.498% | 2.973% | 7.294% | 3.932% | 1.604% | 3.381% | 3.505% | 2.411% | 3.934% | 15.24% | 4.837% | 27.41% | 26.64% | 13.34% | 11.82% | 5.746% | 11.12% | 1.724% | 1.255% | 0.8426% | 0.909% | 0.3572% | 0.3128% | 2.212% | 1.801% | 1.523% | 2.384% | 1.338% | 2.37%
28 | 0.1 STD |  | 6/16/2023 2:59:12 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\0.1 STD.677 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.956% | 2.885% | 5.184% | 409.2% | 21.44% | 5.267% | 9.125% | 3.315% | 5.261% | 4.287% | 6.323% | 2.438% | 455.1% | 3.648% | 16.39% | 18.92% | 2.833% | 15.68% | 10.96% | 17.92% | 4.645% | 69.69% | 18.66% | 37.47% | 21.06% | 72.37% | 63.61% | 1.499% | 6.435% | 2.147% | 0.2205% | 0.8267% | 0.4524% | 0.9052% | 0.3296% | 0.3992% | 3.027% | 4.737% | 3.121%
29 | CCV 10 ppb MAJORS |  | 6/16/2023 3:05:12 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.678 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.4342% | 2.684% | 9.996% | 0.6815% | 2.317% | 1.358% | 1.589% | 0.5377% | 0.7941% | 1.013% | 1.235% | 0.8408% | 1.068% | 1.479% | 1.743% | 1.758% | 1.334% | 2.05% | 1.984% | 1.771% | 0.9794% | 5.42% | 3.02% | 1.871% | 0.3719% | 3.44% | 0.9664% | 0.3917% | 0.6337% | 1.111% | 0.4492% | 1.105% | 1.027% | 0.6158% | 0.921% | 0.9161% | 1.446% | 0.8602% | 1.518%
30 | 9_13_21_F_11_DUP |  | 6/16/2023 3:11:13 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_F_11_DUP.679 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.083% | 1.475% | 7.428% | 3.901% | 1.029% | 1.41% | 0.7642% | 3.137% | 7.779% | 5.04% | 1.463% | 1.763% | 0.4826% | 13.27% | 5.159% | 1.832% | 1.406% | 0.3607% | 0.5973% | 1.304% | 1.879% | 69.3% | 41.8% | 26.1% | 3.373% | 6.653% | 40.72% | 1.804% | 0.824% | 0.5127% | 1.097% | 0.2767% | 0.5001% | 0.9798% | 0.2759% | 16.37% | 7.551% | 2.939% | 2.07%
31 | 9_13_21_F_11_DUP_LSP |  | 6/16/2023 3:17:13 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_F_11_DUP_LSP.680 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.057% | 2.296% | 4.461% | 0.9324% | 1.936% | 3.825% | 2.039% | 2.55% | 2.389% | 0.4909% | 0.9916% | 0.2349% | 2.535% | 0.1819% | 0.9012% | 1.016% | 0.9045% | 0.6708% | 2.056% | 6.489% | 0.3312% | 0.5962% | 4.781% | 0.3163% | 0.7206% | 2.309% | 4.209% | 1.605% | 0.6594% | 1.517% | 0.4821% | 0.4831% | 1.866% | 1.838% | 1.934% | 1.613% | 0.6048% | 0.4703% | 1.512%
32 | 9_13_21_UF_1 |  | 6/16/2023 3:23:13 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_UF_1.681 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.206% | 0.7033% | 6.906% | 0.4392% | 2.502% | 1.408% | 0.8761% | 3.144% | 19.82% | 6.907% | 4.324% | 0.6782% | 10.18% | 10.79% | 4.006% | 0.9946% | 1.475% | 5.023% | 2.326% | 9.136% | 1.379% | 45.24% | 8.068% | 7.369% | 13.35% | 3.422% | 37.77% | 0.7244% | 0.3864% | 0.281% | 0.7427% | 0.2023% | 1.003% | 2.241% | 0.5457% | 20.67% | 21.92% | 5.487% | 0.3349%
33 | 9_13_21_UF_6 |  | 6/16/2023 3:29:13 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_UF_6.682 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.4603% | 2.541% | 229.9% | 5.336% | 1.825% | 1.241% | 1.547% | 2.275% | 12.85% | 6.567% | 10.78% | 1.201% | 3.45% | 7.026% | 8.844% | 2.229% | 1.734% | 1.676% | 2.659% | 1.968% | 0.5397% | 26.88% | 21.17% | 25.66% | 24.85% | 14.09% | 31.84% | 0.6884% | 1.121% | 1.183% | 0.7426% | 0.9056% | 2.44% | 0.7638% | 0.704% | 12.38% | 7.798% | 6.094% | 4.128%
34 | 9_13_21_UF_6_DUP |  | 6/16/2023 3:35:13 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_UF_6_DUP.683 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.437% | 2.573% | 451.1% | 1.31% | 1.886% | 2.417% | 1.27% | 2.367% | 3.603% | 2.597% | 10.26% | 1.127% | 3.437% | 9.829% | 12.53% | 1.431% | 1.642% | 0.6958% | 0.2308% | 6.723% | 0.2505% | 7.895% | 33.09% | 19.83% | 0.6633% | 7.282% | 38.76% | 0.9972% | 0.9991% | 1.278% | 0.5713% | 0.4614% | 0.5448% | 1.12% | 1.845% | 23.21% | 10.45% | 3.074% | 0.8708%
35 | 9_13_21_UF_8 |  | 6/16/2023 3:41:13 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_UF_8.684 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.658% | 3.495% | 762.9% | 0.9986% | 2.538% | 1.979% | 1.417% | 3.873% | 11.19% | 2.603% | 80.21% | 0.941% | 3.104% | 6.941% | 15.95% | 1.701% | 1.045% | 0.1088% | 1.776% | 7.043% | 1.142% | 43.25% | 54.12% | 16.46% | 20.85% | 7.015% | 70.12% | 1.606% | 0.397% | 1.747% | 0.2641% | 1.159% | 0.9784% | 1.742% | 0.5843% | 26.45% | 24.07% | 5.974% | 1.933%
36 | 9_13_21_UF_9 |  | 6/16/2023 3:47:13 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_UF_9.685 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.855% | 1.867% | 77.4% | 3.175% | 3.374% | 1.505% | 2.126% | 5.43% | 4.422% | 11.49% | 3.677% | 0.2041% | 11.41% | 8.136% | 5.263% | 3.46% | 2.028% | 3.323% | 3.132% | 5.587% | 2.224% | 36.17% | 33.66% | 9.695% | 40.8% | 11.24% | 44.79% | 1.552% | 1.909% | 0.7681% | 1.131% | 1.681% | 1.3% | 1.13% | 1.956% | 29.85% | 28.78% | 4.652% | 2.828%
37 | LBLANK |  | 6/16/2023 3:53:13 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.686 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.6004% | 1.931e-06% | 4.527% | 0.7008% | 94.93% | 3.195% | 2.811% | 4200% | 18.58% | 118.4% | 1.636% | 401.1% | 8.929% | 3.062% | 1372% | 28.11% | 4.851% | 1.874% | 36.62% | 42.15% | 8.227% | 174% | 0% | 36.62% | 20.37% | 412.8% | 0% | 32.69% | 8.984% | 1.583% | 0.9188% | 0.8614% | 1.071% | 0.3249% | 1.125% | 28.56% | 8.648% | 4.425% | 5.954%
38 | CCV 1ppb TRACE |  | 6/16/2023 3:59:13 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.687 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9.244% | 12.12% | 7.756% | 12.76% | 24.58% | 6.452% | 8.567% | 11.03% | 13.44% | 1.388% | 4.959% | 1.547% | 28.22% | 13.75% | 14.99% | 9.754% | 0.7964% | 2.601% | 3.992% | 20.16% | 1.184% | 11.66% | 16.09% | 4.293% | 3.464% | 4.174% | 6.546% | 2.131% | 3.184% | 0.2891% | 1.007% | 1.056% | 1.099% | 1.01% | 7.83% | 1.536% | 1.676% | 1.483% | 0.3473%
39 | CCV 1ppb TRACE |  | 6/16/2023 4:05:32 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.688 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.9564% | 4.59% | 8.92% | 13.08% | 5.726% | 3.291% | 2.154% | 2.986% | 0.9137% | 1.687% | 7.241% | 0.1876% | 8.576% | 1.994% | 1.884% | 4.795% | 1.384% | 1.335% | 3.378% | 11.06% | 1.27% | 10.28% | 13.57% | 2.937% | 0.3634% | 4.741% | 7.456% | 1.143% | 2.405% | 2.002% | 0.2311% | 1.067% | 0.8012% | 1.069% | 2.357% | 3.667% | 0.9035% | 1.452% | 0.9987%
40 | CCV 10 ppb MAJORS |  | 6/16/2023 4:11:33 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.689 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.535% | 2.383% | 3.285% | 1.847% | 4.852% | 4.303% | 1.925% | 2.621% | 1.894% | 0.6639% | 1.571% | 1.647% | 1.602% | 2.015% | 2.75% | 1.97% | 0.3064% | 0.6809% | 1.125% | 4.044% | 0.7572% | 2.395% | 5.466% | 3.089% | 1.737% | 4.882% | 2.827% | 0.6091% | 0.6696% | 0.9888% | 0.5536% | 0.3194% | 0.3864% | 0.835% | 1.568% | 1.235% | 0.6602% | 0.9531% | 0.9137%
41 | CCV 10 ppb MAJORS |  | 6/16/2023 4:17:52 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.690 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.112% | 1.095% | 5.305% | 2.297% | 4.474% | 1.888% | 0.9133% | 2.681% | 1.103% | 1.539% | 1.238% | 0.9717% | 1.512% | 1.66% | 1.621% | 1.297% | 1.353% | 0.5107% | 1.301% | 2.171% | 1.621% | 0.2758% | 7.471% | 1.714% | 2.408% | 2.509% | 3.996% | 1.615% | 1.114% | 1.178% | 0.76% | 0.7053% | 1.42% | 0.6062% | 0.9535% | 1.158% | 0.8829% | 1.157% | 2.216%
42 | 9_13_21_UF_10 |  | 6/16/2023 4:23:53 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_UF_10.691 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.777% | 0.7523% | 23.88% | 1.619% | 2.257% | 3.887% | 0.986% | 3.939% | 14% | 7.311% | 6.845% | 0.6607% | 7.509% | 23.55% | 7.231% | 1.527% | 2.189% | 4.578% | 1.78% | 4.655% | 0.4589% | 33.07% | 40.15% | 23.69% | 20.09% | 10.69% | 51.93% | 1.219% | 1.083% | 1.462% | 1.066% | 0.7566% | 0.5826% | 0.2239% | 1.46% | 17.35% | 28.49% | 2.924% | 3.204%
43 | 9_13_21_UF_11 |  | 6/16/2023 4:29:52 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_UF_11.692 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.6343% | 1.598% | 145% | 0.6378% | 0.659% | 0.8598% | 0.5283% | 3.004% | 6.7% | 1.48% | 6.796% | 1.167% | 8.538% | 47.72% | 9.306% | 1.732% | 1.681% | 1.598% | 0.5256% | 2.802% | 1.709% | 96.94% | 22.4% | 28.29% | 15.16% | 13.18% | 19.5% | 0.9303% | 2.921% | 1.227% | 1.29% | 0.6789% | 3.544% | 0.4046% | 0.4105% | 11.74% | 17.37% | 4.139% | 4.726%
44 | 9_13_21_LAB_A |  | 6/16/2023 4:35:53 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_LAB_A.693 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.575% | 1.931e-06% | 210% | 1.248% | 2.833% | 5.633% | 5.287% | 78667% | 54.53% | 15.2% | 1.846% | 2.402% | 24.68% | 12.52% | 20.86% | 1.907% | 1.444% | 7.562% | 53.02% | 98.9% | 14.54% | 189% | 74.14% | 192.9% | 22.74% | 42.8% | 0% | 29.19% | 18.87% | 1.869% | 0.8749% | 0.654% | 1.593% | 0.6699% | 0.573% | 28.88% | 20.14% | 5.459% | 2.9%
45 | 9_13_21_BB |  | 6/16/2023 4:41:52 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_BB.694 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.822% | 1.931e-06% | 191.4% | 1.924% | 82.45% | 8.446% | 9.066% | 83.17% | 53% | 33.35% | 1.429% | 16.77% | 31.72% | 6.376% | 29.15% | 11.1% | 5.14% | 25.39% | 50.32% | 385.4% | 4.301% | 134.6% | 0% | 102.3% | 83.81% | 14.69% | 45.27% | 15.64% | 17.56% | 0.2672% | 0.4443% | 0.6803% | 1.461% | 0.6236% | 1.267% | 37.97% | 17.97% | 8.012% | 7.975%
46 | 9_13_21_FB |  | 6/16/2023 4:47:52 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_FB.695 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4.808% | 15.7% | 3.66% | 1.694% | 19.77% | 12.88% | 4.71% | 29.05% | 13.6% | 9.245% | 5.703% | 1.592% | 21.39% | 42.83% | 6.679% | 6.963% | 2.536% | 0.525% | 121% | 59.14% | 8.143% | 80.95% | 56.99% | 37.38% | 92.27% | 1165% | 61.9% | 4.441% | 3.593% | 0.5812% | 0.8864% | 0.3722% | 0.5201% | 2.576% | 4.93% | 16.71% | 24.48% | 1.614% | 2.339%
47 | 5_18_23_MFB |  | 6/16/2023 4:53:52 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_MFB.696 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.286% | 2.93% | 3.718% | 1.313% | 3.604% | 1.218% | 6.649% | 0.5871% | 2.33% | 1.457% | 2.09% | 1.25% | 1.112% | 0.6851% | 1.631% | 1.764% | 0.5915% | 0.7894% | 11.97% | 2.389% | 1.115% | 19.67% | 4.206% | 4.582% | 1.843% | 14.33% | 4.224% | 0.5199% | 0.7296% | 10.33% | 0.4279% | 0.3683% | 0.6024% | 0.7247% | 0.5646% | 0.9068% | 0.486% | 13.45% | 1.001%
48 | 5_18_23_DB |  | 6/16/2023 4:59:52 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_DB.697 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.8861% | 33.37% | 13.63% | 1.225% | 15.91% | 3.749% | 8.076% | 38.62% | 40.5% | 9.533% | 5.718% | 21.52% | 89.09% | 10.9% | 71.32% | 7.596% | 3.9% | 9.933% | 91.43% | 173.9% | 17.21% | 225.8% | 106.8% | 36.56% | 21.36% | 1158% | 7187% | 29.56% | 1.479% | 1.234% | 1.118% | 1.078% | 1.644% | 0.7032% | 0.7146% | 9.723% | 18.49% | 5.285% | 5.689%
49 | 5_18_23_T_239 |  | 6/16/2023 5:05:52 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_T_239.698 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.128% | 3.192% | 3.592% | 1.376% | 17.85% | 4.395% | 3.344% | 3.954% | 7.655% | 2.36% | 1.708% | 0.849% | 3.255% | 3.964% | 1.369% | 4.11% | 2.639% | 0.9069% | 9.651% | 9.052% | 1.652% | 11.9% | 9.044% | 5.019% | 1.809% | 9.944% | 6.29% | 2.993% | 0.8827% | 6.404% | 0.5446% | 1.433% | 1.146% | 0.595% | 2.293% | 0.9836% | 2.158% | 8.366% | 2.481%
50 | 5_18_23_T_239 |  | 6/16/2023 5:11:52 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_T_239.699 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4.969% | 4.422% | 6.031% | 1.089% | 10.51% | 5.222% | 5.267% | 7.893% | 3.951% | 2.45% | 0.7432% | 1.084% | 6.149% | 7.738% | 5.331% | 6.973% | 1.727% | 0.656% | 2.103% | 10.71% | 0.9178% | 1.866% | 11.88% | 6.081% | 4.122% | 5.775% | 14.57% | 1.901% | 3.266% | 1.134% | 1.015% | 0.3858% | 0.5237% | 0.1556% | 4.989% | 0.503% | 0.6953% | 0.3787% | 0.6957%
51 | 5_18_23_M_224 |  | 6/16/2023 5:17:52 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_M_224.700 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.323% | 7.993% | 0.8617% | 1.556% | 5.196% | 4.002% | 3.179% | 2.615% | 8.033% | 10.58% | 2.695% | 1.373% | 73.13% | 35.58% | 29.32% | 10.61% | 6.93% | 0.9961% | 15.35% | 89.81% | 12.19% | 54.83% | 14.39% | 98.17% | 31.61% | 5.931% | 40.29% | 0.9332% | 2.097% | 0.8471% | 0.2931% | 0.3809% | 2.088% | 0.4269% | 1.043% | 11.47% | 138.2% | 26.24% | 6.065%
52 | LBLANK |  | 6/16/2023 5:23:52 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.701 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.183% | 115.5% | 3.344% | 12.97% | 132.1% | 0.5142% | 4.855% | 95.17% | 35.61% | 69.45% | 1.623% | 36.43% | 24.41% | 8.682% | 55.23% | 32.75% | 12% | 1.693% | 331.9% | 75.9% | 523.3% | 62.32% | 0% | 101.8% | 33.66% | 53.53% | 0% | 10.92% | 13.74% | 7.484% | 1.298% | 3.372% | 1.101% | 0.909% | 1.297% | 0% | 8.047% | 2.579% | 1.69%
53 | CCV 1ppb TRACE |  | 6/16/2023 5:29:53 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.702 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.189% | 1.846% | 17.05% | 22.45% | 8.339% | 1.117% | 5.192% | 5.702% | 2.625% | 1.227% | 5.785% | 2.453% | 6.339% | 0.7688% | 10.51% | 2.307% | 1.999% | 2.043% | 0.6475% | 3.817% | 1.091% | 5.197% | 16.56% | 6.575% | 2.721% | 7.868% | 12.5% | 1.022% | 0.4834% | 0.6882% | 0.8524% | 0.7412% | 1.292% | 1.373% | 0.2669% | 1.603% | 1.888% | 1.652% | 1.242%
54 | CCV 1ppb TRACE |  | 6/16/2023 5:36:12 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.703 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.032% | 1.162% | 34.99% | 16.5% | 11.46% | 1.632% | 7.228% | 3.555% | 5.405% | 3.113% | 2.821% | 0.9124% | 5.243% | 2.67% | 4.247% | 1.102% | 2.056% | 2.961% | 0.2777% | 16.81% | 1.457% | 20.36% | 4.645% | 5.793% | 7.429% | 7.346% | 14.4% | 2.296% | 2.98% | 1.197% | 0.7841% | 0.5756% | 0.9823% | 0.9679% | 2.215% | 1.761% | 1.767% | 1.644% | 0.3649%
55 | LFB |  | 6/16/2023 5:42:12 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.704 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.369% | 2.521% | 2.386% | 0.7684% | 3.189% | 1.411% | 1.501% | 1.374% | 0.5744% | 0.582% | 0.7025% | 1.102% | 1.538% | 0.5445% | 0.9103% | 0.6605% | 0.7689% | 0.6492% | 1.751% | 3.019% | 0.9623% | 4.602% | 3.837% | 3.355% | 0.3693% | 4.766% | 3.396% | 0.8523% | 0.4636% | 0.8486% | 0.7544% | 0.7418% | 0.4742% | 1.207% | 1.048% | 1.629% | 1.534% | 0.7322% | 1.589%
56 | CCV 10 ppb MAJORS |  | 6/16/2023 5:48:11 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.705 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.9674% | 0.7618% | 3.171% | 2.339% | 0.5388% | 1.966% | 2.246% | 1.596% | 1.96% | 0.8779% | 1.475% | 0.1265% | 0.8655% | 1.808% | 1.36% | 1.066% | 0.6891% | 0.1179% | 1.27% | 1.529% | 0.6102% | 2.065% | 2.425% | 1.878% | 1.028% | 2.458% | 2.368% | 1.008% | 0.314% | 1.081% | 0.4757% | 0.6096% | 0.7384% | 0.7501% | 1.019% | 0.3958% | 0.655% | 1.588% | 2.124%
57 | CCV 10 ppb MAJORS |  | 6/16/2023 5:54:30 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.706 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.143% | 2.367% | 4.505% | 0.8531% | 1.112% | 1.116% | 2.768% | 0.6755% | 1.611% | 1.134% | 1.451% | 2.294% | 1.667% | 1.117% | 2.696% | 0.9469% | 0.4459% | 0.6287% | 2.948% | 0.2651% | 1.583% | 3.728% | 5.637% | 3.164% | 1.791% | 1.494% | 3.231% | 0.893% | 1.623% | 1.002% | 1.19% | 0.2564% | 1.13% | 1.691% | 1.553% | 1.033% | 0.42% | 2.247% | 1.692%
58 | 5_18_23_MFB |  | 6/16/2023 6:00:31 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_MFB.707 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.9222% | 2.941% | 2.541% | 0.6466% | 1.922% | 2.418% | 0.5558% | 1.912% | 3.355% | 13.23% | 14.48% | 0.3013% | 2.41% | 2.534% | 3.843% | 1.819% | 12.98% | 13.8% | 1.064% | 10.93% | 1.906% | 1.507% | 3.189% | 2.358% | 2.014% | 3.991% | 8.15% | 12.74% | 0.3503% | 1.004% | 0.5323% | 10.78% | 1.046% | 1.368% | 1.773% | 13.1% | 13.48% | 0.55% | 0.3545%
59 | 10_1_21_W_1 |  | 6/16/2023 6:06:31 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_W_1.708 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4.592% | 2.694% | 6.581% | 2.689% | 0.6843% | 6.075% | 1.198% | 17.45% | 29.85% | 8.849% | 1.182% | 1.201% | 4.494% | 110.7% | 4.375% | 1.847% | 2.381% | 8.934% | 3.113% | 3.799% | 1.53% | 101.4% | 76.74% | 82.41% | 280.6% | 34.07% | 506.3% | 0.3415% | 1.241% | 0.8712% | 1.276% | 0.6219% | 0.2237% | 0.9858% | 2.059% | 27.67% | 1802% | 5.144% | 8.714%
60 | 10_1_21_W_6 |  | 6/16/2023 6:12:31 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_W_6.709 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.258% | 1.092% | 5.285% | 1.774% | 2.283% | 3.997% | 1.652% | 5.578% | 35.63% | 18.43% | 3.44% | 0.576% | 2.95% | 70.43% | 18.31% | 4.494% | 4.623% | 8.456% | 5.081% | 17.62% | 1.79% | 303.3% | 86.48% | 22.71% | 56.53% | 55.16% | 617.8% | 3.723% | 0.5652% | 0.9056% | 0.8356% | 2.591% | 1.434% | 1.621% | 0.7154% | 6.791% | 40.27% | 4.126% | 5.615%
61 | 10_1_21_W_8 |  | 6/16/2023 6:18:32 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_W_8.710 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.302% | 4.253% | 5.175% | 1.832% | 6.674% | 2.682% | 0.8373% | 15.48% | 42.62% | 4.606% | 1.153% | 0.3948% | 0.8411% | 18.4% | 33.85% | 2.746% | 1.353% | 2.437% | 8.258% | 8.202% | 1.005% | 270.4% | 57.59% | 53.7% | 80.92% | 105.5% | 202.8% | 0.5921% | 1.471% | 1.643% | 0.409% | 0.09719% | 0.5937% | 0.6828% | 1.014% | 25.07% | 1261% | 13.2% | 4.588%
62 | 10_1_21_W_9 |  | 6/16/2023 6:24:32 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_W_9.711 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.24% | 1.63% | 5.625% | 3.144% | 1.869% | 2.174% | 0.5162% | 14.97% | 13.41% | 5.445% | 2.45% | 0.8102% | 4.884% | 26.08% | 12.99% | 6.545% | 0.7599% | 7.025% | 10.61% | 9.235% | 2.389% | 4.268% | 57.98% | 59.29% | 45.31% | 62.06% | 979.7% | 1.654% | 1.3% | 0.8161% | 0.1883% | 1.016% | 0.9502% | 1.183% | 1.215% | 45.49% | 150.8% | 6.374% | 5.025%
63 | 10_1_21_W_10 |  | 6/16/2023 6:30:32 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_W_10.712 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.05% | 0.409% | 3.403% | 0.3176% | 3.888% | 12.37% | 0.4439% | 7.265% | 3.289% | 11.73% | 1.455% | 1.02% | 1.197% | 40.9% | 17.83% | 6.786% | 0.4859% | 4.37% | 5.39% | 4.339% | 1.936% | 20.59% | 58.24% | 21.38% | 207% | 123.7% | 211.8% | 2.764% | 1.838% | 1.056% | 1.327% | 1.253% | 0.1956% | 1.093% | 0.7761% | 29.68% | 22.6% | 3.151% | 0.8914%
64 | 10_1_21_W_11 |  | 6/16/2023 6:36:31 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_W_11.713 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.349% | 2.909% | 3.481% | 0.8718% | 5.596% | 3.867% | 4.114% | 7.037% | 11.05% | 0.6688% | 0.3902% | 0.4668% | 1.001% | 23.36% | 19.07% | 5.762% | 1.41% | 4.407% | 6.795% | 13.8% | 1.549% | 18.51% | 91.92% | 45.19% | 140% | 29028% | 138.6% | 1.613% | 1.822% | 0.8303% | 0.9175% | 0.5402% | 0.6778% | 0.5215% | 1.719% | 25.16% | 36.6% | 2.372% | 1.478%
65 | 10_1_21_W_11_LD |  | 6/16/2023 6:42:31 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_W_11_LD.714 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.9121% | 1.931e-06% | 1.662% | 1.601% | 1.476% | 1.09% | 2.204% | 63.84% | 14.68% | 32.64% | 0.9969% | 31.73% | 4.976% | 5.231% | 19.99% | 4.221% | 11.23% | 3.317% | 2.221% | 177.3% | 24.79% | 1.912% | 0% | 0% | 19.62% | 87.03% | 0% | 25.98% | 10.42% | 3.855% | 2.406% | 1.592% | 7.767% | 3.211% | 5.794% | 25.65% | 10.06% | 1.26% | 1.668%
66 | 10_1_21_W_11_LD |  | 6/16/2023 6:48:48 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_W_11_LD.715 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.232% | 2.826% | 2.09% | 3.475% | 2.696% | 1.985% | 5.938% | 12.99% | 17.9% | 2.708% | 1.9% | 0.8981% | 1.595% | 79.6% | 25.74% | 2.225% | 2.664% | 1.259% | 2.128% | 10.49% | 2.768% | 47.37% | 76.92% | 51% | 5163% | 8.183% | 71.73% | 4.117% | 1.128% | 1.642% | 0.2073% | 0.7535% | 0.251% | 1.196% | 0.2052% | 35.53% | 33.11% | 3.959% | 1.815%
67 | 10_1_21_W_11_DUP |  | 6/16/2023 6:54:48 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_W_11_DUP.716 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.875% | 4.336% | 3.317% | 4.492% | 4.72% | 5.661% | 2.47% | 7.261% | 30.6% | 4.597% | 0.685% | 1.389% | 1.154% | 78.03% | 10.23% | 6.476% | 2.497% | 4.653% | 6.122% | 8.064% | 0.6606% | 859.3% | 77.27% | 51.43% | 54.9% | 11.5% | 556% | 2.717% | 2.137% | 0.7506% | 0.3305% | 0.0735% | 0.6337% | 1.047% | 1.247% | 17% | 18.48% | 0.4642% | 0.7058%
68 | LBLANK |  | 6/16/2023 7:00:48 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.717 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.6792% | 118.3% | 4.104% | 2.456% | 26.97% | 9.697% | 0% | 88.99% | 12.2% | 184.3% | 0.9397% | 16.49% | 18.62% | 3.249% | 22.69% | 167.8% | 14.22% | 1.155% | 52.34% | 1.777e-06% | 38.54% | 632.6% | 0% | 240.3% | 13.22% | 611.8% | 757.8% | 14.66% | 26.87% | 1.322% | 1.149% | 1.42% | 0.7713% | 0.2646% | 1.73% | 30.45% | 6.327% | 1.206% | 3.831%
69 | CCV 1ppb TRACE |  | 6/16/2023 7:06:48 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.718 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 5.941% | 4.715% | 12.57% | 19.5% | 7.534% | 1.829% | 5.893% | 3.17% | 2.048% | 4.374% | 1.637% | 0.3137% | 6.492% | 4.61% | 2.016% | 1.226% | 2.132% | 1.535% | 5.295% | 13.36% | 0.8629% | 11.5% | 17.95% | 2.058% | 6.593% | 10.53% | 24.01% | 1.853% | 1.619% | 0.2421% | 0.6822% | 0.3343% | 0.3424% | 1.82% | 0.8832% | 2.24% | 0.8264% | 1.26% | 1.286%
70 | CCV 1ppb TRACE |  | 6/16/2023 7:13:07 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.719 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.321% | 3.327% | 1.642% | 43.72% | 6.907% | 4.167% | 2.494% | 2.45% | 3.28% | 3.373% | 2.677% | 1.67% | 5.76% | 1.57% | 4.558% | 1.378% | 0.3724% | 2.059% | 2.452% | 12.25% | 1.414% | 9.303% | 19.12% | 6.749% | 8.405% | 11.85% | 14.62% | 2.291% | 0.941% | 0.07654% | 0.387% | 1.265% | 1.57% | 1.173% | 0.4386% | 1.74% | 0.7212% | 0.8518% | 2.101%
71 | CCV 10 ppb MAJORS |  | 6/16/2023 7:19:08 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.720 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.2673% | 1.216% | 2.956% | 1.233% | 2.981% | 1.18% | 1.701% | 0.8202% | 1.177% | 0.3951% | 0.4945% | 0.7408% | 2.243% | 1.373% | 1.642% | 1.687% | 0.3238% | 0.8813% | 1.695% | 3.015% | 0.5142% | 5.825% | 4.268% | 2.688% | 2.16% | 4.038% | 3.417% | 1.514% | 0.1612% | 0.9196% | 0.458% | 1.032% | 0.6228% | 1.176% | 0.2518% | 1.223% | 1.082% | 2.325% | 0.8533%
72 | CCV 10 ppb MAJORS |  | 6/16/2023 7:25:27 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.721 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.963% | 3.026% | 6.918% | 1.823% | 1.906% | 2.18% | 2.537% | 0.6022% | 2.094% | 1.828% | 1.666% | 0.7612% | 2.671% | 1.309% | 1.735% | 0.9542% | 0.3518% | 0.7785% | 1.726% | 1.137% | 0.9207% | 1.092% | 2.749% | 2.122% | 1.33% | 2.723% | 4.202% | 0.5111% | 1.202% | 0.5563% | 0.8365% | 0.739% | 1.099% | 0.7696% | 1.348% | 1.634% | 0.3777% | 2.443% | 0.9257%
73 | 10_1_21_W_11_DUP_LSP |  | 6/16/2023 7:31:27 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_W_11_DUP_LSP.722 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.395% | 1.279% | 2.057% | 1.331% | 0.8865% | 1.236% | 0.9673% | 1.695% | 0.3675% | 1.604% | 0.658% | 0.8618% | 0.3925% | 0.7276% | 0.9391% | 1.27% | 0.9038% | 0.6368% | 1.851% | 2.105% | 0.4514% | 1.081% | 3.615% | 3.033% | 2.281% | 2.79% | 0.4531% | 0.5078% | 0.9732% | 0.5045% | 0.874% | 0.3329% | 0.8878% | 0.4816% | 0.6117% | 0.2999% | 0.7454% | 1.213% | 1.903%
74 | 10_1_21_F_1 |  | 6/16/2023 7:37:27 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_F_1.723 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.581% | 0.9222% | 1.445% | 2.821% | 0.3314% | 3.335% | 1.933% | 3.085% | 10.16% | 4.942% | 0.5488% | 0.9396% | 1.867% | 2.646% | 5.746% | 5.95% | 1.131% | 0.7225% | 3.581% | 9.161% | 0.6837% | 28.14% | 72.53% | 46.34% | 55.51% | 1.587% | 39.16% | 1.515% | 0.9953% | 2.205% | 0.382% | 1.085% | 0.3892% | 0.6383% | 0.8666% | 22.79% | 9.843% | 3.164% | 3.973%
75 | 10_1_21_F_1_DUP |  | 6/16/2023 7:43:27 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_F_1_DUP.724 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.952% | 0.9928% | 5.932% | 1.107% | 1.745% | 4.242% | 1.862% | 3.938% | 9.397% | 4.903% | 2.084% | 1.085% | 3.532% | 17.61% | 7.206% | 2.526% | 1.384% | 0.6007% | 1.058% | 7.568% | 1.185% | 37.23% | 70.89% | 15.13% | 16.81% | 19.57% | 30% | 0.91% | 1.375% | 1.078% | 0.9386% | 0.4764% | 0.7746% | 0.9433% | 0.8129% | 16.14% | 33.86% | 2.591% | 1.269%
76 | 10_1_21_F_6 |  | 6/16/2023 7:49:26 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_F_6.725 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.5915% | 1.663% | 0.9312% | 3.212% | 0.7952% | 1.186% | 1.083% | 2.976% | 15.39% | 5.676% | 1.422% | 0.8357% | 0.3287% | 23.87% | 9.75% | 2.965% | 0.6066% | 0.9744% | 0.6817% | 7.252% | 1.12% | 64.07% | 40.53% | 11.22% | 36.14% | 6.804% | 67.22% | 1.805% | 1.802% | 0.4811% | 0.5873% | 1.195% | 0.7977% | 1.239% | 0.628% | 13.35% | 12.9% | 7.37% | 1.667%
77 | 10_1_21_F_6_DUP |  | 6/16/2023 7:55:26 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_F_6_DUP.726 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.161% | 3.065% | 14.19% | 1.369% | 1.806% | 2.294% | 1.328% | 7.138% | 11.03% | 4.361% | 0.6507% | 0.5288% | 0.5335% | 26.12% | 9.239% | 2.983% | 1.165% | 0.3657% | 1.562% | 4.408% | 0.8254% | 43.06% | 95.11% | 26.14% | 33.78% | 8.234% | 73.55% | 0.2191% | 0.6636% | 1.425% | 0.2286% | 0.498% | 0.5496% | 0.5948% | 0.9555% | 18.18% | 10.51% | 6.94% | 3.424%
78 | 10_1_21_F_8 |  | 6/16/2023 8:01:27 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_F_8.727 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.323% | 2.941% | 6.894% | 4.185% | 3.107% | 3.755% | 2.099% | 3.351% | 9.871% | 8.277% | 1.008% | 1.039% | 2.8% | 18.51% | 8.269% | 3.141% | 1.652% | 0.6771% | 1.947% | 3.038% | 2.2% | 12.18% | 50.88% | 16.56% | 13.38% | 16.97% | 29.47% | 1.898% | 0.6346% | 0.8553% | 0.4319% | 0.2814% | 0.9658% | 0.8528% | 2.239% | 42.59% | 4.444% | 3.193% | 1.346%
79 | 10_1_21_F_9 |  | 6/16/2023 8:07:27 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_F_9.728 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.085% | 0.9938% | 10.28% | 1.462% | 1.416% | 0.2691% | 2.542% | 5.792% | 9.799% | 4.411% | 0.5082% | 0.4159% | 1.508% | 22.22% | 11.79% | 2.281% | 0.225% | 1.471% | 2.493% | 7.321% | 2.52% | 7.546% | 30.75% | 6.506% | 23.86% | 6% | 46.66% | 0.5244% | 0.5451% | 1.202% | 0.5019% | 0.6072% | 0.4501% | 1.3% | 1.393% | 13.6% | 9.945% | 1.024% | 1.723%
80 | 10_1_21_F_9_DUP |  | 6/16/2023 8:13:27 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_F_9_DUP.729 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.686% | 1.338% | 4.899% | 4.592% | 1.369% | 4.696% | 1.092% | 3.466% | 11.53% | 3.291% | 1.278% | 0.07186% | 1.476% | 8.679% | 4.854% | 2.764% | 1.752% | 0.4832% | 1.001% | 7.72% | 0.1189% | 17.74% | 51.33% | 36.46% | 26.28% | 10.22% | 94.77% | 1.67% | 1.055% | 1.132% | 0.1082% | 0.6324% | 0.2692% | 1.257% | 0.7% | 16.33% | 4.682% | 3.794% | 3.317%
81 | 10_1_21_F_10 |  | 6/16/2023 8:19:27 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_F_10.730 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.039% | 3.074% | 1.873% | 3.369% | 0.8935% | 2.004% | 1.331% | 4.212% | 6.829% | 9.528% | 0.5782% | 0.8441% | 4.68% | 21.33% | 6.029% | 4.514% | 1.343% | 0.7602% | 0.6292% | 9.681% | 0.4016% | 10.92% | 37.98% | 30.81% | 27.53% | 10.7% | 78.43% | 0.7147% | 0.9909% | 0.8002% | 0.7045% | 0.2221% | 0.7774% | 0.7656% | 1.287% | 5.217% | 18.56% | 1.243% | 0.5704%
82 | 10_1_21_F_11 |  | 6/16/2023 8:25:27 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_F_11.731 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.9746% | 1.223% | 4.987% | 2.671% | 1.005% | 1.223% | 0.6206% | 3.48% | 7.432% | 1.691% | 0.594% | 0.6694% | 1.247% | 16.36% | 10.02% | 1.611% | 1.69% | 1.331% | 2.677% | 2.567% | 1.224% | 33.37% | 30.7% | 19.3% | 28.42% | 15.03% | 39.73% | 0.4962% | 0.09067% | 1.547% | 0.4645% | 1.041% | 0.7159% | 0.54% | 1.854% | 8.088% | 14.27% | 0.8534% | 3.365%
83 | LBLANK |  | 6/16/2023 8:31:27 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.732 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.701% | 117% | 3.149% | 1.54% | 26.02% | 7.512% | 3.07% | 101.2% | 21.47% | 104.7% | 0.1188% | 50.99% | 33.47% | 0% | 205% | 23.14% | 15.49% | 2.875% | 390.4% | 1.777e-06% | 11.8% | 4104% | 0% | 111.4% | 0% | 340.4% | 49.66% | 15.05% | 4.126% | 1.747% | 1.257% | 0.8401% | 0.8553% | 0.4416% | 1.881% | 0% | 5.86% | 5.514% | 4.42%
84 | CCV 1ppb TRACE |  | 6/16/2023 8:37:27 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.733 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 6.14% | 3.165% | 15.63% | 38.32% | 6.88% | 3.622% | 5.239% | 2.323% | 2.712% | 3.73% | 3.595% | 0.1425% | 7.556% | 0.8772% | 2.902% | 1.672% | 0.8743% | 4.562% | 9.709% | 7.038% | 1.523% | 8.726% | 3.234% | 5.331% | 1.788% | 11.69% | 5.065% | 1.08% | 0.6747% | 1.783% | 0.6754% | 0.5846% | 0.6097% | 1.136% | 0.8885% | 1.335% | 2.067% | 1.72% | 1.267%
85 | CCV 1ppb TRACE |  | 6/16/2023 8:43:46 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.734 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.327% | 2.188% | 11.68% | 31.67% | 4.662% | 2.774% | 7.857% | 5.006% | 2.265% | 2.541% | 0.6777% | 0.8635% | 0.7958% | 0.8717% | 0.4694% | 5.092% | 0.8842% | 1.767% | 7.92% | 5.988% | 2.199% | 20.66% | 6.747% | 10.87% | 2.698% | 16.82% | 10.44% | 1.116% | 1.894% | 2.358% | 0.458% | 0.7415% | 1.106% | 0.4806% | 1.06% | 0.5628% | 0.1755% | 3.368% | 0.7512%
86 | LFB |  | 6/16/2023 8:49:46 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.735 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.1591% | 1.854% | 6.106% | 0.8385% | 2.574% | 3.269% | 1.119% | 0.3248% | 0.5305% | 0.8769% | 2.719% | 1.788% | 0.9701% | 0.8488% | 0.1847% | 1.197% | 1.338% | 1.257% | 4.444% | 1.391% | 1.756% | 4.664% | 4.444% | 3.118% | 2.126% | 5.945% | 3.36% | 1.475% | 1.294% | 2.448% | 0.8675% | 0.9842% | 0.7748% | 0.3062% | 0.6025% | 1.642% | 1.409% | 1.869% | 0.3545%
87 | CCV 10 ppb MAJORS |  | 6/16/2023 8:55:46 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.736 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.986% | 2.077% | 5.835% | 0.2603% | 2.208% | 1.176% | 1.583% | 0.6887% | 0.9248% | 1.415% | 0.7486% | 0.7985% | 0.8694% | 0.579% | 0.9823% | 0.5914% | 0.9213% | 0.375% | 1.526% | 1.63% | 0.5155% | 5.455% | 2.69% | 2.557% | 1.252% | 3.559% | 3.112% | 1.465% | 0.8863% | 0.9861% | 0.4187% | 0.4929% | 1.466% | 0.1154% | 0.2436% | 0.4793% | 0.3063% | 0.6561% | 2.009%
88 | 10_1_21_UF_1 |  | 6/16/2023 9:01:46 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_UF_1.737 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.71% | 2.527% | 50.16% | 1.939% | 0.8867% | 5.19% | 0.6352% | 10.17% | 14.43% | 13.16% | 9.368% | 1.953% | 7.015% | 12.12% | 4.398% | 5.352% | 1.948% | 0.9156% | 2.813% | 5.508% | 1.967% | 19.95% | 80.46% | 12.89% | 32.78% | 2.637% | 85.78% | 0.3539% | 1.205% | 0.6437% | 0.8076% | 0.3625% | 0.9473% | 1.633% | 1.092% | 5.273% | 21.84% | 7.777% | 3.456%
89 | 10_1_21_UF_6 |  | 6/16/2023 9:07:46 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_UF_6.738 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.1% | 2.261% | 18.63% | 1.241% | 3.956% | 2.334% | 1.334% | 3.134% | 9.888% | 9.066% | 5.446% | 0.692% | 4.012% | 6.212% | 2.589% | 2.23% | 0.3896% | 2.539% | 5.133% | 0.7796% | 2.223% | 60.57% | 31.47% | 23.9% | 17.45% | 5.523% | 68.45% | 0.7366% | 1.616% | 3.668% | 0.9077% | 0.2122% | 0.3631% | 0.7168% | 1.548% | 21.04% | 13.23% | 5.472% | 5.218%
90 | 10_1_21_UF_8 |  | 6/16/2023 9:13:46 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_UF_8.739 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.928% | 2.327% | 39.04% | 0.9212% | 1.873% | 4.367% | 0.762% | 1.365% | 11.37% | 2.819% | 1.426% | 0.9919% | 1.529% | 16.84% | 4.434% | 3.293% | 0.3511% | 1.301% | 0.428% | 5.761% | 0.8106% | 7.117% | 106.5% | 19.97% | 14.74% | 25.56% | 36.07% | 0.1034% | 1.045% | 1.618% | 0.6212% | 0.3365% | 1.295% | 0.3257% | 1.388% | 17.86% | 18.38% | 11.21% | 2.41%
91 | 10_1_21_UF_8_LD |  | 6/16/2023 9:19:45 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_UF_8_LD.740 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.5223% | 1.959% | 3.664% | 0.5968% | 0.9695% | 1.494% | 1.916% | 6.937% | 4.11% | 1.638% | 5.06% | 0.9425% | 2.242% | 39.54% | 8.497% | 4.813% | 1.954% | 0.7836% | 2.241% | 5.632% | 1.23% | 14.24% | 60.6% | 26.88% | 24.44% | 14.71% | 38.95% | 0.8876% | 0.2117% | 0.1703% | 0.3394% | 0.2906% | 2.277% | 0.7436% | 0.7012% | 21.21% | 15.46% | 6.271% | 4.966%
92 | 10_1_21_UF_8_LD |  | 6/16/2023 9:26:03 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_UF_8_LD.741 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.195% | 2.301% | 20.68% | 2.148% | 1.267% | 2.295% | 0.4945% | 3.02% | 4.003% | 2.158% | 8.427% | 1.099% | 0.7913% | 10.39% | 3.567% | 2.252% | 2.269% | 2.207% | 1.532% | 6.611% | 2.177% | 15.9% | 70.25% | 28.87% | 23.3% | 8.819% | 83.41% | 2.673% | 0.8063% | 0.7377% | 0.7907% | 2.171% | 2.481% | 1.714% | 0.2614% | 15.17% | 14.31% | 4.573% | 4.846%
93 | 10_1_21_UF_8_DUP |  | 6/16/2023 9:32:02 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_UF_8_DUP.742 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.98% | 2.278% | 24.37% | 2.857% | 0.9305% | 3.6% | 1.208% | 5.807% | 1.512% | 1.718% | 2.644% | 0.9763% | 5.016% | 30.77% | 10.27% | 0.5015% | 1.522% | 1.651% | 1.838% | 4.634% | 0.71% | 47.4% | 37.52% | 24.66% | 18.91% | 6.081% | 49.77% | 1.861% | 0.3896% | 0.5614% | 0.3978% | 0.9743% | 1.165% | 1.111% | 0.7187% | 14.86% | 12.77% | 6.852% | 2.941%
94 | 10_1_21_UF_8_DUP_LSP |  | 6/16/2023 9:38:02 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_UF_8_DUP_LSP.743 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.164% | 1.184% | 5.2% | 1.503% | 0.8913% | 0.9713% | 2.757% | 0.885% | 1.587% | 1.898% | 1.593% | 0.4342% | 0.8565% | 1.963% | 1.444% | 1.759% | 0.9367% | 1.224% | 0.8762% | 3.934% | 0.8422% | 1.849% | 6.415% | 3.741% | 0.9899% | 1.749% | 2.092% | 0.5316% | 0.7295% | 0.5751% | 0.3661% | 0.8009% | 0.1717% | 1.2% | 0.8372% | 1.542% | 1.122% | 1.32% | 1.141%
95 | 10_1_21_UF_9 |  | 6/16/2023 9:44:02 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_UF_9.744 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.6236% | 2.372% | 116% | 1.389% | 1.428% | 2.668% | 1.607% | 1.891% | 7.433% | 13.71% | 4.77% | 2.162% | 3.018% | 11.55% | 2.911% | 4.574% | 1.138% | 1.002% | 1.554% | 3.71% | 2.933% | 35.8% | 28.59% | 25.28% | 35.64% | 6.574% | 23.5% | 0.494% | 1.63% | 1.32% | 0.8741% | 0.6536% | 0.8234% | 0.3847% | 1.231% | 7.92% | 21.91% | 8.884% | 8.355%
96 | 10_1_21_UF_10 |  | 6/16/2023 9:50:03 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_UF_10.745 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.503% | 0.9931% | 843.9% | 1.931% | 2.026% | 3.257% | 1.154% | 5.726% | 6.432% | 4.087% | 2.46% | 2.168% | 1.719% | 29.03% | 21.84% | 5.171% | 0.8627% | 1.176% | 5.948% | 5.607% | 1.505% | 46.53% | 12.04% | 34.33% | 32.34% | 10.52% | 32.22% | 0.9214% | 1.823% | 3.721% | 1.277% | 0.9107% | 1.123% | 0.3988% | 0.4589% | 37.76% | 22.97% | 12.23% | 1.818%
97 | 10_1_21_UF_11 |  | 6/16/2023 9:56:03 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_UF_11.746 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.275% | 0.3752% | 249.7% | 2.238% | 1.96% | 3.87% | 2.292% | 3.149% | 15.34% | 14.73% | 1.356% | 0.6564% | 3.73% | 10.8% | 5.32% | 2.513% | 0.788% | 0.4783% | 6.311% | 8.003% | 1.194% | 29.46% | 42.6% | 7.222% | 13.6% | 25.14% | 21.34% | 0.3978% | 1.03% | 4.674% | 0.5039% | 0.4325% | 1.185% | 0.9599% | 1.006% | 20.94% | 14.61% | 0.2721% | 2.581%
98 | LBLANK |  | 6/16/2023 10:02:04 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.747 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.77% | 77.66% | 9.079% | 1.213% | 28.78% | 7.243% | 12.03% | 279.7% | 22.14% | 56.57% | 1.343% | 24.06% | 4.925% | 6.85% | 78.24% | 3.788% | 7.414% | 1.42% | 109% | 345.2% | 11.02% | 1911% | 89.14% | 37.09% | 19.87% | 66.26% | 0% | 12.23% | 33.08% | 2.235% | 0.6251% | 0.5635% | 0.4907% | 0.4222% | 0.8836% | 40.07% | 24.76% | 0.4127% | 1.166%
99 | CCV 1ppb TRACE |  | 6/16/2023 10:08:04 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.748 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.726% | 4.301% | 15.26% | 43.26% | 3.91% | 1.492% | 1.496% | 5.865% | 3.372% | 1.553% | 5.099% | 1.877% | 12.93% | 3.515% | 4.984% | 1.382% | 0.9038% | 0.7363% | 5.823% | 7.787% | 0.4798% | 9.298% | 22.81% | 1.933% | 5.627% | 9.493% | 21.61% | 1.793% | 0.82% | 1.193% | 0.5393% | 0.7116% | 0.9686% | 1.285% | 1.066% | 1.569% | 0.7933% | 1.766% | 0.5941%
100 | CCV 1ppb TRACE |  | 6/16/2023 10:14:23 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.749 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.545% | 2.481% | 18.89% | 52.26% | 5.14% | 1.176% | 7.358% | 5.01% | 1.535% | 0.4907% | 4.858% | 2.408% | 5.69% | 0.2907% | 5.68% | 0.7116% | 1.306% | 1.07% | 7.694% | 11.91% | 1.921% | 6.745% | 8.417% | 2.647% | 3.983% | 7.341% | 11.56% | 0.7815% | 0.1719% | 0.8387% | 0.3328% | 0.6684% | 3.154% | 0.9213% | 1.786% | 1.949% | 0.1715% | 0.7089% | 3.317%
101 | CCV 10 ppb MAJORS |  | 6/16/2023 10:20:23 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.750 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.429% | 3.628% | 11.06% | 0.309% | 1.244% | 2.828% | 2.14% | 3.612% | 1.086% | 0.1597% | 2.146% | 1.286% | 3.59% | 2% | 0.978% | 2.29% | 0.3346% | 1.08% | 0.8617% | 5.029% | 0.7% | 0.6303% | 2.068% | 2.901% | 2.523% | 1.705% | 4.471% | 0.6719% | 1.543% | 0.3927% | 0.8377% | 0.3682% | 0.5189% | 0.7274% | 1.204% | 0.882% | 0.3755% | 1.594% | 0.8672%
102 | 10_1_21_BB |  | 6/16/2023 10:26:24 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_BB.751 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.309% | 139.8% | 56.31% | 3.991% | 21.75% | 3.863% | 4.856% | 937.9% | 46.12% | 10.58% | 12.75% | 18.57% | 18.37% | 17.44% | 151.9% | 2.426% | 8.286% | 21.05% | 90.03% | 179.6% | 17.58% | 85.89% | 89.49% | 123.2% | 49.07% | 484% | 53.34% | 45.46% | 8.111% | 1.103% | 0.5277% | 1.434% | 1.109% | 0.3852% | 1.471% | 40.7% | 5.734% | 3.199% | 3.657%
103 | 10_1_21_FB |  | 6/16/2023 10:32:24 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_FB.752 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 6.214% | 70.5% | 2.743% | 4.175% | 20.7% | 2.325% | 19.65% | 81.41% | 24.45% | 0.4203% | 3.809% | 3.667% | 9.936% | 26.94% | 15.34% | 7.915% | 5.88% | 2.242% | 79.78% | 341.9% | 17.73% | 265% | 90.51% | 110% | 72.29% | 303% | 0% | 233.3% | 19.06% | 1.537% | 0.631% | 0.6416% | 0.3558% | 0.3348% | 1.151% | 40.21% | 47.76% | 9.517% | 7.391%
104 | 10_1_21_LAB_A |  | 6/16/2023 10:38:24 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_LAB_A.753 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.73% | 1.931e-06% | 74.13% | 3.469% | 24.56% | 12.78% | 8.265% | 297% | 12.37% | 21.67% | 24.5% | 12.53% | 45.4% | 11.47% | 68.39% | 14.53% | 2.8% | 3.881% | 96.69% | 18242% | 76.19% | 1178% | 91.06% | 104.4% | 49.79% | 511.3% | 54.79% | 4.483% | 7.482% | 0.9147% | 0.2723% | 0.6468% | 0.4915% | 2.048% | 0.8414% | 0% | 21.82% | 8.795% | 1.28%
105 | 10_1_21_LAB_B |  | 6/16/2023 10:44:24 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_LAB_B.754 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.972% | 119.3% | 17.64% | 5.656% | 23.09% | 6.028% | 2.878% | 314.8% | 20.77% | 3.162% | 7.639% | 17.39% | 149.2% | 6.274% | 98.44% | 9.285% | 1.778% | 8.835% | 24.67% | 87.69% | 229.4% | 59.87% | 0% | 234.4% | 42.84% | 73.71% | 55.54% | 6.873% | 14.27% | 1.494% | 0.3412% | 0.3601% | 0.1852% | 1.299% | 2.03% | 31.69% | 22.75% | 3.486% | 1.369%
106 | 5_18_23_DB |  | 6/16/2023 10:50:24 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_DB.755 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.8144% | 0.7885% | 1.035% | 1.158% | 16.16% | 1.369% | 0.7182% | 2.035% | 6.131% | 0.9507% | 0.832% | 1.584% | 0.7303% | 1.206% | 1.794% | 6.246% | 1.433% | 0.7074% | 2.533% | 10.05% | 1.036% | 5.246% | 8.458% | 6.659% | 5.733% | 12.96% | 3.879% | 1.533% | 0.6418% | 1.37% | 0.3016% | 0.7725% | 0.9521% | 1.032% | 0.8361% | 2.014% | 1.66% | 2.553% | 0.24%
107 | 5_18_23_DB |  | 6/16/2023 10:56:42 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_DB.756 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.388% | 0.6112% | 1.091% | 1.281% | 11.15% | 2.966% | 2.118% | 1.2% | 2.905% | 0.566% | 1.131% | 1.617% | 1.741% | 1.131% | 2.027% | 2.118% | 1.804% | 0.5923% | 5.558% | 4.089% | 1.392% | 2.326% | 16.55% | 8.355% | 1.628% | 4.035% | 0.8809% | 0.5967% | 1.6% | 0.2917% | 0.6116% | 0.6478% | 0.4366% | 0.3877% | 2.18% | 1.603% | 0.4922% | 1.905% | 0.7208%
108 | 5_18_23_T_239 |  | 6/16/2023 11:02:42 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_T_239.757 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.703% | 1.907% | 1.187% | 1.371% | 14.72% | 0.339% | 0.7527% | 2.13% | 1.178% | 1.061% | 1.007% | 0.8387% | 1.57% | 2.339% | 0.9694% | 2.035% | 2.806% | 1.704% | 1.202% | 5.975% | 0.03144% | 1.402% | 3.753% | 1.617% | 4.667% | 6.882% | 16.25% | 0.8764% | 0.9302% | 0.3745% | 0.618% | 0.6688% | 0.9669% | 0.4168% | 0.5022% | 0.8258% | 1.934% | 1.459% | 1.801%
109 | 5_18_23_T_239 |  | 6/16/2023 11:09:00 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_T_239.758 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.4131% | 1.535% | 0.2866% | 0.6331% | 9.522% | 3.363% | 1.463% | 2.404% | 2.238% | 1.682% | 0.1375% | 0.5156% | 0.7733% | 2.193% | 1.64% | 1.236% | 2.467% | 0.9773% | 3.093% | 11.62% | 1.018% | 3.955% | 10.27% | 8.25% | 5.833% | 1.681% | 2.183% | 2.47% | 0.6734% | 0.8346% | 1.095% | 0.3458% | 1.506% | 0.8793% | 0.7705% | 0.4858% | 0.6729% | 0.6073% | 2.037%
110 | 5_18_23_T_239 |  | 6/16/2023 11:15:00 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_T_239.759 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4.523% | 4.352% | 1.068% | 1.85% | 4.143% | 3.843% | 5.673% | 2.538% | 23.8% | 8.893% | 2.843% | 0.4812% | 231.9% | 4.428% | 25.47% | 19.4% | 5.832% | 3.964% | 20.88% | 70.91% | 16.84% | 14.38% | 61.17% | 34.65% | 115.9% | 3.216% | 237.2% | 1.044% | 3.283% | 0.7488% | 0.4336% | 0.58% | 4.1% | 1.307% | 3.672% | 49.16% | 325.2% | 17.54% | 8.779%
111 | 5_18_23_M_224 |  | 6/16/2023 11:20:59 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_M_224.760 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.7549% | 3.332% | 9.231% | 1.806% | 1.381% | 0.8738% | 0.959% | 1.443% | 1.039% | 0.5981% | 1.564% | 1.095% | 3.31% | 2.026% | 1.408% | 2.03% | 0.6717% | 1.233% | 8.734% | 0.9106% | 1.535% | 5.728% | 7.924% | 2.946% | 1.544% | 12.33% | 5.583% | 0.8525% | 0.7338% | 6.87% | 0.5318% | 0.1772% | 0.921% | 0.99% | 0.4839% | 1.599% | 0.2288% | 7.69% | 2.163%
112 | LBLANK |  | 6/16/2023 11:27:00 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.761 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.062% | 119.7% | 5.736% | 6.098% | 10.29% | 4.71% | 6.472% | 155.4% | 102.7% | 51.93% | 0.1276% | 603.3% | 9.155% | 16.63% | 38.83% | 33.96% | 11.36% | 1.59% | 196.2% | 1.777e-06% | 27.04% | 92.07% | 120.1% | 109.6% | 14.3% | 1036% | 163% | 8.19% | 9.743% | 0.3356% | 0.8375% | 0.8353% | 0.9035% | 0.728% | 1.007% | 31.54% | 4.779% | 3.318% | 0.432%
113 | LBLANK |  | 6/16/2023 11:33:18 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.762 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.215% | 119.3% | 0.9354% | 1.228% | 21.88% | 7.053% | 3.226% | 315.4% | 101.4% | 25.67% | 0.4751% | 217.3% | 7.884% | 7.343% | 39.84% | 6.272% | 2.917% | 4.883% | 11.75% | 90.23% | 31.01% | 176.7% | 0% | 198.6% | 14.56% | 42.38% | 0% | 33.8% | 18.04% | 1.577% | 0.4858% | 2.63% | 0.5922% | 1.335% | 0.9195% | 51.93% | 14.67% | 2.715% | 1.653%
114 | CCV 1ppb TRACE |  | 6/16/2023 11:39:19 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.763 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.484% | 2.505% | 20.68% | 7.479% | 7.951% | 5.76% | 2.49% | 1.702% | 3.253% | 4.81% | 3.456% | 0.3748% | 4.739% | 1.23% | 5.97% | 1.777% | 2.581% | 1.831% | 2.854% | 11.71% | 4.622% | 5.976% | 9.831% | 4.258% | 5.978% | 11.98% | 3.943% | 1.865% | 1.286% | 1.072% | 1.182% | 1.366% | 0.5358% | 0.6714% | 0.9509% | 1.363% | 1.446% | 2.634% | 0.9985%
115 | CCV 1ppb TRACE |  | 6/16/2023 11:45:38 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.764 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.799% | 5.46% | 16.76% | 27.88% | 8.49% | 5.15% | 11.6% | 2.833% | 3.202% | 0.8382% | 3.195% | 1.701% | 3.589% | 2.514% | 1.288% | 3.665% | 0.7634% | 2.123% | 3.028% | 19.6% | 2.576% | 8.16% | 32.22% | 4.627% | 5.183% | 4.943% | 21.7% | 1.669% | 0.9604% | 5.296% | 0.9594% | 0.7673% | 0.5262% | 0.5698% | 1.376% | 4.821% | 0.8933% | 6.068% | 0.8664%
116 | LFB |  | 6/16/2023 11:51:38 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.765 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.594% | 0.7547% | 11.11% | 1.733% | 1.505% | 1.78% | 2.146% | 2.31% | 2.211% | 1.168% | 2.72% | 1.044% | 1.817% | 1.172% | 1.045% | 0.9967% | 0.2913% | 0.8767% | 2.868% | 2.997% | 1.596% | 3.222% | 4.437% | 2.106% | 1.531% | 4.891% | 2.991% | 0.3969% | 0.2004% | 2.44% | 0.4442% | 0.8114% | 0.965% | 1.074% | 1.004% | 1.158% | 0.354% | 3.092% | 0.7532%
117 | LFB |  | 6/16/2023 11:57:55 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.766 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.365% | 2.48% | 5.128% | 0.4739% | 1.547% | 2.123% | 1.58% | 2.39% | 1.702% | 0.5856% | 0.8188% | 0.3211% | 1.497% | 0.5277% | 3.956% | 1.208% | 0.5225% | 0.2959% | 0.2409% | 4.096% | 0.6044% | 2.839% | 2.64% | 1.459% | 2.424% | 4.216% | 0.5887% | 1.449% | 0.6604% | 0.6494% | 0.9099% | 0.7759% | 0.3577% | 0.5826% | 1.404% | 0.8615% | 0.7795% | 1.529% | 1.334%
118 | CCV 10 ppb MAJORS |  | 6/17/2023 12:03:55 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.767 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.9527% | 0.9423% | 6.881% | 0.8498% | 3.458% | 1.298% | 1.486% | 0.9653% | 1.204% | 0.3609% | 1.36% | 0.3994% | 0.8611% | 1.662% | 1.682% | 2.276% | 0.1869% | 1.132% | 2.46% | 3.53% | 0.8935% | 2.816% | 3.105% | 0.5264% | 2.644% | 3.972% | 3.936% | 0.2169% | 0.9715% | 1.176% | 0.3713% | 0.6955% | 0.888% | 0.2153% | 0.2875% | 0.3733% | 0.7945% | 0.7553% | 1.26%
119 | 5_18_23_MFB |  | 6/17/2023 12:09:55 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_MFB.768 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.6512% | 1.931e-06% | 22.08% | 2.085% | 34.33% | 13.14% | 9.581% | 89.01% | 76.99% | 21.63% | 1.213% | 22.49% | 64.88% | 3.829% | 489.1% | 11.2% | 6.899% | 6.704% | 60.51% | 84.42% | 55.94% | 43.47% | 90.37% | 96.82% | 103.4% | 124.2% | 0% | 16.7% | 10.56% | 0.5741% | 0.343% | 0.7403% | 0.3774% | 0.4238% | 0.3288% | 53.22% | 31.73% | 12.44% | 6.374%
120 | 10_15_21_W_1 |  | 6/17/2023 12:15:55 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_W_1.769 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.1% | 3.326% | 8.539% | 1.612% | 2.398% | 5.711% | 1.567% | 9.391% | 29.98% | 1.872% | 1.23% | 0.6361% | 1.225% | 85.87% | 17.16% | 2.546% | 0.8928% | 7.556% | 2.197% | 9.112% | 3.268% | 63.66% | 71.56% | 58.8% | 132.8% | 60.66% | 93.22% | 1.639% | 1.449% | 0.902% | 0.4892% | 0.5454% | 0.07775% | 0.8708% | 1.325% | 19.17% | 1676% | 7.112% | 2.386%
121 | 10_15_21_W_6 |  | 6/17/2023 12:21:56 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_W_6.770 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.461% | 0.9226% | 3.687% | 3.56% | 5.132% | 0.1982% | 2.023% | 3.117% | 40.01% | 57.4% | 1.344% | 1.654% | 2.192% | 65.83% | 8.058% | 6.017% | 2.625% | 2.585% | 7.261% | 14.7% | 1.325% | 134.3% | 30.04% | 80.52% | 20.65% | 95.44% | 63.15% | 0.7703% | 1.545% | 1.057% | 1.391% | 1.007% | 1.223% | 0.1285% | 1.332% | 26.15% | 70.11% | 2.326% | 1.209%
122 | 10_15_21_W_8 |  | 6/17/2023 12:27:56 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_W_8.771 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.115% | 8.425% | 2.894% | 4.707% | 1.588% | 1.215% | 0.8374% | 7.312% | 9.373% | 4.887% | 2.827% | 1.137% | 0.9122% | 121.8% | 6.391% | 6.218% | 1.828% | 2.297% | 3.858% | 16.63% | 2.125% | 81.04% | 35.8% | 103% | 139.6% | 20.67% | 1.53% | 3.302% | 1.456% | 0.5462% | 0.5638% | 0.9321% | 0.6423% | 0.1904% | 0.6107% | 37.37% | 156.4% | 4.665% | 2.559%
123 | 10_15_21_W_9 |  | 6/17/2023 12:33:56 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_W_9.772 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.576% | 1.214% | 1.084% | 4.619% | 3.909% | 3.102% | 2.778% | 19.36% | 16.54% | 7.227% | 0.6629% | 2% | 2.815% | 138.8% | 11.63% | 4.619% | 3.004% | 7.523% | 7.955% | 16.27% | 2.315% | 59.78% | 61.42% | 0.7803% | 128.9% | 101.3% | 367.6% | 1.977% | 2.145% | 15% | 0.7616% | 0.54% | 0.269% | 0.9237% | 1.057% | 16.52% | 217.5% | 16.49% | 1.638%
124 | 10_15_21_W_9 |  | 6/17/2023 12:40:15 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_W_9.773 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.118% | 3.263% | 5.987% | 1.911% | 8.385% | 7.008% | 3.319% | 17.01% | 21.78% | 6.161% | 0.5776% | 2.051% | 1.102% | 77.39% | 10.45% | 1.16% | 0.8699% | 1.174% | 13.59% | 28.89% | 6.769% | 34.5% | 95.16% | 51.7% | 369.9% | 80.51% | 273.9% | 0.8042% | 1.829% | 0.2768% | 1.192% | 1.315% | 1.345% | 0.6338% | 1.182% | 34.83% | 302.7% | 1.345% | 2.457%
125 | 10_15_21_W_10 |  | 6/17/2023 12:46:15 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_W_10.774 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.143% | 2.868% | 0.9412% | 3.248% | 3.239% | 2.739% | 3.762% | 3.979% | 11.65% | 24.21% | 1.136% | 1.511% | 0.7344% | 23.09% | 15.33% | 3.807% | 2.392% | 1.247% | 2.751% | 16.55% | 4.952% | 64.04% | 36.27% | 42.46% | 92.36% | 67.01% | 63.75% | 2.451% | 0.6243% | 1.148% | 0.2914% | 0.4235% | 0.5708% | 0.6537% | 0.5292% | 43.44% | 19.17% | 2.126% | 2.221%
126 | 10_15_21_W_10 |  | 6/17/2023 12:52:33 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_W_10.775 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.625% | 2.843% | 1.779% | 3.298% | 4.221% | 6.841% | 2.782% | 15.92% | 23.93% | 14.77% | 0.2636% | 0.7195% | 0.7277% | 18.01% | 10.81% | 1.488% | 1.085% | 2.229% | 9.9% | 3.803% | 3.839% | 102.1% | 71.98% | 59.06% | 58.19% | 41.4% | 398.6% | 0.7919% | 2.675% | 0.7711% | 0.5148% | 0.3204% | 0.5286% | 0.2867% | 1.01% | 22.96% | 13.81% | 2.18% | 2.327%
127 | 10_15_21_W_10_DUP |  | 6/17/2023 12:58:33 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_W_10_DUP.776 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.87% | 1.47% | 0.5589% | 1.721% | 7.798% | 8.331% | 1.13% | 6.245% | 1.952% | 6.738% | 3.353% | 0.4858% | 2.8% | 23.05% | 28.81% | 1.332% | 2.474% | 1.431% | 12.74% | 19.14% | 1.558% | 31.48% | 0% | 73.3% | 665.5% | 18.96% | 695.4% | 1.778% | 2.27% | 0.5889% | 0.9428% | 1.56% | 0.8012% | 0.5468% | 1.441% | 4.268% | 26.21% | 0.6943% | 0.6915%
128 | 10_15_21_W_10_DUP |  | 6/17/2023 1:04:52 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_W_10_DUP.777 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4.556% | 2.151% | 4.333% | 1.396% | 2.657% | 3.472% | 1.834% | 8.36% | 14.02% | 6.172% | 2.494% | 1.456% | 1.536% | 55.61% | 29.96% | 1.496% | 1.285% | 0.5251% | 6.921% | 16.08% | 3.611% | 52.62% | 0% | 33.93% | 25.63% | 19.16% | 203% | 1.378% | 1.028% | 0.5219% | 0.3854% | 0.5241% | 1.382% | 0.4935% | 1.045% | 8.603% | 46.4% | 2.724% | 4.494%
129 | LBLANK |  | 6/17/2023 1:10:52 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.778 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.286% | 122.2% | 2.009% | 2.233% | 20.69% | 7.833% | 3.438% | 279.1% | 62.29% | 49.21% | 0.2181% | 87.19% | 13.57% | 3.73% | 2029% | 9.165% | 3.823% | 1.528% | 1079% | 1.777e-06% | 8.611% | 231.3% | 0% | 196.7% | 15.65% | 158.5% | 61.39% | 72.86% | 5.046% | 0.3605% | 0.9115% | 0.5611% | 0.5531% | 0.133% | 1.395% | 0% | 14.38% | 2.717% | 5.443%
130 | LBLANK |  | 6/17/2023 1:17:12 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.779 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.7685% | 1.931e-06% | 2.921% | 1.922% | 20.84% | 3.166% | 3.481% | 126.7% | 55.43% | 37.03% | 0.8333% | 115.8% | 18.18% | 6.733% | 382.7% | 25.43% | 5.618% | 1.297% | 75.05% | 1.777e-06% | 14.23% | 74.65% | 0% | 227.8% | 50.35% | 4464% | 0% | 17.54% | 4.898% | 2.965% | 0.5999% | 0.3541% | 0.7981% | 0.9525% | 1.186% | 33.61% | 3.816% | 0.3946% | 1.733%
131 | CCV 1ppb TRACE |  | 6/17/2023 1:23:12 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.780 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.137% | 9.593% | 13.29% | 44.97% | 12.72% | 6.057% | 6.493% | 2.452% | 3.847% | 2.886% | 0.814% | 1.999% | 1.415% | 0.9743% | 0.669% | 4.122% | 1.943% | 4.973% | 3.072% | 24.09% | 2.718% | 14.74% | 13.98% | 5.246% | 6.991% | 10.78% | 12.08% | 2.397% | 1.091% | 0.5928% | 0.4277% | 0.8454% | 0.6498% | 0.3416% | 0.6929% | 0.8193% | 2.562% | 0.6317% | 1.417%
132 | CCV 1ppb TRACE |  | 6/17/2023 1:29:31 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.781 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.535% | 4.614% | 29.2% | 42.78% | 13.69% | 4.022% | 6.054% | 0.864% | 2.484% | 1.587% | 2.173% | 1.235% | 7.545% | 3.224% | 7.949% | 5.348% | 1.037% | 0.9791% | 5.805% | 13.54% | 1.592% | 10.9% | 15.46% | 5.374% | 2.593% | 20.57% | 6.598% | 0.7189% | 1.084% | 0.7887% | 1.005% | 0.8591% | 0.5559% | 0.7531% | 1.07% | 1.165% | 0.5035% | 1.596% | 1.348%
133 | CCV 10 ppb MAJORS |  | 6/17/2023 1:35:32 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.782 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.002% | 2.445% | 7.702% | 1.204% | 2.307% | 1.067% | 2.269% | 1.838% | 1.383% | 0.4253% | 1.793% | 1.522% | 1.676% | 0.4678% | 2.687% | 2.348% | 0.4027% | 1.332% | 2.144% | 1.954% | 0.2037% | 2.164% | 4.884% | 4.564% | 1.115% | 0.4638% | 2.866% | 0.3259% | 0.4193% | 1.054% | 0.3696% | 1.044% | 0.2332% | 0.4327% | 0.5378% | 0.6582% | 0.7845% | 2.857% | 0.5708%
134 | CCV 10 ppb MAJORS |  | 6/17/2023 1:41:51 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.783 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.866% | 0.4532% | 4.581% | 0.8886% | 0.7322% | 0.2673% | 0.9153% | 0.3192% | 1.01% | 3.602% | 4.058% | 0.7115% | 0.669% | 0.8262% | 2.204% | 1.517% | 3.469% | 4.398% | 1.709% | 5.412% | 0.8446% | 2.946% | 4.203% | 3.207% | 1.788% | 3.753% | 5.236% | 3.818% | 1.194% | 0.593% | 0.9523% | 2.859% | 0.9469% | 0.4514% | 0.6567% | 5.009% | 4.018% | 0.774% | 1.25%
135 | 10_15_21_W_11 |  | 6/17/2023 1:47:51 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_W_11.784 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.404% | 3.088% | 2.403% | 0.4959% | 5.386% | 7.659% | 0.911% | 3.483% | 31.77% | 4.762% | 1.081% | 2.364% | 1.163% | 11.52% | 30.26% | 5.395% | 1.232% | 1.175% | 5.567% | 27.07% | 4.072% | 143.4% | 122.9% | 108.1% | 84.1% | 226.8% | 400.3% | 3.125% | 0.8717% | 2.839% | 0.7473% | 0.2272% | 0.6943% | 1.043% | 0.6431% | 27.49% | 9.276% | 5.949% | 0.8111%
136 | 10_15_21_W_11 |  | 6/17/2023 1:54:10 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_W_11.785 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.122% | 2.723% | 2.325% | 2.079% | 2.796% | 4.277% | 0.5009% | 3.441% | 6.511% | 4.177% | 1.216% | 1.086% | 0.2539% | 25.68% | 49.51% | 4.019% | 0.3721% | 1.18% | 5.135% | 13.97% | 0.9733% | 91.05% | 72.23% | 57.94% | 334.7% | 69.34% | 452.4% | 1.856% | 0.4357% | 0.6382% | 0.3568% | 0.8426% | 1.27% | 1.042% | 0.5285% | 15.51% | 24.87% | 2.895% | 0.7207%
137 | 10_15_21_F_1 |  | 6/17/2023 2:00:09 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_1.786 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.344% | 2.966% | 10.62% | 0.5722% | 1.513% | 1.136% | 0.06202% | 10.32% | 23.29% | 7.68% | 0.2374% | 0.4545% | 1.895% | 31.36% | 2.514% | 3.965% | 0.7786% | 0.63% | 1.225% | 1.71% | 1.974% | 7.267% | 24.07% | 9.657% | 6.5% | 18.07% | 100.7% | 0.9987% | 0.5171% | 0.5382% | 0.6462% | 0.6477% | 3.372% | 1.247% | 0.7562% | 41.2% | 16.63% | 0.9568% | 4.771%
138 | 10_15_21_F_1 |  | 6/17/2023 2:06:28 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_1.787 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.703% | 1.065% | 12.54% | 2.745% | 0.4802% | 2.536% | 2.415% | 2.998% | 17.07% | 2.383% | 0.5763% | 0.1365% | 0.9274% | 21.17% | 13.05% | 1.181% | 1.596% | 0.8654% | 2.184% | 1.297% | 0.6864% | 11.95% | 38.98% | 29.02% | 13.14% | 6.851% | 142.1% | 0.2832% | 0.8433% | 1.082% | 0.1814% | 0.2782% | 0.1876% | 1.107% | 0.8423% | 19.49% | 31.85% | 1.039% | 0.6199%
139 | 10_15_21_F_1_DUP |  | 6/17/2023 2:12:28 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_1_DUP.788 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.9813% | 1.758% | 23.35% | 1.354% | 0.8837% | 3.488% | 1.101% | 10.29% | 20.24% | 10.15% | 0.2304% | 1.154% | 1.472% | 8.91% | 1.611% | 4.219% | 1.952% | 1.996% | 1.485% | 7.813% | 0.3801% | 20.08% | 52.35% | 28.96% | 16.85% | 10.98% | 61.31% | 1.104% | 0.6435% | 0.8147% | 0.7142% | 0.5063% | 0.2845% | 0.3679% | 0.535% | 3.675% | 23.35% | 3.096% | 1.846%
140 | 10_15_21_F_1_DUP |  | 6/17/2023 2:18:46 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_1_DUP.789 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.848% | 2.069% | 23.22% | 3.607% | 2.219% | 1.754% | 2.242% | 3.99% | 17.84% | 9.168% | 2.261% | 0.4631% | 3.704% | 18.26% | 8.393% | 3% | 1.213% | 0.3319% | 3.188% | 5.917% | 0.9862% | 29.89% | 43.9% | 14.42% | 18.2% | 21.8% | 25% | 0.792% | 1.511% | 1.767% | 1.071% | 0.5785% | 0.5509% | 0.4484% | 1.734% | 41.54% | 20.2% | 4.694% | 1.743%
141 | 10_15_21_F_6 |  | 6/17/2023 2:24:45 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_6.790 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.5248% | 2.162% | 5.927% | 3.05% | 3.797% | 2.396% | 1.125% | 10.76% | 10.55% | 4.952% | 1.185% | 0.9176% | 2.258% | 13.82% | 4.624% | 2.985% | 2.253% | 0.9331% | 1.253% | 2.717% | 1.55% | 45.3% | 38.23% | 32.42% | 20.54% | 7.415% | 56.31% | 1.102% | 1.002% | 0.6614% | 0.647% | 1.256% | 4.023% | 0.812% | 1.467% | 23.21% | 5.649% | 2.852% | 9.872%
142 | 10_15_21_F_6 |  | 6/17/2023 2:31:03 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_6.791 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.4645% | 1.797% | 0.6857% | 1.649% | 0.8095% | 1.321% | 1.162% | 3.03% | 6.772% | 6.497% | 0.5189% | 1.107% | 0.4554% | 11.88% | 8.864% | 2.054% | 1.009% | 2.114% | 1.205% | 8.803% | 0.8962% | 16.1% | 19.03% | 42.67% | 26.32% | 5.948% | 28.12% | 0.32% | 0.6178% | 0.7932% | 0.2577% | 0.2251% | 0.2746% | 0.8431% | 0.5364% | 14.16% | 8.159% | 3.439% | 2.717%
143 | 10_15_21_F_6_LD |  | 6/17/2023 2:37:02 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_6_LD.792 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.864% | 2.66% | 1.264% | 3.454% | 1.216% | 2.969% | 1.841% | 0.5956% | 8.145% | 6.652% | 0.7759% | 1.48% | 2.198% | 23.42% | 1.437% | 0.872% | 0.9542% | 1.512% | 1.105% | 7.084% | 2.887% | 9.725% | 61.89% | 9.295% | 7.607% | 17.07% | 19.47% | 0.7401% | 1.539% | 0.7848% | 1.222% | 0.9164% | 0.7052% | 0.2116% | 1.097% | 18.23% | 2.317% | 4.248% | 1.924%
144 | 10_15_21_F_6_LD |  | 6/17/2023 2:43:20 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_6_LD.793 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.16% | 2.972% | 4.698% | 2.481% | 2.074% | 1.211% | 0.5957% | 1.76% | 14.91% | 3.631% | 5.187% | 0.7059% | 3.05% | 13.18% | 12.67% | 1.252% | 3.712% | 5.021% | 6.775% | 12.99% | 0.85% | 64.21% | 40.43% | 13.08% | 3.965% | 4.617% | 55.84% | 4.462% | 0.612% | 5.774% | 0.3839% | 3.872% | 0.1897% | 1.387% | 1.006% | 5.86% | 12.46% | 5.61% | 0.3456%
145 | LBLANK |  | 6/17/2023 2:49:20 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.794 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.518% | 1.931e-06% | 3.764% | 15.78% | 27.38% | 2.961% | 0.1102% | 929.9% | 9.466% | 136.6% | 1.026% | 95.71% | 13.06% | 3.702% | 157.4% | 65.56% | 6.247% | 2.438% | 82.25% | 214.4% | 28.91% | 344.1% | 0% | 108.8% | 16.62% | 101.4% | 59.8% | 7.99% | 16.41% | 5.908% | 0.6351% | 0.924% | 1.533% | 0.5081% | 1.933% | 0% | 10.26% | 1.533% | 1.823%
146 | CCV 1ppb TRACE |  | 6/17/2023 2:55:21 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.795 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.597% | 1.901% | 9.279% | 13.56% | 8.48% | 4.05% | 4.879% | 6.297% | 3.054% | 0.2875% | 3.888% | 1.544% | 7.57% | 2.562% | 3.201% | 0.7694% | 0.3752% | 2.309% | 3.518% | 9.339% | 2.141% | 9.73% | 10.9% | 13.41% | 2.576% | 17.03% | 2.496% | 1.661% | 1.076% | 1.788% | 0.1805% | 0.9617% | 0.7834% | 0.5332% | 0.4923% | 2.107% | 2.613% | 0.7597% | 1.314%
147 | CCV 1ppb TRACE |  | 6/17/2023 3:01:39 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.796 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.841% | 5.26% | 13.08% | 36.85% | 11.93% | 2.431% | 2.668% | 2.498% | 1.866% | 1.128% | 3.26% | 1.057% | 7.064% | 3.156% | 6.268% | 5.396% | 1.474% | 3.91% | 2.767% | 22.26% | 2.805% | 24.57% | 32.45% | 3.321% | 3.593% | 11.6% | 0.832% | 1.666% | 1.098% | 7.55% | 0.8115% | 0.5285% | 0.6279% | 0.6499% | 0.6774% | 0.6272% | 1.18% | 9.121% | 0.7657%
148 | LFB |  | 6/17/2023 3:07:39 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.797 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.164% | 1.353% | 9.975% | 0.7427% | 3.509% | 1.359% | 3.178% | 1.218% | 3.29% | 1.362% | 0.2341% | 0.3317% | 1.898% | 2.1% | 1.477% | 1.677% | 1.568% | 0.5384% | 1.627% | 1.982% | 1.311% | 3.945% | 6.429% | 1.691% | 3.432% | 4.106% | 3.154% | 0.7472% | 0.7683% | 1.999% | 1.025% | 0.5075% | 3.364% | 0.3256% | 1.357% | 1.39% | 0.6683% | 1.19% | 2.759%
149 | LFB |  | 6/17/2023 3:13:57 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.798 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.2334% | 2.946% | 3.196% | 0.7212% | 2.494% | 1.163% | 2.823% | 2.842% | 1.315% | 1.48% | 0.8276% | 1.551% | 1.226% | 1.584% | 2.099% | 1.827% | 0.6081% | 0.8279% | 1.708% | 3.987% | 0.3363% | 1.376% | 6.936% | 0.4469% | 0.3965% | 5.398% | 4.839% | 1.274% | 0.6709% | 0.6623% | 0.3639% | 0.5003% | 0.7924% | 0.8585% | 0.9356% | 1.862% | 1.167% | 1.193% | 1.674%
150 | CCV 10 ppb MAJORS |  | 6/17/2023 3:19:56 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.799 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.333% | 1.876% | 3.872% | 1.544% | 5.465% | 1.088% | 4.909% | 1.963% | 2.921% | 0.7549% | 0.9334% | 0.0879% | 1.681% | 3.266% | 1.874% | 2.25% | 0.5504% | 0.221% | 1.462% | 8.065% | 0.3989% | 1.753% | 9.072% | 2.095% | 2.063% | 2.149% | 1.081% | 0.3528% | 0.7392% | 0.2922% | 0.5046% | 0.4959% | 0.7286% | 0.1791% | 2.148% | 0.2161% | 0.9853% | 0.8047% | 1.446%
151 | CCV 10 ppb MAJORS |  | 6/17/2023 3:26:14 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.800 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.109% | 4.693% | 5.385% | 2.507% | 1.149% | 1.808% | 2.967% | 1.002% | 1.221% | 0.3388% | 1.731% | 1.1% | 1.744% | 3.286% | 2.689% | 2.031% | 0.204% | 1.486% | 3.749% | 1.685% | 1.079% | 6.684% | 5.421% | 2.696% | 3.142% | 6.179% | 0.6847% | 0.3929% | 0.6422% | 3.224% | 0.144% | 0.6114% | 0.323% | 0.3116% | 1.771% | 0.6004% | 1.088% | 4.757% | 0.6702%
152 | 10_15_21_F_6_DUP |  | 6/17/2023 3:32:14 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_6_DUP.801 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.425% | 1.751% | 3.5% | 1.255% | 1.344% | 5.187% | 1.174% | 7.516% | 7.753% | 10.08% | 0.8054% | 0.4112% | 0.4426% | 19.24% | 10.55% | 2.759% | 0.3559% | 0.7344% | 2.1% | 2.865% | 1.15% | 33.95% | 11.88% | 49.11% | 58.25% | 16.99% | 93.78% | 0.391% | 0.8335% | 1.558% | 0.292% | 0.5471% | 0.4767% | 0.8061% | 0.7152% | 1.541% | 10.72% | 1.957% | 1.563%
153 | 10_15_21_F_6_DUP |  | 6/17/2023 3:38:32 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_6_DUP.802 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.233% | 0.7876% | 7.093% | 2.521% | 2.016% | 0.5349% | 1.516% | 9.699% | 20.2% | 7.825% | 1.465% | 1.302% | 1.548% | 14.58% | 11.22% | 2.247% | 1.246% | 0.8059% | 3.039% | 8.414% | 0.6545% | 17.93% | 80.57% | 9.891% | 21.77% | 27.2% | 70.46% | 0.4461% | 2.105% | 2.836% | 1.254% | 0.4233% | 2.772% | 0.4706% | 0.6008% | 7.332% | 4.947% | 3.262% | 4.12%
154 | 10_15_21_F_6_DUP_LSP |  | 6/17/2023 3:44:32 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_6_DUP_LSP.803 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.099% | 2.515% | 1.038% | 1.354% | 0.2235% | 2.128% | 1.777% | 1.915% | 0.3929% | 0.5381% | 0.3757% | 0.3351% | 1.48% | 2.237% | 2.143% | 0.7129% | 1.089% | 0.6671% | 0.919% | 4.198% | 0.8583% | 2.519% | 0.1631% | 0.3215% | 2.501% | 2.8% | 7.324% | 0.6911% | 0.3114% | 0.8807% | 0.1632% | 0.4564% | 0.9774% | 1.092% | 1.435% | 0.3568% | 0.4268% | 1.302% | 1.237%
155 | 10_15_21_F_6_DUP_LSP |  | 6/17/2023 3:50:52 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_6_DUP_LSP.804 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4.357% | 3.896% | 4.9% | 0.7528% | 3.685% | 2% | 3.463% | 3.468% | 3.774% | 0.9826% | 1.383% | 0.7914% | 3.558% | 3.352% | 3.884% | 3.386% | 0.7585% | 1.72% | 1.173% | 6.843% | 1.194% | 3.297% | 5.93% | 3.002% | 1.916% | 1.602% | 5.337% | 0.1655% | 0.5199% | 0.6434% | 0.7323% | 0.6684% | 0.6743% | 0.5631% | 3.299% | 0.4971% | 1.167% | 1.337% | 1.424%
156 | 10_15_21_F_9 |  | 6/17/2023 3:56:52 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_9.805 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.431% | 1.434% | 5.073% | 1.009% | 2.775% | 2.071% | 2.234% | 4.207% | 7.485% | 5.77% | 0.7352% | 0.9241% | 0.9594% | 10.63% | 9.258% | 2.703% | 1.173% | 1.684% | 4.477% | 10.3% | 1.345% | 49.76% | 53.34% | 38.49% | 21.09% | 24.66% | 61.39% | 1.781% | 0.9163% | 1.257% | 0.7266% | 0.6754% | 0.232% | 0.7402% | 0.3964% | 13.94% | 18.74% | 4.923% | 1.716%
157 | 10_15_21_F_9 |  | 6/17/2023 4:03:12 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_9.806 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.8581% | 2.562% | 6.301% | 2.207% | 1.907% | 5.371% | 0.8316% | 3.627% | 6.469% | 6.99% | 1.096% | 0.3914% | 2.052% | 20.75% | 13.98% | 3.261% | 1.568% | 1.136% | 1.271% | 8.356% | 0.8498% | 4.403% | 31.14% | 19.89% | 27.96% | 14.97% | 14.42% | 1.017% | 1.706% | 0.6497% | 0.4503% | 0.2038% | 0.6271% | 0.5051% | 1.061% | 12.04% | 19.08% | 3.257% | 1.615%
158 | 10_15_21_F_10 |  | 6/17/2023 4:09:13 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_10.807 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.298% | 0.06311% | 4.801% | 0.9876% | 2.729% | 1.399% | 1.073% | 4.535% | 11.2% | 5.072% | 0.2581% | 5.352% | 2.786% | 27.69% | 1.137% | 2.369% | 1.544% | 1.737% | 1.683% | 7.367% | 7.03% | 22.99% | 32.72% | 29.37% | 27.01% | 5.791% | 193.1% | 0.9782% | 6.912% | 2.2% | 4.762% | 0.5503% | 0.9162% | 0.5534% | 0.5655% | 34.49% | 33.6% | 4.151% | 2.166%
159 | 10_15_21_F_10 |  | 6/17/2023 4:15:32 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_10.808 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.017% | 1.784% | 5.283% | 0.8593% | 1.37% | 5.366% | 2.386% | 3.792% | 12.2% | 14.94% | 18.68% | 1.388% | 3.241% | 60.87% | 8.627% | 4.543% | 19.3% | 17.19% | 2.053% | 2.355% | 2.395% | 24.4% | 45.56% | 18.19% | 16.56% | 3.947% | 68.29% | 18.08% | 1.681% | 0.1382% | 0.7617% | 13.17% | 0.7009% | 0.663% | 2.296% | 35.62% | 37.08% | 1.977% | 1.637%
160 | 10_15_21_F_10_DUP |  | 6/17/2023 4:21:32 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_10_DUP.809 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.544% | 2.954% | 7.45% | 0.4536% | 3.687% | 0.8091% | 1.133% | 6.352% | 15.93% | 4.766% | 0.7851% | 2.147% | 3.415% | 17.81% | 2.779% | 3.94% | 0.8292% | 1.488% | 1.405% | 10.23% | 2.494% | 24.01% | 76.27% | 43.36% | 33.53% | 13.75% | 198.6% | 0.5743% | 1.521% | 1.174% | 1.045% | 1.064% | 0.9918% | 0.909% | 0.6035% | 35.17% | 33.26% | 3.26% | 1.381%
161 | 10_15_21_F_10_DUP |  | 6/17/2023 4:27:51 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_10_DUP.810 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.179% | 1.729% | 10.49% | 1.24% | 0.1674% | 1.963% | 2.601% | 3.954% | 25.76% | 5.628% | 1.494% | 0.7548% | 3.059% | 67.37% | 6.946% | 2.249% | 1.701% | 2% | 1.67% | 4.884% | 0.8772% | 29.52% | 77.04% | 27.28% | 36.21% | 14.11% | 94.42% | 1.792% | 0.9033% | 2.058% | 0.6494% | 0.7846% | 0.9728% | 1.412% | 0.5801% | 25.69% | 37.26% | 3.461% | 0.805%
162 | LBLANK |  | 6/17/2023 4:33:51 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.811 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.39% | 123.9% | 1.948% | 0.4317% | 5.774% | 6.455% | 8.707% | 0% | 25.24% | 60.13% | 1.631% | 629.6% | 12.05% | 3.782% | 111.6% | 37.04% | 64.86% | 1.948% | 580.5% | 1.777e-06% | 21.66% | 175.3% | 96.43% | 121.1% | 23.31% | 143.4% | 0% | 12.61% | 27.46% | 1.042% | 0.6218% | 0.2723% | 1.28% | 0.4278% | 2.298% | 34.59% | 15.56% | 2.991% | 4.056%
163 | LBLANK |  | 6/17/2023 4:40:11 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.812 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.786% | 124.6% | 3.544% | 1.895% | 13.82% | 7.161% | 3.771% | 647.4% | 44.14% | 68.11% | 0.9672% | 22.79% | 13.93% | 4.015% | 224% | 28.66% | 14.33% | 0.6318% | 14.8% | 101.5% | 46.74% | 52.79% | 0% | 0% | 0% | 177.6% | 0% | 19.57% | 26.49% | 3.627% | 1.124% | 0.9879% | 0.4514% | 0.4492% | 2.821% | 0% | 6.423% | 0.3439% | 0.7476%
164 | CCV 1ppb TRACE |  | 6/17/2023 4:46:11 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.813 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.958% | 2.781% | 13.87% | 61.35% | 9.688% | 7.081% | 4.114% | 2.058% | 1.33% | 1.631% | 8.959% | 1.6% | 11.87% | 4.369% | 8.94% | 3.622% | 1.63% | 1.646% | 5.999% | 10.5% | 2.176% | 7.189% | 25.73% | 6.199% | 8.038% | 18.1% | 15.74% | 0.6848% | 1.214% | 1.316% | 0.5856% | 0.763% | 0.7265% | 0.7338% | 1.232% | 2.11% | 2.307% | 1.949% | 1.568%
165 | CCV 1ppb TRACE |  | 6/17/2023 4:52:31 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.814 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4.595% | 3.341% | 16.92% | 45.02% | 9.943% | 3.78% | 4.012% | 5.504% | 5.441% | 0.7732% | 3.573% | 1.057% | 12.73% | 3.602% | 6.475% | 5.072% | 1.491% | 1.633% | 3.002% | 20.13% | 1.219% | 26.62% | 14.33% | 9.811% | 5.054% | 29.44% | 5.042% | 1.026% | 1.573% | 7.182% | 0.6419% | 0.9373% | 0.7019% | 0.02218% | 2.09% | 2.427% | 1.26% | 10.79% | 1.481%
166 | CCV 10 ppb MAJORS |  | 6/17/2023 4:58:31 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.815 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.277% | 1.348% | 8.512% | 1.552% | 3.514% | 1.806% | 4.008% | 2.343% | 2.551% | 1.316% | 1.272% | 0.4426% | 1.211% | 0.2009% | 2.314% | 0.5512% | 1.106% | 0.8839% | 0.6866% | 3.939% | 0.3463% | 1.515% | 3.727% | 4.339% | 3.421% | 2.597% | 6.439% | 0.383% | 0.8137% | 0.9511% | 0.3233% | 0.7959% | 0.6926% | 0.3525% | 0.6979% | 0.6313% | 0.7514% | 1.769% | 0.996%
167 | CCV 10 ppb MAJORS |  | 6/17/2023 5:04:50 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.816 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.833% | 3.413% | 1.964% | 2.001% | 2.701% | 2.558% | 1.564% | 2.562% | 1.531% | 1.134% | 0.9041% | 0.7042% | 2.8% | 3.879% | 5.023% | 2.24% | 0.9808% | 1.264% | 1.439% | 4.418% | 1.043% | 4.172% | 8.022% | 0.6407% | 1.412% | 1.656% | 4.412% | 0.9118% | 0.7109% | 2.251% | 0.5143% | 0.8808% | 0.516% | 0.8302% | 2.247% | 0.9409% | 0.7783% | 3.307% | 0.4813%
168 | 10_15_21_F_11 |  | 6/17/2023 5:10:50 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_11.817 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.264% | 2.241% | 8.472% | 0.9493% | 3.843% | 2.161% | 1.936% | 4.223% | 24.92% | 10.78% | 2.3% | 0.5314% | 3.464% | 25.9% | 19.72% | 3.572% | 0.852% | 1.172% | 2.614% | 9.601% | 1.669% | 25.65% | 45.78% | 37.72% | 21.13% | 18.91% | 68.19% | 1.615% | 0.3512% | 0.07403% | 0.4769% | 0.5942% | 0.1247% | 0.2647% | 1.043% | 14.27% | 10.46% | 3.101% | 1.132%
169 | 10_15_21_F_11 |  | 6/17/2023 5:17:09 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_11.818 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.364% | 0.2476% | 3.732% | 4.178% | 2.286% | 2.602% | 1.869% | 2.379% | 14.05% | 5.327% | 0.8801% | 1.6% | 2.218% | 43.13% | 12.97% | 3.327% | 1.849% | 1.053% | 3.297% | 6.843% | 0.9334% | 28.93% | 21.62% | 44.28% | 36.68% | 8.871% | 24.31% | 1.547% | 0.8387% | 0.4129% | 0.9654% | 0.5893% | 3.357% | 0.7831% | 0.9304% | 17.49% | 4.917% | 3.104% | 6.176%
170 | 10_15_21_UF_1 |  | 6/17/2023 5:23:09 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_1.819 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.468% | 2.191% | 151.8% | 1.702% | 0.8603% | 1.83% | 0.7695% | 6.803% | 6.566% | 5.162% | 4.515% | 0.5548% | 3.472% | 20.68% | 10.35% | 1.178% | 2.676% | 1.314% | 1.926% | 10.19% | 0.9397% | 24.31% | 73.57% | 49.07% | 36.78% | 12.15% | 458.8% | 1.781% | 0.7251% | 0.9167% | 0.4293% | 0.508% | 1.344% | 0.3541% | 0.7725% | 19.55% | 23.58% | 4.942% | 9.292%
171 | 10_15_21_UF_1 |  | 6/17/2023 5:29:28 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_1.820 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.711% | 0.1893% | 252.8% | 1.995% | 0.256% | 1.55% | 0.8389% | 6.119% | 4.594% | 3.794% | 3.123% | 1.597% | 2.502% | 9.278% | 3.345% | 4.454% | 2.059% | 2.1% | 2.241% | 8% | 2.275% | 13.79% | 79.64% | 13.45% | 15.14% | 8.975% | 159.6% | 1.189% | 1.11% | 0.1832% | 1.218% | 0.2944% | 0.4764% | 0.9472% | 0.7133% | 11.64% | 29.17% | 7.853% | 8%
172 | 10_15_21_UF_6 |  | 6/17/2023 5:35:27 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_6.821 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.18% | 3.343% | 220% | 2.715% | 3.836% | 2.275% | 2.563% | 0.4401% | 7.59% | 13.67% | 4.688% | 0.4486% | 5.939% | 35.19% | 15.13% | 1.706% | 2.453% | 1.638% | 2.474% | 4.836% | 2.895% | 9.856% | 24.55% | 35.03% | 25.11% | 18.61% | 33.74% | 0.9581% | 0.8412% | 1.196% | 1.086% | 0.1733% | 0.6111% | 0.8935% | 1.225% | 12.25% | 6.596% | 6.095% | 3.257%
173 | 10_15_21_UF_6 |  | 6/17/2023 5:41:46 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_6.822 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.9649% | 0.5068% | 94.29% | 5.335% | 2.144% | 4.567% | 0.1558% | 5.623% | 13.46% | 7.026% | 3.218% | 0.1559% | 3.237% | 32.37% | 4.324% | 2.837% | 0.9331% | 0.6934% | 4.719% | 1.657% | 0.7632% | 34.99% | 0.2632% | 43.23% | 14.99% | 11.32% | 139.7% | 1.126% | 0.2292% | 2.371% | 0.2845% | 0.7375% | 0.683% | 0.8663% | 0.3766% | 7.524% | 9.38% | 2.142% | 1.53%
174 | 10_15_21_UF_8 |  | 6/17/2023 5:47:46 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_8.823 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.707% | 1.284% | 14.59% | 2.033% | 0.6255% | 0.3793% | 0.9042% | 1.534% | 3.423% | 10.93% | 1.639% | 0.3297% | 5.609% | 29.69% | 10.31% | 3.86% | 1.117% | 0.4475% | 2.24% | 17.3% | 1.084% | 28.86% | 32.74% | 36.62% | 31.44% | 12.22% | 64.8% | 0.9742% | 0.4214% | 0.1926% | 0.3072% | 0.5153% | 0.5438% | 0.6148% | 0.7879% | 12.39% | 6.719% | 2.118% | 4.462%
175 | 10_15_21_UF_8 |  | 6/17/2023 5:54:04 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_8.824 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.965% | 2.433% | 5.567% | 2.77% | 1.548% | 3.293% | 2.203% | 2.82% | 11.68% | 2.197% | 2.917% | 1.33% | 6.825% | 20.44% | 11.16% | 4.33% | 1.149% | 1.625% | 3.617% | 6.008% | 0.3015% | 32.8% | 39.38% | 48.49% | 15.63% | 5.22% | 24.76% | 0.3212% | 1.39% | 1.203% | 0.6943% | 0.8069% | 2.57% | 0.7964% | 1.393% | 16.43% | 29.46% | 5.711% | 2.792%
176 | 10_15_21_UF_8_LD |  | 6/17/2023 6:00:03 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_8_LD.825 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.289% | 2.606% | 7.792% | 2.549% | 2.287% | 4.334% | 1.665% | 3.305% | 10.14% | 3.466% | 2.688% | 7.151% | 5.24% | 5.241% | 5.269% | 3.371% | 1.97% | 0.2286% | 0.6498% | 4.862% | 5.222% | 11.61% | 33.64% | 48.29% | 16.47% | 11.86% | 23.92% | 0.8487% | 7.331% | 0.5062% | 5.726% | 0.41% | 0.1099% | 0.5577% | 1.264% | 13.76% | 17.1% | 1.271% | 2.783%
177 | 10_15_21_UF_8_LD |  | 6/17/2023 6:06:21 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_8_LD.826 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.5596% | 2.228% | 1.8% | 0.2878% | 1.999% | 1.861% | 2.044% | 1.325% | 4.983% | 10.59% | 3.33% | 0.04406% | 4.708% | 27.4% | 12.66% | 1.477% | 2.988% | 1.415% | 2.751% | 9.8% | 0.4226% | 67.2% | 20.99% | 29.97% | 20.25% | 3.205% | 32.53% | 1.782% | 0.1594% | 0.6671% | 0.3272% | 1.419% | 0.8862% | 0.7562% | 0.8365% | 18.91% | 11.8% | 3.424% | 3.666%
178 | LBLANK |  | 6/17/2023 6:12:21 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.827 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.613% | 72.42% | 1.868% | 5.137% | 29.2% | 8.53% | 6.194% | 386.5% | 26.65% | 25.15% | 1.426% | 30.56% | 19.16% | 0% | 115.2% | 44.67% | 2.706% | 5.028% | 353.5% | 242.8% | 29.62% | 119.8% | 0% | 222.3% | 57.31% | 256% | 69.15% | 15.59% | 14.24% | 4.406% | 0.3825% | 0.1172% | 0.5691% | 0.8644% | 0.7997% | 43.28% | 11.01% | 2.676% | 5.131%
179 | LBLANK |  | 6/17/2023 6:18:40 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.828 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.359% | 124.2% | 4.624% | 2.844% | 44.05% | 10.63% | 3.651% | 668% | 24.49% | 47.67% | 0.7441% | 147.6% | 9.754% | 8.644% | 538.6% | 7.903% | 4.022% | 0.7551% | 62% | 113.7% | 62.49% | 44.79% | 0% | 222% | 146% | 200.1% | 0% | 11.78% | 16.24% | 1.018% | 0.2777% | 0.4676% | 2.53% | 0.6979% | 1.457% | 35.36% | 6.4% | 2.991% | 9.36%
180 | CCV 1ppb TRACE |  | 6/17/2023 6:24:40 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.829 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4.493% | 7.163% | 4.655% | 32.87% | 10.08% | 2.715% | 4.607% | 3.435% | 2.681% | 1.978% | 6.169% | 0.9752% | 17.58% | 3.161% | 2.511% | 3.574% | 2.007% | 2.674% | 7.287% | 5.914% | 1.492% | 10.51% | 4.901% | 6.954% | 6.842% | 13.1% | 19.47% | 0.6027% | 1.95% | 1.844% | 0.3839% | 0.2344% | 0.6877% | 0.6185% | 0.492% | 1.709% | 1.029% | 2.049% | 1.107%
181 | CCV 1ppb TRACE |  | 6/17/2023 6:31:00 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.830 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.489% | 6.111% | 16.9% | 27.18% | 3.177% | 3.327% | 7.005% | 1.788% | 4.922% | 1.289% | 4.497% | 2.709% | 1.758% | 1.738% | 1.399% | 2.226% | 0.7559% | 4.276% | 4.562% | 5.872% | 1.743% | 13.24% | 9.178% | 12.75% | 5.96% | 9.397% | 16.9% | 1.764% | 0.7212% | 1.818% | 0.9932% | 0.4423% | 0.5217% | 1.183% | 1.402% | 1.145% | 2.025% | 3.829% | 0.881%
182 | LFB |  | 6/17/2023 6:37:00 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.831 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.3726% | 1.303% | 8.569% | 2.075% | 2.091% | 1.642% | 0.7999% | 1.807% | 1.107% | 2.625% | 1.495% | 0.7119% | 2.885% | 0.5268% | 1.769% | 2.134% | 1.664% | 3.782% | 1.944% | 3.049% | 1.05% | 1.178% | 3.36% | 1.31% | 0.7824% | 2.271% | 2.5% | 1.138% | 0.99% | 0.2936% | 0.09303% | 1.796% | 0.6008% | 0.2231% | 1.394% | 2.027% | 1.94% | 1.144% | 0.7658%
183 | LFB |  | 6/17/2023 6:43:18 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.832 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.955% | 1.206% | 1.339% | 1.127% | 1.746% | 3.002% | 2.003% | 3.285% | 1.444% | 0.1806% | 0.5698% | 0.5927% | 2.184% | 1.29% | 1.443% | 2.79% | 0.5672% | 0.4506% | 1.061% | 4.665% | 1.291% | 2.641% | 5.212% | 3.663% | 0.5591% | 1.544% | 3.427% | 0.4208% | 0.697% | 0.2805% | 0.7076% | 0.7562% | 3.879% | 1.099% | 1.651% | 0.6314% | 0.4678% | 0.3102% | 5.154%
184 | CCV 10 ppb MAJORS |  | 6/17/2023 6:49:17 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.833 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.247% | 1.262% | 5.364% | 0.677% | 1.978% | 1.487% | 2.696% | 1.147% | 0.4319% | 1.366% | 9.036% | 1.02% | 3.459% | 0.912% | 1.451% | 1.109% | 0.6885% | 0.1334% | 3.512% | 1.165% | 0.5967% | 3.82% | 5.935% | 2.004% | 0.9336% | 4.081% | 3.311% | 0.347% | 0.3777% | 2.554% | 0.4529% | 0.4295% | 0.9437% | 0.3084% | 0.8975% | 1.672% | 0.5653% | 2.366% | 1.481%
185 | CCV 10 ppb MAJORS |  | 6/17/2023 6:55:37 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.834 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.606% | 5.241% | 7.575% | 2.58% | 3.388% | 1.252% | 2.365% | 2.42% | 2.161% | 0.1906% | 2.071% | 1.025% | 3.942% | 1.828% | 1.642% | 1.254% | 0.4354% | 1.196% | 2.376% | 8.08% | 0.985% | 1.883% | 4.695% | 2.268% | 0.3783% | 1.494% | 5.263% | 0.5068% | 0.3829% | 0.6864% | 0.1965% | 1.142% | 0.3485% | 0.7907% | 1.439% | 0.7469% | 1.158% | 1.09% | 1.055%
186 | 10_15_21_UF_8_DUP |  | 6/17/2023 7:01:37 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_8_DUP.835 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.855% | 3.992% | 1097% | 3.736% | 3.092% | 2.321% | 3.682% | 3.838% | 15.1% | 1.895% | 0.6143% | 0.5931% | 6.223% | 29.37% | 9.614% | 2.524% | 1.309% | 1.002% | 3.263% | 7.596% | 2.573% | 33.08% | 62.01% | 31.01% | 14.84% | 17.5% | 17.76% | 0.516% | 0.8113% | 0.1749% | 0.2182% | 1.236% | 1.021% | 0.4684% | 2.809% | 34.11% | 7.455% | 2.164% | 3.003%
187 | 10_15_21_UF_8_DUP |  | 6/17/2023 7:07:56 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_8_DUP.836 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.71% | 2.83% | 41.93% | 3.883% | 1.353% | 3.154% | 1.648% | 5.957% | 9.958% | 4.271% | 3.917% | 1.992% | 2.37% | 20.63% | 10.96% | 3.942% | 1.161% | 3.14% | 0.7473% | 5.635% | 1.756% | 32.49% | 0.4778% | 40.27% | 22.67% | 27.73% | 63.59% | 0.4942% | 1.553% | 0.5484% | 0.5125% | 1.34% | 0.9152% | 0.5244% | 1.625% | 10.35% | 29.38% | 4.322% | 3.52%
188 | 10_15_21_UF_8_DUP_LSP |  | 6/17/2023 7:13:55 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_8_DUP_LSP.837 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.772% | 0.2693% | 6.093% | 5.682% | 0.191% | 1.553% | 1.741% | 0.5517% | 1.673% | 0.758% | 0.2707% | 0.9826% | 1.445% | 1.886% | 1.689% | 0.794% | 0.6089% | 0.497% | 0.5844% | 0.8744% | 0.4299% | 1.254% | 7.383% | 1.783% | 2.239% | 2.232% | 2.221% | 0.8135% | 0.4757% | 0.4237% | 0.3315% | 0.4178% | 1.069% | 0.8426% | 1.13% | 0.3111% | 0.8818% | 1.511% | 0.3767%
189 | 10_15_21_UF_8_DUP_LSP |  | 6/17/2023 7:20:13 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_8_DUP_LSP.838 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.851% | 2.039% | 3.382% | 1.264% | 1.211% | 1.744% | 1.72% | 1.905% | 1.93% | 1.483% | 0.6524% | 0.6074% | 0.5302% | 1.571% | 2.005% | 1.207% | 0.57% | 1.167% | 3.119% | 5.389% | 1.164% | 5.378% | 3.179% | 3.568% | 0.6463% | 4.098% | 3.333% | 1.151% | 0.05412% | 2.649% | 0.4814% | 0.8073% | 0.515% | 0.4411% | 1.224% | 0.3916% | 0.5218% | 3.734% | 1.137%
190 | 10_15_21_UF_9 |  | 6/17/2023 7:26:13 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_9.839 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.027% | 2.016% | 3.737% | 1.085% | 1.76% | 2.617% | 0.3647% | 2.709% | 5.39% | 1.216% | 2.956% | 0.6749% | 2.665% | 44.81% | 3.403% | 0.7483% | 2.289% | 0.3365% | 0.6632% | 14.89% | 1.228% | 23.15% | 20.39% | 30.13% | 31.19% | 22.91% | 70.6% | 1.288% | 1.045% | 0.2182% | 0.6003% | 0.7352% | 0.3603% | 1.181% | 0.8849% | 28.98% | 23.06% | 2.058% | 4.369%
191 | 10_15_21_UF_9 |  | 6/17/2023 7:32:31 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_9.840 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.402% | 2.225% | 1.422% | 2.478% | 2.998% | 2.425% | 2.313% | 7.057% | 17.78% | 3.422% | 4.916% | 1.768% | 4.578% | 34.95% | 19.26% | 1.71% | 1.606% | 1.819% | 3.619% | 7.09% | 1.948% | 16.31% | 60.57% | 21.36% | 28.7% | 29.98% | 84.2% | 0.3901% | 1.215% | 0.9742% | 0.5266% | 1.202% | 0.3178% | 0.5594% | 1.763% | 11.61% | 8.309% | 12.22% | 4.277%
192 | 10_15_21_UF_10 |  | 6/17/2023 7:38:31 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_10.841 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.9982% | 0.8704% | 286% | 2.224% | 1.604% | 0.2545% | 1.188% | 4.562% | 9.503% | 1.609% | 8.203% | 2.113% | 8.125% | 45.64% | 8.169% | 3.936% | 2.085% | 0.357% | 2.218% | 6.205% | 1.635% | 23.3% | 53.93% | 35.96% | 6.471% | 1.799% | 33.87% | 2.524% | 2.253% | 0.1584% | 1.186% | 0.9719% | 14.82% | 0.1394% | 0.3658% | 17.27% | 9.394% | 3.316% | 25.39%
193 | 10_15_21_UF_10 |  | 6/17/2023 7:44:50 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_10.842 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.9659% | 1.605% | 83.32% | 3.265% | 2.624% | 3.125% | 2.516% | 4.673% | 10.26% | 7.207% | 4.487% | 1.321% | 7.212% | 19.59% | 9.21% | 1.539% | 0.5672% | 1.107% | 1.214% | 5.026% | 3.131% | 13.2% | 20.31% | 19.25% | 39.92% | 44.78% | 89.42% | 1.159% | 0.729% | 0.7467% | 0.8777% | 0.2577% | 2.528% | 0.6488% | 1.17% | 6.817% | 39.82% | 2.622% | 2.291%
194 | 10_15_21_UF_11 |  | 6/17/2023 7:50:50 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_11.843 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.837% | 1.84% | 78.91% | 2.196% | 3.591% | 4.1% | 1.248% | 2.821% | 0.3321% | 6.19% | 5.795% | 1.186% | 4.212% | 40.53% | 10.05% | 0.4831% | 0.3757% | 0.246% | 0.9556% | 2.681% | 1.421% | 29.02% | 110.1% | 27.37% | 13.31% | 18.12% | 87.54% | 0.9595% | 1.414% | 0.505% | 0.2873% | 0.1994% | 3.137% | 0.8055% | 0.9374% | 47.08% | 16.53% | 5.987% | 5.437%
195 | 10_15_21_UF_11 |  | 6/17/2023 7:57:09 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_11.844 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.598% | 3.456% | 41.94% | 1.807% | 2.113% | 2.003% | 2.093% | 8.529% | 4.335% | 18.08% | 0.9818% | 0.3526% | 4.138% | 23.65% | 21.11% | 3.054% | 1.867% | 1.404% | 2.014% | 7.943% | 1.051% | 12.41% | 71.77% | 27.67% | 22.36% | 30.29% | 96.6% | 1.032% | 1.215% | 0.1615% | 0.3486% | 0.1205% | 0.1712% | 1.083% | 1.883% | 17.35% | 19.37% | 2.734% | 1.882%
196 | LBLANK |  | 6/17/2023 8:03:10 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.845 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.8609% | 126.6% | 3.443% | 1.569% | 18.93% | 4.935% | 0% | 160.7% | 56.89% | 144.7% | 1.319% | 187.1% | 9.706% | 0% | 405.6% | 0.5017% | 3.972% | 2.451% | 279.2% | 122% | 39.45% | 419.5% | 0% | 95.18% | 53.86% | 217.8% | 133.4% | 11.96% | 27.73% | 3.274% | 0.9434% | 4.597% | 1.051% | 0.982% | 0.5099% | 0% | 22.86% | 1.106% | 5.17%
197 | LBLANK |  | 6/17/2023 8:09:28 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.846 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.8552% | 1.931e-06% | 5.271% | 1.309% | 14.11% | 8.457% | 3.737% | 154.7% | 54.19% | 34.67% | 1.609% | 228.4% | 15.55% | 8.993% | 103.7% | 23.16% | 12.49% | 0.6659% | 238.9% | 372.8% | 19.94% | 183.4% | 101.2% | 107.2% | 64.41% | 92.12% | 75.97% | 5.203% | 5.607% | 11.6% | 0.3508% | 3.078% | 0.775% | 1.091% | 0.5998% | 23.9% | 9.193% | 4.618% | 5.103%
198 | CCV 1ppb TRACE |  | 6/17/2023 8:15:29 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.847 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.422% | 4.661% | 2.433% | 126.1% | 11.18% | 2.691% | 2.509% | 1.022% | 3.387% | 0.6674% | 4.715% | 2.9% | 7.26% | 2.104% | 8.394% | 2.614% | 3.612% | 0.427% | 5.331% | 19.47% | 2.91% | 8.005% | 15.34% | 8.316% | 1.686% | 21.05% | 14.93% | 2.611% | 0.7784% | 0.4053% | 0.1566% | 0.143% | 0.3961% | 0.2567% | 0.4616% | 1.368% | 1.674% | 0.5683% | 1.321%
199 | CCV 1ppb TRACE |  | 6/17/2023 8:21:48 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.848 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.377% | 3.153% | 25.47% | 95.66% | 12.91% | 5.263% | 2.179% | 4.362% | 3.893% | 1.665% | 3.777% | 1.149% | 3.655% | 0.9383% | 1.07% | 3.476% | 1.927% | 3.438% | 2.992% | 5.118% | 1.192% | 12.72% | 25.19% | 6.283% | 9.482% | 7.592% | 20.04% | 0.7068% | 1.536% | 0.7409% | 0.5037% | 0.6126% | 0.645% | 0.4516% | 0.822% | 0.3406% | 2.284% | 0.3957% | 1.561%
200 | CCV 10 ppb MAJORS |  | 6/17/2023 8:27:48 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.849 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.2681% | 1.064% | 10.09% | 1.69% | 1.779% | 1.603% | 3.494% | 0.8177% | 1.589% | 0.4368% | 1.107% | 0.3435% | 1.668% | 0.5773% | 2.167% | 0.6747% | 0.3925% | 0.1129% | 1.784% | 6.14% | 1.169% | 0.4529% | 5.329% | 1.069% | 2.031% | 1.759% | 0.5253% | 0.3159% | 1.73% | 0.09716% | 0.8362% | 0.7998% | 0.1872% | 1.181% | 0.7117% | 0.5169% | 0.3038% | 1.528% | 0.8682%
201 | CCV 10 ppb MAJORS |  | 6/17/2023 8:34:07 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.850 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.339% | 1.933% | 4.406% | 2.834% | 0.9372% | 1.104% | 1.375% | 1.505% | 1.339% | 1.003% | 2.275% | 1.669% | 0.9572% | 2.216% | 3.986% | 0.7517% | 0.9434% | 1.361% | 2.723% | 3.078% | 1.196% | 4.579% | 3.112% | 5.701% | 1.852% | 3.691% | 3.079% | 0.7197% | 1.974% | 2.453% | 0.9137% | 0.7203% | 0.4736% | 0.2292% | 0.7842% | 0.3928% | 0.8562% | 4.919% | 1.363%
202 | 10_15_21_FB |  | 6/17/2023 8:40:08 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_FB.851 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9.015% | 31.62% | 7.971% | 9.802% | 12.43% | 14.39% | 12.08% | 344.8% | 6.652% | 6.073% | 1.333% | 11.95% | 2.204% | 16.62% | 49.6% | 10.69% | 7.311% | 0.6995% | 116.6% | 116.8% | 32.56% | 92.79% | 0% | 97% | 74.69% | 418.2% | 0% | 70.9% | 4.849% | 0.1586% | 11.15% | 0.5913% | 0.3092% | 0.5303% | 1.478% | 52.69% | 54.57% | 4.832% | 5.13%
203 | 10_15_21_FB |  | 6/17/2023 8:46:26 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_FB.852 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.93% | 33.41% | 1.202% | 5.522% | 20.36% | 15.61% | 18.49% | 58.54% | 10.9% | 7.519% | 2.129% | 6.98% | 6.016% | 27.67% | 61.07% | 19.93% | 7.222% | 0.96% | 106.9% | 123.2% | 14.11% | 60.12% | 102.5% | 192.6% | 41.91% | 833.6% | 0% | 189.4% | 20.52% | 0.9108% | 0.6744% | 0.6422% | 0.3913% | 0.1658% | 0.6935% | 30.89% | 35.37% | 8.546% | 2.847%
204 | 10_15_21_BB |  | 6/17/2023 8:52:26 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_BB.853 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.838% | 73.27% | 15.01% | 1.556% | 8.1% | 10.96% | 6.041% | 17768% | 53.37% | 15.49% | 8.012% | 69.62% | 38.22% | 9.594% | 48.51% | 17.03% | 3.569% | 6.071% | 1419% | 445.5% | 75.92% | 160.6% | 102% | 96.3% | 55.69% | 966.3% | 0% | 6.872% | 42.25% | 0.1579% | 0.6144% | 1.012% | 0.5925% | 1.232% | 1.993% | 37.7% | 13.29% | 11.05% | 8.294%
205 | 10_15_21_BB |  | 6/17/2023 8:58:45 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_BB.854 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.432% | 1.931e-06% | 55.07% | 0.3062% | 5.446% | 1.604% | 12.54% | 0% | 50.83% | 13.66% | 0.6141% | 40.87% | 43.05% | 4.311% | 100.3% | 7.164% | 1.898% | 21.83% | 367% | 1.777e-06% | 30.8% | 46048% | 0% | 107.8% | 66.27% | 385.6% | 80.45% | 10.76% | 35.21% | 0.4912% | 1.114% | 1.255% | 1.123% | 0.5711% | 0.7145% | 0% | 4.76% | 7.674% | 4.335%
206 | 10_15_21_LAB_A |  | 6/17/2023 9:04:45 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_LAB_A.855 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.456% | 1.931e-06% | 21.09% | 1.234% | 22.87% | 2.482% | 6.631% | 0% | 26.92% | 7.052% | 9.858% | 16.14% | 147.3% | 4.272% | 2908% | 4.885% | 5.132% | 60.02% | 149.2% | 1.777e-06% | 91.55% | 279.9% | 0% | 0% | 19.12% | 165.9% | 0% | 31.65% | 34.51% | 0.5217% | 0.2406% | 4.992% | 2.44% | 1.429% | 1.157% | 61.3% | 17.22% | 11.9% | 7.141%
207 | 10_15_21_LAB_A |  | 6/17/2023 9:11:03 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_LAB_A.856 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.007% | 1.931e-06% | 19.24% | 2.276% | 21.69% | 12.76% | 3.764% | 321.6% | 10.63% | 1.495% | 2.343% | 12.11% | 102.6% | 7.765% | 278.6% | 12.93% | 7.501% | 42.78% | 188.8% | 122.5% | 45.33% | 24.21% | 0% | 216.9% | 0.2341% | 73.05% | 0% | 70.33% | 10.85% | 0.4016% | 0.5544% | 0.3027% | 0.7753% | 0.5973% | 0.2759% | 61.35% | 30.04% | 12.96% | 5.551%
208 | 10_15_21_LAB_B |  | 6/17/2023 9:17:03 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_LAB_B.857 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.247% | 127.7% | 39.86% | 1.6% | 8.056% | 11.12% | 6.805% | 0% | 18.54% | 18.62% | 3.651% | 46.41% | 56.93% | 7.86% | 88.31% | 8.795% | 56.26% | 11.67% | 144.5% | 482.7% | 40.86% | 184% | 0% | 192.2% | 43.2% | 59.68% | 82.63% | 8.327% | 33.05% | 0.09203% | 0.7893% | 0.1261% | 0.6512% | 1.254% | 1.026% | 37.98% | 10.75% | 3.896% | 5.554%
209 | 10_15_21_LAB_B |  | 6/17/2023 9:23:20 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_LAB_B.858 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.777% | 1.931e-06% | 53.93% | 4.447% | 46.78% | 3.041% | 11.49% | 4.192% | 26.7% | 13.89% | 4.107% | 26.28% | 18.7% | 9.38% | 242.8% | 23.29% | 7.986% | 13.49% | 112.9% | 1.777e-06% | 12.67% | 16.85% | 103.1% | 0% | 0% | 17.36% | 83% | 1.148% | 32.28% | 0.238% | 0.5297% | 1.005% | 0.773% | 0.5101% | 1.359% | 61.63% | 13.61% | 10.17% | 2.5%
210 | 5_18_23_DB |  | 6/17/2023 9:29:20 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_DB.859 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.524% | 1.931e-06% | 20.04% | 4.707% | 2.453% | 5.179% | 3.874% | 0% | 32.04% | 30.96% | 1.157% | 214.7% | 8.329% | 4.646% | 93.4% | 6.362% | 0.7397% | 0.3606% | 132.1% | 1.777e-06% | 40.75% | 3090% | 0% | 0% | 21.84% | 108.4% | 0% | 4.551% | 28.05% | 1.367% | 0.8639% | 0.2967% | 0.9445% | 1.284% | 0.6405% | 62.61% | 3.721% | 12.41% | 2.518%
211 | 5_18_23_DB |  | 6/17/2023 9:35:39 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_DB.860 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.532% | 1.931e-06% | 28.04% | 2.046% | 30.51% | 1.058% | 10.32% | 3.157% | 70.6% | 67.95% | 0.6656% | 52.24% | 13.34% | 7.979% | 285.8% | 11.21% | 7.027% | 4.341% | 70.49% | 127.9% | 140.1% | 106.2% | 0% | 215.2% | 19.55% | 12.19% | 306.2% | 9.477% | 12.26% | 1.684% | 1.583% | 0.691% | 0.3408% | 1.09% | 1.047% | 45.58% | 11.95% | 7.938% | 1.305%
212 | LBLANK |  | 6/17/2023 9:41:39 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.861 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.9164% | 1.931e-06% | 4.299% | 13.2% | 44.04% | 15.46% | 4.075% | 36189% | 38.64% | 94.07% | 1.809% | 177.5% | 7.487% | 14.23% | 52.38% | 19.8% | 11.12% | 3.391% | 331.5% | 134.2% | 102.5% | 301.3% | 0% | 191.7% | 22.27% | 1278% | 83.95% | 8.306% | 18.41% | 6.235% | 0.6584% | 0.5743% | 0.1421% | 0.631% | 0.6248% | 31.56% | 8.155% | 1.746% | 7.856%
213 | LBLANK |  | 6/17/2023 9:47:59 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.862 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.8834% | 1.931e-06% | 1.977% | 1.598% | 68.46% | 9.528% | 4.125% | 0% | 58.56% | 175.7% | 0.6676% | 66.62% | 27.96% | 4.534% | 152.9% | 14.46% | 14.74% | 1.917% | 7806% | 1.777e-06% | 64.16% | 458.9% | 103.6% | 0% | 0% | 34.56% | 0% | 7.403% | 6.748% | 3.775% | 0.2978% | 0.3872% | 0.5834% | 0.4001% | 2.111% | 38.4% | 7.958% | 4.96% | 4.296%
214 | CCV 1ppb TRACE |  | 6/17/2023 9:53:59 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.863 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.66% | 2.61% | 36.14% | 65.94% | 10.48% | 4.711% | 7.311% | 1.055% | 5.555% | 1.413% | 3.711% | 0.9289% | 12% | 2.018% | 7.334% | 2.901% | 2.257% | 1.299% | 7.765% | 15.33% | 1.528% | 5.255% | 16.93% | 14.25% | 4.225% | 6.621% | 24.22% | 1.71% | 2.664% | 0.5838% | 1.076% | 1.098% | 0.8227% | 0.8006% | 2.185% | 2.472% | 2.293% | 1.308% | 1.641%
215 | CCV 1ppb TRACE |  | 6/17/2023 10:00:19 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.864 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.814% | 1.495% | 24.82% | 100.5% | 9.039% | 6.399% | 7.726% | 2.496% | 1.238% | 7.464% | 16.57% | 0.9959% | 8.641% | 0.368% | 5.361% | 2.692% | 8.233% | 11.84% | 2.116% | 12.65% | 2.493% | 14.83% | 14.85% | 19.06% | 3.013% | 19.52% | 13.46% | 11.53% | 2.43% | 0.6072% | 0.7237% | 7.333% | 0.6219% | 0.8248% | 1.009% | 9.965% | 9.688% | 1.859% | 0.3808%
216 | LFB |  | 6/17/2023 10:06:18 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.865 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.546% | 2.686% | 8.492% | 0.3324% | 1.646% | 2.507% | 4.836% | 1.761% | 3.712% | 0.9378% | 1.141% | 0.6715% | 2.837% | 1.542% | 1.777% | 2.433% | 0.3666% | 0.7172% | 1.662% | 4.356% | 0.6179% | 5.393% | 9.387% | 2.174% | 2.604% | 1.503% | 7.005% | 1.206% | 0.6866% | 0.5999% | 0.4423% | 1.127% | 0.5146% | 0.2978% | 1.436% | 0.6128% | 0.1444% | 0.8397% | 0.5003%
217 | LFB |  | 6/17/2023 10:12:36 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.866 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.53% | 0.5538% | 10.36% | 1.963% | 5.573% | 1.134% | 1.6% | 2.04% | 1.389% | 3.166% | 4.739% | 1.192% | 1.613% | 0.7677% | 1.404% | 1.389% | 3.713% | 3.257% | 0.503% | 3.49% | 0.4911% | 4.786% | 9.76% | 4.339% | 3.261% | 3.503% | 6.01% | 3.565% | 0.4086% | 1.821% | 1.218% | 2.141% | 0.6679% | 0.7912% | 0.5973% | 4.553% | 4.469% | 2.098% | 0.5975%
218 | CCV 10 ppb MAJORS |  | 6/17/2023 10:18:36 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.867 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.687% | 0.8758% | 0.6871% | 1.208% | 2.751% | 1.488% | 2.113% | 2.2% | 2.158% | 1.167% | 1.262% | 1.536% | 2.854% | 1.104% | 1.042% | 2.547% | 0.6427% | 1.155% | 0.5745% | 4.628% | 1.095% | 0.8475% | 2.878% | 1.346% | 2.301% | 1.847% | 2.357% | 0.6816% | 1.139% | 0.8282% | 0.7479% | 1.438% | 0.6097% | 0.7542% | 2.072% | 1.437% | 1.729% | 1.443% | 1.189%
219 | CCV 10 ppb MAJORS |  | 6/17/2023 10:24:55 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.868 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.446% | 2.865% | 4.48% | 1.508% | 3.839% | 1.97% | 3.176% | 2.503% | 1.561% | 1.41% | 1.197% | 0.3845% | 2.583% | 3.213% | 2.028% | 2.125% | 0.8129% | 0.9016% | 1.077% | 1.68% | 0.974% | 1.822% | 5.334% | 3.203% | 1.157% | 0.8085% | 6.222% | 1.003% | 0.4736% | 0.5097% | 0.1366% | 0.1719% | 0.6183% | 0.7821% | 2.089% | 0.9213% | 0.3083% | 1.426% | 1.045%
220 | 5_18_23_DB |  | 6/17/2023 10:30:55 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_DB.869 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.453% | 1.931e-06% | 33% | 2.102% | 43.73% | 9.954% | 9.988% | 327.3% | 13980% | 38.12% | 3.773% | 25.67% | 12.42% | 5.933% | 49.9% | 17.21% | 4.498% | 7.245% | 307.6% | 124.4% | 45.97% | 166.2% | 0% | 120.6% | 55.6% | 395.7% | 1360% | 38.51% | 21.8% | 3.284% | 0.3298% | 1.188% | 2.003% | 0.9762% | 1.486% | 0% | 14.71% | 15.59% | 4.099%
221 | 5_18_23_DB |  | 6/17/2023 10:37:14 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_DB.870 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.367% | 126.7% | 2.638% | 2.162% | 4.033% | 11.81% | 3.856% | 161.6% | 167.9% | 38.31% | 3.294% | 32.15% | 18.38% | 13.45% | 67.57% | 16.19% | 6.523% | 6.406% | 217.9% | 389.7% | 127.8% | 106.1% | 101.6% | 216.8% | 18.81% | 30.4% | 144.2% | 147.5% | 15.26% | 1.134% | 0.6438% | 0.9004% | 0.3227% | 0.3216% | 0.9927% | 36.86% | 7.783% | 14.28% | 8.205%
222 | 5_18_23_DB |  | 6/17/2023 10:43:14 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_DB.871 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.698% | 1.931e-06% | 20.56% | 2.586% | 12.6% | 5.832% | 15.02% | 3902% | 52.5% | 12.43% | 5.711% | 24.48% | 146% | 9.794% | 168.6% | 15.03% | 5.875% | 4.877% | 59.78% | 416.3% | 71.4% | 40.43% | 0% | 96.17% | 42.16% | 696.6% | 80.22% | 17.35% | 23.01% | 15.75% | 1.103% | 0.1436% | 0.6102% | 0.5082% | 0.5605% | 30.75% | 7.585% | 27.09% | 2.63%
223 | 5_18_23_DB |  | 6/17/2023 10:49:33 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_DB.872 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.709% | 1.931e-06% | 13.82% | 0.3475% | 33.4% | 4.782% | 0.06262% | 7540% | 158.6% | 25.65% | 1.494% | 46% | 41.28% | 10.61% | 100.8% | 13.53% | 6.35% | 6.042% | 334.6% | 1.777e-06% | 127.2% | 48.7% | 73.96% | 216.2% | 28.42% | 348.5% | 0% | 18.73% | 16.37% | 1.045% | 0.8654% | 0.355% | 11.84% | 1.151% | 0.9395% | 45.35% | 9.918% | 1.077% | 21.03%
224 | 5_18_23_DB |  | 6/17/2023 10:55:33 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_DB.873 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 3.078% | 1.931e-06% | 30.16% | 1.857% | 40.65% | 6.171% | 11.61% | 167.7% | 248.6% | 28.74% | 4.787% | 34.51% | 87.43% | 4.635% | 461.8% | 4.192% | 12.91% | 7.008% | 756.9% | 1.777e-06% | 36.6% | 30.88% | 0% | 0% | 132.6% | 334.3% | 0% | 24.64% | 22.88% | 1.29% | 0.4664% | 0.6213% | 0.4251% | 1.018% | 0.8267% | 45.82% | 16.63% | 16.73% | 12.61%
225 | 5_18_23_DB |  | 6/17/2023 11:01:51 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_DB.874 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.497% | 1.931e-06% | 26.54% | 3.959% | 28.51% | 7.256% | 3.818% | 0% | 70.24% | 13088% | 4.91% | 12.41% | 23.36% | 14% | 208.9% | 15.84% | 5.116% | 16.66% | 75.94% | 464.6% | 15.04% | 516.4% | 0% | 130% | 95.05% | 168.6% | 0% | 15.62% | 23.84% | 0.4389% | 0.1387% | 1.361% | 0.1499% | 0.3278% | 1.918% | 38.57% | 18.97% | 7.353% | 7.583%
226 | 5_18_23_MFB |  | 6/17/2023 11:07:51 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_MFB.875 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.403% | 3.422% | 7.453% | 1.294% | 2.839% | 1.912% | 1.242% | 2.341% | 2.503% | 0.8728% | 0.8931% | 1.97% | 0.8765% | 1.308% | 2.166% | 1.987% | 0.7466% | 0.3956% | 2.789% | 7.619% | 0.8321% | 3.715% | 4.501% | 2.062% | 4.971% | 1.31% | 2.689% | 0.5383% | 2.682% | 0.8757% | 2.025% | 1.277% | 0.4159% | 1.774% | 0.576% | 1.229% | 1.478% | 0.826% | 1.732%
227 | 5_18_23_MFB |  | 6/17/2023 11:14:09 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_MFB.876 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.95% | 1.504% | 7.103% | 1.074% | 5.236% | 2.345% | 0.8781% | 2.283% | 2.002% | 1.299% | 1.024% | 1.635% | 0.5998% | 2.625% | 2.944% | 4.292% | 0.5598% | 0.628% | 1.89% | 6.735% | 0.761% | 5.408% | 14.84% | 0.3806% | 2.308% | 5.898% | 5.173% | 0.1382% | 0.4084% | 1.223% | 1.105% | 0.5077% | 0.9751% | 0.09416% | 0.9342% | 1.245% | 0.3789% | 2.594% | 1.302%
228 | 5_18_23_MFB |  | 6/17/2023 11:20:08 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_MFB.877 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.3924% | 1.626% | 9.594% | 0.6716% | 4.235% | 3.34% | 3.461% | 3.996% | 1.88% | 1.667% | 1.328% | 0.9603% | 2.94% | 0.6175% | 0.2325% | 2.052% | 0.9718% | 0.9213% | 1.318% | 6.087% | 2.909% | 6.9% | 5.773% | 1.181% | 2.969% | 6.062% | 7.944% | 0.7474% | 0.4237% | 0.6761% | 0.8451% | 0.1682% | 5.476% | 0.5712% | 0.8766% | 1.382% | 1.208% | 0.2196% | 5.484%
229 | 5_18_23_MFB |  | 6/17/2023 11:26:26 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_MFB.878 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.251% | 4.854% | 5.927% | 0.9417% | 2.141% | 2.545% | 1.496% | 2.37% | 1.334% | 1.379% | 1.594% | 0.2227% | 0.3518% | 1.66% | 3.9% | 0.5092% | 0.4503% | 0.2917% | 3.753% | 6.499% | 2.473% | 9.624% | 3.464% | 5.631% | 7.676% | 4.317% | 7.358% | 0.7453% | 1.045% | 3.141% | 0.5436% | 0.4914% | 0.5853% | 1.073% | 0.5426% | 0.6244% | 0.4041% | 3.085% | 1.23%
230 | LBLANK |  | 6/17/2023 11:32:26 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.879 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.009801% | 127.3% | 2.49% | 11.08% | 27.35% | 0.2491% | 4.051% | 89.06% | 20.05% | 89.09% | 1.663% | 150.2% | 65.72% | 4.452% | 126% | 30.29% | 19.02% | 3.757% | 451.8% | 1.777e-06% | 134.9% | 294% | 0% | 96.07% | 34.67% | 2703% | 81.2% | 63.13% | 15.38% | 2.088% | 0.5003% | 1.362% | 0.3981% | 0.5041% | 0.6416% | 38.15% | 9.902% | 2.425% | 8.594%
231 | LBLANK |  | 6/17/2023 11:38:51 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.880 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.879% | 1.931e-06% | 1.166% | 3.279% | 38.99% | 2.725% | 4.049% | 7957% | 44.4% | 0.2811% | 0.4508% | 38.21% | 25.57% | 4.393% | 212.5% | 20.61% | 4.26% | 2.487% | 406.8% | 1.777e-06% | 22.26% | 61.61% | 0% | 0% | 77.44% | 393.3% | 0% | 107.3% | 6.948% | 1.885% | 0.4801% | 0.1343% | 0.5602% | 0.4477% | 0.5455% | 0% | 7.107% | 1.431% | 4.85%
232 | CCV 1ppb TRACE |  | 6/17/2023 11:44:51 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.881 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2.369% | 5.957% | 14.66% | 9.018% | 6.8% | 3.837% | 1.15% | 1.762% | 8.134% | 2.835% | 7.226% | 0.1711% | 6.948% | 1.223% | 2.454% | 2.481% | 2.15% | 3.362% | 8.398% | 13.73% | 5.205% | 13.23% | 21.49% | 5.993% | 3.093% | 2.166% | 6.782% | 1.823% | 1.399% | 0.7664% | 0.2276% | 1.049% | 0.4844% | 0.9052% | 0.5254% | 1.023% | 0.8539% | 0.3795% | 1.383%
233 | CCV 1ppb TRACE |  | 6/17/2023 11:51:11 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.882 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 0.5947% | 2.679% | 3.768% | 8.819% | 7.594% | 3.916% | 5.848% | 1.114% | 3.019% | 3.512% | 4.177% | 0.9864% | 10.48% | 2.038% | 4.431% | 2.875% | 1.771% | 3.546% | 6.442% | 9.039% | 1.453% | 1.761% | 13.73% | 18.27% | 4.806% | 2.407% | 19.68% | 0.7271% | 1.405% | 0.2555% | 0.2396% | 1.031% | 0.4663% | 0.7229% | 0.5741% | 3.373% | 1.048% | 1.823% | 1.287%
234 | CCV 10 ppb MAJORS |  | 6/17/2023 11:57:11 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.883 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1.275% | 1.424% | 5.272% | 0.5683% | 4.916% | 1.289% | 0.5192% | 1.958% | 0.1688% | 1.364% | 2.84% | 0.7222% | 0.2356% | 0.2678% | 3.044% | 0.8978% | 1.101% | 0.5258% | 2.597% | 4.674% | 1.637% | 2.382% | 7.061% | 3.876% | 2.81% | 4.309% | 6.96% | 0.3073% | 0.6464% | 1.104% | 0.5806% | 0.6938% | 1.186% | 0.5447% | 0.7936% | 1.944% | 0.8456% | 1.898% | 1.628%
235 | CAL BLANK |  | 6/17/2023 12:25:05 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CAL BLANK.884 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth

## Unfactored Concentrations
| Sample Id | R | Acquisition Time | QC Status | Dataset File | Method File | Na KED 23/23 Helium KED_LowLow (ppb) | Mg KED 24/24 Helium KED_LowLow (ppb) | Al KED 27/27 Helium KED_LowLow (ppb) | P Oshift 31/47 Oxygen DRC (ppb) | S (34) -1 34/34 Helium KED_LowLow (ppb) | K KED 39/39 Helium KED_LowLow (ppb) | Ca KED 43/43 Helium KED_LowLow (ppb) | V KED 51/51 Helium KED_LowLow (ppb) | Cr (52) 52/52 Helium KED_LowLow (ppb) | Cr (53) NH3 53/53 Ammonia DRC (ppb) | Fe (54) NH3 54/54 Ammonia DRC (ppb) | Mn Oshift 55/71 Oxygen DRC (ppb) | Fe (56) KED 56/56 Helium KED_LowLow (ppb) | Co KED 59/59 Helium KED_LowLow (ppb) | Ni KED 60/60 Helium KED_LowLow (ppb) | Cu (63) KED 63/63 Helium KED_LowLow (ppb) | Cu (65) NH3 65/65 Ammonia DRC (ppb) | Zn NH3 66/66 Ammonia DRC (ppb) | As 75/75 (ppb) | As KED 75/75 Helium KED_LowLow (ppb) | As Oshift 75/91 Oxygen DRC (ppb) | Se (77) 77/77 (ppb) | Se (77) Oshift 77/93 Oxygen DRC (ppb) | Se (78) Oshift 78/94 Oxygen DRC (ppb) | Se (80) Oshift 80/96 Oxygen DRC (ppb) | Se (82) 82/82 (ppb) | Se (82) Oshift 82/98 Oxygen DRC (ppb) | Mo (95) NH3 95/95 Ammonia DRC (ppb) | Mo (95) OShift 95/127 Oxygen DRC (ppb) | Rh (IS1) 103/103 (IS) | Rh (IS2) 103/103 (IS) Oxygen DRC | Rh (IS4) 103/103 (IS) Ammonia DRC | Rh (IS5) 103/103 (IS) Ammonia DRC | Rh (IS6) 103/103 (IS) Oxygen DRC | Rh (IS7) 103/103 (IS) Helium KED_LowLow | Cd (111) NH3 111/111 Ammonia DRC (ppb) | Cd (114) NH3 114/114 Ammonia DRC (ppb) | Pb 208/208 (ppb) | Pb NH3 208/208 Ammonia DRC (ppb)
1 | sac blank |  | 06/16/2023 11:58:56 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\sac blank .650 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth
2 | sac blank cal |  | 06/16/2023 12:04:57 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\sac blank cal.651 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth
3 | sac blank cal |  | 06/16/2023 12:10:57 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\sac blank cal.652 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth
4 | CAL BLANK |  | 06/16/2023 12:16:57 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CAL BLANK.653 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth
5 | LOW 0.01 |  | 06/16/2023 12:22:58 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LOW 0.01.654 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth
6 | LOW 0.05 |  | 06/16/2023 12:28:58 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LOW 0.05.655 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | [7 empty] | 0.05 | 0.05 | 0.05 |  |  |  | 0.05 | 0.05 | 0.05 | 0.05 |  |  | 0.05 | 0.05 | [14 empty] | 0.05 | 0.05 | 0.05 | 0.05
7 | LOW 0.1 |  | 06/16/2023 12:34:59 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LOW 0.1.656 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | [7 empty] | 0.091491 | 0.10172 | 0.097655 |  | 0.1 |  | 0.099402 | 0.10001 | 0.10345 | 0.10293 |  | 0.1 | 0.11294 | 0.096974 |  | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | [6 empty] | 0.098881 | 0.096188 | 0.10299 | 0.10182
8 | LOW 0.5 |  | 06/16/2023 12:40:59 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LOW 0.5.657 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth |  |  |  | 0.5 |  |  |  | 0.47899 | 0.54495 | 0.51826 |  | 0.49806 |  | 0.51852 | 0.49897 | 0.56111 | 0.56962 | 0.5 | 0.62329 | 0.56584 | 0.49801 | 0.5 | 0.55552 | 0.52402 | 0.51467 | 0.55061 | 0.55207 | 0.51936 | 0.50551 | [6 empty] | 0.50456 | 0.49928 | 0.53645 | 0.5312
9 | STD 1 |  | 06/16/2023 12:46:58 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 1.658 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1000 | 1000 | 1 | 0.523 |  | 1000 | 1000 | 0.96784 | 1.0723 | 1.0004 | 1 | 0.99303 | 1 | 1.0119 | 0.97409 | 1.0383 | 1.0389 | 0.89245 | 1.0612 | 1.0017 | 0.98309 | 1.0627 | 0.99131 | 1.0406 | 1.0155 | 1.0001 | 1.0004 | 1.0047 | 1.0047 | [6 empty] | 0.97751 | 0.97957 | 1.0241 | 1.0424
10 | STD 5 |  | 06/16/2023 12:52:58 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 5.659 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 5021.2304 | 5049.1253 | 5.6267 | 2.9923 | 5000 | 4966.6842 | 5131.8876 | 5.0136 | 5.3783 | 4.9234 | 4.6709 | 5.0395 | 4.7022 | 5.1448 | 4.8947 | 5.1802 | 5.0303 | 4.5954 | 5.5479 | 5.1629 | 4.9865 | 5.0238 | 5.1474 | 5.2097 | 4.9795 | 5.3795 | 5.329 | 5.03 | 4.9866 | [6 empty] | 4.9537 | 4.9369 | 4.9653 | 5.1031
11 | STD 10 |  | 06/16/2023 12:58:58 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 10.660 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9999.5136 | 10083.0832 | 10.3725 | 5.7932 | 9888.1205 | 9980.4144 | 10192.9569 | 10.1267 | 10.5701 | 9.8979 | 9.8012 | 10.2265 | 9.6613 | 10.203 | 9.6861 | 10.1208 | 9.9206 | 9.4306 | 10.8636 | 10.0086 | 9.9641 | 10.6374 | 9.9419 | 10.1848 | 10.0889 | 10.3578 | 10.3331 | 10.0956 | 10.0999 | [6 empty] | 9.8008 | 9.801 | 9.4087 | 9.9173
12 | STD 25 |  | 06/16/2023 13:04:58 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 25.661 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 25949.6012 | 25382.7433 | 26.4382 | 15.7537 | 25263.2589 | 25792.5995 | 25454.0498 | 25.9986 | 26.7776 | 26.1222 | 25.926 | 26.1442 | 25.0156 | 26.2376 | 24.8838 | 25.7058 | 25.6116 | 24.4674 | 27.7521 | 26.3788 | 26.3983 | 25.8426 | 26.8878 | 26.6488 | 26.8519 | 26.6799 | 27.4171 | 26.391 | 26.2628 | [6 empty] | 25.3755 | 25.3304 | 21.5059 | 24.7359
13 | STD 50 |  | 06/16/2023 13:10:58 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 50.662 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 51920.6916 | 49981.3479 | 52.4153 | 35.8084 | 50235.6864 | 52623.4836 | 51975.4887 | 53.2049 | 54.0967 | 52.3163 | 50.8468 | 51.6675 | 52.3849 | 52.2692 | 49.0791 | 50.1304 | 49.6667 | 48.1366 | 54.8391 | 51.1131 | 52.4779 | 51.2563 | 53.1132 | 53.2205 | 52.8092 | 52.6819 | 53.9795 | 52.4782 | 52.5072 | [6 empty] | 49.3176 | 48.9369 | 35.3187 | 46.7403
14 | STD 100 |  | 06/16/2023 13:16:57 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\STD 100.663 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 108940.0898 | 104423.2102 | 109.062 | 76.9405 | 102894.999 | 111261.013 | 108386.8988 | 116.1759 | 116.2272 | 112.535 | 102.9693 | 106.2473 | 103.1562 | 109.9238 | 102.8191 | 102.6854 | 100.5096 | 96.8635 | 117.6454 | 106.6672 | 105.7244 | 107.4687 | 107.6878 | 109.1336 | 108.317 | 114.3442 | 111.5977 | 110.7929 | 108.8356 | [6 empty] | 98.5794 | 97.2169 | 57.984 | 102.7341
15 | Calibration Curves |  | 06/16/2023 13:16:57
16 | LBLANK |  | 06/16/2023 13:22:58 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.664 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 60.4509 | 4.5053 | -0.89613 | -0.23606 | 324.0337 | 71.0745 | 42.4779 | 0.0012092 | -0.0036826 | 0.00097579 | -0.59692 | 0.00078078 | -0.28103 | -0.0045444 | 0.034784 | -0.015084 | -0.014802 | 0.45922 | 0.032511 | -0.011484 | -0.0019304 | -0.079577 | 0.018001 | 0.02332 | -0.006465 | 0.0037711 | -0.021474 | 0.20341 | 0.012281 | [6 empty] | 0.0012708 | -0.0008358 | -0.011131 | -0.007241
17 | LBLANK |  | 06/16/2023 13:41:50 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.665 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 56.0325 | 0.92206 | -0.78861 | -0.82037 | -463.6058 | 66.3065 | 39.6685 | 0.0014031 | -0.0072756 | -0.00021147 | -0.57632 | -0.00038413 | -0.59158 | -0.0069095 | 0.006604 | -0.012116 | -0.0099493 | -0.10807 | 0.046967 | -0.01135 | -0.0022639 | -0.16067 | 0.014389 | 0.016408 | -0.0074087 | 0.010201 | -0.021474 | 0.05543 | 0.0072977 | [6 empty] | 0.00068284 | -0.0015896 | -0.010754 | -0.0043189
18 | ICV PE |  | 06/16/2023 13:47:50 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\ICV PE.666 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10056.5045 | 10399.9312 | 9.8698 | 18.7224 | 9864.5338 | 9803.376 | 9913.8204 | 10.0101 | 10.5126 | 10.1321 | 10.071 | 10.2031 | 9.9179 | 10.0963 | 9.7904 | 10.0569 | 10.0276 | 9.713 | 9.7284 | 9.0755 | 10.0024 | 6.9848 | 9.9015 | 10.1402 | 10.1402 | 9.8368 | 9.9585 | 9.5255 | 9.5507 | [6 empty] | 9.5729 | 9.6206 | 9.2728 | 9.5202
19 | CCV 1ppb TRACE |  | 06/16/2023 13:53:50 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.667 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 55.6556 | 0.91518 | -1.0076 | -0.99296 | -2402.6525 | 50.0019 | 40.5457 | 7.111e-06 | -0.013837 | -0.0030719 | -0.6073 | -0.0012092 | -2.2335 | -0.00661 | -0.0051573 | -0.01323 | -0.01183 | -0.21218 | 0.71528 | -0.0089217 | -0.004222 | -1.9308 | 0.0055366 | 0.0096123 | -0.0090549 | -0.090629 | -0.021474 | 0.035588 | 0.0014145 | [6 empty] | 0.00074945 | -0.0022094 | -0.010929 | -0.0047576
20 | CCV 1ppb TRACE |  | 06/16/2023 14:00:07 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.668 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1051.1896 | 1076.671 | 1.2775 | 0.14651 | 1270.6449 | 1104.0053 | 989.4019 | 1.0496 | 1.09 | 1.0467 | 0.8539 | 1.0223 | 1.167 | 1.0494 | 1.0433 | 1.0774 | 1.0596 | 0.82662 | 0.96458 | 0.93718 | 0.9793 | 1.1859 | 0.9275 | 0.99008 | 1.0313 | 0.96014 | 0.9736 | 1.0233 | 0.98635 | [6 empty] | 0.97624 | 0.97117 | 1.0298 | 0.98783
21 | LFB |  | 06/16/2023 14:06:07 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.669 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10631.0355 | 10937.2198 | 10.7055 | 11.5311 | 10903.3394 | 10615.1967 | 10398.4198 | 10.7744 | 11.3367 | 9.9786 | 9.5522 | 9.9937 | 10.6854 | 10.9986 | 10.5865 | 10.8235 | 9.914 | 10.0895 | 9.9048 | 11.207 | 9.9616 | 10.3388 | 10.1419 | 10.1303 | 9.9546 | 10.4514 | 10.5175 | 9.7079 | 9.7943 | [6 empty] | 9.669 | 9.5549 | 9.1443 | 9.426
22 | LLOQ 0.01 |  | 06/16/2023 14:12:05 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LLOQ 0.01.670 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 68.7627 | 11.998 | 0.31844 | 1.2213 | 404.9005 | 74.1205 | 49.5998 | 0.011602 | 0.0069612 | 0.013291 | -0.064712 | 0.029726 | 0.26583 | 0.0096876 | 0.02016 | 0.01962 | 0.02022 | -0.0037104 | 0.05627 | 0.0034775 | 0.0077831 | 0.028379 | 0.019141 | 0.023973 | 9.0899e-05 | 0.031049 | 0.001602 | 0.050454 | 0.021506 | [6 empty] | 0.010088 | 0.0088365 | 0.0077222 | 0.013587
23 | CCV 10 ppb MAJORS |  | 06/16/2023 14:18:05 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.671 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10122.8889 | 10275.35 | 9.7477 | 9.1705 | 10255.1656 | 9972.5124 | 9698.068 | 10.1342 | 10.5512 | 10.4501 | 10.0215 | 10.4576 | 9.83 | 10.4451 | 10.2808 | 10.3056 | 10.327 | 10.4275 | 9.9744 | 10.2396 | 10.2403 | 10.5057 | 9.9803 | 9.9928 | 10.1894 | 10.2047 | 10.2849 | 10.1979 | 10.2561 | [6 empty] | 9.8422 | 9.8537 | 9.3505 | 9.6711
24 | LLOQ 0.1 |  | 06/16/2023 14:24:05 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LLOQ 0.1.672 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 157.5781 | 102.3591 | -0.73809 | -0.51602 | 462.677 | 164.6774 | 129.5481 | 0.10451 | 0.088081 | 0.10195 | -0.2596 | 0.10732 | 0.13125 | 0.09552 | 0.10538 | 0.10931 | 0.10692 | 0.014436 | 0.13764 | 0.094082 | 0.094888 | 0.13829 | 0.096396 | 0.10776 | 0.091667 | 0.056754 | 0.13188 | 0.13578 | 0.10093 | [6 empty] | 0.097918 | 0.09334 | 0.089389 | 0.093548
25 | 9_13_21_F_11_DUP |  | 06/16/2023 14:30:04 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_F_11_DUP.673 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 6749.2116 | 8683.4952 | 8.7958 | 2.7535 | 14706.9096 | 1924.6586 | 33015.485 | 0.76408 | 0.11728 | 0.11583 | 4.4686 | 5.0477 | 5.0885 | 0.016785 | 0.24271 | 2.2292 | 2.293 | 3.555 | 4.1806 | 4.0028 | 4.1759 | 0.29654 | 0.10333 | 0.11553 | 0.093087 | 0.43196 | 0.065493 | 2.3338 | 2.2977 | [6 empty] | 0.012929 | 0.011039 | 0.11831 | 0.11785
26 | LBLANK |  | 06/16/2023 14:47:08 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.675 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 55.7227 | 1.5596 | -0.86348 | -0.86945 | 243.5678 | 64.2909 | 40.9968 | 0.0012129 | -0.0026043 | -0.00061988 | -0.57775 | -0.00058098 | -0.27505 | -0.0062026 | 0.00013072 | -0.011968 | -0.010147 | -0.1891 | 0.029022 | -0.001768 | -0.0025415 | -0.10492 | 0.010392 | 0.0056793 | -0.0070266 | -0.017754 | -0.017342 | 0.020267 | 0.0038164 | [6 empty] | 0.0005995 | -0.0016388 | -0.01057 | -0.0043334
27 | 0.5 STD |  | 06/16/2023 14:53:09 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\0.5 STD.676 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 567.1011 | 543.0869 | 0.48963 | 1.4303 | 822.519 | 587.9992 | 554.3056 | 0.5089 | 0.53802 | 0.52638 | 0.57417 | 0.51448 | 0.74484 | 0.51448 | 0.51989 | 0.56375 | 0.5828 | 0.55951 | 0.5425 | 0.47652 | 0.50708 | 0.42628 | 0.51372 | 0.53534 | 0.50474 | 0.45577 | 0.56183 | 0.51778 | 0.50407 | [6 empty] | 0.4982 | 0.50077 | 0.5187 | 0.51064
28 | 0.1 STD |  | 06/16/2023 14:59:12 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\0.1 STD.677 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 155.8817 | 102.4158 | -0.81077 | -0.012616 | 408.199 | 159.4332 | 148.202 | 0.088929 | 0.097049 | 0.099312 | -0.25647 | 0.10293 | -0.028854 | 0.09906 | 0.10391 | 0.097358 | 0.10278 | -0.032258 | 0.17035 | 0.12607 | 0.099857 | -0.085497 | 0.13845 | 0.10755 | 0.096673 | 0.068767 | 0.045051 | 0.11071 | 0.10245 | [6 empty] | 0.10497 | 0.10023 | 0.096042 | 0.094023
29 | CCV 10 ppb MAJORS |  | 06/16/2023 15:05:12 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.678 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9998.191 | 10263.8003 | 8.857 | 9.0615 | 10259.6438 | 10055.9356 | 9931.5792 | 10.1704 | 10.489 | 10.3977 | 10.0399 | 10.3779 | 9.8225 | 10.2719 | 10.0946 | 10.1712 | 10.319 | 10.349 | 10.0115 | 10.2913 | 10.2486 | 10.1398 | 10.4099 | 10.2028 | 10.292 | 10.3317 | 10.3477 | 10.1689 | 10.3237 | [6 empty] | 9.8926 | 9.9207 | 9.3714 | 9.774
30 | 9_13_21_F_11_DUP |  | 06/16/2023 15:11:13 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_F_11_DUP.679 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 6661.3322 | 8884.6197 | 9.2147 | 2.6296 | 14635.657 | 1901.3119 | 32311.5268 | 0.7198 | 0.11733 | 0.11484 | 4.4854 | 5.0046 | 5.0335 | 0.016368 | 0.24765 | 2.1914 | 2.2814 | 3.5103 | 4.0506 | 3.7434 | 4.1249 | 0.20292 | 0.1016 | 0.094639 | 0.076293 | 0.37117 | 0.036995 | 2.316 | 2.3 | [6 empty] | 0.010934 | 0.0095419 | 0.11362 | 0.11771
31 | 9_13_21_F_11_DUP_LSP |  | 06/16/2023 15:17:13 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_F_11_DUP_LSP.680 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 16372.2944 | 18220.7526 | 18.303 | 13.5298 | 23548.6851 | 11957.1968 | 40450.9585 | 10.8758 | 10.5394 | 10.4667 | 14.8681 | 14.7905 | 14.809 | 10.2056 | 9.9521 | 11.8482 | 11.8968 | 13.1843 | 13.7685 | 13.8019 | 13.7087 | 10.1245 | 9.7167 | 9.8906 | 9.5655 | 10.7426 | 9.7061 | 12.1925 | 12.2044 | [6 empty] | 9.76 | 9.7339 | 9.3837 | 9.6345
32 | 9_13_21_UF_1 |  | 06/16/2023 15:23:13 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_UF_1.681 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 8600.8984 | 6172.4502 | 1.1442 | 7.7774 | 21031.1995 | 2115.0093 | 33630.2459 | 0.31841 | 0.070932 | 0.056883 | 0.39907 | 13.1481 | 1.3661 | 0.026365 | 0.49956 | 0.59207 | 0.6441 | 1.1866 | 3.9346 | 3.6791 | 3.9304 | 0.32996 | 0.10727 | 0.036967 | 0.039942 | 0.3644 | 0.055555 | 15.4627 | 15.6921 | [6 empty] | 0.0076473 | 0.005176 | 0.041612 | 0.04545
33 | 9_13_21_UF_6 |  | 06/16/2023 15:29:13 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_UF_6.682 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9119.8134 | 7168.0192 | 0.040541 | 2.4342 | 15318.6539 | 1909.3491 | 33678.1192 | 0.64699 | 0.083507 | 0.061443 | 0.12175 | 12.0435 | 1.1733 | 0.016883 | 0.25807 | 1.8047 | 1.868 | 2.2027 | 4.0122 | 3.8794 | 4.1069 | 0.39017 | 0.11304 | 0.095389 | 0.07704 | 0.47657 | 0.074028 | 5.9911 | 6.0122 | [6 empty] | 0.015542 | 0.013462 | 0.058878 | 0.062798
34 | 9_13_21_UF_6_DUP |  | 06/16/2023 15:35:13 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_UF_6_DUP.683 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9054.7298 | 7120.8688 | 0.026518 | 2.4373 | 15410.9294 | 1870.7874 | 33131.715 | 0.65816 | 0.10579 | 0.091806 | 0.2807 | 11.9175 | 1.1758 | 0.01591 | 0.26493 | 1.8118 | 1.8602 | 1.7857 | 4.1471 | 3.9179 | 4.1674 | 0.3714 | 0.077033 | 0.065904 | 0.074097 | 0.462 | 0.062133 | 5.9135 | 5.9981 | [6 empty] | 0.016588 | 0.011743 | 0.062784 | 0.063414
35 | 9_13_21_UF_8 |  | 06/16/2023 15:41:13 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_UF_8.684 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9413.63 | 6987.0859 | -0.0090554 | 2.914 | 13899.2824 | 1983.0529 | 32356.6161 | 0.84496 | 0.074651 | 0.057237 | 0.028408 | 10.6395 | 0.88943 | 0.011977 | 0.25932 | 1.6028 | 1.6104 | 0.98706 | 4.5497 | 3.9851 | 4.5172 | 0.20632 | 0.077473 | 0.080379 | 0.072963 | 0.35976 | 0.058022 | 4.9521 | 5.0425 | [6 empty] | 0.011199 | 0.0070041 | 0.049838 | 0.051821
36 | 9_13_21_UF_9 |  | 06/16/2023 15:47:13 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_UF_9.685 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 8672.9482 | 7152.9978 | 0.2647 | 2.6899 | 12879.5084 | 1879.952 | 29543.5492 | 0.73641 | 0.10046 | 0.077492 | 0.19619 | 3.8521 | 1.1171 | 0.010035 | 0.20771 | 1.4859 | 1.479 | 0.73751 | 3.7713 | 3.6218 | 3.875 | 0.41215 | 0.088717 | 0.071676 | 0.063054 | 0.38646 | 0.08164 | 3.7576 | 3.8449 | [6 empty] | 0.006981 | 0.0034572 | 0.050297 | 0.051291
37 | LBLANK |  | 06/16/2023 15:53:13 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.686 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 54.4446 | 0.23633 | -0.96263 | -0.87372 | 236.4523 | 59.288 | 40.4154 | 2.2133e-05 | -0.011409 | -0.0013358 | -0.57434 | 0.00021345 | -0.40485 | -0.0076191 | 0.00016619 | -0.010467 | -0.011086 | -0.22461 | 0.09778 | 0.035163 | 0.038651 | -0.021279 | 0.0055366 | 0.0076658 | -0.0066793 | -0.0046485 | -0.021474 | -0.0033911 | 0.0030236 | [6 empty] | 0.00060382 | -0.0025156 | -0.010272 | -0.0044217
38 | CCV 1ppb TRACE |  | 06/16/2023 15:59:13 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.687 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1190.5987 | 1169.6923 | 1.2344 | 0.06955 | 1723.162 | 1147.5669 | 1120.4383 | 1.1169 | 1.1415 | 1.0426 | 0.85164 | 1.0106 | 1.3864 | 1.1093 | 1.1316 | 1.1304 | 1.0666 | 0.84399 | 0.96898 | 1.1141 | 1.0042 | 1.0799 | 0.98937 | 1.0329 | 0.95564 | 1.0227 | 0.99694 | 0.97943 | 1.0019 | [6 empty] | 1.0047 | 0.9778 | 1.0507 | 1.0158
39 | CCV 1ppb TRACE |  | 06/16/2023 16:05:32 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.688 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1057.7035 | 1073.3082 | 1.2211 | 0.097238 | 1427.3953 | 1096.0919 | 1040.0608 | 1.0547 | 1.0688 | 1.0731 | 0.92835 | 1.0313 | 1.2401 | 1.0406 | 1.1001 | 1.044 | 1.068 | 0.83757 | 0.92826 | 1.0796 | 1.0157 | 1.1541 | 0.94276 | 1.0118 | 0.97498 | 0.95912 | 1.1056 | 0.99301 | 1.0103 | [6 empty] | 0.99726 | 0.98164 | 1.0647 | 1.0064
40 | CCV 10 ppb MAJORS |  | 06/16/2023 16:11:33 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.689 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9961.7153 | 10510.6335 | 9.7507 | 8.8572 | 10582.1499 | 10164.7328 | 9939.21 | 10.1587 | 10.4579 | 10.3163 | 10.2467 | 10.2249 | 9.8969 | 10.2916 | 9.9266 | 10.1153 | 10.131 | 10.3464 | 9.7164 | 10.2678 | 9.9965 | 10.2444 | 9.977 | 10.1643 | 10.0885 | 10.0647 | 10.1743 | 9.9111 | 10.0907 | [6 empty] | 9.8342 | 9.8947 | 9.6594 | 9.8891
41 | CCV 10 ppb MAJORS |  | 06/16/2023 16:17:52 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.690 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10132.4338 | 10727.9424 | 9.4693 | 8.9241 | 10627.4135 | 10161.8632 | 9907.9138 | 10.2231 | 10.6301 | 10.4494 | 10.0884 | 10.3631 | 10.0249 | 10.3471 | 9.9207 | 10.238 | 10.3375 | 10.5247 | 9.926 | 9.9607 | 10.1212 | 10.1338 | 9.9371 | 10.052 | 9.8966 | 10.2714 | 10.1326 | 10.0575 | 10.1465 | [6 empty] | 9.7552 | 9.8221 | 9.5135 | 9.8035
42 | 9_13_21_UF_10 |  | 06/16/2023 16:23:53 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_UF_10.691 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 7990.4275 | 8816.9371 | 0.52768 | 3.6403 | 14848.7104 | 2008.2639 | 33636.0982 | 0.77593 | 0.075592 | 0.056051 | -0.10492 | 4.3968 | 1.0208 | 0.0087647 | 0.19679 | 1.18 | 1.2209 | 0.85804 | 4.2053 | 3.9632 | 4.2088 | 0.14297 | 0.071487 | 0.07471 | 0.072311 | 0.4533 | 0.057613 | 2.4596 | 2.4743 | [6 empty] | 0.0097692 | 0.0042645 | 0.05325 | 0.057264
43 | 9_13_21_UF_11 |  | 06/16/2023 16:29:52 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_UF_11.692 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 7994.5118 | 9055.3802 | 0.10884 | 3.5752 | 14604.9877 | 2018.1785 | 33979.6864 | 0.76714 | 0.16724 | 0.14546 | 0.82035 | 4.8666 | 1.7196 | 0.0074467 | 0.19257 | 1.4241 | 1.4436 | 1.1555 | 4.1652 | 3.7122 | 4.1654 | 0.2023 | 0.073047 | 0.094714 | 0.073834 | 0.40225 | 0.078638 | 2.2978 | 2.3288 | [6 empty] | 0.0080234 | 0.0069188 | 0.064531 | 0.066184
44 | 9_13_21_LAB_A |  | 06/16/2023 16:35:53 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_LAB_A.693 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 63.8156 | 0.23633 | -0.12208 | -0.62349 | 285.1991 | 59.7555 | 42.3034 | -2.0077e-06 | 0.010343 | 0.014273 | -0.34594 | 0.038506 | -0.25411 | -0.0065759 | 0.027704 | 0.58511 | 0.59532 | 0.12296 | 0.093665 | 0.011442 | 0.023397 | -0.061019 | 0.021136 | 0.0044888 | -0.0058471 | 0.032777 | -0.021474 | -0.0044156 | 0.0019025 | [6 empty] | 0.00060519 | -0.0019527 | 0.018159 | 0.02355
45 | 9_13_21_BB |  | 06/16/2023 16:41:52 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_BB.694 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 56.6196 | 0.23633 | 0.09474 | -0.63072 | 229.6033 | 62.8683 | 45.752 | 0.0012963 | 0.0071925 | 0.016015 | -0.36435 | 0.0026256 | -0.21291 | -0.0073475 | 0.0066767 | 0.12189 | 0.12284 | -0.04338 | 0.056527 | 0.0054457 | 0.013177 | -0.1341 | 0.0055366 | 0.0028146 | -0.0058124 | -0.03456 | -0.017025 | -0.0055398 | 0.0031128 | [6 empty] | 0.00081119 | -0.001438 | 0.019392 | 0.024504
46 | 9_13_21_FB |  | 06/16/2023 16:47:52 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\9_13_21_FB.695 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 236.6815 | 43.7084 | 14.4759 | 0.6 | 637.7343 | 126.7295 | 270.1861 | 0.015376 | 0.075294 | 0.080747 | 0.76598 | 0.30634 | 1.1792 | 0.0046164 | 0.048366 | 0.2146 | 0.25594 | 18.6562 | 0.017427 | 0.05652 | 0.04947 | 0.074959 | 0.016188 | 0.0079667 | 0.0070413 | 0.0023331 | -0.012523 | 0.096043 | 0.10539 | [6 empty] | 0.0097213 | 0.0067231 | 0.089345 | 0.092484
47 | 5_18_23_MFB |  | 06/16/2023 16:53:52 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_MFB.696 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4836.2346 | 5082.0378 | 4.6388 | 8.5802 | 5184.0438 | 4837.1276 | 4882.1192 | 4.9302 | 5.2641 | 5.2 | 4.6128 | 5.0652 | 4.6849 | 5.037 | 5.0001 | 5.1198 | 5.1607 | 4.8593 | 4.3888 | 5.0084 | 4.8807 | 4.4037 | 5.1275 | 4.8968 | 4.9452 | 4.5838 | 5.1712 | 4.8166 | 4.8727 | [6 empty] | 4.8335 | 4.8479 | 4.658 | 5.0361
48 | 5_18_23_DB |  | 06/16/2023 16:59:52 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_DB.697 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 69.2273 | 12.5222 | -0.27238 | 4.1911 | 484.1611 | 72.9149 | 59.3009 | 0.013734 | 0.010734 | 0.018163 | -0.31381 | 0.01657 | -0.025073 | 0.0073802 | 0.010834 | 0.073513 | 0.074039 | -0.11093 | 0.014335 | 0.012004 | 0.015188 | 0.069414 | 0.052083 | 0.016129 | 0.012074 | 0.0042037 | 0.00037767 | 0.0061381 | 0.014442 | [6 empty] | 0.014689 | 0.011044 | 0.024244 | 0.030415
49 | 5_18_23_T_239 |  | 06/16/2023 17:05:52 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_T_239.698 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 46409.2332 | 13062.6293 | 115.9872 | 12.2818 | 1173.2444 | 3943.0834 | 49867.5428 | 1.9186 | 1.4536 | 1.4751 | 43.2952 | 3.69 | 41.8006 | 1.4258 | 1.7129 | 0.77367 | 0.85401 | 6.0938 | 2.5039 | 2.2333 | 2.4468 | 5.416 | 2.0766 | 1.97 | 1.9029 | 2.3068 | 1.9684 | 1.0927 | 1.1055 | [6 empty] | 3.3754 | 3.3183 | 1.4041 | 1.2647
50 | 5_18_23_T_239 |  | 06/16/2023 17:11:52 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_T_239.699 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 48071.2603 | 13505.3814 | 121.1505 | 12.8964 | 1362.1759 | 4062.0503 | 53018.5233 | 2.0132 | 1.602 | 1.4479 | 42.6029 | 3.6219 | 44.1082 | 1.5019 | 1.8124 | 0.85556 | 0.87471 | 6.0507 | 2.434 | 2.4273 | 2.3709 | 4.9176 | 1.9015 | 1.7989 | 1.8994 | 2.1095 | 1.938 | 1.1136 | 1.0972 | [6 empty] | 3.3567 | 3.2663 | 1.2908 | 1.2357
51 | 5_18_23_M_224 |  | 06/16/2023 17:17:52 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_M_224.700 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4052.0246 | 642.9194 | 7.3955 | 17.0722 | 916.4976 | 519.1015 | 2963.3336 | 1.0225 | 0.031492 | 0.031926 | -0.31224 | 5.5129 | -0.091135 | 0.0065512 | 0.051456 | 0.064557 | 0.062569 | 0.69528 | 0.23479 | 0.040533 | 0.059438 | 0.23278 | 0.0679 | 0.019298 | 0.016721 | 0.76179 | 0.036066 | 0.2415 | 0.20725 | [6 empty] | 0.0031649 | 0.00078166 | 0.0089421 | 0.012469
52 | LBLANK |  | 06/16/2023 17:23:52 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.701 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 55.8376 | 0.70899 | -1.0141 | -0.18597 | 157.8543 | 64.6363 | 41.096 | 0.0010043 | -0.0086365 | -0.0021681 | -0.58445 | -0.00097667 | -0.38029 | -0.007202 | 0.0047496 | -0.0094054 | -0.010432 | -0.22261 | -0.020689 | -0.0078919 | -0.00011531 | 0.21802 | 0.0055366 | 0.0028999 | -0.0080639 | -0.014407 | -0.021474 | 0.017648 | 0.0020502 | [6 empty] | 0.00050427 | -0.0021097 | -0.010416 | -0.004357
53 | CCV 1ppb TRACE |  | 06/16/2023 17:29:53 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.702 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 987.8331 | 1010.777 | 1.2439 | 0.23021 | 1094.2251 | 1046.5265 | 971.0796 | 1.0204 | 1.0526 | 1.0616 | 0.90075 | 1.0129 | 1.035 | 1.0071 | 0.96643 | 1.0699 | 1.0575 | 0.8381 | 0.96922 | 0.9777 | 0.97344 | 1.046 | 1.1939 | 1.0451 | 0.97312 | 0.98967 | 0.95556 | 1.0014 | 1.0082 | [6 empty] | 0.9601 | 0.97936 | 1.0383 | 1.0069
54 | CCV 1ppb TRACE |  | 06/16/2023 17:36:12 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.703 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1009.6197 | 1032.865 | 1.0845 | 0.13398 | 1218.9479 | 1035.6432 | 996.5243 | 0.99993 | 1.0281 | 1.0567 | 0.85323 | 1.0119 | 1.2074 | 1.005 | 1.0396 | 1.0711 | 1.0542 | 0.83458 | 0.90204 | 1.0035 | 0.99199 | 1.1499 | 0.89965 | 0.94656 | 0.99771 | 1.0148 | 0.96448 | 0.99645 | 1.0146 | [6 empty] | 0.97104 | 0.98178 | 1.0613 | 1.0075
55 | LFB |  | 06/16/2023 17:42:12 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.704 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9318.2332 | 9688.4626 | 9.7361 | 11.4387 | 10212.8116 | 9450.1073 | 9627.478 | 9.6244 | 10.0298 | 9.8195 | 9.6886 | 9.8477 | 9.3763 | 9.8275 | 9.545 | 9.8513 | 9.8229 | 9.7185 | 9.5215 | 10.0351 | 9.8384 | 10.3697 | 10.479 | 9.8586 | 9.9861 | 10.1133 | 9.8555 | 9.5679 | 9.8013 | [6 empty] | 9.4964 | 9.4689 | 9.4318 | 9.6216
56 | CCV 10 ppb MAJORS |  | 06/16/2023 17:48:11 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.705 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9642.9197 | 10168.3327 | 9.0177 | 8.9859 | 10240.0218 | 9932.3643 | 9801.3678 | 9.9058 | 10.2795 | 10.4564 | 10.2546 | 10.287 | 9.766 | 10.216 | 9.7392 | 10.0267 | 10.2615 | 10.4433 | 9.5325 | 9.8829 | 9.9566 | 10.3203 | 9.9852 | 9.9118 | 10.1161 | 9.7431 | 9.7936 | 10.0474 | 10.2082 | [6 empty] | 9.7785 | 9.8097 | 9.5589 | 9.9094
57 | CCV 10 ppb MAJORS |  | 06/16/2023 17:54:30 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.706 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9842.1164 | 10306.0446 | 9.6058 | 9.0394 | 10768.3658 | 9919.952 | 10003.6006 | 10.227 | 10.4489 | 10.5586 | 10.1865 | 10.2453 | 9.9996 | 10.4423 | 10.1498 | 10.3751 | 10.3184 | 10.385 | 9.8983 | 9.9591 | 9.8456 | 10.486 | 9.6585 | 10.0605 | 9.9058 | 10.3007 | 10.396 | 10.0906 | 10.1703 | [6 empty] | 9.6967 | 9.7855 | 9.6044 | 9.8314
58 | 5_18_23_MFB |  | 06/16/2023 18:00:31 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_MFB.707 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4976.5158 | 5147.4868 | 4.9909 | 8.543 | 5454.665 | 4914.5063 | 5048.8721 | 5.0602 | 5.2755 | 5.6289 | 5.309 | 5.1463 | 4.9703 | 5.2693 | 5.1257 | 5.25 | 5.5752 | 5.242 | 4.7392 | 5.1705 | 4.9592 | 4.9333 | 4.7759 | 4.9131 | 4.8191 | 5.0688 | 5.0131 | 5.2372 | 4.9371 | [6 empty] | 5.1679 | 5.1922 | 4.9494 | 4.8554
59 | 10_1_21_W_1 |  | 06/16/2023 18:06:31 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_W_1.708 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1809.7496 | 1366.4408 | 5.0964 | 3.1972 | 5506.4698 | 502.6889 | 7806.1746 | 0.053229 | 0.0098102 | 0.016988 | 7.5556 | 5.789 | 7.7203 | 0.0014734 | 0.079968 | 0.46544 | 0.46547 | 0.16074 | 0.69481 | 0.66979 | 0.67812 | 0.12079 | 0.032968 | 0.022083 | -0.0022877 | 0.077872 | 0.0015327 | 3.5298 | 3.5381 | [6 empty] | 0.0026279 | 6.599e-05 | 0.022528 | 0.026239
60 | 10_1_21_W_6 |  | 06/16/2023 18:12:31 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_W_6.709 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1829.2879 | 1606.9129 | 8.9897 | 1.7914 | 4261.9224 | 455.8822 | 7633.578 | 0.14421 | 0.013414 | 0.017521 | 9.3953 | 3.2925 | 9.5866 | 0.00098277 | 0.075184 | 0.84881 | 0.83956 | 0.64956 | 0.7535 | 0.65615 | 0.78522 | 0.040281 | 0.044644 | 0.01374 | 0.013881 | 0.056784 | -0.0026399 | 1.5877 | 1.6668 | [6 empty] | 0.0057574 | 0.0015786 | 0.067478 | 0.07909
61 | 10_1_21_W_8 |  | 06/16/2023 18:18:32 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_W_8.710 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1861.2963 | 1519.0654 | 3.4101 | 1.9807 | 3849.6145 | 458.502 | 7373.0043 | 0.16039 | 0.019854 | 0.020336 | 5.1233 | 2.4486 | 5.2766 | -0.0034375 | 0.045822 | 0.82341 | 0.82485 | 0.51531 | 0.80668 | 0.69827 | 0.85654 | 0.10225 | 0.033808 | 0.026485 | 0.006546 | 0.061177 | -0.0070859 | 1.4505 | 1.4482 | [6 empty] | 0.0030228 | 6.2098e-05 | 0.026713 | 0.0328
62 | 10_1_21_W_9 |  | 06/16/2023 18:24:32 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_W_9.711 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1568.3974 | 1380.7452 | 5.1934 | 1.7039 | 3104.5776 | 418.9389 | 6409.642 | 0.13384 | 0.043758 | 0.045602 | 5.3406 | 1.8014 | 5.4711 | -0.0028335 | 0.033099 | 0.56246 | 0.56041 | 0.1864 | 0.64656 | 0.67178 | 0.66101 | 0.10495 | 0.016749 | 0.013764 | 0.0055303 | 0.052236 | 0.0022145 | 0.99206 | 1.0037 | [6 empty] | 0.0043542 | 0.00058309 | 0.025359 | 0.031942
63 | 10_1_21_W_10 |  | 06/16/2023 18:30:32 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_W_10.712 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1389.6206 | 1711.7712 | 30.1756 | 3.435 | 3384.5393 | 420.0594 | 6934.8936 | 0.14515 | 0.027355 | 0.03211 | 14.9574 | 3.7925 | 14.4594 | 0.0030418 | 0.051527 | 0.61913 | 0.63275 | 0.67065 | 0.73793 | 0.62318 | 0.75786 | 0.18985 | 0.0169 | 0.030373 | 0.0032508 | 0.085612 | -0.00689 | 0.58556 | 0.61175 | [6 empty] | 0.0040781 | 0.002379 | 0.10042 | 0.10142
64 | 10_1_21_W_11 |  | 06/16/2023 18:36:31 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_W_11.713 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1334.4543 | 1653.3531 | 21.1115 | 3.12 | 3350.5535 | 419.7848 | 6718.1624 | 0.16741 | 0.047703 | 0.040418 | 17.5419 | 3.8911 | 16.7815 | 0.0035638 | 0.055717 | 0.78855 | 0.81081 | 0.84412 | 0.77576 | 0.71771 | 0.7898 | 0.17462 | 0.027946 | 0.024552 | 0.0030542 | -0.00010942 | 0.011738 | 0.55225 | 0.56414 | [6 empty] | 0.0047143 | 0.002778 | 0.21187 | 0.20855
65 | 10_1_21_W_11_LD |  | 06/16/2023 18:42:31 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_W_11_LD.714 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 54.4392 | 0.23633 | -1.0383 | -0.92404 | -2381.0441 | 46.3568 | 39.6077 | -0.00066774 | -0.016067 | -0.0033762 | -0.61044 | -0.0009563 | -2.3984 | -0.0076572 | -0.0085046 | -0.013579 | -0.01327 | -0.23067 | 0.73465 | -0.0056089 | -0.0041382 | -1.8771 | 0.0055366 | -0.00050913 | -0.010059 | -0.062944 | -0.021474 | -0.0012166 | 0.0013969 | [6 empty] | 0.00059194 | -0.0027606 | -0.011182 | -0.0047675
66 | 10_1_21_W_11_LD |  | 06/16/2023 18:48:48 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_W_11_LD.715 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1339.1526 | 1688.0961 | 22.0867 | 2.483 | 3244.2172 | 424.2943 | 6864.056 | 0.15839 | 0.047825 | 0.052847 | 17.8769 | 4.057 | 16.8064 | 0.0058586 | 0.050277 | 0.76652 | 0.81322 | 0.8674 | 0.80021 | 0.87521 | 0.84892 | 0.16215 | 0.034303 | 0.021451 | -2.8715e-05 | 0.031092 | -0.011745 | 0.57455 | 0.58916 | [6 empty] | 0.0059181 | 0.0033105 | 0.21835 | 0.21327
67 | 10_1_21_W_11_DUP |  | 06/16/2023 18:54:48 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_W_11_DUP.716 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1391.7828 | 1697.8068 | 23.0783 | 2.4225 | 3279.9744 | 401.5499 | 6786.8219 | 0.15423 | 0.026191 | 0.027036 | 18.466 | 4.2314 | 17.8155 | 0.0023811 | 0.045374 | 0.80496 | 0.81933 | 0.87544 | 0.78012 | 0.74155 | 0.81046 | 0.0074674 | 0.034973 | 0.031283 | 0.014287 | 0.11711 | -0.0015615 | 0.55047 | 0.57562 | [6 empty] | 0.0049392 | 0.0041534 | 0.13131 | 0.12775
68 | LBLANK |  | 06/16/2023 19:00:48 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.717 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 54.1008 | 0.74518 | -0.99077 | -0.58503 | 373.8525 | 57.4128 | 39.1036 | 0.0011378 | -0.0072258 | -0.0012028 | -0.58948 | -0.00056814 | -0.25972 | -0.0077432 | 0.0035129 | -0.00485 | -0.010419 | -0.22553 | 0.077648 | -0.01135 | -0.0034741 | -0.026405 | 0.0055366 | 0.0013136 | -0.011205 | 0.0074769 | -0.0022024 | -0.0050221 | 0.0024088 | [6 empty] | 0.00061183 | -0.0020735 | -0.010642 | -0.0044551
69 | CCV 1ppb TRACE |  | 06/16/2023 19:06:48 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.718 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1002.6883 | 1004.4991 | 1.3914 | 0.16075 | 1337.7963 | 1006.692 | 987.1159 | 1.0297 | 1.0642 | 1.0889 | 0.92501 | 0.99538 | 1.0695 | 1.0281 | 1.0015 | 1.0313 | 1.0631 | 0.85139 | 0.91214 | 1.0631 | 0.9722 | 1.0056 | 1.0267 | 0.98058 | 0.97723 | 0.95232 | 1.0954 | 1.0007 | 1.0144 | [6 empty] | 0.983 | 0.98388 | 1.087 | 1.0288
70 | CCV 1ppb TRACE |  | 06/16/2023 19:13:07 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.719 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1013.9286 | 1074.6779 | 1.2329 | 0.11892 | 1416.6104 | 1024.4187 | 1009.3335 | 1.0007 | 1.0725 | 1.0779 | 0.94758 | 1.0146 | 1.0703 | 1.009 | 0.98826 | 1.0691 | 1.0609 | 0.84908 | 0.96223 | 0.93815 | 0.98826 | 1.0093 | 1.0598 | 0.96533 | 0.97458 | 1.0173 | 1.0478 | 1.0036 | 1.0065 | [6 empty] | 0.97373 | 1.0071 | 1.0945 | 1.0464
71 | CCV 10 ppb MAJORS |  | 06/16/2023 19:19:08 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.720 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9635.2676 | 10151.3372 | 9.4319 | 8.8801 | 10594.1633 | 9744.4391 | 9907.7371 | 9.8795 | 10.1958 | 10.4225 | 10.0754 | 10.2562 | 9.6368 | 10.0051 | 9.6712 | 10.0261 | 10.1498 | 10.3049 | 9.8162 | 9.9649 | 9.9857 | 10.6706 | 9.934 | 10.0132 | 10.0792 | 10.4874 | 9.8281 | 9.9381 | 10.214 | [6 empty] | 9.6553 | 9.7762 | 10.0501 | 10.1056
72 | CCV 10 ppb MAJORS |  | 06/16/2023 19:25:27 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.721 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9689.1906 | 9953.3571 | 9.403 | 8.9625 | 10606.1927 | 9782.6699 | 9988.6639 | 10.0685 | 10.5402 | 10.6579 | 10.2485 | 10.4787 | 9.9744 | 10.2245 | 9.8293 | 10.151 | 10.2338 | 10.3994 | 9.7787 | 10.1096 | 10.0639 | 10.5556 | 10.3708 | 10.2393 | 10.1633 | 10.5873 | 10.6943 | 10.0407 | 10.2687 | [6 empty] | 9.7836 | 9.7555 | 9.7335 | 9.9283
73 | 10_1_21_W_11_DUP_LSP |  | 06/16/2023 19:31:27 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_W_11_DUP_LSP.722 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 21099.4095 | 21681.7785 | 50.5648 | 23.655 | 23812.2807 | 20403.4455 | 27095.5336 | 20.7957 | 21.2822 | 20.8528 | 40.2941 | 25.1748 | 39.462 | 20.7043 | 19.8801 | 20.7994 | 20.4969 | 20.9631 | 20.7963 | 20.5839 | 21.1385 | 20.3535 | 20.5938 | 20.0012 | 20.2836 | 21.0746 | 20.9631 | 20.5648 | 21.1591 | [6 empty] | 19.0252 | 19.1029 | 17.47 | 19.1188
74 | 10_1_21_F_1 |  | 06/16/2023 19:37:27 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_F_1.723 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 8742.7159 | 6405.1866 | 3.167 | 6.0707 | 25855.2236 | 2306.763 | 37692.4243 | 0.23152 | 0.073985 | 0.073037 | 11.4893 | 18.1682 | 12.0924 | 0.022389 | 0.42714 | 0.84493 | 0.90022 | 14.784 | 2.9138 | 2.704 | 2.9412 | 0.44004 | 0.071566 | 0.043453 | 0.02738 | 0.47343 | 0.039489 | 17.0741 | 17.6609 | [6 empty] | 0.010785 | 0.0063999 | 0.05286 | 0.056564
75 | 10_1_21_F_1_DUP |  | 06/16/2023 19:43:27 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_F_1_DUP.724 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9263.3257 | 6778.5892 | 2.9322 | 5.5897 | 26615.3499 | 2410.8315 | 39017.5819 | 0.22279 | 0.062003 | 0.072469 | 8.2742 | 18.699 | 8.941 | 0.028896 | 0.48295 | 0.79325 | 0.88599 | 14.4302 | 2.9967 | 2.7682 | 3.0565 | 0.46945 | 0.054116 | 0.061047 | 0.040309 | 0.41476 | 0.060321 | 17.7716 | 18.1895 | [6 empty] | 0.010208 | 0.0058715 | 0.051917 | 0.054907
76 | 10_1_21_F_6 |  | 06/16/2023 19:49:26 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_F_6.725 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9298.8645 | 7693.4769 | 2.9702 | 2.1459 | 19871.8072 | 2124.291 | 39223.3973 | 0.61224 | 0.068168 | 0.073008 | 7.3023 | 11.2861 | 8.1134 | 0.014791 | 0.28575 | 2.1475 | 2.1909 | 5.2438 | 3.4724 | 3.1372 | 3.5601 | 0.32161 | 0.077464 | 0.090957 | 0.058505 | 0.41388 | 0.079956 | 8.0802 | 8.2927 | [6 empty] | 0.014942 | 0.012828 | 0.093 | 0.096764
77 | 10_1_21_F_6_DUP |  | 06/16/2023 19:55:26 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_F_6_DUP.726 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9073.1547 | 7643.3788 | 3.1145 | 2.0632 | 19827.4014 | 2064.2721 | 37819.6465 | 0.6252 | 0.0774 | 0.071969 | 7.3302 | 11.3088 | 8.1017 | 0.015033 | 0.24745 | 2.1239 | 2.2132 | 5.949 | 3.5128 | 3.1918 | 3.5796 | 0.38132 | 0.090731 | 0.067165 | 0.074158 | 0.51767 | 0.055653 | 7.9611 | 8.3101 | [6 empty] | 0.013986 | 0.0092485 | 0.10072 | 0.10339
78 | 10_1_21_F_8 |  | 06/16/2023 20:01:27 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_F_8.727 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9802.4587 | 7873.6674 | 2.9216 | 2.1788 | 18423.7304 | 2146.8071 | 38184.1905 | 0.80314 | 0.059471 | 0.057574 | 8.5179 | 10.7656 | 9.2095 | 0.012868 | 0.21698 | 2.3688 | 2.4325 | 5.689 | 4.0224 | 3.7941 | 4.0989 | 0.53463 | 0.090242 | 0.087958 | 0.083625 | 0.45329 | 0.060293 | 7.1881 | 7.4649 | [6 empty] | 0.0090979 | 0.0088977 | 0.20797 | 0.2049
79 | 10_1_21_F_9 |  | 06/16/2023 20:07:27 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_F_9.728 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 8271.2277 | 7374.2568 | 3.0834 | 2.8137 | 15606.1718 | 1905.2783 | 33829.3697 | 0.64553 | 0.11516 | 0.12303 | 5.8875 | 5.2989 | 6.6274 | 0.011546 | 0.2179 | 2.2008 | 2.2179 | 6.9589 | 3.1205 | 2.9042 | 3.1696 | 0.57604 | 0.089734 | 0.098979 | 0.068709 | 0.4393 | 0.10568 | 4.9655 | 5.1406 | [6 empty] | 0.01079 | 0.009053 | 0.56867 | 0.54358
80 | 10_1_21_F_9_DUP |  | 06/16/2023 20:13:27 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_F_9_DUP.729 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 8050.5994 | 7277.0856 | 3.9511 | 2.8481 | 15276.1024 | 1878.9129 | 33176.8647 | 0.66988 | 0.096024 | 0.10924 | 5.5729 | 5.2829 | 6.3201 | 0.026412 | 0.24651 | 2.2178 | 2.1847 | 6.9416 | 3.0468 | 2.9396 | 3.1595 | 0.492 | 0.053814 | 0.05712 | 0.071754 | 0.39993 | 0.070376 | 4.9542 | 5.0875 | [6 empty] | 0.010424 | 0.008658 | 0.090634 | 0.092873
81 | 10_1_21_F_10 |  | 06/16/2023 20:19:27 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_F_10.730 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 7215.9152 | 8769.5475 | 7.2734 | 3.7305 | 15496.8067 | 1938.2461 | 35354.5845 | 0.71378 | 0.053591 | 0.058758 | 5.449 | 5.7516 | 6.1681 | 0.013086 | 0.18023 | 1.642 | 1.6867 | 7.6751 | 3.6392 | 3.2996 | 3.6945 | 0.62804 | 0.072647 | 0.061787 | 0.092025 | 0.43305 | 0.045708 | 3.0536 | 3.1571 | [6 empty] | 0.0081257 | 0.0063848 | 18.7896 | 20.6667
82 | 10_1_21_F_11 |  | 06/16/2023 20:25:27 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_F_11.731 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 7110.057 | 8822.8049 | 5.9823 | 2.7911 | 15956.1457 | 1911.5726 | 35337.4308 | 0.73616 | 0.090888 | 0.099734 | 6.8677 | 5.0239 | 7.4942 | 0.013871 | 0.19325 | 1.5776 | 1.6761 | 7.2731 | 3.6826 | 3.3321 | 3.8473 | 0.48399 | 0.090918 | 0.08483 | 0.078173 | 0.3447 | 0.045618 | 2.918 | 2.9617 | [6 empty] | 0.0087624 | 0.0053262 | 0.10438 | 0.10478
83 | LBLANK |  | 06/16/2023 20:31:27 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.732 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 55.3543 | 0.72782 | -1.0117 | -0.91666 | 447.6718 | 60.514 | 39.8091 | -0.00057673 | -0.0069438 | -0.0012405 | -0.57586 | -0.00052007 | -0.26616 | -0.0078885 | 0.0018702 | -0.011305 | -0.0096492 | -0.22954 | -0.016738 | -0.01135 | -0.0037458 | 0.0040065 | 0.0055366 | 0.0048521 | -0.01206 | -0.013452 | -0.016689 | -0.006875 | 0.0028274 | [6 empty] | 0.00050427 | -0.0023379 | -0.010438 | -0.004459
84 | CCV 1ppb TRACE |  | 06/16/2023 20:37:27 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.733 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1036.9693 | 1078.0255 | 1.2511 | 0.12445 | 1274.8891 | 1063.1539 | 952.9247 | 1.058 | 1.0708 | 1.104 | 0.95774 | 1.0155 | 1.1228 | 1.0253 | 0.97374 | 1.0594 | 1.0654 | 0.85663 | 0.93263 | 1.0259 | 0.98818 | 1.0653 | 0.97813 | 0.96768 | 0.95858 | 0.99066 | 0.90136 | 0.97415 | 0.99372 | [6 empty] | 0.96191 | 0.99069 | 1.0722 | 1.0157
85 | CCV 1ppb TRACE |  | 06/16/2023 20:43:46 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.734 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1040.4587 | 1103.9276 | 1.193 | 0.10895 | 1444.0459 | 1062.3671 | 1072.8531 | 1.0397 | 1.0729 | 1.0873 | 0.96254 | 1.0395 | 1.2168 | 1.029 | 1.0613 | 1.0792 | 1.0569 | 0.84771 | 0.9166 | 0.94519 | 0.9981 | 1.0994 | 1.105 | 0.94652 | 0.94985 | 0.97181 | 0.9856 | 0.98237 | 1.0205 | [6 empty] | 0.97908 | 0.97651 | 1.0438 | 1.0203
86 | LFB |  | 06/16/2023 20:49:46 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.735 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9466.6541 | 9904.8327 | 9.4772 | 11.9438 | 10401.5438 | 9535.8912 | 9640.5655 | 9.5913 | 10.0779 | 10.0633 | 9.7932 | 9.9504 | 9.2658 | 9.741 | 9.4367 | 9.7016 | 9.8291 | 9.7731 | 9.8013 | 10.0316 | 9.739 | 10.624 | 9.8981 | 9.9175 | 9.9831 | 10.3206 | 9.9436 | 9.5256 | 9.8087 | [6 empty] | 9.4004 | 9.4604 | 9.7719 | 9.6345
87 | CCV 10 ppb MAJORS |  | 06/16/2023 20:55:46 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.736 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9907.5254 | 10350.9786 | 9.9565 | 9.377 | 10682.0945 | 9945.1814 | 10117.4077 | 10.2223 | 10.5432 | 10.7419 | 10.3071 | 10.4417 | 9.7082 | 10.1628 | 9.7941 | 10.0052 | 10.3036 | 10.5344 | 9.7441 | 10.2478 | 9.9968 | 10.3172 | 10.0712 | 10.1348 | 10.1377 | 10.582 | 10.2023 | 9.9813 | 10.2245 | [6 empty] | 9.7546 | 9.7844 | 9.7097 | 9.9626
88 | 10_1_21_UF_1 |  | 06/16/2023 21:01:46 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_UF_1.737 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9565.9098 | 6356.3857 | 0.34999 | 4.9861 | 25052.5414 | 2310.0364 | 37921.4774 | 0.23106 | 0.081269 | 0.064605 | 0.37956 | 17.867 | 1.4258 | 0.021832 | 0.38227 | 0.57217 | 0.62899 | 12.3857 | 2.7295 | 2.6363 | 2.8453 | 0.43792 | 0.060909 | 0.046443 | 0.037455 | 0.42122 | 0.046161 | 17.0459 | 17.3986 | [6 empty] | 0.011255 | 0.0054113 | 0.022377 | 0.029689
89 | 10_1_21_UF_6 |  | 06/16/2023 21:07:46 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_UF_6.738 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10567.788 | 7987.2439 | 0.52271 | 3.1185 | 20518.5617 | 2195.1584 | 40248.9315 | 0.62622 | 0.095584 | 0.081501 | 0.28123 | 11.4151 | 1.4518 | 0.014599 | 0.26024 | 1.4437 | 1.5485 | 4.4138 | 3.3061 | 3.3731 | 3.5376 | 0.42696 | 0.067389 | 0.076407 | 0.080424 | 0.42684 | 0.10384 | 8.1382 | 8.2675 | [6 empty] | 0.012277 | 0.0088445 | 0.046355 | 0.054753
90 | 10_1_21_UF_8 |  | 06/16/2023 21:13:46 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_UF_8.739 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10287.4416 | 7656.4409 | 0.52817 | 4.573 | 18040.6633 | 2133.4152 | 37895.7799 | 0.8091 | 0.16492 | 0.16073 | 0.88635 | 10.322 | 2.0103 | 0.017304 | 0.21873 | 1.2939 | 1.3682 | 4.395 | 3.8363 | 3.8934 | 3.9829 | 0.52551 | 0.080604 | 0.077414 | 0.074223 | 0.46841 | 0.068935 | 6.9984 | 7.159 | [6 empty] | 0.01133 | 0.0056188 | 0.046361 | 0.051983
91 | 10_1_21_UF_8_LD |  | 06/16/2023 21:19:45 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_UF_8_LD.740 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10423.8202 | 7773.2355 | 1.4202 | 7.7731 | 17452.2012 | 2117.3708 | 36943.8326 | 0.79313 | 0.16307 | 0.14829 | 0.88894 | 10.1712 | 1.8357 | 0.011939 | 0.23289 | 1.2906 | 1.3532 | 4.3519 | 3.865 | 3.6391 | 3.9593 | 0.70811 | 0.061687 | 0.090947 | 0.071889 | 0.45243 | 0.084089 | 7.0082 | 7.1914 | [6 empty] | 0.010528 | 0.0060094 | 0.044366 | 0.052771
92 | 10_1_21_UF_8_LD |  | 06/16/2023 21:26:03 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_UF_8_LD.741 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10273.4334 | 7508.2639 | 1.4141 | 8.1758 | 17632.898 | 2164.6772 | 36960.6201 | 0.82172 | 0.17465 | 0.15192 | 0.86339 | 10.2554 | 1.758 | 0.012183 | 0.24061 | 1.2665 | 1.36 | 4.3385 | 3.8132 | 3.5223 | 3.9473 | 0.52993 | 0.05594 | 0.087908 | 0.075636 | 0.48323 | 0.047787 | 6.9729 | 7.2246 | [6 empty] | 0.011505 | 0.0077663 | 0.045513 | 0.050741
93 | 10_1_21_UF_8_DUP |  | 06/16/2023 21:32:02 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_UF_8_DUP.742 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10395.9339 | 7712.0563 | 0.49512 | 5.3695 | 17918.3544 | 2159.0861 | 37599.3298 | 0.79388 | 0.28663 | 0.2676 | 1.699 | 10.2653 | 2.6554 | 0.015724 | 0.23428 | 1.3991 | 1.406 | 4.8656 | 3.9155 | 3.4748 | 3.9794 | 0.42577 | 0.10131 | 0.084821 | 0.089093 | 0.51932 | 0.081039 | 7.0821 | 7.2285 | [6 empty] | 0.0096322 | 0.0088035 | 0.053256 | 0.055655
94 | 10_1_21_UF_8_DUP_LSP |  | 06/16/2023 21:38:02 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_UF_8_DUP_LSP.743 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 19395.3071 | 16899.0275 | 11.1808 | 16.4541 | 26417.594 | 12033.9032 | 43587.2902 | 11.0948 | 10.6658 | 10.7603 | 12.718 | 19.5298 | 12.6553 | 10.179 | 9.692 | 10.9061 | 11.1057 | 14.1191 | 13.1521 | 13.3973 | 13.3559 | 10.7136 | 9.8977 | 10.2027 | 9.9796 | 10.9439 | 9.791 | 16.3919 | 16.5672 | [6 empty] | 9.5021 | 9.5257 | 9.4286 | 9.751
95 | 10_1_21_UF_9 |  | 06/16/2023 21:44:02 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_UF_9.744 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9181.1223 | 7265.0967 | 0.090679 | 5.4856 | 15397.3048 | 1878.5572 | 34058.2355 | 0.65891 | 0.07963 | 0.05955 | 0.56798 | 5.1421 | 1.5824 | 0.011856 | 0.20379 | 1.3465 | 1.3943 | 6.0107 | 3.016 | 3.0197 | 3.158 | 0.43566 | 0.07489 | 0.10003 | 0.073556 | 0.38209 | 0.079943 | 4.9891 | 5.1125 | [6 empty] | 0.0099357 | 0.0089202 | 0.040606 | 0.047614
96 | 10_1_21_UF_10 |  | 06/16/2023 21:50:03 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_UF_10.745 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 8285.4207 | 8770.3211 | 0.012551 | 4.5596 | 15678.3237 | 1938.0638 | 35826.3709 | 0.73358 | 0.29517 | 0.28494 | 1.6836 | 5.4361 | 2.5709 | 0.010188 | 0.19433 | 1.0301 | 1.0762 | 5.0537 | 3.276 | 3.2359 | 3.6643 | 0.3434 | 0.094532 | 0.094293 | 0.074442 | 0.38877 | 0.080798 | 3.038 | 3.1485 | [6 empty] | 0.0069081 | 0.0042631 | 0.057783 | 0.063994
97 | 10_1_21_UF_11 |  | 06/16/2023 21:56:03 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_UF_11.746 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 7796.2775 | 8747.7439 | -0.045743 | 3.8452 | 15584.5364 | 1913.1878 | 35413.076 | 0.71627 | 0.067322 | 0.060274 | 1.6133 | 4.6069 | 2.4692 | 0.010103 | 0.16813 | 0.94025 | 1.0007 | 5.2103 | 3.3255 | 3.3699 | 3.7391 | 0.38255 | 0.11298 | 0.091824 | 0.085954 | 0.31002 | 0.074532 | 2.8191 | 2.873 | [6 empty] | 0.0073545 | 0.0044562 | 0.045802 | 0.054742
98 | LBLANK |  | 06/16/2023 22:02:04 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.747 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 55.3996 | 2.2875 | -0.90892 | -0.86876 | 710.7681 | 58.875 | 42.0219 | -0.00021605 | -0.006893 | -0.0011939 | -0.57746 | -0.00077284 | -0.32539 | -0.0075884 | 0.0085943 | -0.012158 | -0.010057 | -0.22863 | 0.037665 | -0.0037923 | 0.01697 | 0.0048278 | 0.011408 | 0.0089106 | -0.0076485 | -0.040154 | -0.021474 | -0.0075587 | 0.003048 | [6 empty] | 0.00083952 | -0.0021333 | -0.010903 | -0.0044755
99 | CCV 1ppb TRACE |  | 06/16/2023 22:08:04 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.748 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1025.6458 | 1070.4949 | 1.142 | 0.089874 | 1758.7854 | 1038.7734 | 1043.2635 | 1.0473 | 1.0654 | 1.1002 | 0.95648 | 1.0125 | 0.99979 | 1.0335 | 1.0321 | 1.0676 | 1.0825 | 0.83832 | 0.94036 | 1.0034 | 0.98286 | 1.1848 | 1.1063 | 0.996 | 0.96327 | 0.93271 | 0.98896 | 0.97105 | 1.0138 | [6 empty] | 0.98231 | 0.96731 | 1.1049 | 1.0381
100 | CCV 1ppb TRACE |  | 06/16/2023 22:14:23 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.749 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1042.9753 | 1077.1368 | 1.0946 | 0.055224 | 1588.5975 | 1046.2009 | 1073.8464 | 1.0625 | 1.0754 | 1.08 | 0.90942 | 1.0248 | 1.1486 | 1.0834 | 1.0595 | 1.065 | 1.0643 | 0.82146 | 0.89691 | 1.0153 | 0.99907 | 1.2416 | 1.0274 | 0.96342 | 0.97571 | 0.95118 | 1.1049 | 0.97055 | 1.017 | [6 empty] | 0.97399 | 0.98445 | 1.0887 | 0.93946
101 | CCV 10 ppb MAJORS |  | 06/16/2023 22:20:23 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.750 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9730.104 | 10148.7696 | 9.5046 | 9.0768 | 10884.6744 | 9980.2188 | 9967.1371 | 10.1639 | 10.6508 | 10.6538 | 10.304 | 10.5466 | 9.7354 | 10.2861 | 9.7647 | 10.1525 | 10.228 | 10.429 | 9.6983 | 9.8855 | 10.1663 | 10.4224 | 10.3729 | 10.0838 | 9.9256 | 10.2341 | 10.3914 | 9.995 | 10.3779 | [6 empty] | 9.7695 | 9.8099 | 9.8755 | 10.1874
102 | 10_1_21_BB |  | 06/16/2023 22:26:24 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_BB.751 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 55.1546 | 1.2239 | -0.25314 | -0.61745 | 692.0038 | 58.1697 | 43.4113 | 0.00010834 | 0.012387 | 0.024695 | 0.080283 | 0.0045032 | 0.32342 | -0.0062728 | 0.0030991 | 0.079735 | 0.081493 | -0.031276 | 0.023456 | 0.0068852 | 0.0058004 | -0.075634 | 0.011455 | 0.0070818 | -0.009398 | -0.0091788 | -0.016418 | -0.0046728 | 0.0035826 | [6 empty] | 0.00095124 | -0.0017145 | 0.020115 | 0.02603
103 | 10_1_21_FB |  | 06/16/2023 22:32:24 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_FB.752 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 248.2807 | 1.2699 | 4.3483 | 0.32434 | 485.9773 | 63.7857 | 52.5276 | 0.0022868 | 0.047236 | 0.056925 | 0.47676 | 0.040819 | 0.65784 | -0.0052927 | 0.052968 | 0.22894 | 0.23379 | 4.1399 | -0.031358 | 0.0038511 | 0.0057717 | 0.090935 | 0.011596 | 0.0053161 | -0.0057177 | -0.015839 | -0.021474 | -0.00084081 | 0.0082393 | [6 empty] | 0.00084424 | -0.0016373 | 0.019352 | 0.024384
104 | 10_1_21_LAB_A |  | 06/16/2023 22:38:24 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_LAB_A.753 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 56.5068 | 0.23633 | -0.21045 | -0.67577 | 481.8848 | 59.104 | 45.7502 | -0.00020634 | 0.0084091 | 0.01495 | -0.11956 | 0.0076566 | -0.073696 | -0.0069724 | 0.0065624 | 0.083533 | 0.078229 | 0.12564 | 0.061907 | -6.1539e-05 | 0.0012466 | 0.0036966 | 0.011675 | 0.011181 | -0.0084087 | 0.0031808 | -0.016313 | -0.0075698 | 0.0030503 | [6 empty] | 0.00050427 | -0.0021193 | 0.018135 | 0.02193
105 | 10_1_21_LAB_B |  | 06/16/2023 22:44:24 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_1_21_LAB_B.754 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 56.7511 | 0.75978 | -0.33986 | -0.60508 | 356.0908 | 54.3773 | 42.0522 | -0.00019717 | 0.043227 | 0.046668 | -0.073799 | 0.010293 | 0.035383 | -0.0069807 | 0.0084083 | 0.092568 | 0.091436 | 0.089311 | 0.054174 | -0.0075353 | 0.000731 | -0.18223 | 0.0055366 | 0.0014405 | -0.0074628 | -0.035078 | -0.01626 | -0.0082757 | 0.0021832 | [6 empty] | 0.00061721 | -0.0018803 | 0.040749 | 0.042802
106 | 5_18_23_DB |  | 06/16/2023 22:50:24 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_DB.755 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 43769.9619 | 12318.5837 | 113.1533 | 10.8613 | 1207.6708 | 3716.8541 | 49307.7328 | 1.9543 | 1.4534 | 1.4731 | 42.6719 | 3.5835 | 39.5921 | 1.3396 | 1.6309 | 0.77295 | 0.86221 | 5.8381 | 2.2272 | 2.1675 | 2.3597 | 4.8171 | 2.0437 | 1.8249 | 1.8442 | 1.9192 | 2.0587 | 1.0445 | 1.08 | [6 empty] | 3.267 | 3.2022 | 1.3019 | 1.2691
107 | 5_18_23_DB |  | 06/16/2023 22:56:42 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_DB.756 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 43362.4959 | 12410.459 | 108.812 | 11.8107 | 961.5035 | 3682.17 | 49526.3751 | 1.8454 | 1.4003 | 1.4197 | 41.8412 | 3.573 | 38.854 | 1.357 | 1.5545 | 0.7449 | 0.85922 | 5.795 | 2.2864 | 2.1086 | 2.3171 | 5.0834 | 1.9291 | 1.8769 | 1.853 | 2.1861 | 1.9397 | 1.0555 | 1.0916 | [6 empty] | 3.2344 | 3.1803 | 1.3582 | 1.2892
108 | 5_18_23_T_239 |  | 06/16/2023 23:02:42 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_T_239.757 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 43818.6967 | 12401.9544 | 112.5601 | 12.1886 | 1049.0914 | 3786.4273 | 50300.4608 | 1.8823 | 1.4043 | 1.4782 | 42.5428 | 3.6191 | 39.7527 | 1.3633 | 1.5978 | 0.75129 | 0.84652 | 5.9218 | 2.2436 | 2.1529 | 2.3504 | 5.183 | 1.8762 | 1.911 | 1.8645 | 2.1089 | 1.7647 | 1.1012 | 1.0772 | [6 empty] | 3.2441 | 3.2062 | 1.3426 | 1.2786
109 | 5_18_23_T_239 |  | 06/16/2023 23:09:00 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_T_239.758 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 43638.5231 | 12318.0966 | 114.4458 | 11.7129 | 1252.4876 | 3804.2649 | 50684.1641 | 1.886 | 1.4278 | 1.4217 | 41.8146 | 3.451 | 39.7999 | 1.3533 | 1.7091 | 0.74796 | 0.83342 | 5.7601 | 2.2777 | 2.1467 | 2.2225 | 5.0414 | 1.6799 | 1.8186 | 1.7585 | 2.1142 | 1.806 | 1.0965 | 1.0362 | [6 empty] | 3.2232 | 3.183 | 1.3481 | 1.3047
110 | 5_18_23_T_239 |  | 06/16/2023 23:15:00 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_T_239.759 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4013.3513 | 652.6814 | 7.8734 | 16.8415 | 1351.5605 | 524.8745 | 3247.7153 | 1.0284 | 0.029052 | 0.037332 | -0.29472 | 5.436 | -0.050935 | 0.0093901 | 0.049561 | 0.052033 | 0.070904 | 0.59276 | 0.20613 | 0.033469 | 0.053173 | 0.26488 | 0.018852 | 0.018528 | 0.012771 | 0.82724 | 0.017952 | 0.23726 | 0.19843 | [6 empty] | 0.0044874 | -0.00018506 | 0.0082028 | 0.013524
111 | 5_18_23_M_224 |  | 06/16/2023 23:20:59 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_M_224.760 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4417.0847 | 4583.8643 | 4.3343 | 9.4264 | 5249.3873 | 4611.0071 | 4862.5717 | 4.8126 | 4.9756 | 5.1286 | 4.7419 | 5.0416 | 4.4366 | 4.8558 | 4.6759 | 4.9378 | 5.063 | 4.9123 | 4.9229 | 4.7641 | 4.8185 | 5.9733 | 5.0512 | 4.5848 | 4.7783 | 5.345 | 5.0127 | 4.8095 | 4.942 | [6 empty] | 4.7627 | 4.7858 | 5.9992 | 5.2479
112 | LBLANK |  | 06/16/2023 23:27:00 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.761 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 54.9966 | 0.765 | -0.98765 | 0.27807 | 453.1504 | 57.124 | 40.6215 | 0.0015915 | -0.0054195 | -0.0026971 | -0.57853 | -0.00011699 | -0.28595 | -0.00588 | -0.0025338 | -0.012831 | -0.010844 | -0.22593 | 0.014198 | -0.01135 | -0.0014082 | -0.099117 | 0.018063 | 0.0054514 | -0.011141 | -0.0027456 | -0.011061 | 0.015293 | 0.0032846 | [6 empty] | 0.00061655 | -0.0017854 | -0.01059 | -0.004417
113 | LBLANK |  | 06/16/2023 23:33:18 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.762 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 55.6865 | 0.75871 | -0.99821 | -0.70648 | 531.5513 | 57.2547 | 40.6163 | 0.00052196 | -0.0056343 | -0.0020015 | -0.57838 | -0.00023853 | -0.28665 | -0.0072719 | -0.0024492 | -0.013827 | -0.011843 | -0.22239 | 0.054473 | -0.0074625 | -0.0027578 | -0.084283 | 0.0055366 | 0.0034757 | -0.011125 | -0.045514 | -0.021474 | 0.0079127 | 0.0024118 | [6 empty] | 0.00072017 | -0.0025324 | -0.010044 | -0.0041253
114 | CCV 1ppb TRACE |  | 06/16/2023 23:39:19 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.763 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 973.5156 | 996.1043 | 1.0589 | 0.14864 | 1343.6848 | 993.3618 | 1012.2498 | 1.0111 | 1.0245 | 1.0526 | 0.85519 | 1.0001 | 1.0524 | 0.97807 | 0.94431 | 1.0144 | 1.0229 | 0.78053 | 0.92841 | 0.99216 | 0.95211 | 1.2711 | 0.88487 | 0.95893 | 0.97227 | 0.95235 | 0.9513 | 0.97658 | 0.99264 | [6 empty] | 0.95932 | 0.96684 | 1.1404 | 1.0837
115 | CCV 1ppb TRACE |  | 06/16/2023 23:45:38 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.764 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 988.5811 | 943.7088 | 1.1674 | 0.089159 | 1565.2944 | 990.3717 | 1077.1318 | 0.99564 | 1.0509 | 1.0646 | 0.93036 | 0.99354 | 1.0632 | 0.99528 | 0.99444 | 1.0052 | 1.0578 | 0.82988 | 0.99061 | 0.90047 | 0.97965 | 1.6046 | 1.0735 | 1.0641 | 0.97289 | 1.148 | 0.93666 | 1.0003 | 1.0057 | [6 empty] | 0.95725 | 0.97728 | 1.3077 | 1.0615
116 | LFB |  | 06/16/2023 23:51:38 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.765 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 8856.5512 | 9233.4125 | 9.0278 | 11.5031 | 10225.7146 | 8971.6175 | 9523.1364 | 9.4157 | 9.8454 | 10.0143 | 9.6841 | 9.7627 | 8.711 | 9.4044 | 9.2996 | 9.5611 | 9.7698 | 9.8099 | 9.0941 | 9.6085 | 9.6807 | 10.0371 | 9.4857 | 9.6289 | 9.9364 | 9.6183 | 9.8764 | 9.6503 | 9.8539 | [6 empty] | 9.5391 | 9.5837 | 9.4655 | 10.0337
117 | LFB |  | 06/16/2023 23:57:55 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.766 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 8924.9738 | 9503.9294 | 9.1653 | 11.5659 | 10370.3317 | 9199.315 | 9709.2367 | 9.5653 | 10.0675 | 10.037 | 9.7676 | 9.8653 | 8.7551 | 9.6751 | 9.3953 | 9.6844 | 9.9154 | 9.7938 | 9.2828 | 9.5281 | 9.8623 | 10.2711 | 9.9673 | 10.1907 | 10.0648 | 9.8316 | 10.3033 | 9.6507 | 9.8937 | [6 empty] | 9.5445 | 9.6102 | 9.6924 | 9.9435
118 | CCV 10 ppb MAJORS |  | 06/17/2023 00:03:55 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.767 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9146.4813 | 9732.3602 | 9.1596 | 9.2258 | 10290.665 | 9419.1562 | 9950.6236 | 9.838 | 10.2238 | 10.3956 | 10.111 | 10.221 | 9.3725 | 9.9119 | 9.504 | 9.8709 | 10.1029 | 10.2163 | 9.8929 | 9.2593 | 9.9075 | 10.9237 | 9.476 | 9.8004 | 10.0228 | 10.2829 | 10.5384 | 9.9149 | 10.2118 | [6 empty] | 9.6348 | 9.7148 | 10.4492 | 10.3041
119 | 5_18_23_MFB |  | 06/17/2023 00:09:55 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_MFB.768 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 56.6753 | 0.23633 | -0.59115 | 2.7087 | 406.8146 | 61.5672 | 47.099 | 0.0011718 | 0.0025575 | 0.0076277 | -0.43906 | 0.0022547 | -0.074802 | -0.0066936 | 0.0010111 | 0.066099 | 0.065083 | 0.062457 | 0.05115 | -0.0076308 | -0.0021745 | 0.063854 | 0.011577 | 0.0090925 | -0.0075515 | -0.040273 | -0.021474 | 0.01862 | 0.0028276 | [6 empty] | 0.00094041 | -0.0021917 | 0.0081761 | 0.013714
120 | 10_15_21_W_1 |  | 06/17/2023 00:15:55 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_W_1.769 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 2005.5658 | 1413.4402 | 4.2652 | 3.7352 | 6577.8116 | 561.4068 | 8786.5903 | 0.047893 | 0.027796 | 0.028673 | 11.1255 | 5.9621 | 10.6707 | 0.0026552 | 0.079964 | 0.29786 | 0.32803 | 0.15963 | 0.61128 | 0.62592 | 0.67172 | 0.28938 | 0.031875 | 0.018321 | 0.004674 | 0.11832 | -0.010342 | 4.1547 | 4.2854 | [6 empty] | 0.0039643 | 3.6953e-05 | 0.027694 | 0.029919
121 | 10_15_21_W_6 |  | 06/17/2023 00:21:56 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_W_6.770 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1820.2365 | 1468.9528 | 14.72 | 2.0336 | 4891.6762 | 491.3656 | 8091.1895 | 0.12697 | 0.024065 | 0.051147 | 17.161 | 3.8165 | 15.9952 | 0.0028552 | 0.056043 | 1.1194 | 1.1462 | 0.73276 | 0.64307 | 0.62888 | 0.70034 | 0.10601 | 0.03898 | 0.03133 | 0.0099013 | 0.09291 | -0.015736 | 2.3016 | 2.3506 | [6 empty] | 0.0081055 | 0.0029772 | 0.12268 | 0.12325
122 | 10_15_21_W_8 |  | 06/17/2023 00:27:56 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_W_8.771 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1833.5767 | 1389.3738 | 5.9307 | 1.9088 | 4423.7611 | 464.205 | 7863.5214 | 0.13799 | 0.015746 | 0.017335 | 8.2087 | 2.5864 | 7.6748 | 0.00036112 | 0.053662 | 0.67381 | 0.72028 | 0.27556 | 0.72553 | 0.66762 | 0.72841 | 0.20391 | 0.032373 | 0.018776 | 0.0070127 | 0.06854 | 0.012562 | 2.0589 | 2.0931 | [6 empty] | 0.0029558 | 0.00054757 | 0.052617 | 0.054571
123 | 10_15_21_W_9 |  | 06/17/2023 00:33:56 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_W_9.772 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1572.1129 | 1292.7546 | 9.5485 | 1.8094 | 3728.0764 | 403.3684 | 6875.8952 | 0.11606 | 0.030872 | 0.034331 | 9.0725 | 2.3189 | 8.4067 | -0.0014588 | 0.056117 | 0.58915 | 0.62044 | 0.25387 | 0.5951 | 0.54605 | 0.58819 | 0.78538 | 0.019037 | 0.018882 | 0.006106 | 0.16521 | 0.0071705 | 1.4202 | 1.4846 | [6 empty] | 0.0044981 | 0.00048436 | 0.06522 | 0.054579
124 | 10_15_21_W_9 |  | 06/17/2023 00:40:15 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_W_9.773 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1590.7698 | 1274.7171 | 8.7663 | 1.8495 | 3698.9078 | 410.1781 | 6898.2629 | 0.1177 | 0.029519 | 0.037816 | 9.099 | 2.28 | 8.5242 | -0.0010056 | 0.036501 | 0.62082 | 0.63553 | 0.26782 | 0.57476 | 0.47825 | 0.59738 | 0.19708 | 0.012288 | 0.012467 | 0.0042995 | 0.11182 | 0.00718 | 1.4431 | 1.4681 | [6 empty] | 0.0033199 | 0.00028658 | 0.051986 | 0.05653
125 | 10_15_21_W_10 |  | 06/17/2023 00:46:15 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_W_10.774 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1285.8078 | 1505.7855 | 30.7228 | 2.9233 | 3222.2775 | 403.2486 | 6651.2359 | 0.15196 | 0.033668 | 0.034283 | 25.7102 | 5.1996 | 23.6581 | 0.0072383 | 0.048663 | 0.86917 | 0.88971 | 1.0798 | 0.65321 | 0.6412 | 0.6658 | 0.16078 | 0.032907 | 0.023412 | 0.0083637 | 0.046454 | -0.015696 | 0.63991 | 0.66836 | [6 empty] | 0.0067906 | 0.0042742 | 0.1793 | 0.16979
126 | 10_15_21_W_10 |  | 06/17/2023 00:52:33 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_W_10.775 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1268.8095 | 1473.1627 | 31.2012 | 2.7824 | 3205.0898 | 367.1392 | 6757.8141 | 0.15041 | 0.025625 | 0.031782 | 25.6163 | 5.2486 | 23.8756 | 0.0071848 | 0.040801 | 0.86183 | 0.88084 | 1.0925 | 0.60909 | 0.48603 | 0.68398 | 0.075173 | 0.032777 | 0.0168 | 0.0092596 | 0.060801 | -0.0043024 | 0.63699 | 0.66042 | [6 empty] | 0.0064709 | 0.0032196 | 0.17381 | 0.17129
127 | 10_15_21_W_10_DUP |  | 06/17/2023 00:58:33 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_W_10_DUP.776 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1229.5166 | 1420.7281 | 28.7666 | 2.8389 | 3184.337 | 382.2393 | 6571.0237 | 0.1441 | 0.057515 | 0.063047 | 24.983 | 5.0238 | 22.7382 | 0.0059246 | 0.044269 | 0.85086 | 0.8821 | 0.96952 | 0.59848 | 0.54385 | 0.67152 | 0.24149 | 0.0055366 | 0.010246 | -0.00095537 | 0.071134 | -0.0042822 | 0.62367 | 0.63807 | [6 empty] | 0.007231 | 0.0034866 | 0.16196 | 0.16484
128 | 10_15_21_W_10_DUP |  | 06/17/2023 01:04:52 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_W_10_DUP.777 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1247.1092 | 1458.0539 | 27.6383 | 2.85 | 3208.9612 | 387.2612 | 6409.5866 | 0.14541 | 0.055372 | 0.057049 | 25.2636 | 5.0287 | 22.9964 | 0.0059069 | 0.048256 | 0.86047 | 0.87087 | 0.95387 | 0.62072 | 0.60296 | 0.69031 | 0.16657 | 0.0055366 | 0.019184 | 0.013621 | 0.12264 | -0.0098875 | 0.61297 | 0.65128 | [6 empty] | 0.0075176 | 0.0030841 | 0.16071 | 0.15901
129 | LBLANK |  | 06/17/2023 01:10:52 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.778 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 54.3812 | 0.80206 | -0.99466 | -0.53709 | 462.2207 | 59.9633 | 39.8954 | 0.00063728 | -0.0098968 | -0.0023999 | -0.58696 | -0.00071708 | -0.45669 | -0.0077222 | -0.00023988 | -0.010848 | -0.012549 | -0.23468 | 0.0044616 | -0.01135 | -0.0031847 | -0.035661 | 0.0055366 | 0.0037498 | -0.011061 | -0.02219 | -0.015854 | -0.0025302 | 0.0025004 | [6 empty] | 0.00050427 | -0.0021725 | -0.010437 | -0.0044141
130 | LBLANK |  | 06/17/2023 01:17:12 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.779 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 54.8716 | 0.23633 | -0.95262 | -0.80843 | 506.4749 | 61.3376 | 39.9056 | -0.00052778 | -0.0079865 | -0.0020988 | -0.58614 | -0.00064534 | -0.34749 | -0.007392 | 0.0018527 | -0.011356 | -0.011045 | -0.22706 | 0.066471 | -0.01135 | -0.0036289 | -0.21459 | 0.0055366 | 0.0016159 | -0.0060409 | -0.0019992 | -0.021474 | -0.0033717 | 0.002501 | [6 empty] | 0.00062566 | -0.0023092 | -0.01067 | -0.0043961
131 | CCV 1ppb TRACE |  | 06/17/2023 01:23:12 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.780 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 972.3873 | 1032.7939 | 1.2642 | 0.080488 | 1454.5502 | 1054.1315 | 1036.1287 | 1.0181 | 1.0324 | 1.0541 | 0.89712 | 0.99441 | 0.96018 | 1.0006 | 0.96316 | 1.0235 | 1.043 | 0.80002 | 0.89649 | 1.0668 | 0.95919 | 1.105 | 0.86769 | 1.0299 | 0.95419 | 0.95715 | 1.0801 | 0.98211 | 1.0186 | [6 empty] | 0.97143 | 0.99687 | 1.1638 | 1.0874
132 | CCV 1ppb TRACE |  | 06/17/2023 01:29:31 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.781 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 965.4585 | 962.0715 | 1.0941 | 0.10977 | 1521.441 | 1017.0475 | 1037.8145 | 1.022 | 1.0321 | 1.0666 | 0.93879 | 1.0066 | 0.95728 | 0.98594 | 0.98441 | 1.0047 | 1.034 | 0.83014 | 0.91458 | 1.0204 | 0.97087 | 1.174 | 0.95901 | 1.029 | 0.98943 | 0.95493 | 1.0794 | 0.98612 | 1.0012 | [6 empty] | 0.97472 | 0.97825 | 1.1556 | 1.1375
133 | CCV 10 ppb MAJORS |  | 06/17/2023 01:35:32 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.782 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9198.8992 | 9642.5842 | 8.8068 | 8.8042 | 10567.5571 | 9467.1743 | 10040.5209 | 9.9471 | 10.2946 | 10.3894 | 10.2326 | 10.2764 | 9.1999 | 9.9296 | 9.4652 | 9.813 | 9.9928 | 10.2096 | 9.5101 | 10.0143 | 9.8402 | 10.1058 | 9.9663 | 9.8738 | 9.9824 | 10.0215 | 9.9297 | 9.8692 | 10.3203 | [6 empty] | 9.7451 | 9.8086 | 10.3905 | 10.5557
134 | CCV 10 ppb MAJORS |  | 06/17/2023 01:41:51 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.783 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9210.9318 | 9666.5118 | 9.5342 | 8.785 | 10531.8412 | 9347.6247 | 10086.3066 | 9.9852 | 10.1448 | 10.6126 | 10.2737 | 10.2175 | 9.2229 | 9.8789 | 9.5724 | 9.6853 | 10.1415 | 10.2698 | 9.4948 | 9.719 | 9.8399 | 10.1636 | 10.0204 | 9.9062 | 9.97 | 10.3088 | 10.1556 | 9.9226 | 10.2306 | [6 empty] | 9.7404 | 9.771 | 10.2194 | 10.4249
135 | 10_15_21_W_11 |  | 06/17/2023 01:47:51 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_W_11.784 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1202.0649 | 1500.0398 | 41.9447 | 2.8231 | 3194.1838 | 398.3913 | 6621.6961 | 0.17238 | 0.043371 | 0.046943 | 33.9716 | 5.7171 | 31.3393 | 0.010605 | 0.057631 | 1.0295 | 1.0937 | 1.7004 | 0.62582 | 0.60114 | 0.70033 | -0.049296 | 0.019075 | 0.0059015 | 0.011138 | 0.014905 | -0.0042934 | 0.56267 | 0.55949 | [6 empty] | 0.0061421 | 0.0065104 | 0.18482 | 0.20569
136 | 10_15_21_W_11 |  | 06/17/2023 01:54:10 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_W_11.785 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1185.5387 | 1477.3299 | 40.491 | 3.0542 | 2985.4201 | 371.5053 | 6594.0206 | 0.17399 | 0.040188 | 0.045111 | 34.2345 | 5.6832 | 30.7074 | 0.01021 | 0.055387 | 0.99603 | 1.0887 | 1.6934 | 0.736 | 0.74405 | 0.7189 | 0.099548 | 0.033358 | 0.023794 | 0.0014201 | 0.044907 | -0.0038902 | 0.55994 | 0.57723 | [6 empty] | 0.0093589 | 0.0053224 | 0.21451 | 0.20386
137 | 10_15_21_F_1 |  | 06/17/2023 02:00:09 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_1.786 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10219.3731 | 6847.8694 | 1.9268 | 6.4141 | 32544.0982 | 2787.0104 | 45317.1903 | 0.19546 | 0.067471 | 0.075495 | 12.227 | 19.9237 | 12.1574 | 0.024947 | 0.39748 | 0.77675 | 0.83846 | 1.2518 | 2.7852 | 2.5831 | 2.8186 | 0.61552 | 0.092999 | 0.068957 | 0.046631 | 0.52158 | 0.046131 | 21.3393 | 21.8333 | [6 empty] | 0.0086021 | 0.0075661 | 3.6005 | 3.1027
138 | 10_15_21_F_1 |  | 06/17/2023 02:06:28 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_1.787 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10386.9913 | 7032.3456 | 1.9451 | 6.3023 | 32560.42 | 2844.9986 | 45525.2282 | 0.20668 | 0.067385 | 0.074497 | 12.4622 | 20.0952 | 12.4585 | 0.024122 | 0.34814 | 0.76734 | 0.8441 | 1.2882 | 2.7229 | 2.5591 | 2.8237 | 0.6889 | 0.082766 | 0.070961 | 0.047669 | 0.45425 | 0.07357 | 21.2365 | 21.9279 | [6 empty] | 0.010252 | 0.0079057 | 3.497 | 3.4967
139 | 10_15_21_F_1_DUP |  | 06/17/2023 02:12:28 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_1_DUP.788 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9901.0171 | 6587.3786 | 1.1049 | 5.1146 | 30527.5174 | 2650.462 | 42640.337 | 0.203 | 0.053164 | 0.050694 | 11.6102 | 19.3078 | 11.6988 | 0.023372 | 0.35659 | 0.70677 | 0.77966 | 1.2538 | 2.6899 | 2.4528 | 2.7299 | 0.69758 | 0.081364 | 0.056555 | 0.050794 | 0.5058 | 0.071988 | 20.3201 | 20.8434 | [6 empty] | 0.010912 | 0.006162 | 0.05058 | 0.053585
140 | 10_15_21_F_1_DUP |  | 06/17/2023 02:18:46 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_1_DUP.789 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10138.9519 | 6771.2102 | 1.1569 | 5.1716 | 30874.0792 | 2702.4277 | 43277.5899 | 0.17748 | 0.04635 | 0.050235 | 11.6423 | 19.1878 | 11.8251 | 0.021165 | 0.35103 | 0.72759 | 0.77051 | 1.2365 | 2.6951 | 2.2757 | 2.7049 | 0.66134 | 0.053359 | 0.084417 | 0.039096 | 0.47696 | 0.042306 | 20.1097 | 20.7729 | [6 empty] | 0.0092028 | 0.0066434 | 0.04779 | 0.049862
141 | 10_15_21_F_6 |  | 06/17/2023 02:24:45 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_6.790 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9629.9114 | 7669.1188 | 11.2546 | 2.3267 | 24141.9699 | 2245.8268 | 41752.8675 | 0.5675 | 0.080044 | 0.086269 | 18.7791 | 8.2352 | 18.7771 | 0.015592 | 0.28296 | 2.0075 | 2.1282 | 4.9782 | 2.8731 | 2.6881 | 2.8923 | 0.69087 | 0.13604 | 0.099793 | 0.099692 | 0.50131 | 0.030706 | 11.4307 | 11.707 | [6 empty] | 0.015011 | 0.012246 | 0.091037 | 0.079212
142 | 10_15_21_F_6 |  | 06/17/2023 02:31:03 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_6.791 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9430.877 | 7611.3164 | 11.0525 | 2.3269 | 24142.7901 | 2294.784 | 41845.4663 | 0.54157 | 0.08211 | 0.089646 | 18.656 | 8.2356 | 18.2018 | 0.020376 | 0.31111 | 2.0248 | 2.083 | 5.0895 | 2.7877 | 2.5437 | 2.964 | 0.53509 | 0.1666 | 0.077532 | 0.068499 | 0.46481 | 0.096984 | 11.4216 | 11.8225 | [6 empty] | 0.015578 | 0.014996 | 0.092431 | 0.09269
143 | 10_15_21_F_6_LD |  | 06/17/2023 02:37:02 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_6_LD.792 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9320.0319 | 7348.1645 | 11.938 | 4.6263 | 23009.2172 | 2170.4842 | 40625.0249 | 0.52264 | 0.08457 | 0.094454 | 15.5315 | 8.2293 | 15.1796 | 0.023452 | 0.30503 | 1.994 | 2.0703 | 4.9787 | 2.8201 | 2.5914 | 2.913 | 0.6383 | 0.089652 | 0.097468 | 0.064297 | 0.4884 | 0.084981 | 11.4579 | 11.7377 | [6 empty] | 0.016067 | 0.014251 | 0.089412 | 0.090078
144 | 10_15_21_F_6_LD |  | 06/17/2023 02:43:20 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_6_LD.793 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9444.1291 | 7475.209 | 11.822 | 5.3002 | 23798.8203 | 2237.406 | 41932.4511 | 0.5491 | 0.083496 | 0.089356 | 15.8424 | 8.2016 | 15.5376 | 0.026349 | 0.30831 | 2.0273 | 2.1363 | 5.0521 | 2.6901 | 2.6837 | 2.9146 | 0.43153 | 0.089894 | 0.086774 | 0.090903 | 0.40693 | 0.073686 | 11.7513 | 11.7456 | [6 empty] | 0.017802 | 0.013349 | 0.085059 | 0.091604
145 | LBLANK |  | 06/17/2023 02:49:20 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.794 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 54.1513 | 0.23633 | -0.98894 | -0.12799 | 465.4333 | 59.8055 | 41.4958 | -0.00014348 | -0.0075797 | -0.00058329 | -0.5883 | -0.00033068 | -0.23142 | -0.0072372 | 0.0046778 | -0.0084419 | -0.010987 | -0.22901 | 0.055995 | -0.0032668 | -0.0029648 | -0.040557 | 0.0055366 | 0.0056923 | -0.010118 | 0.0074721 | -0.015962 | -0.0057121 | 0.0044433 | [6 empty] | 0.00050427 | -0.0025665 | -0.010739 | -0.0043062
146 | CCV 1ppb TRACE |  | 06/17/2023 02:55:21 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.795 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1005.7882 | 1048.4076 | 1.304 | 0.3476 | 1727.1377 | 1054.9772 | 1057.608 | 1.086 | 1.0813 | 1.1036 | 0.97469 | 1.0112 | 1.0803 | 0.99081 | 1.0656 | 1.041 | 1.0678 | 0.83716 | 0.86468 | 1.013 | 0.98192 | 1.2262 | 0.89459 | 0.96245 | 0.92098 | 0.96661 | 0.94546 | 0.98616 | 1.0248 | [6 empty] | 0.99923 | 0.98593 | 1.1311 | 1.0721
147 | CCV 1ppb TRACE |  | 06/17/2023 03:01:39 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.796 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1039.1883 | 1046.4968 | 1.1749 | 0.1367 | 1691.0827 | 1014.6489 | 1058.595 | 1.0321 | 1.0505 | 1.1139 | 0.95184 | 1.0209 | 1.2062 | 1.0493 | 1.0383 | 1.0472 | 1.0801 | 0.82147 | 0.92262 | 0.97322 | 0.99683 | 1.0198 | 0.83139 | 1.0285 | 0.99269 | 0.82037 | 0.95962 | 0.99313 | 1.0076 | [6 empty] | 0.95779 | 0.96766 | 1.0514 | 1.0706
148 | LFB |  | 06/17/2023 03:07:39 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.797 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9123.3373 | 9345.9849 | 9.3734 | 11.6738 | 10722.5977 | 9204.8392 | 9932.1496 | 9.6151 | 10.052 | 10.2338 | 9.729 | 9.9726 | 9.1071 | 9.643 | 9.3195 | 9.6476 | 9.8048 | 9.8163 | 9.5112 | 9.7902 | 9.8163 | 10.2543 | 10.7167 | 10.0953 | 10.0937 | 10.3666 | 10.3098 | 9.5825 | 9.8779 | [6 empty] | 9.5213 | 9.5604 | 9.9174 | 10.4426
149 | LFB |  | 06/17/2023 03:13:57 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.798 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 8972.1227 | 9656.0419 | 9.6568 | 11.9956 | 10536.2556 | 9338.0058 | 9596.5639 | 9.7802 | 10.0446 | 10.2039 | 9.831 | 9.9603 | 9.1934 | 9.7123 | 9.1212 | 9.6313 | 9.8862 | 9.8124 | 9.3311 | 9.4447 | 9.7939 | 10.1977 | 9.7979 | 10.1906 | 10.1112 | 10.0706 | 10.0154 | 9.5099 | 9.9398 | [6 empty] | 9.5094 | 9.6022 | 9.6689 | 10.0364
150 | CCV 10 ppb MAJORS |  | 06/17/2023 03:19:56 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.799 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9501.9728 | 10048.5273 | 9.3208 | 9.5016 | 10864.3158 | 9851.0134 | 10385.0432 | 10.3208 | 10.5674 | 10.6308 | 10.5023 | 10.3868 | 9.8696 | 10.1599 | 9.681 | 10.0892 | 10.2151 | 10.3869 | 9.4912 | 10.348 | 9.9173 | 10.5936 | 9.8965 | 10.0221 | 10.1009 | 10.2473 | 10.755 | 9.902 | 10.3375 | [6 empty] | 9.7112 | 9.7188 | 10.031 | 10.3503
151 | CCV 10 ppb MAJORS |  | 06/17/2023 03:26:14 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.800 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9456.6663 | 9965.1394 | 9.5303 | 9.2573 | 10804.9655 | 9722.0743 | 10241.1771 | 10.3568 | 10.5133 | 10.7993 | 10.4783 | 10.5147 | 9.6045 | 10.1885 | 9.6873 | 9.9861 | 10.2508 | 10.472 | 9.3425 | 9.7767 | 9.9689 | 10.1017 | 9.7904 | 9.905 | 10.1544 | 9.6905 | 10.3999 | 10.0468 | 10.3589 | [6 empty] | 9.7324 | 9.8083 | 9.7941 | 10.3318
152 | 10_15_21_F_6_DUP |  | 06/17/2023 03:32:14 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_6_DUP.801 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 8811.1153 | 6992.3724 | 6.8005 | 2.3788 | 22820.0661 | 2141.8205 | 39844.1552 | 0.50363 | 0.081727 | 0.093434 | 10.2046 | 7.7409 | 10.3714 | 0.024951 | 0.31427 | 1.9886 | 2.0722 | 4.8828 | 2.766 | 2.6448 | 2.8368 | 0.59003 | 0.10713 | 0.066391 | 0.072447 | 0.46184 | 0.052217 | 11.2417 | 11.4655 | [6 empty] | 0.019252 | 0.014072 | 0.10432 | 0.10827
153 | 10_15_21_F_6_DUP |  | 06/17/2023 03:38:32 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_6_DUP.802 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 8910.3035 | 7051.8656 | 7.1826 | 2.3587 | 22781.2926 | 2139.8691 | 40020.173 | 0.50234 | 0.093405 | 0.097135 | 9.9386 | 7.8319 | 10.4932 | 0.026948 | 0.29225 | 1.9905 | 2.0379 | 4.8686 | 3.0144 | 2.65 | 2.864 | 0.68447 | 0.057486 | 0.079187 | 0.094594 | 0.53331 | 0.065473 | 10.9494 | 11.57 | [6 empty] | 0.018388 | 0.015016 | 0.11671 | 0.11496
154 | 10_15_21_F_6_DUP_LSP |  | 06/17/2023 03:44:32 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_6_DUP_LSP.803 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 17563.3886 | 16273.8317 | 16.946 | 14.3103 | 31736.4955 | 11646.9161 | 46706.9872 | 10.6931 | 10.4294 | 10.5597 | 20.3091 | 17.5141 | 19.4848 | 9.9816 | 9.5785 | 11.2966 | 11.593 | 14.3953 | 12.1378 | 11.9845 | 12.4204 | 10.5632 | 10.0777 | 10.0092 | 10.0159 | 10.5568 | 9.4897 | 20.1646 | 20.9537 | [6 empty] | 9.3567 | 9.4372 | 9.8518 | 10.1426
155 | 10_15_21_F_6_DUP_LSP |  | 06/17/2023 03:50:52 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_6_DUP_LSP.804 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 18749.7413 | 17093.0077 | 18.5224 | 15.1751 | 33563.1567 | 12318.1988 | 49989.2252 | 11.3525 | 11.2101 | 10.702 | 20.5879 | 17.8444 | 20.7693 | 10.6252 | 10.1331 | 11.9968 | 11.7464 | 14.5267 | 12.069 | 12.398 | 12.3779 | 10.5394 | 10.3708 | 10.065 | 9.8342 | 10.5624 | 10.2297 | 20.3677 | 20.9185 | [6 empty] | 9.5494 | 9.5547 | 9.7951 | 10.1567
156 | 10_15_21_F_9 |  | 06/17/2023 03:56:52 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_9.805 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 8102.7062 | 6954.4507 | 4.8346 | 2.9816 | 18217.2736 | 1955.6885 | 36413.1928 | 0.57781 | 0.10968 | 0.12751 | 7.9751 | 4.7308 | 8.3261 | 0.012742 | 0.23535 | 2.363 | 2.4114 | 4.1252 | 2.6276 | 2.6289 | 2.7552 | 0.42187 | 0.070854 | 0.082572 | 0.065844 | 0.38652 | 0.052124 | 7.3815 | 7.5648 | [6 empty] | 0.013028 | 0.0081197 | 0.086792 | 0.090247
157 | 10_15_21_F_9 |  | 06/17/2023 04:03:12 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_9.806 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 8150.1238 | 6942.8337 | 4.8298 | 2.4293 | 18057.9904 | 1921.0344 | 36289.5407 | 0.55985 | 0.10642 | 0.12468 | 7.9007 | 4.7679 | 8.0582 | 0.0095882 | 0.22446 | 2.3313 | 2.4272 | 4.0293 | 2.6614 | 2.5905 | 2.7507 | 0.62515 | 0.07126 | 0.071393 | 0.047871 | 0.43153 | 0.077226 | 7.333 | 7.547 | [6 empty] | 0.015327 | 0.0085878 | 0.088836 | 0.092073
158 | 10_15_21_F_10 |  | 06/17/2023 04:09:13 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_10.807 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 6299.0039 | 7642.8526 | 5.376 | 3.9528 | 14518.378 | 1736.9165 | 34055.4244 | 0.63737 | 0.13138 | 0.14164 | 7.1129 | 5.3923 | 7.6122 | 0.011977 | 0.33376 | 1.3945 | 1.4571 | 3.226 | 2.9094 | 2.7114 | 3.1781 | 0.53466 | 0.14439 | 0.080377 | 0.066856 | 0.36566 | 0.017569 | 3.2005 | 3.5127 | [6 empty] | 0.0084272 | 0.0047624 | 0.23764 | 0.22821
159 | 10_15_21_F_10 |  | 06/17/2023 04:15:32 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_10.808 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 6195.3918 | 7319.3991 | 5.6836 | 3.7716 | 14253.6366 | 1690.3526 | 33381.6621 | 0.60657 | 0.12683 | 0.16107 | 8.5265 | 5.0844 | 7.3288 | 0.0087652 | 0.32234 | 1.3962 | 1.7344 | 3.8172 | 2.9324 | 2.6849 | 2.9969 | 0.4144 | 0.10164 | 0.077081 | 0.074064 | 0.3276 | 0.072154 | 3.777 | 3.3264 | [6 empty] | 0.0078382 | 0.0058589 | 0.23614 | 0.23336
160 | 10_15_21_F_10_DUP |  | 06/17/2023 04:21:32 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_10_DUP.809 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 6217.952 | 7506.678 | 6.1406 | 3.5061 | 14037.0184 | 1770.1345 | 34553.1907 | 0.64102 | 0.099026 | 0.095613 | 6.8738 | 4.9928 | 7.4375 | 0.012175 | 0.15514 | 1.3159 | 1.3214 | 2.582 | 2.4829 | 2.6549 | 2.6722 | 0.65078 | 0.050031 | 0.058871 | 0.044605 | 0.30788 | 0.029009 | 3.0638 | 3.2789 | [6 empty] | 0.0085942 | 0.0035429 | 0.38288 | 0.35871
161 | 10_15_21_F_10_DUP |  | 06/17/2023 04:27:51 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_10_DUP.810 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 6073.8187 | 7368.8303 | 5.6486 | 3.4927 | 13772.809 | 1723.6082 | 33934.5645 | 0.61669 | 0.10096 | 0.097561 | 7.0339 | 4.9706 | 7.3089 | 0.0073058 | 0.16144 | 1.2764 | 1.3309 | 2.6238 | 2.421 | 2.3436 | 2.6484 | 0.3797 | 0.050117 | 0.058691 | 0.045758 | 0.38427 | 0.041455 | 3.1715 | 3.2682 | [6 empty] | 0.0086233 | 0.0038039 | 0.36084 | 0.36895
162 | LBLANK |  | 06/17/2023 04:33:51 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.811 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 55.1512 | 0.82977 | -0.99851 | -0.9124 | 474.7171 | 66.3204 | 43.2566 | -0.00091385 | -0.0086804 | -0.00088664 | -0.58225 | -0.00014399 | -0.37868 | -0.0077199 | 0.011284 | -0.010375 | -0.0073362 | -0.23426 | 0.0091799 | -0.01135 | -0.0037193 | -0.079842 | 0.01249 | 0.0083662 | -0.007874 | 0.021953 | -0.021474 | -0.0057551 | 0.0029491 | [6 empty] | 0.00063013 | -0.0023352 | -0.010447 | -0.0041987
163 | LBLANK |  | 06/17/2023 04:40:11 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.812 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 54.1682 | 0.84263 | -1.0169 | -0.93954 | 495.8656 | 57.3309 | 39.974 | 0.00033378 | -0.0077697 | -0.00062604 | -0.58195 | -0.0011218 | -0.40732 | -0.0077098 | 0.0020553 | -0.010273 | -0.011342 | -0.22944 | 0.032275 | -0.0071572 | -0.0024656 | -0.091799 | 0.0055366 | -0.00050913 | -0.01206 | -0.010603 | -0.021474 | -0.0068265 | 0.0022609 | [6 empty] | 0.00050427 | -0.0025427 | -0.010484 | -0.0040756
164 | CCV 1ppb TRACE |  | 06/17/2023 04:46:11 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.813 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 983.7392 | 1072.1924 | 1.1539 | 0.055486 | 1537.7798 | 1029.8857 | 1079.6466 | 1.0459 | 1.0545 | 1.07 | 0.89737 | 1.0298 | 0.96944 | 0.97811 | 1.0569 | 1.0311 | 1.0457 | 0.82232 | 0.9227 | 0.99993 | 0.96635 | 1.3222 | 0.92116 | 0.996 | 0.98525 | 1.0635 | 0.96342 | 0.98122 | 1.0053 | [6 empty] | 0.97088 | 0.96993 | 1.1709 | 1.1087
165 | CCV 1ppb TRACE |  | 06/17/2023 04:52:31 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.814 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 984.0166 | 996.0293 | 1.0865 | 0.068408 | 1722.9088 | 1051.2473 | 1043.3635 | 1.0329 | 1.0552 | 1.0901 | 0.96717 | 1.0441 | 1.0858 | 1.0268 | 0.98463 | 1.0304 | 1.0665 | 0.83325 | 0.85805 | 0.89092 | 0.98626 | 1.5568 | 0.94617 | 1.0005 | 0.99509 | 0.87096 | 1.0018 | 0.98554 | 1.0256 | [6 empty] | 0.97611 | 0.98955 | 1.2189 | 1.1084
166 | CCV 10 ppb MAJORS |  | 06/17/2023 04:58:31 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.815 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9082.5039 | 9580.7748 | 8.9554 | 9.0381 | 10710.8149 | 9556.8428 | 10161.4654 | 9.8986 | 10.3285 | 10.4954 | 10.2133 | 10.198 | 8.9676 | 9.7084 | 9.3707 | 9.7398 | 10.0391 | 10.1715 | 9.4579 | 9.7517 | 9.8535 | 10.3911 | 9.875 | 9.7636 | 9.9566 | 10.2759 | 9.9948 | 9.8994 | 10.2608 | [6 empty] | 9.734 | 9.7572 | 10.5532 | 10.797
167 | CCV 10 ppb MAJORS |  | 06/17/2023 05:04:50 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.816 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9381.6047 | 9735.7882 | 9.5173 | 8.783 | 10954.4274 | 9589.6856 | 10173.9004 | 10.0587 | 10.435 | 10.4507 | 10.1238 | 10.2891 | 9.4533 | 9.9703 | 9.6684 | 9.7898 | 10.0039 | 10.2206 | 9.5555 | 10.1011 | 9.9143 | 10.3912 | 10.362 | 9.79 | 9.9271 | 10.0439 | 10.2064 | 9.9061 | 10.2198 | [6 empty] | 9.702 | 9.7842 | 10.4492 | 10.6178
168 | 10_15_21_F_11 |  | 06/17/2023 05:10:50 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_11.817 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 5592.9865 | 7242.6501 | 5.3809 | 2.8956 | 13472.9781 | 1637.2603 | 32214.5355 | 0.59902 | 0.045408 | 0.052248 | 6.4671 | 4.3891 | 6.6122 | 0.010852 | 0.16026 | 1.3174 | 1.3805 | 3.8596 | 2.9805 | 2.7283 | 3.0592 | 0.48999 | 0.10636 | 0.06849 | 0.075806 | 0.48096 | 0.083389 | 2.7031 | 2.8195 | [6 empty] | 0.006384 | 0.0037648 | 0.074611 | 0.075565
169 | 10_15_21_F_11 |  | 06/17/2023 05:17:09 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_F_11.818 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 5709.2532 | 7276.4977 | 6.0489 | 2.9814 | 14036.4157 | 1622.1786 | 32411.0066 | 0.63879 | 0.047943 | 0.056146 | 6.4569 | 4.3605 | 6.5465 | 0.0062589 | 0.16589 | 1.3106 | 1.388 | 3.9062 | 2.9168 | 2.8841 | 3.0314 | 0.42432 | 0.060216 | 0.093855 | 0.066084 | 0.44742 | 0.044583 | 2.7347 | 2.8179 | [6 empty] | 0.0081176 | 0.0047789 | 0.073105 | 0.059691
170 | 10_15_21_UF_1 |  | 06/17/2023 05:23:09 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_1.819 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10677.3627 | 6508.6069 | -0.091374 | 4.4228 | 30497.0149 | 2696.2837 | 43622.0162 | 0.17706 | 0.14795 | 0.13663 | 0.90798 | 18.9121 | 1.8741 | 0.020279 | 0.32587 | 0.51999 | 0.59483 | 1.7301 | 2.5021 | 2.4878 | 2.5515 | 0.66694 | 0.036776 | 0.046769 | 0.043952 | 0.54831 | 0.0049962 | 20.2131 | 20.74 | [6 empty] | 0.0094326 | 0.0058988 | 0.027038 | 0.032148
171 | 10_15_21_UF_1 |  | 06/17/2023 05:29:28 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_1.820 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10912.8776 | 6553.5644 | 0.036679 | 4.675 | 31010.5811 | 2777.1165 | 44445.534 | 0.1867 | 0.15768 | 0.13581 | 0.88232 | 19.2145 | 2.0467 | 0.022778 | 0.3689 | 0.52478 | 0.60167 | 1.7154 | 2.416 | 2.3614 | 2.5725 | 0.66903 | 0.043898 | 0.050948 | 0.037463 | 0.48918 | 0.030998 | 20.2482 | 20.9722 | [6 empty] | 0.0091044 | 0.0062453 | 0.026177 | 0.030681
172 | 10_15_21_UF_6 |  | 06/17/2023 05:35:27 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_6.821 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9778.1275 | 6907.1205 | 0.0462 | 2.8905 | 22982.8535 | 2177.2132 | 40607.0865 | 0.50769 | 0.069257 | 0.068044 | 1.7509 | 7.575 | 2.737 | 0.014942 | 0.2576 | 1.227 | 1.3191 | 3.8486 | 2.5823 | 2.4143 | 2.6986 | 0.63864 | 0.096684 | 0.081297 | 0.071793 | 0.44489 | 0.055334 | 10.8516 | 11.2565 | [6 empty] | 0.013609 | 0.013152 | 0.044547 | 0.044508
173 | 10_15_21_UF_6 |  | 06/17/2023 05:41:46 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_6.822 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9509.7211 | 6878.5393 | 0.13504 | 2.8128 | 22388.3122 | 2150.7365 | 39413.0655 | 0.47552 | 0.071657 | 0.065158 | 1.7019 | 7.5941 | 2.5708 | 0.013491 | 0.23676 | 1.264 | 1.3155 | 3.8334 | 2.6357 | 2.7498 | 2.6374 | 0.62927 | 0.074189 | 0.099048 | 0.072203 | 0.51086 | 0.036481 | 10.7183 | 11.203 | [6 empty] | 0.018228 | 0.012559 | 0.044285 | 0.046527
174 | 10_15_21_UF_8 |  | 06/17/2023 05:47:46 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_8.823 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10167.1898 | 7133.965 | 1.4165 | 2.6727 | 22209.4721 | 2276.7367 | 40130.2502 | 0.65756 | 0.10186 | 0.079606 | 0.76854 | 7.8594 | 1.8432 | 0.014035 | 0.256 | 1.1536 | 1.2398 | 3.4563 | 3.272 | 2.9149 | 3.2465 | 0.59386 | 0.12179 | 0.080902 | 0.06668 | 0.49962 | 0.070402 | 9.9962 | 10.4041 | [6 empty] | 0.0099515 | 0.0074862 | 0.03899 | 0.04215
175 | 10_15_21_UF_8 |  | 06/17/2023 05:54:04 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_8.824 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10308.5276 | 7292.6146 | 1.4987 | 2.7549 | 22366.0507 | 2222.1842 | 40438.484 | 0.69611 | 0.10715 | 0.08496 | 0.74773 | 7.8134 | 1.8393 | 0.014904 | 0.24387 | 1.1724 | 1.2325 | 3.4805 | 3.1527 | 2.9835 | 3.2151 | 0.70476 | 0.091005 | 0.075842 | 0.084091 | 0.35646 | 0.044069 | 10.0139 | 10.3562 | [6 empty] | 0.0093285 | 0.0067423 | 0.041126 | 0.042234
176 | 10_15_21_UF_8_LD |  | 06/17/2023 06:00:03 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_8_LD.825 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10279.8382 | 7278.2515 | 3.003 | 5.1627 | 22462.4541 | 2243.611 | 40395.2592 | 0.70522 | 0.1198 | 0.097023 | 1.1432 | 7.3959 | 2.2455 | 0.018584 | 0.26318 | 1.2075 | 1.3062 | 7.5269 | 3.1388 | 2.8911 | 3.1269 | 0.61915 | 0.043555 | 0.072032 | 0.060847 | 0.56655 | 0.06132 | 9.9496 | 9.8832 | [6 empty] | 0.010438 | 0.007625 | 0.047093 | 0.04997
177 | 10_15_21_UF_8_LD |  | 06/17/2023 06:06:21 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_8_LD.826 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10118.7739 | 7023.569 | 3.1024 | 5.6801 | 21919.3357 | 2272.9515 | 39953.4492 | 0.69466 | 0.10874 | 0.099702 | 1.162 | 7.7619 | 2.0974 | 0.020192 | 0.26859 | 1.2615 | 1.3177 | 7.5997 | 3.0957 | 2.9964 | 3.2702 | 0.55472 | 0.1301 | 0.10105 | 0.076206 | 0.59828 | 0.070642 | 9.9627 | 10.3167 | [6 empty] | 0.010205 | 0.0070167 | 0.044442 | 0.04924
178 | LBLANK |  | 06/17/2023 06:12:21 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.827 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 54.6956 | 1.442 | -0.99661 | -0.49387 | 622.9496 | 55.1933 | 41.6973 | 0.00031926 | -0.0091733 | -0.0024377 | -0.58138 | -0.00087292 | -0.30531 | -0.0078885 | 0.0037548 | -0.009191 | -0.011487 | -0.22872 | 0.022123 | 0.0062932 | 0.010213 | -0.063336 | 0.0055366 | 0.0017967 | -0.0056003 | 0.014232 | -0.015347 | -0.0066375 | 0.0044315 | [6 empty] | 0.0008888 | -0.0025361 | -0.010858 | -0.0043132
179 | LBLANK |  | 06/17/2023 06:18:40 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.828 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 55.203 | 0.83572 | -0.90719 | -0.87994 | 515.7658 | 57.3635 | 40.8248 | 0.00031991 | -0.0091201 | -0.001297 | -0.57984 | -0.00024085 | -0.31988 | -0.0071727 | -0.00096481 | -0.013315 | -0.01122 | -0.22905 | -0.046507 | -0.0068514 | 0.0038727 | 0.1286 | 0.0055366 | 0.0018063 | -0.0033927 | -0.011958 | -0.021474 | -0.00815 | 0.0030957 | [6 empty] | 0.00063363 | -0.0025293 | -0.010509 | -0.0042858
180 | CCV 1ppb TRACE |  | 06/17/2023 06:24:40 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.829 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 968.0848 | 959.1468 | 1.2952 | 0.059317 | 1514.0892 | 1011.1718 | 996.4636 | 1.0541 | 1.0725 | 1.0959 | 0.95095 | 1.0273 | 0.94525 | 0.9819 | 0.97449 | 0.98646 | 1.0628 | 0.84725 | 0.88584 | 1.0034 | 1.0054 | 1.1615 | 1.1205 | 1.0212 | 0.92404 | 0.90827 | 1.1525 | 1.0126 | 0.99527 | [6 empty] | 0.98473 | 0.98321 | 1.1529 | 1.1166
181 | CCV 1ppb TRACE |  | 06/17/2023 06:31:00 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.830 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 989.569 | 1004.8621 | 1.1962 | 0.085023 | 1760.5622 | 1005.7186 | 1055.8563 | 1.0353 | 1.0936 | 1.0879 | 0.92351 | 1.0383 | 1.0326 | 1.0053 | 0.9927 | 1.0485 | 1.0576 | 0.83569 | 0.83491 | 1.0503 | 0.9721 | 1.1549 | 0.86982 | 1.0134 | 0.98219 | 0.97531 | 1.0308 | 0.99013 | 1.0449 | [6 empty] | 0.96264 | 0.97011 | 1.1216 | 1.1282
182 | LFB |  | 06/17/2023 06:37:00 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.831 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 8970.3074 | 9183.3327 | 9.2986 | 11.5407 | 10881.8268 | 9176.0433 | 9770.5853 | 9.6765 | 10.1655 | 10.2248 | 9.854 | 9.8512 | 8.8593 | 9.5357 | 9.1662 | 9.5941 | 9.9232 | 10.0133 | 9.4709 | 9.8337 | 9.8297 | 9.9138 | 9.7686 | 9.9008 | 9.8871 | 10.1287 | 10.0335 | 9.6339 | 9.9816 | [6 empty] | 9.6126 | 9.6824 | 10.1483 | 10.3675
183 | LFB |  | 06/17/2023 06:43:18 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.832 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 8867.5677 | 9124.608 | 9.0696 | 11.6353 | 10358.6309 | 9136.001 | 9792.1015 | 9.651 | 9.9827 | 9.9947 | 9.5324 | 9.8232 | 8.8429 | 9.5589 | 9.2102 | 9.5548 | 9.6558 | 9.4595 | 9.0273 | 9.728 | 9.669 | 9.5589 | 10.1294 | 9.8792 | 10.122 | 9.8282 | 9.9723 | 9.3787 | 9.762 | [6 empty] | 9.3795 | 9.4313 | 9.8528 | 11.2168
184 | CCV 10 ppb MAJORS |  | 06/17/2023 06:49:17 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.833 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9306.831 | 9713.4271 | 9.1386 | 9.2248 | 10647.9564 | 9569.9926 | 10321.301 | 10.2197 | 10.4124 | 10.5272 | 10.534 | 10.2872 | 9.3711 | 9.8855 | 9.4574 | 9.9306 | 10.1239 | 10.3161 | 9.2763 | 9.8277 | 9.9325 | 10.1599 | 9.835 | 10.2077 | 9.7385 | 10.0393 | 10.3327 | 9.9262 | 10.3355 | [6 empty] | 9.7562 | 9.8718 | 10.3118 | 10.5736
185 | CCV 10 ppb MAJORS |  | 06/17/2023 06:55:37 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.834 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9321.9399 | 9621.4959 | 9.6796 | 9.0552 | 10832.7867 | 9651.5489 | 10077.7649 | 10.1472 | 10.5289 | 10.5721 | 10.2148 | 10.5022 | 9.5828 | 9.9356 | 9.5294 | 9.8754 | 10.1215 | 10.1735 | 9.5486 | 10.0409 | 9.8948 | 10.1073 | 10.3434 | 10.1782 | 9.9266 | 10.2306 | 10.2963 | 9.9969 | 10.3937 | [6 empty] | 9.6756 | 9.772 | 10.4837 | 10.7976
186 | 10_15_21_UF_8_DUP |  | 06/17/2023 07:01:37 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_8_DUP.835 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 10384.274 | 7292.7311 | 0.011584 | 2.6336 | 22873.8333 | 2289.1992 | 41227.5619 | 0.69471 | 0.12987 | 0.10074 | 1.5959 | 7.6724 | 2.7848 | 0.013787 | 0.25072 | 1.2456 | 1.2584 | 2.6476 | 3.1381 | 2.9242 | 3.2768 | 0.61622 | 0.078082 | 0.13022 | 0.089055 | 0.50312 | 0.067117 | 10.07 | 10.3259 | [6 empty] | 0.0097243 | 0.0070001 | 0.041976 | 0.046897
187 | 10_15_21_UF_8_DUP |  | 06/17/2023 07:07:56 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_8_DUP.836 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9850.2514 | 6887.3176 | -0.048696 | 2.8155 | 22047.339 | 2217.6119 | 39826.8979 | 0.68307 | 0.11885 | 0.10378 | 1.6106 | 7.6629 | 2.5062 | 0.011616 | 0.24294 | 1.1865 | 1.2277 | 2.6478 | 3.1327 | 2.9767 | 3.1707 | 0.55869 | 0.078707 | 0.079626 | 0.086299 | 0.39406 | 0.067957 | 9.8998 | 10.2526 | [6 empty] | 0.012761 | 0.007404 | 0.043746 | 0.046002
188 | 10_15_21_UF_8_DUP_LSP |  | 06/17/2023 07:13:55 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_8_DUP_LSP.837 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 18258.2928 | 15957.0251 | 11.0449 | 15.7559 | 30341.579 | 11662.155 | 46464.476 | 10.7604 | 10.5319 | 10.3998 | 12.0851 | 16.998 | 11.9649 | 9.7962 | 9.4143 | 10.5352 | 10.681 | 11.525 | 12.0351 | 12.7318 | 12.5689 | 10.1763 | 10.1885 | 9.8168 | 9.7248 | 10.3647 | 9.7098 | 18.8764 | 19.5579 | [6 empty] | 9.4169 | 9.424 | 10.2377 | 10.4079
189 | 10_15_21_UF_8_DUP_LSP |  | 06/17/2023 07:20:13 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_8_DUP_LSP.838 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 18290.0209 | 15456.0348 | 10.8726 | 15.6574 | 29759.8218 | 11437.4433 | 46012.4646 | 10.7682 | 10.3547 | 10.4974 | 11.9713 | 16.9751 | 11.7553 | 9.7419 | 9.3663 | 10.2005 | 10.663 | 11.5272 | 12.353 | 12.1812 | 12.4703 | 10.5643 | 10.1203 | 9.74 | 10.0045 | 10.7222 | 9.9553 | 18.9821 | 19.5368 | [6 empty] | 9.4894 | 9.5252 | 10.3605 | 10.4457
190 | 10_15_21_UF_9 |  | 06/17/2023 07:26:13 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_9.839 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 8681.2319 | 6514.3072 | 12.7994 | 6.7919 | 17805.5469 | 1935.3941 | 35857.3928 | 0.55291 | 0.11357 | 0.11179 | 0.52013 | 4.3694 | 1.4831 | 0.0095453 | 0.19127 | 1.0147 | 1.0846 | 2.8383 | 2.5441 | 2.4415 | 2.596 | 0.69525 | 0.069559 | 0.080838 | 0.061884 | 0.42996 | 0.10037 | 7.0353 | 7.2486 | [6 empty] | 0.011369 | 0.0087918 | 0.02821 | 0.028778
191 | 10_15_21_UF_9 |  | 06/17/2023 07:32:31 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_9.840 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 8598.2121 | 6573.2885 | 12.3818 | 6.0986 | 18024.4361 | 1882.2505 | 35627.2489 | 0.5333 | 0.11919 | 0.10858 | 0.48428 | 4.4017 | 1.4087 | 0.0092074 | 0.1751 | 0.99183 | 1.0864 | 2.8711 | 2.453 | 2.3753 | 2.632 | 0.50747 | 0.046004 | 0.08197 | 0.050878 | 0.36164 | 0.074505 | 7.0407 | 7.275 | [6 empty] | 0.010457 | 0.0080746 | 0.026571 | 0.030637
192 | 10_15_21_UF_10 |  | 06/17/2023 07:38:31 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_10.841 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 6908.1782 | 7277.051 | -0.047034 | 4.3194 | 14283.6831 | 1647.0466 | 33688.9073 | 0.61551 | 0.10345 | 0.086187 | 0.28748 | 4.5774 | 1.2897 | 0.0072488 | 0.17595 | 0.80388 | 0.8599 | 2.1185 | 2.7805 | 2.8519 | 2.8633 | 0.54562 | 0.14189 | 0.076286 | 0.070962 | 0.36588 | 0.094407 | 3.1106 | 3.1919 | [6 empty] | 0.0078488 | 0.0029637 | 0.028569 | 0.033127
193 | 10_15_21_UF_10 |  | 06/17/2023 07:44:50 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_10.842 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 6816.0919 | 7000.3872 | -0.18804 | 4.0854 | 14019.0841 | 1695.7911 | 33118.622 | 0.59768 | 0.098621 | 0.087848 | 0.26931 | 4.5917 | 1.1492 | 0.0065144 | 0.13791 | 0.79565 | 0.85683 | 2.1546 | 2.7302 | 2.7956 | 2.8861 | 0.61787 | 0.070913 | 0.098112 | 0.059739 | 0.40034 | 0.026922 | 3.0991 | 3.2061 | [6 empty] | 0.0075895 | 0.0031747 | 0.030267 | 0.027758
194 | 10_15_21_UF_11 |  | 06/17/2023 07:50:50 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_11.843 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 6563.6549 | 7219.8262 | -0.040966 | 3.3337 | 13955.4355 | 1598.3116 | 32909.7825 | 0.58637 | 0.073131 | 0.056315 | 1.4423 | 4.0029 | 2.1636 | 0.0069465 | 0.12095 | 0.78428 | 0.85014 | 3.0595 | 2.9087 | 2.7622 | 2.9895 | 0.34193 | 0.046214 | 0.071946 | 0.053482 | 0.38851 | 0.095383 | 2.6352 | 2.7668 | [6 empty] | 0.0054095 | 0.0046742 | 0.032175 | 0.032871
195 | 10_15_21_UF_11 |  | 06/17/2023 07:57:09 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_UF_11.844 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 6694.7831 | 7168.185 | 0.095456 | 3.5003 | 14096.4545 | 1566.5985 | 32446.2584 | 0.58537 | 0.077728 | 0.063078 | 1.4567 | 4.0649 | 2.1564 | 0.0075975 | 0.16379 | 0.80875 | 0.86873 | 3.0912 | 2.8091 | 2.7826 | 3.0339 | 0.52947 | 0.071978 | 0.099996 | 0.067411 | 0.38781 | 0.069767 | 2.6288 | 2.7605 | [6 empty] | 0.0066366 | 0.0043425 | 0.03311 | 0.036126
196 | LBLANK |  | 06/17/2023 08:03:10 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.845 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 55.035 | 0.87745 | -1.013 | -0.87717 | 568.7107 | 58.9971 | 39.1036 | -0.00047409 | -0.0055501 | -0.0021796 | -0.57735 | -0.00023199 | -0.31439 | -0.0078885 | 0.00083711 | -0.0086675 | -0.011112 | -0.23141 | 0.014612 | -0.0066593 | 0.0085284 | 0.022527 | 0.0055366 | 0.011718 | -0.0074358 | -0.0091548 | -0.0084538 | -0.0067453 | 0.0027929 | [6 empty] | 0.00050427 | -0.0024531 | -0.010606 | -0.0041835
197 | LBLANK |  | 06/17/2023 08:09:28 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.846 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 53.943 | 0.23633 | -0.99374 | -0.94322 | 454.9155 | 62.6287 | 42.7364 | -0.00048271 | -0.007886 | -0.0021642 | -0.58084 | -0.00022223 | -0.37327 | -0.0071464 | 0.0083564 | -0.012219 | -0.011096 | -0.23312 | -0.067417 | -0.0021395 | 0.0089333 | 0.23586 | 0.013317 | 0.0068941 | -0.0062653 | -0.01877 | -0.014927 | -0.0084082 | 0.0025485 | [6 empty] | 0.0010413 | -0.0022897 | -0.010089 | -0.0042789
198 | CCV 1ppb TRACE |  | 06/17/2023 08:15:29 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.847 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 1002.3381 | 996.7207 | 1.3114 | 0.063046 | 1835.2714 | 1027.1575 | 1055.0865 | 1.0733 | 1.0884 | 1.0695 | 0.95485 | 1.0151 | 1.0531 | 1.0411 | 1.0045 | 1.0406 | 1.0572 | 0.81105 | 0.8506 | 1.0771 | 0.99088 | 1.287 | 0.90792 | 0.9629 | 0.98127 | 0.93088 | 1.0329 | 0.99225 | 1.0192 | [6 empty] | 1.0105 | 0.99141 | 1.2018 | 1.1506
199 | CCV 1ppb TRACE |  | 06/17/2023 08:21:48 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.848 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 997.0932 | 995.0847 | 1.4091 | 0.051982 | 1719.7104 | 1061.1018 | 978.5833 | 1.0568 | 1.0815 | 1.0769 | 0.93359 | 1.0394 | 1.0595 | 1.0099 | 0.94677 | 1.0184 | 1.0648 | 0.81691 | 0.90047 | 0.92465 | 0.98077 | 1.2687 | 1.067 | 1.0128 | 1.0302 | 0.95362 | 0.97098 | 0.99058 | 1.0264 | [6 empty] | 0.99939 | 0.99145 | 1.1946 | 1.146
200 | CCV 10 ppb MAJORS |  | 06/17/2023 08:27:48 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.849 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9196.0967 | 9579.6142 | 9.6821 | 8.8967 | 11025.4809 | 9642.3472 | 10209.7829 | 9.8544 | 10.2666 | 10.5153 | 9.984 | 10.2443 | 9.1805 | 9.7764 | 9.3451 | 9.8159 | 9.9135 | 10.1591 | 9.4528 | 9.8617 | 9.9315 | 10.4755 | 10.0874 | 10.0847 | 10.1008 | 10.2158 | 10.2704 | 10.0021 | 10.3519 | [6 empty] | 9.8043 | 9.8706 | 10.8231 | 11.1097
201 | CCV 10 ppb MAJORS |  | 06/17/2023 08:34:07 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.850 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9343.0282 | 9610.9488 | 9.238 | 8.8704 | 10996.6849 | 9457.8814 | 10283.2552 | 10.1219 | 10.5213 | 10.4927 | 10.1373 | 10.3406 | 9.3419 | 9.8617 | 9.4333 | 9.8444 | 10.0404 | 10.1828 | 8.909 | 10.0913 | 9.973 | 9.5723 | 10.678 | 9.9592 | 10.05 | 10.1431 | 10.2223 | 9.9624 | 10.4288 | [6 empty] | 9.759 | 9.8639 | 10.1266 | 11.0322
202 | 10_15_21_FB |  | 06/17/2023 08:40:08 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_FB.851 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 309.5635 | 5.9134 | 8.7846 | 0.64889 | 660.5355 | 73.1704 | 50.8432 | 0.00037281 | 0.081431 | 0.073486 | 1.4713 | 0.058921 | 1.571 | -0.0056663 | 0.0097588 | 0.10545 | 0.11742 | 7.5315 | -0.052131 | -0.00678 | 0.0090817 | 0.13273 | 0.0055366 | 0.0051056 | -0.0070843 | -0.021988 | -0.021474 | 0.00096373 | 0.012359 | [6 empty] | 0.0011888 | -0.00094353 | 0.050446 | 0.053852
203 | 10_15_21_FB |  | 06/17/2023 08:46:26 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_FB.852 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 300.7921 | 6.6407 | 8.8997 | 0.585 | 577.2082 | 80.9715 | 51.9309 | 0.0026006 | 0.079389 | 0.076795 | 1.478 | 0.050685 | 1.4772 | -0.0061818 | 0.014396 | 0.09774 | 0.11869 | 7.5547 | -0.035538 | -0.006632 | 0.0082709 | 0.16972 | 0.013568 | 0.0045383 | -0.0084892 | 0.0132 | -0.021474 | -0.00030517 | 0.010018 | [6 empty] | 0.00078386 | -0.0014934 | 0.047406 | 0.052525
204 | 10_15_21_BB |  | 06/17/2023 08:52:26 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_BB.853 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 56.1217 | 1.5306 | -0.36638 | -0.72755 | 443.6407 | 59.1756 | 44.6656 | -8.8222e-06 | 0.01434 | 0.019423 | -0.35299 | 0.002982 | -0.23378 | -0.0061632 | 0.011964 | 0.075239 | 0.085513 | -0.06074 | 0.0029923 | 0.0072197 | 0.00084329 | 0.10906 | 0.013477 | 0.0045487 | -0.0073402 | 0.0022803 | -0.021474 | -0.007672 | 0.0025852 | [6 empty] | 0.00064456 | -0.0020845 | 0.013018 | 0.019339
205 | 10_15_21_BB |  | 06/17/2023 08:58:45 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_BB.854 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 56.1369 | 0.23633 | -0.19821 | -0.72812 | 487.6436 | 56.437 | 46.5609 | -0.00091385 | 0.012905 | 0.020062 | -0.3422 | 0.0043811 | -0.16623 | -0.0076969 | 0.0053235 | 0.08384 | 0.082649 | -0.061945 | 0.014109 | -0.01135 | -0.0010676 | 0.00016846 | 0.0055366 | 0.0071314 | -0.0061376 | -0.022243 | -0.014663 | -0.0064993 | 0.0025805 | [6 empty] | 0.00050427 | -0.0022672 | 0.012778 | 0.019384
206 | 10_15_21_LAB_A |  | 06/17/2023 09:04:45 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_LAB_A.855 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 56.9618 | 0.23633 | -0.39347 | -0.7624 | 312.4661 | 55.5559 | 41.9002 | -0.00091385 | 0.038638 | 0.043247 | -0.18434 | 0.0070936 | -0.07752 | -0.0076986 | 0.00026748 | 0.088188 | 0.08877 | -0.039874 | -0.018785 | -0.01135 | -0.0012161 | 0.023675 | 0.0055366 | -0.00050913 | -0.010861 | 0.032175 | -0.021474 | -0.0030283 | 0.0064621 | [6 empty] | 0.00078049 | -0.0022502 | 0.015723 | 0.020626
207 | 10_15_21_LAB_A |  | 06/17/2023 09:11:03 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_LAB_A.856 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 56.3986 | 0.23633 | -0.38434 | -0.75607 | 619.7993 | 55.4549 | 42.808 | 0.00041334 | 0.040328 | 0.046871 | -0.16125 | 0.0061968 | -0.053997 | -0.0073192 | 0.0023391 | 0.097948 | 0.093321 | -0.025752 | -0.019415 | -0.0066474 | -0.0017623 | 0.19994 | 0.0055366 | 0.0020185 | -0.0084854 | -0.023321 | -0.021474 | -0.001479 | 0.0073253 | [6 empty] | 0.00078087 | -0.0022045 | 0.017173 | 0.020528
208 | 10_15_21_LAB_B |  | 06/17/2023 09:17:03 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_LAB_B.857 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 55.8835 | 0.89899 | -0.32662 | -0.69175 | 463.7233 | 57.0996 | 41.9431 | -0.00091385 | 0.012211 | 0.016121 | -0.36697 | 0.0028359 | -0.1743 | -0.0073101 | 0.007829 | 0.098107 | 0.14728 | 0.098667 | -0.020084 | -0.0017266 | -0.0027726 | 0.11412 | 0.0055366 | 0.0046528 | -0.0084127 | -0.085032 | -0.014538 | -0.0089123 | 0.0021771 | [6 empty] | 0.00064589 | -0.0025438 | 0.016229 | 0.020504
209 | 10_15_21_LAB_B |  | 06/17/2023 09:23:20 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\10_15_21_LAB_B.858 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 54.2331 | 0.23633 | -0.18095 | -0.6595 | 232.4913 | 61.6954 | 41.8829 | 0.00043462 | 0.0085678 | 0.015795 | -0.3636 | 0.0021673 | -0.23348 | -0.0071177 | 0.0019464 | 0.11115 | 0.10095 | 0.097623 | 0.029191 | -0.01135 | -0.0025741 | 0.074651 | 0.013686 | -0.00050913 | -0.01206 | -0.081376 | -0.014517 | -0.0086601 | 0.0020907 | [6 empty] | 0.00078281 | -0.002016 | 0.017091 | 0.020987
210 | 5_18_23_DB |  | 06/17/2023 09:29:20 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_DB.859 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 54.7914 | 0.23633 | -0.5834 | 2.7988 | 400.8324 | 54.8712 | 42.8917 | -0.00091385 | -0.0069955 | -0.0014792 | -0.51844 | 0.00065848 | -0.3335 | -0.0071132 | -0.005392 | 0.07844 | 0.083187 | -0.14136 | -0.02964 | -0.01135 | -0.0020394 | -0.0062654 | 0.0055366 | -0.00050913 | -0.0096315 | -0.045747 | -0.021474 | -0.0094574 | 0.0014597 | [6 empty] | 0.00078973 | -0.0027816 | 0.017861 | 0.022752
211 | 5_18_23_DB |  | 06/17/2023 09:35:39 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_DB.860 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 54.5141 | 0.23633 | -0.59238 | 3.9776 | 284.2138 | 60.6181 | 42.9465 | 0.00045495 | -0.0039211 | -0.0013658 | -0.50863 | 0.00031322 | -0.41929 | -0.0073018 | -0.0015406 | 0.088481 | 0.08197 | -0.14344 | 0.024899 | -0.0065295 | -0.0018156 | 0.047586 | 0.0055366 | 0.0020976 | -0.010837 | -0.053863 | -0.007759 | -0.0089791 | 0.0012842 | [6 empty] | 0.00092955 | -0.0022477 | 0.018101 | 0.022976
212 | LBLANK |  | 06/17/2023 09:41:39 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.861 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 53.9622 | 0.23633 | -1.0181 | 0.27629 | 380.1367 | 56.2211 | 41.0345 | 4.3949e-06 | -0.0085363 | -0.0022115 | -0.58825 | -0.00039326 | -0.3623 | -0.0072896 | -0.0088998 | -0.010558 | -0.010593 | -0.23139 | -0.014607 | -0.0063945 | -0.0010943 | -0.06612 | 0.0055366 | 0.004779 | -0.0095937 | 0.0057133 | -0.014463 | -0.0078214 | 0.0022834 | [6 empty] | 0.00079341 | -0.0025369 | -0.010532 | -0.0040134
213 | LBLANK |  | 06/17/2023 09:47:59 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.862 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 53.9519 | 0.23633 | -0.99086 | -0.76234 | 211.4654 | 63.8271 | 40.0575 | -0.00091385 | -0.0061699 | -0.00078087 | -0.58438 | -0.00083815 | -0.39507 | -0.0074963 | -0.0044315 | -0.013919 | -0.010237 | -0.2302 | 0.00056601 | -0.01135 | -0.0027413 | -0.032854 | 0.013787 | -0.00050913 | -0.01206 | -0.11708 | -0.021474 | -0.0086093 | 0.0023704 | [6 empty] | 0.00064791 | -0.0022976 | -0.010158 | -0.0039429
214 | CCV 1ppb TRACE |  | 06/17/2023 09:53:59 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.863 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 941.3352 | 971.9387 | 0.8378 | 0.11132 | 1197.1077 | 1024.1048 | 1066.1094 | 1.0225 | 1.0422 | 1.0843 | 0.88773 | 1.0488 | 0.97888 | 0.98638 | 0.96001 | 1.0284 | 1.0613 | 0.83248 | 0.79129 | 0.88526 | 0.96246 | 1.2541 | 1.0028 | 1.0659 | 0.96182 | 0.90306 | 0.94047 | 0.97769 | 1.0342 | [6 empty] | 0.9842 | 0.99871 | 1.2063 | 1.1878
215 | CCV 1ppb TRACE |  | 06/17/2023 10:00:19 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.864 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 970.6048 | 967.9932 | 0.98345 | 0.050434 | 1617.0556 | 1048.122 | 1118.9109 | 1.0582 | 1.0757 | 1.095 | 0.94119 | 1.0547 | 0.98358 | 0.99119 | 0.98077 | 0.99756 | 1.0699 | 0.81686 | 0.85719 | 0.99734 | 0.99774 | 1.1592 | 1.0502 | 0.93461 | 0.92341 | 0.94712 | 1.08 | 1.003 | 1.0383 | [6 empty] | 1.0128 | 1.0111 | 1.1968 | 1.1765
216 | LFB |  | 06/17/2023 10:06:18 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.865 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 8604.3312 | 8848.0206 | 9.4861 | 10.8847 | 10437.6528 | 9135.7801 | 9755.5838 | 9.5967 | 9.8197 | 9.7091 | 9.2974 | 9.6264 | 8.5047 | 9.2071 | 8.9238 | 9.2904 | 9.4477 | 9.3339 | 9.072 | 9.2887 | 9.5888 | 9.8521 | 9.3578 | 9.4377 | 9.9006 | 9.5877 | 10.1963 | 9.3895 | 9.8947 | [6 empty] | 9.4288 | 9.514 | 10.5896 | 10.8361
217 | LFB |  | 06/17/2023 10:12:36 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LFB.866 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 8488.4552 | 8863.8757 | 9.0709 | 11.3655 | 10232.5317 | 8918.4416 | 9541.5482 | 9.341 | 9.5082 | 10.0825 | 9.539 | 9.7991 | 8.522 | 9.1402 | 8.9403 | 9.3112 | 9.6675 | 9.5192 | 9.166 | 9.9727 | 9.709 | 10.0455 | 9.4721 | 9.5825 | 9.8053 | 10.0227 | 10.4228 | 9.6155 | 10.0292 | [6 empty] | 9.588 | 9.728 | 10.5883 | 10.7092
218 | CCV 10 ppb MAJORS |  | 06/17/2023 10:18:36 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.867 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9088.8882 | 9458.5212 | 9.7384 | 9.1108 | 11032.9994 | 9584.4373 | 10268.8251 | 10.0607 | 10.3768 | 10.457 | 10.1439 | 10.2826 | 9.2199 | 9.7878 | 9.2477 | 9.7816 | 10.0346 | 9.9822 | 9.4649 | 9.426 | 9.8989 | 10.2056 | 10.015 | 10.1258 | 9.9487 | 10.0483 | 10.2024 | 9.9111 | 10.431 | [6 empty] | 9.6437 | 9.7748 | 10.8962 | 11.1018
219 | CCV 10 ppb MAJORS |  | 06/17/2023 10:24:55 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.868 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9136.6389 | 9344.4324 | 9.3527 | 8.8583 | 11013.0283 | 9452.938 | 10278.224 | 10.0455 | 10.1677 | 10.581 | 10.0385 | 10.2585 | 9.2147 | 9.7741 | 9.4132 | 9.7052 | 10.0953 | 10.1275 | 9.4815 | 9.6238 | 9.8198 | 10.1195 | 10.0042 | 9.8117 | 9.8013 | 10.4542 | 10.049 | 9.9984 | 10.4653 | [6 empty] | 9.7188 | 9.8832 | 10.7595 | 10.8331
220 | 5_18_23_DB |  | 06/17/2023 10:30:55 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_DB.869 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 56.0874 | 0.23633 | -0.34413 | 3.0299 | 483.6795 | 56.5471 | 42.803 | 0.00039696 | 1.6621e-05 | 0.0065821 | 1.1027 | 0.0076569 | 1.2403 | -0.0060061 | -0.0043181 | 0.081987 | 0.075413 | -0.13935 | -0.0097701 | -0.0066049 | -0.0019681 | 0.12984 | 0.0055366 | 0.0095666 | -0.0073449 | -0.0050493 | -0.0014681 | 0.0034523 | 0.0037905 | [6 empty] | 0.00050427 | -0.0019947 | 0.017713 | 0.021232
221 | 5_18_23_DB |  | 06/17/2023 10:37:14 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_DB.870 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 55.5641 | 0.87935 | -0.45883 | 4.2547 | 414.7607 | 59.2883 | 40.9256 | -0.00047279 | 0.00071555 | 0.0043316 | 1.0919 | 0.0080144 | 1.1945 | -0.0073199 | -0.0050174 | 0.082157 | 0.080496 | -0.14309 | -0.032794 | -0.0020638 | -0.0012618 | 0.084607 | 0.013394 | 0.0020228 | -0.010879 | -0.064205 | -0.0080586 | -0.00069779 | 0.0019706 | [6 empty] | 0.00064062 | -0.002559 | 0.015932 | 0.022555
222 | 5_18_23_DB |  | 06/17/2023 10:43:14 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_DB.871 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 55.0292 | 0.23633 | -0.41535 | 4.4563 | 526.0722 | 63.4354 | 46.4853 | -3.8842e-05 | 0.007231 | 0.013075 | -0.38083 | 0.0034809 | -0.10681 | 0.06763 | 0.0058478 | 0.089958 | 0.08793 | -0.09427 | -0.091237 | -0.0019545 | -0.0011934 | 0.43327 | 0.0055366 | 0.004616 | -0.0084655 | 0.013129 | -0.014676 | -0.004036 | 0.0023358 | [6 empty] | 0.00078193 | -0.0023205 | 0.025718 | 0.025116
223 | 5_18_23_DB |  | 06/17/2023 10:49:33 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_DB.872 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 55.5987 | 0.23633 | -0.542 | 4.249 | 541.8147 | 54.1982 | 41.9035 | -2.0521e-05 | 0.0043401 | 0.0089419 | -0.35624 | 0.0030895 | -0.10843 | 0.063046 | -0.0055082 | 0.086176 | 0.090599 | -0.10312 | 0.0080738 | -0.01135 | -0.0015096 | 0.18504 | 0.037918 | 0.0020514 | -0.007219 | -0.022544 | -0.021474 | -0.0054082 | 0.0025307 | [6 empty] | 0.00092317 | -0.0027262 | 0.020813 | 0.020228
224 | 5_18_23_DB |  | 06/17/2023 10:55:33 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_DB.873 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 56.4577 | 0.23633 | -0.49485 | 4.087 | 584.6872 | 57.1195 | 41.9141 | -0.00046435 | -0.0015507 | 0.0015378 | -0.46333 | 0.0056896 | -0.10679 | -0.0071152 | 0.00051067 | 0.065643 | 0.060561 | -0.09366 | 0.0048129 | -0.01135 | -0.0031009 | 0.03863 | 0.0055366 | -0.00050913 | -0.0058545 | -0.035078 | -0.021474 | -0.00518 | 0.0018303 | [6 empty] | 0.0010707 | -0.0022452 | 0.0097389 | 0.014763
225 | 5_18_23_DB |  | 06/17/2023 11:01:51 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_DB.874 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 56.1408 | 0.23633 | -0.58117 | 3.9238 | 489.0645 | 57.4775 | 43.8027 | -0.00091385 | -0.0046186 | 1.0946e-05 | -0.46647 | 0.0044928 | -0.19377 | -0.0072986 | 0.0047816 | 0.05814 | 0.06201 | -0.096595 | -0.031544 | 0.003152 | -0.0042019 | -0.023642 | 0.0055366 | 0.012573 | -0.0059201 | -0.035507 | -0.021474 | -0.0060373 | 0.0017349 | [6 empty] | 0.00064873 | -0.0021798 | 0.0081603 | 0.016227
226 | 5_18_23_MFB |  | 06/17/2023 11:07:51 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_MFB.875 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4502.9532 | 4575.0836 | 4.35 | 9.6796 | 5801.114 | 4707.6221 | 4994.9937 | 4.9413 | 5.2152 | 5.117 | 4.8635 | 5.1413 | 4.775 | 4.8546 | 4.6213 | 5.0005 | 5.0256 | 4.6334 | 4.5372 | 4.8443 | 4.9364 | 5.1613 | 5.2911 | 5.0694 | 4.9715 | 4.8808 | 4.9016 | 4.7381 | 5.1323 | [6 empty] | 4.6559 | 4.7335 | 5.656 | 5.6578
227 | 5_18_23_MFB |  | 06/17/2023 11:14:09 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_MFB.876 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4463.4639 | 4710.4861 | 4.735 | 9.6626 | 5434.759 | 4631.4346 | 5054.1685 | 4.8814 | 5.0092 | 5.1374 | 4.8842 | 5.0269 | 4.5976 | 4.7959 | 4.5058 | 4.9135 | 4.9987 | 4.6561 | 4.4796 | 4.7587 | 4.7943 | 5.0754 | 5.2604 | 4.7412 | 4.8838 | 5.0667 | 4.7738 | 4.749 | 4.9597 | [6 empty] | 4.7477 | 4.7746 | 5.5546 | 5.638
228 | 5_18_23_MFB |  | 06/17/2023 11:20:08 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_MFB.877 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4428.9243 | 4430.2363 | 4.6763 | 9.9353 | 5721.9346 | 4590.4168 | 4902.0196 | 4.8057 | 5.0277 | 4.9683 | 4.5221 | 4.8921 | 4.2767 | 4.7135 | 4.5537 | 4.871 | 4.9143 | 4.5065 | 4.4756 | 4.3624 | 4.6548 | 5.2603 | 5.1506 | 4.8765 | 4.7152 | 4.8929 | 4.8166 | 4.5883 | 4.8284 | [6 empty] | 4.6554 | 4.6797 | 5.5783 | 6.1916
229 | 5_18_23_MFB |  | 06/17/2023 11:26:26 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\5_18_23_MFB.878 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 4388.0558 | 4647.9886 | 4.4676 | 10.0162 | 5605.7713 | 4604.6739 | 4913.2522 | 4.8383 | 5.1078 | 5.1317 | 4.6075 | 4.9353 | 4.368 | 4.7526 | 4.6859 | 4.9266 | 5.0203 | 4.5944 | 4.1423 | 4.7415 | 4.7736 | 4.6372 | 4.8249 | 4.7784 | 4.6967 | 4.4521 | 4.679 | 4.7558 | 4.9286 | [6 empty] | 4.6784 | 4.704 | 5.0777 | 5.5252
230 | LBLANK |  | 06/17/2023 11:32:26 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.879 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 54.5088 | 0.8925 | -0.97753 | 0.44344 | 471.7096 | 64.0202 | 40.0401 | 0.00089021 | -0.0052172 | -0.0014034 | -0.57981 | -0.0005055 | -0.15477 | -0.0075028 | 0.0056466 | -0.010101 | -0.010062 | -0.23466 | -0.011186 | -0.01135 | -0.00099884 | 0.060126 | 0.0055366 | 0.0046556 | -0.0060193 | 0.0016578 | -0.01462 | 0.0021293 | 0.0029673 | [6 empty] | 0.00064673 | -0.0020761 | -0.010465 | -0.0040721
231 | LBLANK |  | 06/17/2023 11:38:51 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\LBLANK.880 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 54.7736 | 0.23633 | -0.97863 | -0.6603 | 461.2309 | 58.709 | 40.0395 | -1.947e-05 | -0.0050509 | -0.0020565 | -0.59531 | -0.00067734 | -0.28885 | -0.0075076 | -0.0011957 | -0.012398 | -0.012075 | -0.23104 | 0.0064449 | -0.01135 | -0.0027775 | 0.11097 | 0.0055366 | -0.00050913 | -0.0072037 | -0.017012 | -0.021474 | -0.0016192 | 0.0022626 | [6 empty] | 0.00050427 | -0.0024889 | -0.010459 | -0.0041141
232 | CCV 1ppb TRACE |  | 06/17/2023 11:44:51 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.881 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 967.4726 | 1004.7918 | 1.361 | 0.24606 | 1518.8956 | 990.955 | 1104.9739 | 1.019 | 1.0727 | 1.0913 | 0.91663 | 1.0591 | 1.1138 | 0.97472 | 0.98598 | 1.0189 | 1.0193 | 0.81806 | 0.85939 | 0.99708 | 0.99701 | 1.1505 | 0.8848 | 0.99406 | 1.0403 | 1.0933 | 1.002 | 0.98413 | 1.0285 | [6 empty] | 0.96878 | 0.98541 | 1.2091 | 1.1621
233 | CCV 1ppb TRACE |  | 06/17/2023 11:51:11 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 1ppb TRACE.882 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 971.2691 | 1012.0527 | 1.2375 | 0.23216 | 1682.2554 | 1039.9667 | 1059.4263 | 1.0437 | 1.0384 | 1.0541 | 0.89143 | 1.083 | 1.0717 | 0.99625 | 1.0157 | 1.0338 | 1.0378 | 0.79909 | 0.86597 | 1.0611 | 0.96681 | 1.4148 | 0.9827 | 0.9507 | 0.98401 | 1.0236 | 1.0158 | 0.97625 | 1.0435 | [6 empty] | 0.95965 | 0.98965 | 1.1992 | 1.1692
234 | CCV 10 ppb MAJORS |  | 06/17/2023 11:57:11 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CCV 10 ppb MAJORS.883 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth | 9036.4133 | 9290.9252 | 9.1368 | 8.8344 | 10684.1627 | 9550.7685 | 10259.8601 | 10.0001 | 10.1686 | 10.5382 | 10.1855 | 10.3267 | 9.0796 | 9.6226 | 9.3076 | 9.688 | 10.0191 | 10.0356 | 9.2331 | 10.1107 | 9.9015 | 10.1159 | 10.0307 | 9.7954 | 10.1018 | 10.1547 | 9.9832 | 10.0086 | 10.4225 | [6 empty] | 9.6742 | 9.7936 | 10.7344 | 11.149
235 | CAL BLANK |  | 06/17/2023 12:25:05 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230515_BAMRAB\CAL BLANK.884 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\EBL Methods\230607_UCFR_5% _BAMRAB.mth

## Internal Standards
Row Index Number | Rh (IS1) 103/103 | Rh (IS2) 103/103 | Rh (IS4) 103/103 | Rh (IS5) 103/103 | Rh (IS6) 103/103 | Rh (IS7) 103/103
1
2
3
4 | 100% | 100% | 100% | 100% | 100% | 100%
5 | 100.4% | 100.2% | 100.2% | 99.9% | 98.19% | 101.8%
6 | 99.33% | 98.34% | 98.99% | 98.96% | 95.43% | 99.25%
7 | 98.4% | 98.02% | 100.3% | 99.99% | 96.31% | 99.11%
8 | 96.44% | 95.8% | 96.57% | 97.76% | 93.95% | 98.62%
9 | 95.19% | 95.52% | 95.89% | 96.38% | 94.22% | 98.28%
10 | 93.85% | 90.94% | 92.39% | 92.66% | 91.44% | 92.35%
11 | 93.11% | 89.26% | 90.58% | 91.62% | 89.43% | 91.1%
12 | 92.8% | 84.29% | 84.33% | 86.04% | 93.17% | 87.43%
13 | 88.05% | 79.22% | 79.29% | 81.35% | 80.66% | 82.78%
14 | 78.14% | 72.24% | 69.65% | 71.72% | 73.39% | 74.82%
16 | 98.14% | 97.31% | 96% | 97.6% | 95.78% | 95.62%
17 | 94.29% | 90.98% | 91.27% | 94.82% | 89.08% | 99.79%
18 | 90.84% | 84.9% | 86.62% | 88.49% | 84.45% | 87.2%
19 | 102.1% | 101% | 101.5% | 103.3% | 100.1% | 111.2%
20 | 92.39% | 90.3% | 91.8% | 93.04% | 90.48% | 91.64%
21 | 89.22% | 83.36% | 85.25% | 86.67% | 82.9% | 80.88%
22 | 93.13% | 89.42% | 92.11% | 92.22% | 87.86% | 91.82%
23 | 88.31% | 81.2% | 83.3% | 85.97% | 82.26% | 85.11%
24 | 93.22% | 89.09% | 90.83% | 92.57% | 86.35% | 90.45%
25 | 82.33% | 78.59% | 79.73% | 82.14% | 79.72% | 82.75%
26 | 86.29% | 83.37% | 84.58% | 86.4% | 80.13% | 84.33%
27 | 85.29% | 82.06% | 83.44% | 85.04% | 80.46% | 83.16%
28 | 86% | 82.14% | 84.5% | 85.76% | 80.04% | 85.3%
29 | 84.03% | 74.88% | 78.34% | 80.25% | 76.75% | 80.04%
30 | 80.86% | 75.84% | 76.82% | 78.73% | 76.93% | 81%
31 | 79.68% | 73.4% | 73.97% | 75.96% | 73.32% | 76.61%
32 | 80.7% | 75.48% | 76.67% | 79.63% | 76.7% | 78.5%
33 | 80.64% | 75.14% | 74.99% | 79.33% | 76.74% | 78.88%
34 | 78.33% | 73.56% | 75.16% | 76.43% | 74.06% | 77.46%
35 | 77.83% | 73.13% | 74.93% | 76.85% | 74.09% | 76.26%
36 | 78.72% | 73.1% | 74.63% | 76.01% | 73.74% | 76.56%
37 | 84% | 78.59% | 81.52% | 83.52% | 77.53% | 81.84%
38 | 83.14% | 78.55% | 79.84% | 81.62% | 78% | 77.21%
39 | 83.25% | 78.52% | 80.81% | 82.55% | 78.15% | 82.08%
40 | 82.64% | 73.78% | 76.88% | 78.33% | 74.42% | 78.4%
41 | 83.34% | 74.42% | 77.21% | 79.09% | 75.88% | 78.81%
42 | 78.82% | 73.64% | 74.04% | 77.44% | 74.68% | 77.82%
43 | 77.24% | 71.76% | 73.9% | 74.61% | 72.72% | 76.55%
44 | 83.93% | 77.69% | 80.4% | 82.41% | 76.32% | 83.78%
45 | 84.28% | 77.19% | 80.23% | 82% | 75.47% | 80.58%
46 | 81.45% | 75.99% | 78.95% | 80.53% | 75.59% | 75.89%
47 | 88.19% | 74.05% | 77.27% | 79.41% | 73.98% | 78.45%
48 | 85.33% | 77.5% | 81.83% | 82.51% | 76.68% | 81.98%
49 | 72.24% | 65.79% | 69.66% | 72.21% | 65.36% | 71.35%
50 | 75.63% | 65.42% | 69.33% | 71.95% | 65.46% | 68.77%
51 | 81.37% | 71.3% | 76.78% | 79.24% | 72.12% | 79.09%
52 | 74.67% | 75.09% | 82.71% | 82.32% | 74.77% | 79.98%
53 | 82.72% | 75.25% | 78.92% | 79.98% | 73.49% | 79.36%
54 | 83.29% | 75.52% | 79.03% | 81.67% | 74.72% | 81.02%
55 | 81.51% | 70.96% | 75.83% | 77.13% | 71.52% | 76.66%
56 | 82.99% | 71.63% | 76.33% | 77.5% | 72.42% | 78.59%
57 | 83.53% | 73.2% | 76.77% | 77.85% | 73.62% | 78.01%
58 | 84.17% | 74.32% | 75.02% | 80.68% | 74.65% | 80.22%
59 | 82.38% | 74.06% | 76.88% | 79.67% | 73.56% | 79.66%
60 | 81.59% | 72.24% | 77.83% | 71.77% | 72.5% | 79.23%
61 | 80.76% | 71.38% | 74.63% | 78.22% | 71.42% | 76.42%
62 | 81.66% | 72.06% | 76.22% | 78.25% | 72.58% | 78.2%
63 | 81.98% | 70.86% | 75.66% | 78.61% | 71.22% | 77.31%
64 | 81.93% | 71.81% | 75.79% | 78.32% | 72.08% | 76.75%
65 | 96.01% | 90.46% | 91.58% | 96.28% | 87.67% | 101.6%
66 | 79.18% | 70.23% | 73.83% | 75.82% | 69.5% | 77.46%
67 | 79.96% | 68.7% | 73.72% | 76.3% | 69.15% | 74.9%
68 | 81.81% | 71.35% | 76.1% | 77.27% | 70.63% | 74.68%
69 | 79.09% | 70.91% | 74.53% | 76.32% | 69.14% | 75.39%
70 | 80.14% | 70.96% | 75% | 76.9% | 70.67% | 76.66%
71 | 78.21% | 66.14% | 71.46% | 73.35% | 66.94% | 73%
72 | 80.29% | 67.91% | 72.92% | 75.02% | 68.57% | 74.83%
73 | 79.54% | 66.48% | 70.87% | 72.6% | 66.31% | 71.74%
74 | 76% | 67.31% | 70.03% | 72.32% | 68.35% | 72.87%
75 | 75.66% | 66.9% | 69.1% | 71.4% | 68.35% | 72.19%
76 | 76.28% | 67.51% | 69.56% | 71.78% | 68.01% | 72.08%
77 | 75.61% | 66.49% | 69.15% | 70.86% | 68.02% | 72.25%
78 | 75.86% | 66.88% | 69.44% | 72.22% | 68.14% | 71.31%
79 | 77.04% | 67.22% | 69.96% | 72.49% | 68.47% | 72.23%
80 | 75.94% | 66.97% | 69.98% | 71.82% | 68.02% | 71.67%
81 | 80.36% | 66.14% | 68.64% | 71.4% | 67.79% | 72.24%
82 | 75.28% | 66.27% | 68.33% | 70.69% | 66.89% | 71.1%
83 | 80.56% | 72.06% | 75.43% | 76.14% | 70.8% | 74.53%
84 | 79.15% | 71.08% | 73.5% | 75.55% | 70.54% | 75.67%
85 | 80.81% | 70.6% | 74.1% | 76.12% | 69.51% | 75.24%
86 | 76.21% | 66.1% | 70.78% | 72.1% | 66.42% | 72.54%
87 | 79.38% | 66.35% | 70.25% | 72.9% | 67.44% | 72.79%
88 | 76.65% | 65.75% | 68.42% | 70.48% | 66.52% | 72.35%
89 | 78.64% | 65.29% | 68.02% | 70.93% | 66.31% | 69.81%
90 | 75.12% | 64.33% | 67.96% | 69.74% | 66.44% | 70.18%
91 | 75.06% | 64.73% | 67.7% | 68.27% | 65.04% | 70.38%
92 | 74.74% | 64.06% | 67.6% | 70.05% | 64.94% | 70.56%
93 | 73.66% | 63.27% | 66.21% | 68.96% | 64.11% | 68.33%
94 | 74.52% | 62.62% | 65.2% | 66.82% | 63.18% | 67.69%
95 | 76.64% | 64.05% | 67.62% | 68.8% | 65.96% | 70.59%
96 | 81.01% | 63.62% | 67.48% | 69.79% | 64.73% | 69.23%
97 | 79.93% | 64.06% | 67.96% | 70.32% | 64.94% | 70.16%
98 | 80% | 68.37% | 73.37% | 74.96% | 68.27% | 72.86%
99 | 78.75% | 68.28% | 71.13% | 73.65% | 66.5% | 72.75%
100 | 78.85% | 67.64% | 72.11% | 79.46% | 66.12% | 72.15%
101 | 76.72% | 61.69% | 67.72% | 68.64% | 62.94% | 68.68%
102 | 80.84% | 67.97% | 73.66% | 74.56% | 66.73% | 74.27%
103 | 79.01% | 66.48% | 71.82% | 73.41% | 65.65% | 72.11%
104 | 78.87% | 66.06% | 72.49% | 73.44% | 64.68% | 72.2%
105 | 79.39% | 65.72% | 72.29% | 73.68% | 64.35% | 72.14%
106 | 70.23% | 54.54% | 61.13% | 64.17% | 54.65% | 61.01%
107 | 68.53% | 54.89% | 62.38% | 63.77% | 54.65% | 62.32%
108 | 69.01% | 54.9% | 61.65% | 64.91% | 54.28% | 61.45%
109 | 68.16% | 56.3% | 61.92% | 63.18% | 54.43% | 60.83%
110 | 75.54% | 60.86% | 67.78% | 73.71% | 60.62% | 66.04%
111 | 69.37% | 61.15% | 68.63% | 70.21% | 60.93% | 69.05%
112 | 79.88% | 65.1% | 72.29% | 73.47% | 63.15% | 71.26%
113 | 78.97% | 64.45% | 76.36% | 73.61% | 62.69% | 70.84%
114 | 78.33% | 63.95% | 71.94% | 73.62% | 62.34% | 70.62%
115 | 68.17% | 63.54% | 70.08% | 72.26% | 62.14% | 69.83%
116 | 76.6% | 59.15% | 65.61% | 67.66% | 58.8% | 66.09%
117 | 75.47% | 59.81% | 66.63% | 68.64% | 60.52% | 66.87%
118 | 73.91% | 60.71% | 67.92% | 69.11% | 61.08% | 69.03%
119 | 82.84% | 66.85% | 74.76% | 75.53% | 66.3% | 73.48%
120 | 75.91% | 61.36% | 68.55% | 71.07% | 61.78% | 69.65%
121 | 75.43% | 60.52% | 67.62% | 69.47% | 61.33% | 68.51%
122 | 75.42% | 60.24% | 66.51% | 69.72% | 60.63% | 67.85%
123 | 64.24% | 59.69% | 67.57% | 69.72% | 60.6% | 68.45%
124 | 74.86% | 59.47% | 66.53% | 68.83% | 60.02% | 66.94%
125 | 74.27% | 59.13% | 66.39% | 68.91% | 59.31% | 66.02%
126 | 75.35% | 59.5% | 67.14% | 69.75% | 60.2% | 66.31%
127 | 75.79% | 59.93% | 68.01% | 68.51% | 59.87% | 66.78%
128 | 73.61% | 58.74% | 66.43% | 68.29% | 58.67% | 65.15%
129 | 75.47% | 60.31% | 68.86% | 69.07% | 59.2% | 66.46%
130 | 78.66% | 60.28% | 67.57% | 68.82% | 58.69% | 65.99%
131 | 73.11% | 59.55% | 66.73% | 67.87% | 58.29% | 65.02%
132 | 73.05% | 59.32% | 67.15% | 66.97% | 58.04% | 66.05%
133 | 71.84% | 55.53% | 63.03% | 64.79% | 55.81% | 62.55%
134 | 74.38% | 57.44% | 64.92% | 66.46% | 57.54% | 64.95%
135 | 81.75% | 59.94% | 66.66% | 69.27% | 59.93% | 67.66%
136 | 73.51% | 58.22% | 65.48% | 67.98% | 58.67% | 66.76%
137 | 71.71% | 55.53% | 62.58% | 71.65% | 55.64% | 60.92%
138 | 73.09% | 57.6% | 62.86% | 64.85% | 58.08% | 63.5%
139 | 71.5% | 58.55% | 63.61% | 64.86% | 58.9% | 65.68%
140 | 71.84% | 59% | 64% | 65.73% | 59.23% | 65.02%
141 | 71.79% | 58.9% | 63.23% | 72.67% | 59.7% | 64.51%
142 | 70.77% | 57.71% | 62.34% | 64.11% | 58.35% | 63.68%
143 | 70.8% | 57.71% | 62.43% | 65.85% | 58.21% | 64.51%
144 | 73.42% | 57.45% | 60.98% | 63.53% | 57.64% | 63.45%
145 | 83.04% | 62.37% | 68.15% | 69.16% | 61.39% | 67.07%
146 | 73.82% | 61.85% | 67.65% | 70.31% | 60.25% | 67.06%
147 | 77.68% | 60.61% | 66.43% | 68.2% | 60.4% | 66.76%
148 | 71.49% | 55.88% | 62.55% | 61.51% | 56.6% | 63.05%
149 | 72.51% | 56.35% | 62.72% | 63.99% | 56.37% | 63.42%
150 | 73.59% | 56.86% | 63.74% | 65.15% | 57.57% | 63.49%
151 | 75.33% | 57.22% | 63.82% | 65.71% | 58.15% | 65.02%
152 | 69.09% | 55.72% | 61.47% | 63.14% | 55.86% | 62.48%
153 | 64.34% | 54.86% | 61.69% | 60.62% | 56.07% | 61.74%
154 | 69.23% | 53.83% | 59.46% | 60.99% | 55% | 60.05%
155 | 69.95% | 54.34% | 59.62% | 61.31% | 55.48% | 58.36%
156 | 69.92% | 55.76% | 60.72% | 62.6% | 56.63% | 62.25%
157 | 69.62% | 55.42% | 60.65% | 62.32% | 56.18% | 61.59%
158 | 69.08% | 52.96% | 60.85% | 63.27% | 55.44% | 61.37%
159 | 68.79% | 54.69% | 54.14% | 62.08% | 55.7% | 61.54%
160 | 70.18% | 54.25% | 61.67% | 63.43% | 55.29% | 60.18%
161 | 71.4% | 54.25% | 60.34% | 62.5% | 54.7% | 60.65%
162 | 73.86% | 57.76% | 65.06% | 66.13% | 56.79% | 64.46%
163 | 74.08% | 57.24% | 64.76% | 65.52% | 56.33% | 63.18%
164 | 70.53% | 56.91% | 64.34% | 65.48% | 55.44% | 62.27%
165 | 68.59% | 56.76% | 64.06% | 65.48% | 56.04% | 62.69%
166 | 68.97% | 52.07% | 59.53% | 60.56% | 52.49% | 59.64%
167 | 70.01% | 53.02% | 60.58% | 61.85% | 53.53% | 59.43%
168 | 67.48% | 52.18% | 59.74% | 61.58% | 53.35% | 60.27%
169 | 67.05% | 51.69% | 59.01% | 71.43% | 52.47% | 58.92%
170 | 67.15% | 51.76% | 58.55% | 60.23% | 52.18% | 58.42%
171 | 69.17% | 52.44% | 58.93% | 60.53% | 53.04% | 58.71%
172 | 67.7% | 53.32% | 59.25% | 61.12% | 53.65% | 59.68%
173 | 66.32% | 53% | 59.97% | 60.89% | 53.89% | 60.11%
174 | 66.64% | 52.13% | 58.84% | 60.16% | 53.04% | 58.63%
175 | 66.7% | 52.1% | 58.41% | 60.65% | 52.45% | 57.81%
176 | 67.22% | 53.95% | 58.4% | 60% | 52.6% | 58.05%
177 | 66.7% | 51.95% | 58.12% | 59.96% | 51.98% | 58.7%
178 | 74.46% | 55.99% | 63.74% | 64.78% | 55.15% | 61.9%
179 | 70.72% | 55.68% | 62.84% | 64.97% | 54.29% | 61.61%
180 | 70.87% | 54.75% | 61.81% | 63.28% | 53.75% | 61.49%
181 | 72.14% | 54.29% | 61.78% | 63.44% | 53.57% | 61.09%
182 | 67.11% | 50.06% | 57.27% | 59.5% | 51.03% | 56.9%
183 | 69.45% | 50.89% | 59.35% | 56.14% | 51.01% | 57.72%
184 | 69.18% | 51.46% | 59.22% | 61.5% | 52.07% | 58.33%
185 | 69.38% | 51.94% | 60.3% | 61.29% | 52.48% | 59.72%
186 | 65.07% | 50.15% | 56.87% | 58.89% | 50.96% | 55.35%
187 | 64.78% | 49.73% | 56.66% | 58.3% | 50.21% | 56.33%
188 | 65.14% | 48.71% | 55.52% | 56.7% | 48.89% | 54.13%
189 | 64.76% | 49.64% | 56.32% | 57.22% | 50.47% | 56.17%
190 | 65.56% | 50.56% | 57.25% | 58.82% | 51.24% | 57.38%
191 | 65.38% | 49.86% | 56.62% | 58.36% | 50.2% | 56.63%
192 | 65.22% | 50.19% | 56.81% | 60.84% | 50.58% | 55.87%
193 | 65.38% | 49.5% | 56.53% | 65.53% | 49.67% | 56.39%
194 | 64.29% | 49.72% | 56.69% | 60.77% | 50.14% | 56.3%
195 | 64.63% | 48.61% | 55.97% | 57.11% | 48.85% | 55.36%
196 | 70.13% | 52.47% | 61.14% | 61.84% | 51.43% | 57.98%
197 | 63.11% | 52.16% | 61.39% | 61.27% | 50.77% | 58.98%
198 | 66.38% | 51.53% | 58.78% | 60.64% | 50.55% | 57.11%
199 | 65.35% | 50.63% | 58.33% | 59.72% | 50.47% | 57.66%
200 | 64.18% | 47.35% | 55.14% | 56.66% | 47.89% | 54.53%
201 | 69.06% | 48.3% | 56.36% | 57.59% | 48.85% | 55.02%
202 | 67.97% | 46.53% | 59.49% | 61.32% | 51.47% | 59.04%
203 | 67.01% | 50.73% | 58.38% | 60.81% | 50.53% | 58.16%
204 | 67.81% | 51.03% | 58.8% | 60.92% | 49.91% | 57.45%
205 | 66.93% | 50.83% | 59.3% | 60.3% | 49.92% | 57.49%
206 | 66.76% | 50.48% | 61.5% | 60.16% | 49.67% | 57.39%
207 | 66.42% | 50.65% | 58.92% | 59.96% | 49.93% | 57.77%
208 | 65.81% | 49.71% | 57.63% | 59.33% | 48.71% | 56.87%
209 | 65.66% | 49.31% | 58.17% | 59.69% | 48.62% | 56.85%
210 | 65.93% | 49.46% | 57.44% | 59.09% | 48.35% | 56.54%
211 | 65.07% | 49.16% | 57.68% | 59.72% | 48.69% | 56%
212 | 61.76% | 48.88% | 56.7% | 58.42% | 48.03% | 55.26%
213 | 67.39% | 49.01% | 57.02% | 58.57% | 47.88% | 56.03%
214 | 63.74% | 48.27% | 55.93% | 57.54% | 47.68% | 54.99%
215 | 64.21% | 48.39% | 55.3% | 57.93% | 47.52% | 55.42%
216 | 61.95% | 44.92% | 53.11% | 54.51% | 45.72% | 52.89%
217 | 62.44% | 45.73% | 53.05% | 55.54% | 46.1% | 54.38%
218 | 64.34% | 46.69% | 54.84% | 56.92% | 47.11% | 54.4%
219 | 64.84% | 47.31% | 54.99% | 58.39% | 48.17% | 55.62%
220 | 66.85% | 51.14% | 60.21% | 63.36% | 50.31% | 58.43%
221 | 67.86% | 51.12% | 59.47% | 61.18% | 49.74% | 58.44%
222 | 60.27% | 50.22% | 58.86% | 60.28% | 49.53% | 58.14%
223 | 64.43% | 49.76% | 58.4% | 67.1% | 48.91% | 57.3%
224 | 65.9% | 48.82% | 57.57% | 59.14% | 48.25% | 56.7%
225 | 65.64% | 49.11% | 57.07% | 59.01% | 47.6% | 57.08%
226 | 62.84% | 45.61% | 54.22% | 56.11% | 46.34% | 53.1%
227 | 64.46% | 47.19% | 54.88% | 57.1% | 47.12% | 54.81%
228 | 63.85% | 48.18% | 55.96% | 52.41% | 48.54% | 55.94%
229 | 68.61% | 47.39% | 55.16% | 57.05% | 47.94% | 55.29%
230 | 67.84% | 49.94% | 58.18% | 59.6% | 48.74% | 56.74%
231 | 67.3% | 49.73% | 58.1% | 59.9% | 49.02% | 57.48%
232 | 64.74% | 48.75% | 57.11% | 58.92% | 48.03% | 56.74%
233 | 64.34% | 48.42% | 56.66% | 58.22% | 48.19% | 56.4%
234 | 63.41% | 45.29% | 53.25% | 55.18% | 45.69% | 53.64%
235

## QC
QC Standards
QC STD 1
|  | Sample Id | R | Acquisition Time | Na KED 23/23 | Mg KED 24/24 | Al KED 27/27 | P Oshift 31/47 | S (34) -1 34/34 | K KED 39/39 | Ca KED 43/43 | V KED 51/51 | Cr (52) 52/52 | Cr (53) NH3 53/53 | Fe (54) NH3 54/54 | Mn Oshift 55/71 | Fe (56) KED 56/56 | Co KED 59/59 | Ni KED 60/60 | Cu (63) KED 63/63 | Cu (65) NH3 65/65 | Zn NH3 66/66 | As 75/75 | As KED 75/75 | As Oshift 75/91 | Se (77) 77/77 | Se (77) Oshift 77/93 | Se (78) Oshift 78/94 | Se (80) Oshift 80/96 | Se (82) 82/82 | Se (82) Oshift 82/98 | Mo (95) NH3 95/95 | Mo (95) OShift 95/127 | Rh (IS1) 103/103 (IS) | Rh (IS2) 103/103 (IS) | Rh (IS4) 103/103 (IS) | Rh (IS5) 103/103 (IS) | Rh (IS6) 103/103 (IS) | Rh (IS7) 103/103 (IS) | Cd (111) NH3 111/111 | Cd (114) NH3 114/114 | Pb 208/208 | Pb NH3 208/208
| 37 | LBLANK |  | 06/16/2023 15:53:13 | 54.4446 | 0.23633 | -0.96263 | -0.87372 | 236.4523 | 59.288 | 40.4154 | 2.2133e-05 | -0.011409 | -0.0013358 | -0.57434 | 0.00021345 | -0.40485 | -0.0076191 | 0.00016619 | -0.010467 | -0.011086 | -0.22461 | 0.09778 | 0.035163 | 0.038651 | -0.021279 | 0.0055366 | 0.0076658 | -0.0066793 | -0.0046485 | -0.021474 | -0.0033911 | 0.0030236 | [6 empty] | 0.00060382 | -0.0025156 | -0.010272 | -0.0044217
|  | RSD |  |  | 0.6004% | 1.931e-06% | 4.527% | 0.7008% | 94.93% | 3.195% | 2.811% | 4200% | 18.58% | 118.4% | 1.636% | 401.1% | 8.929% | 3.062% | 1372% | 28.11% | 4.851% | 1.874% | 36.62% | 42.15% | 8.227% | 174% | 0% | 36.62% | 20.37% | 412.8% | 0% | 32.69% | 8.984% | [6 empty] | 28.56% | 8.648% | 4.425% | 5.954%
|  | % Recovery | [31 empty] | 84% | 78.59% | 81.52% | 83.52% | 77.53%
| 52 | LBLANK |  | 06/16/2023 17:23:52 | 55.8376 | 0.70899 | -1.0141 | -0.18597 | 157.8543 | 64.6363 | 41.096 | 0.0010043 | -0.0086365 | -0.0021681 | -0.58445 | -0.00097667 | -0.38029 | -0.007202 | 0.0047496 | -0.0094054 | -0.010432 | -0.22261 | -0.020689 | -0.0078919 | -0.00011531 | 0.21802 | 0.0055366 | 0.0028999 | -0.0080639 | -0.014407 | -0.021474 | 0.017648 | 0.0020502 | [6 empty] | 0.00050427 | -0.0021097 | -0.010416 | -0.004357
|  | RSD |  |  | 1.183% | 115.5% | 3.344% | 12.97% | 132.1% | 0.5142% | 4.855% | 95.17% | 35.61% | 69.45% | 1.623% | 36.43% | 24.41% | 8.682% | 55.23% | 32.75% | 12% | 1.693% | 331.9% | 75.9% | 523.3% | 62.32% | 0% | 101.8% | 33.66% | 53.53% | 0% | 10.92% | 13.74% | [6 empty] | 0% | 8.047% | 2.579% | 1.69%
|  | % Recovery | [31 empty] | 74.67% | 75.09% | 82.71% | 82.32% | 74.77%
| 68 | LBLANK |  | 06/16/2023 19:00:48 | 54.1008 | 0.74518 | -0.99077 | -0.58503 | 373.8525 | 57.4128 | 39.1036 | 0.0011378 | -0.0072258 | -0.0012028 | -0.58948 | -0.00056814 | -0.25972 | -0.0077432 | 0.0035129 | -0.00485 | -0.010419 | -0.22553 | 0.077648 | -0.01135 | -0.0034741 | -0.026405 | 0.0055366 | 0.0013136 | -0.011205 | 0.0074769 | -0.0022024 | -0.0050221 | 0.0024088 | [6 empty] | 0.00061183 | -0.0020735 | -0.010642 | -0.0044551
|  | RSD |  |  | 0.6792% | 118.3% | 4.104% | 2.456% | 26.97% | 9.697% | 0% | 88.99% | 12.2% | 184.3% | 0.9397% | 16.49% | 18.62% | 3.249% | 22.69% | 167.8% | 14.22% | 1.155% | 52.34% | 1.777e-06% | 38.54% | 632.6% | 0% | 240.3% | 13.22% | 611.8% | 757.8% | 14.66% | 26.87% | [6 empty] | 30.45% | 6.327% | 1.206% | 3.831%
|  | % Recovery | [31 empty] | 81.81% | 71.35% | 76.1% | 77.27% | 70.63%
| 83 | LBLANK |  | 06/16/2023 20:31:27 | 55.3543 | 0.72782 | -1.0117 | -0.91666 | 447.6718 | 60.514 | 39.8091 | -0.00057673 | -0.0069438 | -0.0012405 | -0.57586 | -0.00052007 | -0.26616 | -0.0078885 | 0.0018702 | -0.011305 | -0.0096492 | -0.22954 | -0.016738 | -0.01135 | -0.0037458 | 0.0040065 | 0.0055366 | 0.0048521 | -0.01206 | -0.013452 | -0.016689 | -0.006875 | 0.0028274 | [6 empty] | 0.00050427 | -0.0023379 | -0.010438 | -0.004459
|  | RSD |  |  | 1.701% | 117% | 3.149% | 1.54% | 26.02% | 7.512% | 3.07% | 101.2% | 21.47% | 104.7% | 0.1188% | 50.99% | 33.47% | 0% | 205% | 23.14% | 15.49% | 2.875% | 390.4% | 1.777e-06% | 11.8% | 4104% | 0% | 111.4% | 0% | 340.4% | 49.66% | 15.05% | 4.126% | [6 empty] | 0% | 5.86% | 5.514% | 4.42%
|  | % Recovery | [31 empty] | 80.56% | 72.06% | 75.43% | 76.14% | 70.8%
| 98 | LBLANK |  | 06/16/2023 22:02:04 | 55.3996 | 2.2875 | -0.90892 | -0.86876 | 710.7681 | 58.875 | 42.0219 | -0.00021605 | -0.006893 | -0.0011939 | -0.57746 | -0.00077284 | -0.32539 | -0.0075884 | 0.0085943 | -0.012158 | -0.010057 | -0.22863 | 0.037665 | -0.0037923 | 0.01697 | 0.0048278 | 0.011408 | 0.0089106 | -0.0076485 | -0.040154 | -0.021474 | -0.0075587 | 0.003048 | [6 empty] | 0.00083952 | -0.0021333 | -0.010903 | -0.0044755
|  | RSD |  |  | 1.77% | 77.66% | 9.079% | 1.213% | 28.78% | 7.243% | 12.03% | 279.7% | 22.14% | 56.57% | 1.343% | 24.06% | 4.925% | 6.85% | 78.24% | 3.788% | 7.414% | 1.42% | 109% | 345.2% | 11.02% | 1911% | 89.14% | 37.09% | 19.87% | 66.26% | 0% | 12.23% | 33.08% | [6 empty] | 40.07% | 24.76% | 0.4127% | 1.166%
|  | % Recovery | [31 empty] | 80% | 68.37% | 73.37% | 74.96% | 68.27%
| 112 | LBLANK |  | 06/16/2023 23:27:00 | 54.9966 | 0.765 | -0.98765 | 0.27807 | 453.1504 | 57.124 | 40.6215 | 0.0015915 | -0.0054195 | -0.0026971 | -0.57853 | -0.00011699 | -0.28595 | -0.00588 | -0.0025338 | -0.012831 | -0.010844 | -0.22593 | 0.014198 | -0.01135 | -0.0014082 | -0.099117 | 0.018063 | 0.0054514 | -0.011141 | -0.0027456 | -0.011061 | 0.015293 | 0.0032846 | [6 empty] | 0.00061655 | -0.0017854 | -0.01059 | -0.004417
|  | RSD |  |  | 2.062% | 119.7% | 5.736% | 6.098% | 10.29% | 4.71% | 6.472% | 155.4% | 102.7% | 51.93% | 0.1276% | 603.3% | 9.155% | 16.63% | 38.83% | 33.96% | 11.36% | 1.59% | 196.2% | 1.777e-06% | 27.04% | 92.07% | 120.1% | 109.6% | 14.3% | 1036% | 163% | 8.19% | 9.743% | [6 empty] | 31.54% | 4.779% | 3.318% | 0.432%
|  | % Recovery | [31 empty] | 79.88% | 65.1% | 72.29% | 73.47% | 63.15%
| 113 | LBLANK |  | 06/16/2023 23:33:18 | 55.6865 | 0.75871 | -0.99821 | -0.70648 | 531.5513 | 57.2547 | 40.6163 | 0.00052196 | -0.0056343 | -0.0020015 | -0.57838 | -0.00023853 | -0.28665 | -0.0072719 | -0.0024492 | -0.013827 | -0.011843 | -0.22239 | 0.054473 | -0.0074625 | -0.0027578 | -0.084283 | 0.0055366 | 0.0034757 | -0.011125 | -0.045514 | -0.021474 | 0.0079127 | 0.0024118 | [6 empty] | 0.00072017 | -0.0025324 | -0.010044 | -0.0041253
|  | RSD |  |  | 3.215% | 119.3% | 0.9354% | 1.228% | 21.88% | 7.053% | 3.226% | 315.4% | 101.4% | 25.67% | 0.4751% | 217.3% | 7.884% | 7.343% | 39.84% | 6.272% | 2.917% | 4.883% | 11.75% | 90.23% | 31.01% | 176.7% | 0% | 198.6% | 14.56% | 42.38% | 0% | 33.8% | 18.04% | [6 empty] | 51.93% | 14.67% | 2.715% | 1.653%
|  | % Recovery | [31 empty] | 78.97% | 64.45% | 76.36% | 73.61% | 62.69%
| 129 | LBLANK |  | 06/17/2023 01:10:52 | 54.3812 | 0.80206 | -0.99466 | -0.53709 | 462.2207 | 59.9633 | 39.8954 | 0.00063728 | -0.0098968 | -0.0023999 | -0.58696 | -0.00071708 | -0.45669 | -0.0077222 | -0.00023988 | -0.010848 | -0.012549 | -0.23468 | 0.0044616 | -0.01135 | -0.0031847 | -0.035661 | 0.0055366 | 0.0037498 | -0.011061 | -0.02219 | -0.015854 | -0.0025302 | 0.0025004 | [6 empty] | 0.00050427 | -0.0021725 | -0.010437 | -0.0044141
|  | RSD |  |  | 1.286% | 122.2% | 2.009% | 2.233% | 20.69% | 7.833% | 3.438% | 279.1% | 62.29% | 49.21% | 0.2181% | 87.19% | 13.57% | 3.73% | 2029% | 9.165% | 3.823% | 1.528% | 1079% | 1.777e-06% | 8.611% | 231.3% | 0% | 196.7% | 15.65% | 158.5% | 61.39% | 72.86% | 5.046% | [6 empty] | 0% | 14.38% | 2.717% | 5.443%
|  | % Recovery | [31 empty] | 75.47% | 60.31% | 68.86% | 69.07% | 59.2%
| 130 | LBLANK |  | 06/17/2023 01:17:12 | 54.8716 | 0.23633 | -0.95262 | -0.80843 | 506.4749 | 61.3376 | 39.9056 | -0.00052778 | -0.0079865 | -0.0020988 | -0.58614 | -0.00064534 | -0.34749 | -0.007392 | 0.0018527 | -0.011356 | -0.011045 | -0.22706 | 0.066471 | -0.01135 | -0.0036289 | -0.21459 | 0.0055366 | 0.0016159 | -0.0060409 | -0.0019992 | -0.021474 | -0.0033717 | 0.002501 | [6 empty] | 0.00062566 | -0.0023092 | -0.01067 | -0.0043961
|  | RSD |  |  | 0.7685% | 1.931e-06% | 2.921% | 1.922% | 20.84% | 3.166% | 3.481% | 126.7% | 55.43% | 37.03% | 0.8333% | 115.8% | 18.18% | 6.733% | 382.7% | 25.43% | 5.618% | 1.297% | 75.05% | 1.777e-06% | 14.23% | 74.65% | 0% | 227.8% | 50.35% | 4464% | 0% | 17.54% | 4.898% | [6 empty] | 33.61% | 3.816% | 0.3946% | 1.733%
|  | % Recovery | [31 empty] | 78.66% | 60.28% | 67.57% | 68.82% | 58.69%
| 145 | LBLANK |  | 06/17/2023 02:49:20 | 54.1513 | 0.23633 | -0.98894 | -0.12799 | 465.4333 | 59.8055 | 41.4958 | -0.00014348 | -0.0075797 | -0.00058329 | -0.5883 | -0.00033068 | -0.23142 | -0.0072372 | 0.0046778 | -0.0084419 | -0.010987 | -0.22901 | 0.055995 | -0.0032668 | -0.0029648 | -0.040557 | 0.0055366 | 0.0056923 | -0.010118 | 0.0074721 | -0.015962 | -0.0057121 | 0.0044433 | [6 empty] | 0.00050427 | -0.0025665 | -0.010739 | -0.0043062
|  | RSD |  |  | 1.518% | 1.931e-06% | 3.764% | 15.78% | 27.38% | 2.961% | 0.1102% | 929.9% | 9.466% | 136.6% | 1.026% | 95.71% | 13.06% | 3.702% | 157.4% | 65.56% | 6.247% | 2.438% | 82.25% | 214.4% | 28.91% | 344.1% | 0% | 108.8% | 16.62% | 101.4% | 59.8% | 7.99% | 16.41% | [6 empty] | 0% | 10.26% | 1.533% | 1.823%
|  | % Recovery | [31 empty] | 83.04% | 62.37% | 68.15% | 69.16% | 61.39%
| 162 | LBLANK |  | 06/17/2023 04:33:51 | 55.1512 | 0.82977 | -0.99851 | -0.9124 | 474.7171 | 66.3204 | 43.2566 | -0.00091385 | -0.0086804 | -0.00088664 | -0.58225 | -0.00014399 | -0.37868 | -0.0077199 | 0.011284 | -0.010375 | -0.0073362 | -0.23426 | 0.0091799 | -0.01135 | -0.0037193 | -0.079842 | 0.01249 | 0.0083662 | -0.007874 | 0.021953 | -0.021474 | -0.0057551 | 0.0029491 | [6 empty] | 0.00063013 | -0.0023352 | -0.010447 | -0.0041987
|  | RSD |  |  | 1.39% | 123.9% | 1.948% | 0.4317% | 5.774% | 6.455% | 8.707% | 0% | 25.24% | 60.13% | 1.631% | 629.6% | 12.05% | 3.782% | 111.6% | 37.04% | 64.86% | 1.948% | 580.5% | 1.777e-06% | 21.66% | 175.3% | 96.43% | 121.1% | 23.31% | 143.4% | 0% | 12.61% | 27.46% | [6 empty] | 34.59% | 15.56% | 2.991% | 4.056%
|  | % Recovery | [31 empty] | 73.86% | 57.76% | 65.06% | 66.13% | 56.79%
| 163 | LBLANK |  | 06/17/2023 04:40:11 | 54.1682 | 0.84263 | -1.0169 | -0.93954 | 495.8656 | 57.3309 | 39.974 | 0.00033378 | -0.0077697 | -0.00062604 | -0.58195 | -0.0011218 | -0.40732 | -0.0077098 | 0.0020553 | -0.010273 | -0.011342 | -0.22944 | 0.032275 | -0.0071572 | -0.0024656 | -0.091799 | 0.0055366 | -0.00050913 | -0.01206 | -0.010603 | -0.021474 | -0.0068265 | 0.0022609 | [6 empty] | 0.00050427 | -0.0025427 | -0.010484 | -0.0040756
|  | RSD |  |  | 0.786% | 124.6% | 3.544% | 1.895% | 13.82% | 7.161% | 3.771% | 647.4% | 44.14% | 68.11% | 0.9672% | 22.79% | 13.93% | 4.015% | 224% | 28.66% | 14.33% | 0.6318% | 14.8% | 101.5% | 46.74% | 52.79% | 0% | 0% | 0% | 177.6% | 0% | 19.57% | 26.49% | [6 empty] | 0% | 6.423% | 0.3439% | 0.7476%
|  | % Recovery | [31 empty] | 74.08% | 57.24% | 64.76% | 65.52% | 56.33%
| 178 | LBLANK |  | 06/17/2023 06:12:21 | 54.6956 | 1.442 | -0.99661 | -0.49387 | 622.9496 | 55.1933 | 41.6973 | 0.00031926 | -0.0091733 | -0.0024377 | -0.58138 | -0.00087292 | -0.30531 | -0.0078885 | 0.0037548 | -0.009191 | -0.011487 | -0.22872 | 0.022123 | 0.0062932 | 0.010213 | -0.063336 | 0.0055366 | 0.0017967 | -0.0056003 | 0.014232 | -0.015347 | -0.0066375 | 0.0044315 | [6 empty] | 0.0008888 | -0.0025361 | -0.010858 | -0.0043132
|  | RSD |  |  | 1.613% | 72.42% | 1.868% | 5.137% | 29.2% | 8.53% | 6.194% | 386.5% | 26.65% | 25.15% | 1.426% | 30.56% | 19.16% | 0% | 115.2% | 44.67% | 2.706% | 5.028% | 353.5% | 242.8% | 29.62% | 119.8% | 0% | 222.3% | 57.31% | 256% | 69.15% | 15.59% | 14.24% | [6 empty] | 43.28% | 11.01% | 2.676% | 5.131%
|  | % Recovery | [31 empty] | 74.46% | 55.99% | 63.74% | 64.78% | 55.15%
| 179 | LBLANK |  | 06/17/2023 06:18:40 | 55.203 | 0.83572 | -0.90719 | -0.87994 | 515.7658 | 57.3635 | 40.8248 | 0.00031991 | -0.0091201 | -0.001297 | -0.57984 | -0.00024085 | -0.31988 | -0.0071727 | -0.00096481 | -0.013315 | -0.01122 | -0.22905 | -0.046507 | -0.0068514 | 0.0038727 | 0.1286 | 0.0055366 | 0.0018063 | -0.0033927 | -0.011958 | -0.021474 | -0.00815 | 0.0030957 | [6 empty] | 0.00063363 | -0.0025293 | -0.010509 | -0.0042858
|  | RSD |  |  | 1.359% | 124.2% | 4.624% | 2.844% | 44.05% | 10.63% | 3.651% | 668% | 24.49% | 47.67% | 0.7441% | 147.6% | 9.754% | 8.644% | 538.6% | 7.903% | 4.022% | 0.7551% | 62% | 113.7% | 62.49% | 44.79% | 0% | 222% | 146% | 200.1% | 0% | 11.78% | 16.24% | [6 empty] | 35.36% | 6.4% | 2.991% | 9.36%
|  | % Recovery | [31 empty] | 70.72% | 55.68% | 62.84% | 64.97% | 54.29%
| 196 | LBLANK |  | 06/17/2023 08:03:10 | 55.035 | 0.87745 | -1.013 | -0.87717 | 568.7107 | 58.9971 | 39.1036 | -0.00047409 | -0.0055501 | -0.0021796 | -0.57735 | -0.00023199 | -0.31439 | -0.0078885 | 0.00083711 | -0.0086675 | -0.011112 | -0.23141 | 0.014612 | -0.0066593 | 0.0085284 | 0.022527 | 0.0055366 | 0.011718 | -0.0074358 | -0.0091548 | -0.0084538 | -0.0067453 | 0.0027929 | [6 empty] | 0.00050427 | -0.0024531 | -0.010606 | -0.0041835
|  | RSD |  |  | 0.8609% | 126.6% | 3.443% | 1.569% | 18.93% | 4.935% | 0% | 160.7% | 56.89% | 144.7% | 1.319% | 187.1% | 9.706% | 0% | 405.6% | 0.5017% | 3.972% | 2.451% | 279.2% | 122% | 39.45% | 419.5% | 0% | 95.18% | 53.86% | 217.8% | 133.4% | 11.96% | 27.73% | [6 empty] | 0% | 22.86% | 1.106% | 5.17%
|  | % Recovery | [31 empty] | 70.13% | 52.47% | 61.14% | 61.84% | 51.43%
| 197 | LBLANK |  | 06/17/2023 08:09:28 | 53.943 | 0.23633 | -0.99374 | -0.94322 | 454.9155 | 62.6287 | 42.7364 | -0.00048271 | -0.007886 | -0.0021642 | -0.58084 | -0.00022223 | -0.37327 | -0.0071464 | 0.0083564 | -0.012219 | -0.011096 | -0.23312 | -0.067417 | -0.0021395 | 0.0089333 | 0.23586 | 0.013317 | 0.0068941 | -0.0062653 | -0.01877 | -0.014927 | -0.0084082 | 0.0025485 | [6 empty] | 0.0010413 | -0.0022897 | -0.010089 | -0.0042789
|  | RSD |  |  | 0.8552% | 1.931e-06% | 5.271% | 1.309% | 14.11% | 8.457% | 3.737% | 154.7% | 54.19% | 34.67% | 1.609% | 228.4% | 15.55% | 8.993% | 103.7% | 23.16% | 12.49% | 0.6659% | 238.9% | 372.8% | 19.94% | 183.4% | 101.2% | 107.2% | 64.41% | 92.12% | 75.97% | 5.203% | 5.607% | [6 empty] | 23.9% | 9.193% | 4.618% | 5.103%
|  | % Recovery | [31 empty] | 63.11% | 52.16% | 61.39% | 61.27% | 50.77%
| 212 | LBLANK |  | 06/17/2023 09:41:39 | 53.9622 | 0.23633 | -1.0181 | 0.27629 | 380.1367 | 56.2211 | 41.0345 | 4.3949e-06 | -0.0085363 | -0.0022115 | -0.58825 | -0.00039326 | -0.3623 | -0.0072896 | -0.0088998 | -0.010558 | -0.010593 | -0.23139 | -0.014607 | -0.0063945 | -0.0010943 | -0.06612 | 0.0055366 | 0.004779 | -0.0095937 | 0.0057133 | -0.014463 | -0.0078214 | 0.0022834 | [6 empty] | 0.00079341 | -0.0025369 | -0.010532 | -0.0040134
|  | RSD |  |  | 0.9164% | 1.931e-06% | 4.299% | 13.2% | 44.04% | 15.46% | 4.075% | 36189% | 38.64% | 94.07% | 1.809% | 177.5% | 7.487% | 14.23% | 52.38% | 19.8% | 11.12% | 3.391% | 331.5% | 134.2% | 102.5% | 301.3% | 0% | 191.7% | 22.27% | 1278% | 83.95% | 8.306% | 18.41% | [6 empty] | 31.56% | 8.155% | 1.746% | 7.856%
|  | % Recovery | [31 empty] | 61.76% | 48.88% | 56.7% | 58.42% | 48.03%
| 213 | LBLANK |  | 06/17/2023 09:47:59 | 53.9519 | 0.23633 | -0.99086 | -0.76234 | 211.4654 | 63.8271 | 40.0575 | -0.00091385 | -0.0061699 | -0.00078087 | -0.58438 | -0.00083815 | -0.39507 | -0.0074963 | -0.0044315 | -0.013919 | -0.010237 | -0.2302 | 0.00056601 | -0.01135 | -0.0027413 | -0.032854 | 0.013787 | -0.00050913 | -0.01206 | -0.11708 | -0.021474 | -0.0086093 | 0.0023704 | [6 empty] | 0.00064791 | -0.0022976 | -0.010158 | -0.0039429
|  | RSD |  |  | 0.8834% | 1.931e-06% | 1.977% | 1.598% | 68.46% | 9.528% | 4.125% | 0% | 58.56% | 175.7% | 0.6676% | 66.62% | 27.96% | 4.534% | 152.9% | 14.46% | 14.74% | 1.917% | 7806% | 1.777e-06% | 64.16% | 458.9% | 103.6% | 0% | 0% | 34.56% | 0% | 7.403% | 6.748% | [6 empty] | 38.4% | 7.958% | 4.96% | 4.296%
|  | % Recovery | [31 empty] | 67.39% | 49.01% | 57.02% | 58.57% | 47.88%
| 230 | LBLANK |  | 06/17/2023 11:32:26 | 54.5088 | 0.8925 | -0.97753 | 0.44344 | 471.7096 | 64.0202 | 40.0401 | 0.00089021 | -0.0052172 | -0.0014034 | -0.57981 | -0.0005055 | -0.15477 | -0.0075028 | 0.0056466 | -0.010101 | -0.010062 | -0.23466 | -0.011186 | -0.01135 | -0.00099884 | 0.060126 | 0.0055366 | 0.0046556 | -0.0060193 | 0.0016578 | -0.01462 | 0.0021293 | 0.0029673 | [6 empty] | 0.00064673 | -0.0020761 | -0.010465 | -0.0040721
|  | RSD |  |  | 0.009801% | 127.3% | 2.49% | 11.08% | 27.35% | 0.2491% | 4.051% | 89.06% | 20.05% | 89.09% | 1.663% | 150.2% | 65.72% | 4.452% | 126% | 30.29% | 19.02% | 3.757% | 451.8% | 1.777e-06% | 134.9% | 294% | 0% | 96.07% | 34.67% | 2703% | 81.2% | 63.13% | 15.38% | [6 empty] | 38.15% | 9.902% | 2.425% | 8.594%
|  | % Recovery | [31 empty] | 67.84% | 49.94% | 58.18% | 59.6% | 48.74%
| 231 | LBLANK |  | 06/17/2023 11:38:51 | 54.7736 | 0.23633 | -0.97863 | -0.6603 | 461.2309 | 58.709 | 40.0395 | -1.947e-05 | -0.0050509 | -0.0020565 | -0.59531 | -0.00067734 | -0.28885 | -0.0075076 | -0.0011957 | -0.012398 | -0.012075 | -0.23104 | 0.0064449 | -0.01135 | -0.0027775 | 0.11097 | 0.0055366 | -0.00050913 | -0.0072037 | -0.017012 | -0.021474 | -0.0016192 | 0.0022626 | [6 empty] | 0.00050427 | -0.0024889 | -0.010459 | -0.0041141
|  | RSD |  |  | 0.879% | 1.931e-06% | 1.166% | 3.279% | 38.99% | 2.725% | 4.049% | 7957% | 44.4% | 0.2811% | 0.4508% | 38.21% | 25.57% | 4.393% | 212.5% | 20.61% | 4.26% | 2.487% | 406.8% | 1.777e-06% | 22.26% | 61.61% | 0% | 0% | 77.44% | 393.3% | 0% | 107.3% | 6.948% | [6 empty] | 0% | 7.107% | 1.431% | 4.85%
|  | % Recovery | [31 empty] | 67.3% | 49.73% | 58.1% | 59.9% | 49.02%
QC STD 2
|  | Sample Id | R | Acquisition Time | Na KED 23/23 | Mg KED 24/24 | Al KED 27/27 | P Oshift 31/47 | S (34) -1 34/34 | K KED 39/39 | Ca KED 43/43 | V KED 51/51 | Cr (52) 52/52 | Cr (53) NH3 53/53 | Fe (54) NH3 54/54 | Mn Oshift 55/71 | Fe (56) KED 56/56 | Co KED 59/59 | Ni KED 60/60 | Cu (63) KED 63/63 | Cu (65) NH3 65/65 | Zn NH3 66/66 | As 75/75 | As KED 75/75 | As Oshift 75/91 | Se (77) 77/77 | Se (77) Oshift 77/93 | Se (78) Oshift 78/94 | Se (80) Oshift 80/96 | Se (82) 82/82 | Se (82) Oshift 82/98 | Mo (95) NH3 95/95 | Mo (95) OShift 95/127 | Rh (IS1) 103/103 (IS) | Rh (IS2) 103/103 (IS) | Rh (IS4) 103/103 (IS) | Rh (IS5) 103/103 (IS) | Rh (IS6) 103/103 (IS) | Rh (IS7) 103/103 (IS) | Cd (111) NH3 111/111 | Cd (114) NH3 114/114 | Pb 208/208 | Pb NH3 208/208
| 18 | ICV PE |  | 06/16/2023 13:47:50 | 10056.5045 | 10399.9312 | 9.8698 | 18.7224 | 9864.5338 | 9803.376 | 9913.8204 | 10.0101 | 10.5126 | 10.1321 | 10.071 | 10.2031 | 9.9179 | 10.0963 | 9.7904 | 10.0569 | 10.0276 | 9.713 | 9.7284 | 9.0755 | 10.0024 | 6.9848 | 9.9015 | 10.1402 | 10.1402 | 9.8368 | 9.9585 | 9.5255 | 9.5507 | [6 empty] | 9.5729 | 9.6206 | 9.2728 | 9.5202
|  | RSD |  |  | 1.54% | 0.4419% | 2.368% | 1.088% | 1.033% | 1.39% | 1.257% | 2.059% | 0.5718% | 1.794% | 1.486% | 0.481% | 1.843% | 0.7227% | 1.62% | 0.3411% | 0.8926% | 0.6228% | 1.085% | 3.691% | 1.317% | 3.432% | 4.596% | 0.6279% | 1.94% | 1.243% | 3.82% | 0.6342% | 1.031% | [6 empty] | 1.364% | 1.291% | 0.3706% | 1.502%
|  | % Recovery |  |  | 100.6% | 104% | 98.7% | 93.61% | 98.65% | 98.03% | 99.14% | 100.1% | 105.1% | 101.3% | 100.7% | 102% | 99.18% | 101% | 97.9% | 100.6% | 100.3% | 97.13% | 97.28% | 90.76% | 100% | 69.85% | 99.02% | 101.4% | 101.4% | 98.37% | 99.58% | 95.25% | 95.51% | 90.84% | 84.9% | 86.62% | 88.49% | 84.45% |  | 95.73% | 96.21% | 92.73% | 95.2%
QC STD 3
|  | Sample Id | R | Acquisition Time | Na KED 23/23 | Mg KED 24/24 | Al KED 27/27 | P Oshift 31/47 | S (34) -1 34/34 | K KED 39/39 | Ca KED 43/43 | V KED 51/51 | Cr (52) 52/52 | Cr (53) NH3 53/53 | Fe (54) NH3 54/54 | Mn Oshift 55/71 | Fe (56) KED 56/56 | Co KED 59/59 | Ni KED 60/60 | Cu (63) KED 63/63 | Cu (65) NH3 65/65 | Zn NH3 66/66 | As 75/75 | As KED 75/75 | As Oshift 75/91 | Se (77) 77/77 | Se (77) Oshift 77/93 | Se (78) Oshift 78/94 | Se (80) Oshift 80/96 | Se (82) 82/82 | Se (82) Oshift 82/98 | Mo (95) NH3 95/95 | Mo (95) OShift 95/127 | Rh (IS1) 103/103 (IS) | Rh (IS2) 103/103 (IS) | Rh (IS4) 103/103 (IS) | Rh (IS5) 103/103 (IS) | Rh (IS6) 103/103 (IS) | Rh (IS7) 103/103 (IS) | Cd (111) NH3 111/111 | Cd (114) NH3 114/114 | Pb 208/208 | Pb NH3 208/208
| 19 | CCV 1ppb TRACE |  | 06/16/2023 13:53:50 | 55.6556 | 0.91518 | -1.0076 | -0.99296 | -2402.6525 | 50.0019 | 40.5457 | 7.111e-06 | -0.013837 | -0.0030719 | -0.6073 | -0.0012092 | -2.2335 | -0.00661 | -0.0051573 | -0.01323 | -0.01183 | -0.21218 | 0.71528 | -0.0089217 | -0.004222 | -1.9308 | 0.0055366 | 0.0096123 | -0.0090549 | -0.090629 | -0.021474 | 0.035588 | 0.0014145 | [6 empty] | 0.00074945 | -0.0022094 | -0.010929 | -0.0047576
|  | RSD |  |  | 0.4014% | 128.5% | 2.185% | 0.6069% | 1.243% | 8.21% | 0.04259% | 11218% | 10.2% | 25.53% | 0.6791% | 11.59% | 5.018% | 9.173% | 48.41% | 6.57% | 5.336% | 3.579% | 2.663% | 47.15% | 9.991% | 2.341% | 0% | 58.73% | 23.41% | 42.25% | 0% | 4.473% | 19.67% | [6 empty] | 33.2% | 11.78% | 2.257% | 3.397%
|  | % Recovery |  |  | 5.566% | 0.09152% | -100.8% | -99.3% | -240.3% | 5% | 4.055% | 0.0007111% | -1.384% | -0.3072% | -60.73% | -0.1209% | -223.3% | -0.661% | -0.5157% | -1.323% | -1.183% | -21.22% | 71.53% | -0.8922% | -0.4222% | -193.1% | 0.5537% | 0.9612% | -0.9055% | -9.063% | -2.147% | 3.559% | 0.1414% | 102.1% | 101% | 101.5% | 103.3% | 100.1% |  | 0.07495% | -0.2209% | -1.093% | -0.4758%
| 20 | CCV 1ppb TRACE |  | 06/16/2023 14:00:07 | 1051.1896 | 1076.671 | 1.2775 | 0.14651 | 1270.6449 | 1104.0053 | 989.4019 | 1.0496 | 1.09 | 1.0467 | 0.8539 | 1.0223 | 1.167 | 1.0494 | 1.0433 | 1.0774 | 1.0596 | 0.82662 | 0.96458 | 0.93718 | 0.9793 | 1.1859 | 0.9275 | 0.99008 | 1.0313 | 0.96014 | 0.9736 | 1.0233 | 0.98635 | [6 empty] | 0.97624 | 0.97117 | 1.0298 | 0.98783
|  | RSD |  |  | 2.402% | 5.783% | 9.168% | 70.55% | 6.41% | 1.564% | 4.403% | 2.875% | 0.8365% | 0.06988% | 5.113% | 1.448% | 2.843% | 0.3966% | 2.618% | 2.672% | 0.8651% | 3.048% | 6.198% | 2.583% | 1.815% | 12.78% | 7.67% | 5.36% | 4.046% | 4.028% | 4.982% | 1.123% | 3.051% | [6 empty] | 2.268% | 2.109% | 2.031% | 0.9436%
|  | % Recovery |  |  | 105.1% | 107.7% | 127.7% | 14.65% | 127.1% | 110.4% | 98.94% | 105% | 109% | 104.7% | 85.39% | 102.2% | 116.7% | 104.9% | 104.3% | 107.7% | 106% | 82.66% | 96.46% | 93.72% | 97.93% | 118.6% | 92.75% | 99.01% | 103.1% | 96.01% | 97.36% | 102.3% | 98.64% | 92.39% | 90.3% | 91.8% | 93.04% | 90.48% |  | 97.62% | 97.12% | 103% | 98.78%
| 38 | CCV 1ppb TRACE |  | 06/16/2023 15:59:13 | 1190.5987 | 1169.6923 | 1.2344 | 0.06955 | 1723.162 | 1147.5669 | 1120.4383 | 1.1169 | 1.1415 | 1.0426 | 0.85164 | 1.0106 | 1.3864 | 1.1093 | 1.1316 | 1.1304 | 1.0666 | 0.84399 | 0.96898 | 1.1141 | 1.0042 | 1.0799 | 0.98937 | 1.0329 | 0.95564 | 1.0227 | 0.99694 | 0.97943 | 1.0019 | [6 empty] | 1.0047 | 0.9778 | 1.0507 | 1.0158
|  | RSD |  |  | 9.244% | 12.12% | 7.756% | 12.76% | 24.58% | 6.452% | 8.567% | 11.03% | 13.44% | 1.388% | 4.959% | 1.547% | 28.22% | 13.75% | 14.99% | 9.754% | 0.7964% | 2.601% | 3.992% | 20.16% | 1.184% | 11.66% | 16.09% | 4.293% | 3.464% | 4.174% | 6.546% | 2.131% | 3.184% | [6 empty] | 1.536% | 1.676% | 1.483% | 0.3473%
|  | % Recovery |  |  | 119.1% | 117% | 123.4% | 6.955% | 172.3% | 114.8% | 112% | 111.7% | 114.1% | 104.3% | 85.16% | 101.1% | 138.6% | 110.9% | 113.2% | 113% | 106.7% | 84.4% | 96.9% | 111.4% | 100.4% | 108% | 98.94% | 103.3% | 95.56% | 102.3% | 99.69% | 97.94% | 100.2% | 83.14% | 78.55% | 79.84% | 81.62% | 78% |  | 100.5% | 97.78% | 105.1% | 101.6%
| 39 | CCV 1ppb TRACE |  | 06/16/2023 16:05:32 | 1057.7035 | 1073.3082 | 1.2211 | 0.097238 | 1427.3953 | 1096.0919 | 1040.0608 | 1.0547 | 1.0688 | 1.0731 | 0.92835 | 1.0313 | 1.2401 | 1.0406 | 1.1001 | 1.044 | 1.068 | 0.83757 | 0.92826 | 1.0796 | 1.0157 | 1.1541 | 0.94276 | 1.0118 | 0.97498 | 0.95912 | 1.1056 | 0.99301 | 1.0103 | [6 empty] | 0.99726 | 0.98164 | 1.0647 | 1.0064
|  | RSD |  |  | 0.9564% | 4.59% | 8.92% | 13.08% | 5.726% | 3.291% | 2.154% | 2.986% | 0.9137% | 1.687% | 7.241% | 0.1876% | 8.576% | 1.994% | 1.884% | 4.795% | 1.384% | 1.335% | 3.378% | 11.06% | 1.27% | 10.28% | 13.57% | 2.937% | 0.3634% | 4.741% | 7.456% | 1.143% | 2.405% | [6 empty] | 3.667% | 0.9035% | 1.452% | 0.9987%
|  | % Recovery |  |  | 105.8% | 107.3% | 122.1% | 9.724% | 142.7% | 109.6% | 104% | 105.5% | 106.9% | 107.3% | 92.84% | 103.1% | 124% | 104.1% | 110% | 104.4% | 106.8% | 83.76% | 92.83% | 108% | 101.6% | 115.4% | 94.28% | 101.2% | 97.5% | 95.91% | 110.6% | 99.3% | 101% | 83.25% | 78.52% | 80.81% | 82.55% | 78.15% |  | 99.73% | 98.16% | 106.5% | 100.6%
| 53 | CCV 1ppb TRACE |  | 06/16/2023 17:29:53 | 987.8331 | 1010.777 | 1.2439 | 0.23021 | 1094.2251 | 1046.5265 | 971.0796 | 1.0204 | 1.0526 | 1.0616 | 0.90075 | 1.0129 | 1.035 | 1.0071 | 0.96643 | 1.0699 | 1.0575 | 0.8381 | 0.96922 | 0.9777 | 0.97344 | 1.046 | 1.1939 | 1.0451 | 0.97312 | 0.98967 | 0.95556 | 1.0014 | 1.0082 | [6 empty] | 0.9601 | 0.97936 | 1.0383 | 1.0069
|  | RSD |  |  | 2.189% | 1.846% | 17.05% | 22.45% | 8.339% | 1.117% | 5.192% | 5.702% | 2.625% | 1.227% | 5.785% | 2.453% | 6.339% | 0.7688% | 10.51% | 2.307% | 1.999% | 2.043% | 0.6475% | 3.817% | 1.091% | 5.197% | 16.56% | 6.575% | 2.721% | 7.868% | 12.5% | 1.022% | 0.4834% | [6 empty] | 1.603% | 1.888% | 1.652% | 1.242%
|  | % Recovery |  |  | 98.78% | 101.1% | 124.4% | 23.02% | 109.4% | 104.7% | 97.11% | 102% | 105.3% | 106.2% | 90.07% | 101.3% | 103.5% | 100.7% | 96.64% | 107% | 105.7% | 83.81% | 96.92% | 97.77% | 97.34% | 104.6% | 119.4% | 104.5% | 97.31% | 98.97% | 95.56% | 100.1% | 100.8% | 82.72% | 75.25% | 78.92% | 79.98% | 73.49% |  | 96.01% | 97.94% | 103.8% | 100.7%
| 54 | CCV 1ppb TRACE |  | 06/16/2023 17:36:12 | 1009.6197 | 1032.865 | 1.0845 | 0.13398 | 1218.9479 | 1035.6432 | 996.5243 | 0.99993 | 1.0281 | 1.0567 | 0.85323 | 1.0119 | 1.2074 | 1.005 | 1.0396 | 1.0711 | 1.0542 | 0.83458 | 0.90204 | 1.0035 | 0.99199 | 1.1499 | 0.89965 | 0.94656 | 0.99771 | 1.0148 | 0.96448 | 0.99645 | 1.0146 | [6 empty] | 0.97104 | 0.98178 | 1.0613 | 1.0075
|  | RSD |  |  | 2.032% | 1.162% | 34.99% | 16.5% | 11.46% | 1.632% | 7.228% | 3.555% | 5.405% | 3.113% | 2.821% | 0.9124% | 5.243% | 2.67% | 4.247% | 1.102% | 2.056% | 2.961% | 0.2777% | 16.81% | 1.457% | 20.36% | 4.645% | 5.793% | 7.429% | 7.346% | 14.4% | 2.296% | 2.98% | [6 empty] | 1.761% | 1.767% | 1.644% | 0.3649%
|  | % Recovery |  |  | 101% | 103.3% | 108.4% | 13.4% | 121.9% | 103.6% | 99.65% | 99.99% | 102.8% | 105.7% | 85.32% | 101.2% | 120.7% | 100.5% | 104% | 107.1% | 105.4% | 83.46% | 90.2% | 100.3% | 99.2% | 115% | 89.97% | 94.66% | 99.77% | 101.5% | 96.45% | 99.64% | 101.5% | 83.29% | 75.52% | 79.03% | 81.67% | 74.72% |  | 97.1% | 98.18% | 106.1% | 100.7%
| 69 | CCV 1ppb TRACE |  | 06/16/2023 19:06:48 | 1002.6883 | 1004.4991 | 1.3914 | 0.16075 | 1337.7963 | 1006.692 | 987.1159 | 1.0297 | 1.0642 | 1.0889 | 0.92501 | 0.99538 | 1.0695 | 1.0281 | 1.0015 | 1.0313 | 1.0631 | 0.85139 | 0.91214 | 1.0631 | 0.9722 | 1.0056 | 1.0267 | 0.98058 | 0.97723 | 0.95232 | 1.0954 | 1.0007 | 1.0144 | [6 empty] | 0.983 | 0.98388 | 1.087 | 1.0288
|  | RSD |  |  | 5.941% | 4.715% | 12.57% | 19.5% | 7.534% | 1.829% | 5.893% | 3.17% | 2.048% | 4.374% | 1.637% | 0.3137% | 6.492% | 4.61% | 2.016% | 1.226% | 2.132% | 1.535% | 5.295% | 13.36% | 0.8629% | 11.5% | 17.95% | 2.058% | 6.593% | 10.53% | 24.01% | 1.853% | 1.619% | [6 empty] | 2.24% | 0.8264% | 1.26% | 1.286%
|  | % Recovery |  |  | 100.3% | 100.4% | 139.1% | 16.07% | 133.8% | 100.7% | 98.71% | 103% | 106.4% | 108.9% | 92.5% | 99.54% | 106.9% | 102.8% | 100.1% | 103.1% | 106.3% | 85.14% | 91.21% | 106.3% | 97.22% | 100.6% | 102.7% | 98.06% | 97.72% | 95.23% | 109.5% | 100.1% | 101.4% | 79.09% | 70.91% | 74.53% | 76.32% | 69.14% |  | 98.3% | 98.39% | 108.7% | 102.9%
| 70 | CCV 1ppb TRACE |  | 06/16/2023 19:13:07 | 1013.9286 | 1074.6779 | 1.2329 | 0.11892 | 1416.6104 | 1024.4187 | 1009.3335 | 1.0007 | 1.0725 | 1.0779 | 0.94758 | 1.0146 | 1.0703 | 1.009 | 0.98826 | 1.0691 | 1.0609 | 0.84908 | 0.96223 | 0.93815 | 0.98826 | 1.0093 | 1.0598 | 0.96533 | 0.97458 | 1.0173 | 1.0478 | 1.0036 | 1.0065 | [6 empty] | 0.97373 | 1.0071 | 1.0945 | 1.0464
|  | RSD |  |  | 2.321% | 3.327% | 1.642% | 43.72% | 6.907% | 4.167% | 2.494% | 2.45% | 3.28% | 3.373% | 2.677% | 1.67% | 5.76% | 1.57% | 4.558% | 1.378% | 0.3724% | 2.059% | 2.452% | 12.25% | 1.414% | 9.303% | 19.12% | 6.749% | 8.405% | 11.85% | 14.62% | 2.291% | 0.941% | [6 empty] | 1.74% | 0.7212% | 0.8518% | 2.101%
|  | % Recovery |  |  | 101.4% | 107.5% | 123.3% | 11.89% | 141.7% | 102.4% | 100.9% | 100.1% | 107.3% | 107.8% | 94.76% | 101.5% | 107% | 100.9% | 98.83% | 106.9% | 106.1% | 84.91% | 96.22% | 93.81% | 98.83% | 100.9% | 106% | 96.53% | 97.46% | 101.7% | 104.8% | 100.4% | 100.6% | 80.14% | 70.96% | 75% | 76.9% | 70.67% |  | 97.37% | 100.7% | 109.5% | 104.6%
| 84 | CCV 1ppb TRACE |  | 06/16/2023 20:37:27 | 1036.9693 | 1078.0255 | 1.2511 | 0.12445 | 1274.8891 | 1063.1539 | 952.9247 | 1.058 | 1.0708 | 1.104 | 0.95774 | 1.0155 | 1.1228 | 1.0253 | 0.97374 | 1.0594 | 1.0654 | 0.85663 | 0.93263 | 1.0259 | 0.98818 | 1.0653 | 0.97813 | 0.96768 | 0.95858 | 0.99066 | 0.90136 | 0.97415 | 0.99372 | [6 empty] | 0.96191 | 0.99069 | 1.0722 | 1.0157
|  | RSD |  |  | 6.14% | 3.165% | 15.63% | 38.32% | 6.88% | 3.622% | 5.239% | 2.323% | 2.712% | 3.73% | 3.595% | 0.1425% | 7.556% | 0.8772% | 2.902% | 1.672% | 0.8743% | 4.562% | 9.709% | 7.038% | 1.523% | 8.726% | 3.234% | 5.331% | 1.788% | 11.69% | 5.065% | 1.08% | 0.6747% | [6 empty] | 1.335% | 2.067% | 1.72% | 1.267%
|  | % Recovery |  |  | 103.7% | 107.8% | 125.1% | 12.45% | 127.5% | 106.3% | 95.29% | 105.8% | 107.1% | 110.4% | 95.77% | 101.5% | 112.3% | 102.5% | 97.37% | 105.9% | 106.5% | 85.66% | 93.26% | 102.6% | 98.82% | 106.5% | 97.81% | 96.77% | 95.86% | 99.07% | 90.14% | 97.42% | 99.37% | 79.15% | 71.08% | 73.5% | 75.55% | 70.54% |  | 96.19% | 99.07% | 107.2% | 101.6%
| 85 | CCV 1ppb TRACE |  | 06/16/2023 20:43:46 | 1040.4587 | 1103.9276 | 1.193 | 0.10895 | 1444.0459 | 1062.3671 | 1072.8531 | 1.0397 | 1.0729 | 1.0873 | 0.96254 | 1.0395 | 1.2168 | 1.029 | 1.0613 | 1.0792 | 1.0569 | 0.84771 | 0.9166 | 0.94519 | 0.9981 | 1.0994 | 1.105 | 0.94652 | 0.94985 | 0.97181 | 0.9856 | 0.98237 | 1.0205 | [6 empty] | 0.97908 | 0.97651 | 1.0438 | 1.0203
|  | RSD |  |  | 1.327% | 2.188% | 11.68% | 31.67% | 4.662% | 2.774% | 7.857% | 5.006% | 2.265% | 2.541% | 0.6777% | 0.8635% | 0.7958% | 0.8717% | 0.4694% | 5.092% | 0.8842% | 1.767% | 7.92% | 5.988% | 2.199% | 20.66% | 6.747% | 10.87% | 2.698% | 16.82% | 10.44% | 1.116% | 1.894% | [6 empty] | 0.5628% | 0.1755% | 3.368% | 0.7512%
|  | % Recovery |  |  | 104% | 110.4% | 119.3% | 10.89% | 144.4% | 106.2% | 107.3% | 104% | 107.3% | 108.7% | 96.25% | 103.9% | 121.7% | 102.9% | 106.1% | 107.9% | 105.7% | 84.77% | 91.66% | 94.52% | 99.81% | 109.9% | 110.5% | 94.65% | 94.98% | 97.18% | 98.56% | 98.24% | 102.1% | 80.81% | 70.6% | 74.1% | 76.12% | 69.51% |  | 97.91% | 97.65% | 104.4% | 102%
| 99 | CCV 1ppb TRACE |  | 06/16/2023 22:08:04 | 1025.6458 | 1070.4949 | 1.142 | 0.089874 | 1758.7854 | 1038.7734 | 1043.2635 | 1.0473 | 1.0654 | 1.1002 | 0.95648 | 1.0125 | 0.99979 | 1.0335 | 1.0321 | 1.0676 | 1.0825 | 0.83832 | 0.94036 | 1.0034 | 0.98286 | 1.1848 | 1.1063 | 0.996 | 0.96327 | 0.93271 | 0.98896 | 0.97105 | 1.0138 | [6 empty] | 0.98231 | 0.96731 | 1.1049 | 1.0381
|  | RSD |  |  | 1.726% | 4.301% | 15.26% | 43.26% | 3.91% | 1.492% | 1.496% | 5.865% | 3.372% | 1.553% | 5.099% | 1.877% | 12.93% | 3.515% | 4.984% | 1.382% | 0.9038% | 0.7363% | 5.823% | 7.787% | 0.4798% | 9.298% | 22.81% | 1.933% | 5.627% | 9.493% | 21.61% | 1.793% | 0.82% | [6 empty] | 1.569% | 0.7933% | 1.766% | 0.5941%
|  | % Recovery |  |  | 102.6% | 107% | 114.2% | 8.987% | 175.9% | 103.9% | 104.3% | 104.7% | 106.5% | 110% | 95.65% | 101.3% | 99.98% | 103.4% | 103.2% | 106.8% | 108.2% | 83.83% | 94.04% | 100.3% | 98.29% | 118.5% | 110.6% | 99.6% | 96.33% | 93.27% | 98.9% | 97.1% | 101.4% | 78.75% | 68.28% | 71.13% | 73.65% | 66.5% |  | 98.23% | 96.73% | 110.5% | 103.8%
| 100 | CCV 1ppb TRACE |  | 06/16/2023 22:14:23 | 1042.9753 | 1077.1368 | 1.0946 | 0.055224 | 1588.5975 | 1046.2009 | 1073.8464 | 1.0625 | 1.0754 | 1.08 | 0.90942 | 1.0248 | 1.1486 | 1.0834 | 1.0595 | 1.065 | 1.0643 | 0.82146 | 0.89691 | 1.0153 | 0.99907 | 1.2416 | 1.0274 | 0.96342 | 0.97571 | 0.95118 | 1.1049 | 0.97055 | 1.017 | [6 empty] | 0.97399 | 0.98445 | 1.0887 | 0.93946
|  | RSD |  |  | 3.545% | 2.481% | 18.89% | 52.26% | 5.14% | 1.176% | 7.358% | 5.01% | 1.535% | 0.4907% | 4.858% | 2.408% | 5.69% | 0.2907% | 5.68% | 0.7116% | 1.306% | 1.07% | 7.694% | 11.91% | 1.921% | 6.745% | 8.417% | 2.647% | 3.983% | 7.341% | 11.56% | 0.7815% | 0.1719% | [6 empty] | 1.949% | 0.1715% | 0.7089% | 3.317%
|  | % Recovery |  |  | 104.3% | 107.7% | 109.5% | 5.522% | 158.9% | 104.6% | 107.4% | 106.3% | 107.5% | 108% | 90.94% | 102.5% | 114.9% | 108.3% | 105.9% | 106.5% | 106.4% | 82.15% | 89.69% | 101.5% | 99.91% | 124.2% | 102.7% | 96.34% | 97.57% | 95.12% | 110.5% | 97.05% | 101.7% | 78.85% | 67.64% | 72.11% | 79.46% | 66.12% |  | 97.4% | 98.44% | 108.9% | 93.95%
| 114 | CCV 1ppb TRACE |  | 06/16/2023 23:39:19 | 973.5156 | 996.1043 | 1.0589 | 0.14864 | 1343.6848 | 993.3618 | 1012.2498 | 1.0111 | 1.0245 | 1.0526 | 0.85519 | 1.0001 | 1.0524 | 0.97807 | 0.94431 | 1.0144 | 1.0229 | 0.78053 | 0.92841 | 0.99216 | 0.95211 | 1.2711 | 0.88487 | 0.95893 | 0.97227 | 0.95235 | 0.9513 | 0.97658 | 0.99264 | [6 empty] | 0.95932 | 0.96684 | 1.1404 | 1.0837
|  | RSD |  |  | 3.484% | 2.505% | 20.68% | 7.479% | 7.951% | 5.76% | 2.49% | 1.702% | 3.253% | 4.81% | 3.456% | 0.3748% | 4.739% | 1.23% | 5.97% | 1.777% | 2.581% | 1.831% | 2.854% | 11.71% | 4.622% | 5.976% | 9.831% | 4.258% | 5.978% | 11.98% | 3.943% | 1.865% | 1.286% | [6 empty] | 1.363% | 1.446% | 2.634% | 0.9985%
|  | % Recovery |  |  | 97.35% | 99.61% | 105.9% | 14.86% | 134.4% | 99.34% | 101.2% | 101.1% | 102.4% | 105.3% | 85.52% | 100% | 105.2% | 97.81% | 94.43% | 101.4% | 102.3% | 78.05% | 92.84% | 99.22% | 95.21% | 127.1% | 88.49% | 95.89% | 97.23% | 95.24% | 95.13% | 97.66% | 99.26% | 78.33% | 63.95% | 71.94% | 73.62% | 62.34% |  | 95.93% | 96.68% | 114% | 108.4%
| 115 | CCV 1ppb TRACE |  | 06/16/2023 23:45:38 | 988.5811 | 943.7088 | 1.1674 | 0.089159 | 1565.2944 | 990.3717 | 1077.1318 | 0.99564 | 1.0509 | 1.0646 | 0.93036 | 0.99354 | 1.0632 | 0.99528 | 0.99444 | 1.0052 | 1.0578 | 0.82988 | 0.99061 | 0.90047 | 0.97965 | 1.6046 | 1.0735 | 1.0641 | 0.97289 | 1.148 | 0.93666 | 1.0003 | 1.0057 | [6 empty] | 0.95725 | 0.97728 | 1.3077 | 1.0615
|  | RSD |  |  | 1.799% | 5.46% | 16.76% | 27.88% | 8.49% | 5.15% | 11.6% | 2.833% | 3.202% | 0.8382% | 3.195% | 1.701% | 3.589% | 2.514% | 1.288% | 3.665% | 0.7634% | 2.123% | 3.028% | 19.6% | 2.576% | 8.16% | 32.22% | 4.627% | 5.183% | 4.943% | 21.7% | 1.669% | 0.9604% | [6 empty] | 4.821% | 0.8933% | 6.068% | 0.8664%
|  | % Recovery |  |  | 98.86% | 94.37% | 116.7% | 8.916% | 156.5% | 99.04% | 107.7% | 99.56% | 105.1% | 106.5% | 93.04% | 99.35% | 106.3% | 99.53% | 99.44% | 100.5% | 105.8% | 82.99% | 99.06% | 90.05% | 97.96% | 160.5% | 107.4% | 106.4% | 97.29% | 114.8% | 93.67% | 100% | 100.6% | 68.17% | 63.54% | 70.08% | 72.26% | 62.14% |  | 95.73% | 97.73% | 130.8% | 106.1%
| 131 | CCV 1ppb TRACE |  | 06/17/2023 01:23:12 | 972.3873 | 1032.7939 | 1.2642 | 0.080488 | 1454.5502 | 1054.1315 | 1036.1287 | 1.0181 | 1.0324 | 1.0541 | 0.89712 | 0.99441 | 0.96018 | 1.0006 | 0.96316 | 1.0235 | 1.043 | 0.80002 | 0.89649 | 1.0668 | 0.95919 | 1.105 | 0.86769 | 1.0299 | 0.95419 | 0.95715 | 1.0801 | 0.98211 | 1.0186 | [6 empty] | 0.97143 | 0.99687 | 1.1638 | 1.0874
|  | RSD |  |  | 2.137% | 9.593% | 13.29% | 44.97% | 12.72% | 6.057% | 6.493% | 2.452% | 3.847% | 2.886% | 0.814% | 1.999% | 1.415% | 0.9743% | 0.669% | 4.122% | 1.943% | 4.973% | 3.072% | 24.09% | 2.718% | 14.74% | 13.98% | 5.246% | 6.991% | 10.78% | 12.08% | 2.397% | 1.091% | [6 empty] | 0.8193% | 2.562% | 0.6317% | 1.417%
|  | % Recovery |  |  | 97.24% | 103.3% | 126.4% | 8.049% | 145.5% | 105.4% | 103.6% | 101.8% | 103.2% | 105.4% | 89.71% | 99.44% | 96.02% | 100.1% | 96.32% | 102.3% | 104.3% | 80% | 89.65% | 106.7% | 95.92% | 110.5% | 86.77% | 103% | 95.42% | 95.71% | 108% | 98.21% | 101.9% | 73.11% | 59.55% | 66.73% | 67.87% | 58.29% |  | 97.14% | 99.69% | 116.4% | 108.7%
| 132 | CCV 1ppb TRACE |  | 06/17/2023 01:29:31 | 965.4585 | 962.0715 | 1.0941 | 0.10977 | 1521.441 | 1017.0475 | 1037.8145 | 1.022 | 1.0321 | 1.0666 | 0.93879 | 1.0066 | 0.95728 | 0.98594 | 0.98441 | 1.0047 | 1.034 | 0.83014 | 0.91458 | 1.0204 | 0.97087 | 1.174 | 0.95901 | 1.029 | 0.98943 | 0.95493 | 1.0794 | 0.98612 | 1.0012 | [6 empty] | 0.97472 | 0.97825 | 1.1556 | 1.1375
|  | RSD |  |  | 2.535% | 4.614% | 29.2% | 42.78% | 13.69% | 4.022% | 6.054% | 0.864% | 2.484% | 1.587% | 2.173% | 1.235% | 7.545% | 3.224% | 7.949% | 5.348% | 1.037% | 0.9791% | 5.805% | 13.54% | 1.592% | 10.9% | 15.46% | 5.374% | 2.593% | 20.57% | 6.598% | 0.7189% | 1.084% | [6 empty] | 1.165% | 0.5035% | 1.596% | 1.348%
|  | % Recovery |  |  | 96.55% | 96.21% | 109.4% | 10.98% | 152.1% | 101.7% | 103.8% | 102.2% | 103.2% | 106.7% | 93.88% | 100.7% | 95.73% | 98.59% | 98.44% | 100.5% | 103.4% | 83.01% | 91.46% | 102% | 97.09% | 117.4% | 95.9% | 102.9% | 98.94% | 95.49% | 107.9% | 98.61% | 100.1% | 73.05% | 59.32% | 67.15% | 66.97% | 58.04% |  | 97.47% | 97.83% | 115.6% | 113.7%
| 146 | CCV 1ppb TRACE |  | 06/17/2023 02:55:21 | 1005.7882 | 1048.4076 | 1.304 | 0.3476 | 1727.1377 | 1054.9772 | 1057.608 | 1.086 | 1.0813 | 1.1036 | 0.97469 | 1.0112 | 1.0803 | 0.99081 | 1.0656 | 1.041 | 1.0678 | 0.83716 | 0.86468 | 1.013 | 0.98192 | 1.2262 | 0.89459 | 0.96245 | 0.92098 | 0.96661 | 0.94546 | 0.98616 | 1.0248 | [6 empty] | 0.99923 | 0.98593 | 1.1311 | 1.0721
|  | RSD |  |  | 2.597% | 1.901% | 9.279% | 13.56% | 8.48% | 4.05% | 4.879% | 6.297% | 3.054% | 0.2875% | 3.888% | 1.544% | 7.57% | 2.562% | 3.201% | 0.7694% | 0.3752% | 2.309% | 3.518% | 9.339% | 2.141% | 9.73% | 10.9% | 13.41% | 2.576% | 17.03% | 2.496% | 1.661% | 1.076% | [6 empty] | 2.107% | 2.613% | 0.7597% | 1.314%
|  | % Recovery |  |  | 100.6% | 104.8% | 130.4% | 34.76% | 172.7% | 105.5% | 105.8% | 108.6% | 108.1% | 110.4% | 97.47% | 101.1% | 108% | 99.08% | 106.6% | 104.1% | 106.8% | 83.72% | 86.47% | 101.3% | 98.19% | 122.6% | 89.46% | 96.24% | 92.1% | 96.66% | 94.55% | 98.62% | 102.5% | 73.82% | 61.85% | 67.65% | 70.31% | 60.25% |  | 99.92% | 98.59% | 113.1% | 107.2%
| 147 | CCV 1ppb TRACE |  | 06/17/2023 03:01:39 | 1039.1883 | 1046.4968 | 1.1749 | 0.1367 | 1691.0827 | 1014.6489 | 1058.595 | 1.0321 | 1.0505 | 1.1139 | 0.95184 | 1.0209 | 1.2062 | 1.0493 | 1.0383 | 1.0472 | 1.0801 | 0.82147 | 0.92262 | 0.97322 | 0.99683 | 1.0198 | 0.83139 | 1.0285 | 0.99269 | 0.82037 | 0.95962 | 0.99313 | 1.0076 | [6 empty] | 0.95779 | 0.96766 | 1.0514 | 1.0706
|  | RSD |  |  | 1.841% | 5.26% | 13.08% | 36.85% | 11.93% | 2.431% | 2.668% | 2.498% | 1.866% | 1.128% | 3.26% | 1.057% | 7.064% | 3.156% | 6.268% | 5.396% | 1.474% | 3.91% | 2.767% | 22.26% | 2.805% | 24.57% | 32.45% | 3.321% | 3.593% | 11.6% | 0.832% | 1.666% | 1.098% | [6 empty] | 0.6272% | 1.18% | 9.121% | 0.7657%
|  | % Recovery |  |  | 103.9% | 104.6% | 117.5% | 13.67% | 169.1% | 101.5% | 105.9% | 103.2% | 105.1% | 111.4% | 95.18% | 102.1% | 120.6% | 104.9% | 103.8% | 104.7% | 108% | 82.15% | 92.26% | 97.32% | 99.68% | 102% | 83.14% | 102.9% | 99.27% | 82.04% | 95.96% | 99.31% | 100.8% | 77.68% | 60.61% | 66.43% | 68.2% | 60.4% |  | 95.78% | 96.77% | 105.1% | 107.1%
| 164 | CCV 1ppb TRACE |  | 06/17/2023 04:46:11 | 983.7392 | 1072.1924 | 1.1539 | 0.055486 | 1537.7798 | 1029.8857 | 1079.6466 | 1.0459 | 1.0545 | 1.07 | 0.89737 | 1.0298 | 0.96944 | 0.97811 | 1.0569 | 1.0311 | 1.0457 | 0.82232 | 0.9227 | 0.99993 | 0.96635 | 1.3222 | 0.92116 | 0.996 | 0.98525 | 1.0635 | 0.96342 | 0.98122 | 1.0053 | [6 empty] | 0.97088 | 0.96993 | 1.1709 | 1.1087
|  | RSD |  |  | 3.958% | 2.781% | 13.87% | 61.35% | 9.688% | 7.081% | 4.114% | 2.058% | 1.33% | 1.631% | 8.959% | 1.6% | 11.87% | 4.369% | 8.94% | 3.622% | 1.63% | 1.646% | 5.999% | 10.5% | 2.176% | 7.189% | 25.73% | 6.199% | 8.038% | 18.1% | 15.74% | 0.6848% | 1.214% | [6 empty] | 2.11% | 2.307% | 1.949% | 1.568%
|  | % Recovery |  |  | 98.37% | 107.2% | 115.4% | 5.549% | 153.8% | 103% | 108% | 104.6% | 105.4% | 107% | 89.74% | 103% | 96.94% | 97.81% | 105.7% | 103.1% | 104.6% | 82.23% | 92.27% | 99.99% | 96.63% | 132.2% | 92.12% | 99.6% | 98.53% | 106.3% | 96.34% | 98.12% | 100.5% | 70.53% | 56.91% | 64.34% | 65.48% | 55.44% |  | 97.09% | 96.99% | 117.1% | 110.9%
| 165 | CCV 1ppb TRACE |  | 06/17/2023 04:52:31 | 984.0166 | 996.0293 | 1.0865 | 0.068408 | 1722.9088 | 1051.2473 | 1043.3635 | 1.0329 | 1.0552 | 1.0901 | 0.96717 | 1.0441 | 1.0858 | 1.0268 | 0.98463 | 1.0304 | 1.0665 | 0.83325 | 0.85805 | 0.89092 | 0.98626 | 1.5568 | 0.94617 | 1.0005 | 0.99509 | 0.87096 | 1.0018 | 0.98554 | 1.0256 | [6 empty] | 0.97611 | 0.98955 | 1.2189 | 1.1084
|  | RSD |  |  | 4.595% | 3.341% | 16.92% | 45.02% | 9.943% | 3.78% | 4.012% | 5.504% | 5.441% | 0.7732% | 3.573% | 1.057% | 12.73% | 3.602% | 6.475% | 5.072% | 1.491% | 1.633% | 3.002% | 20.13% | 1.219% | 26.62% | 14.33% | 9.811% | 5.054% | 29.44% | 5.042% | 1.026% | 1.573% | [6 empty] | 2.427% | 1.26% | 10.79% | 1.481%
|  | % Recovery |  |  | 98.4% | 99.6% | 108.7% | 6.841% | 172.3% | 105.1% | 104.3% | 103.3% | 105.5% | 109% | 96.72% | 104.4% | 108.6% | 102.7% | 98.46% | 103% | 106.6% | 83.33% | 85.8% | 89.09% | 98.63% | 155.7% | 94.62% | 100.1% | 99.51% | 87.1% | 100.2% | 98.55% | 102.6% | 68.59% | 56.76% | 64.06% | 65.48% | 56.04% |  | 97.61% | 98.95% | 121.9% | 110.8%
| 180 | CCV 1ppb TRACE |  | 06/17/2023 06:24:40 | 968.0848 | 959.1468 | 1.2952 | 0.059317 | 1514.0892 | 1011.1718 | 996.4636 | 1.0541 | 1.0725 | 1.0959 | 0.95095 | 1.0273 | 0.94525 | 0.9819 | 0.97449 | 0.98646 | 1.0628 | 0.84725 | 0.88584 | 1.0034 | 1.0054 | 1.1615 | 1.1205 | 1.0212 | 0.92404 | 0.90827 | 1.1525 | 1.0126 | 0.99527 | [6 empty] | 0.98473 | 0.98321 | 1.1529 | 1.1166
|  | RSD |  |  | 4.493% | 7.163% | 4.655% | 32.87% | 10.08% | 2.715% | 4.607% | 3.435% | 2.681% | 1.978% | 6.169% | 0.9752% | 17.58% | 3.161% | 2.511% | 3.574% | 2.007% | 2.674% | 7.287% | 5.914% | 1.492% | 10.51% | 4.901% | 6.954% | 6.842% | 13.1% | 19.47% | 0.6027% | 1.95% | [6 empty] | 1.709% | 1.029% | 2.049% | 1.107%
|  | % Recovery |  |  | 96.81% | 95.91% | 129.5% | 5.932% | 151.4% | 101.1% | 99.65% | 105.4% | 107.2% | 109.6% | 95.09% | 102.7% | 94.52% | 98.19% | 97.45% | 98.65% | 106.3% | 84.72% | 88.58% | 100.3% | 100.5% | 116.2% | 112% | 102.1% | 92.4% | 90.83% | 115.3% | 101.3% | 99.53% | 70.87% | 54.75% | 61.81% | 63.28% | 53.75% |  | 98.47% | 98.32% | 115.3% | 111.7%
| 181 | CCV 1ppb TRACE |  | 06/17/2023 06:31:00 | 989.569 | 1004.8621 | 1.1962 | 0.085023 | 1760.5622 | 1005.7186 | 1055.8563 | 1.0353 | 1.0936 | 1.0879 | 0.92351 | 1.0383 | 1.0326 | 1.0053 | 0.9927 | 1.0485 | 1.0576 | 0.83569 | 0.83491 | 1.0503 | 0.9721 | 1.1549 | 0.86982 | 1.0134 | 0.98219 | 0.97531 | 1.0308 | 0.99013 | 1.0449 | [6 empty] | 0.96264 | 0.97011 | 1.1216 | 1.1282
|  | RSD |  |  | 2.489% | 6.111% | 16.9% | 27.18% | 3.177% | 3.327% | 7.005% | 1.788% | 4.922% | 1.289% | 4.497% | 2.709% | 1.758% | 1.738% | 1.399% | 2.226% | 0.7559% | 4.276% | 4.562% | 5.872% | 1.743% | 13.24% | 9.178% | 12.75% | 5.96% | 9.397% | 16.9% | 1.764% | 0.7212% | [6 empty] | 1.145% | 2.025% | 3.829% | 0.881%
|  | % Recovery |  |  | 98.96% | 100.5% | 119.6% | 8.502% | 176.1% | 100.6% | 105.6% | 103.5% | 109.4% | 108.8% | 92.35% | 103.8% | 103.3% | 100.5% | 99.27% | 104.8% | 105.8% | 83.57% | 83.49% | 105% | 97.21% | 115.5% | 86.98% | 101.3% | 98.22% | 97.53% | 103.1% | 99.01% | 104.5% | 72.14% | 54.29% | 61.78% | 63.44% | 53.57% |  | 96.26% | 97.01% | 112.2% | 112.8%
| 198 | CCV 1ppb TRACE |  | 06/17/2023 08:15:29 | 1002.3381 | 996.7207 | 1.3114 | 0.063046 | 1835.2714 | 1027.1575 | 1055.0865 | 1.0733 | 1.0884 | 1.0695 | 0.95485 | 1.0151 | 1.0531 | 1.0411 | 1.0045 | 1.0406 | 1.0572 | 0.81105 | 0.8506 | 1.0771 | 0.99088 | 1.287 | 0.90792 | 0.9629 | 0.98127 | 0.93088 | 1.0329 | 0.99225 | 1.0192 | [6 empty] | 1.0105 | 0.99141 | 1.2018 | 1.1506
|  | RSD |  |  | 2.422% | 4.661% | 2.433% | 126.1% | 11.18% | 2.691% | 2.509% | 1.022% | 3.387% | 0.6674% | 4.715% | 2.9% | 7.26% | 2.104% | 8.394% | 2.614% | 3.612% | 0.427% | 5.331% | 19.47% | 2.91% | 8.005% | 15.34% | 8.316% | 1.686% | 21.05% | 14.93% | 2.611% | 0.7784% | [6 empty] | 1.368% | 1.674% | 0.5683% | 1.321%
|  | % Recovery |  |  | 100.2% | 99.67% | 131.1% | 6.305% | 183.5% | 102.7% | 105.5% | 107.3% | 108.8% | 106.9% | 95.48% | 101.5% | 105.3% | 104.1% | 100.5% | 104.1% | 105.7% | 81.11% | 85.06% | 107.7% | 99.09% | 128.7% | 90.79% | 96.29% | 98.13% | 93.09% | 103.3% | 99.22% | 101.9% | 66.38% | 51.53% | 58.78% | 60.64% | 50.55% |  | 101% | 99.14% | 120.2% | 115.1%
| 199 | CCV 1ppb TRACE |  | 06/17/2023 08:21:48 | 997.0932 | 995.0847 | 1.4091 | 0.051982 | 1719.7104 | 1061.1018 | 978.5833 | 1.0568 | 1.0815 | 1.0769 | 0.93359 | 1.0394 | 1.0595 | 1.0099 | 0.94677 | 1.0184 | 1.0648 | 0.81691 | 0.90047 | 0.92465 | 0.98077 | 1.2687 | 1.067 | 1.0128 | 1.0302 | 0.95362 | 0.97098 | 0.99058 | 1.0264 | [6 empty] | 0.99939 | 0.99145 | 1.1946 | 1.146
|  | RSD |  |  | 3.377% | 3.153% | 25.47% | 95.66% | 12.91% | 5.263% | 2.179% | 4.362% | 3.893% | 1.665% | 3.777% | 1.149% | 3.655% | 0.9383% | 1.07% | 3.476% | 1.927% | 3.438% | 2.992% | 5.118% | 1.192% | 12.72% | 25.19% | 6.283% | 9.482% | 7.592% | 20.04% | 0.7068% | 1.536% | [6 empty] | 0.3406% | 2.284% | 0.3957% | 1.561%
|  | % Recovery |  |  | 99.71% | 99.51% | 140.9% | 5.198% | 172% | 106.1% | 97.86% | 105.7% | 108.1% | 107.7% | 93.36% | 103.9% | 106% | 101% | 94.68% | 101.8% | 106.5% | 81.69% | 90.05% | 92.47% | 98.08% | 126.9% | 106.7% | 101.3% | 103% | 95.36% | 97.1% | 99.06% | 102.6% | 65.35% | 50.63% | 58.33% | 59.72% | 50.47% |  | 99.94% | 99.15% | 119.5% | 114.6%
| 214 | CCV 1ppb TRACE |  | 06/17/2023 09:53:59 | 941.3352 | 971.9387 | 0.8378 | 0.11132 | 1197.1077 | 1024.1048 | 1066.1094 | 1.0225 | 1.0422 | 1.0843 | 0.88773 | 1.0488 | 0.97888 | 0.98638 | 0.96001 | 1.0284 | 1.0613 | 0.83248 | 0.79129 | 0.88526 | 0.96246 | 1.2541 | 1.0028 | 1.0659 | 0.96182 | 0.90306 | 0.94047 | 0.97769 | 1.0342 | [6 empty] | 0.9842 | 0.99871 | 1.2063 | 1.1878
|  | RSD |  |  | 1.66% | 2.61% | 36.14% | 65.94% | 10.48% | 4.711% | 7.311% | 1.055% | 5.555% | 1.413% | 3.711% | 0.9289% | 12% | 2.018% | 7.334% | 2.901% | 2.257% | 1.299% | 7.765% | 15.33% | 1.528% | 5.255% | 16.93% | 14.25% | 4.225% | 6.621% | 24.22% | 1.71% | 2.664% | [6 empty] | 2.472% | 2.293% | 1.308% | 1.641%
|  | % Recovery |  |  | 94.13% | 97.19% | 83.78% | 11.13% | 119.7% | 102.4% | 106.6% | 102.2% | 104.2% | 108.4% | 88.77% | 104.9% | 97.89% | 98.64% | 96% | 102.8% | 106.1% | 83.25% | 79.13% | 88.53% | 96.25% | 125.4% | 100.3% | 106.6% | 96.18% | 90.31% | 94.05% | 97.77% | 103.4% | 63.74% | 48.27% | 55.93% | 57.54% | 47.68% |  | 98.42% | 99.87% | 120.6% | 118.8%
| 215 | CCV 1ppb TRACE |  | 06/17/2023 10:00:19 | 970.6048 | 967.9932 | 0.98345 | 0.050434 | 1617.0556 | 1048.122 | 1118.9109 | 1.0582 | 1.0757 | 1.095 | 0.94119 | 1.0547 | 0.98358 | 0.99119 | 0.98077 | 0.99756 | 1.0699 | 0.81686 | 0.85719 | 0.99734 | 0.99774 | 1.1592 | 1.0502 | 0.93461 | 0.92341 | 0.94712 | 1.08 | 1.003 | 1.0383 | [6 empty] | 1.0128 | 1.0111 | 1.1968 | 1.1765
|  | RSD |  |  | 2.814% | 1.495% | 24.82% | 100.5% | 9.039% | 6.399% | 7.726% | 2.496% | 1.238% | 7.464% | 16.57% | 0.9959% | 8.641% | 0.368% | 5.361% | 2.692% | 8.233% | 11.84% | 2.116% | 12.65% | 2.493% | 14.83% | 14.85% | 19.06% | 3.013% | 19.52% | 13.46% | 11.53% | 2.43% | [6 empty] | 9.965% | 9.688% | 1.859% | 0.3808%
|  | % Recovery |  |  | 97.06% | 96.8% | 98.34% | 5.043% | 161.7% | 104.8% | 111.9% | 105.8% | 107.6% | 109.5% | 94.12% | 105.5% | 98.36% | 99.12% | 98.08% | 99.76% | 107% | 81.69% | 85.72% | 99.73% | 99.77% | 115.9% | 105% | 93.46% | 92.34% | 94.71% | 108% | 100.3% | 103.8% | 64.21% | 48.39% | 55.3% | 57.93% | 47.52% |  | 101.3% | 101.1% | 119.7% | 117.6%
| 232 | CCV 1ppb TRACE |  | 06/17/2023 11:44:51 | 967.4726 | 1004.7918 | 1.361 | 0.24606 | 1518.8956 | 990.955 | 1104.9739 | 1.019 | 1.0727 | 1.0913 | 0.91663 | 1.0591 | 1.1138 | 0.97472 | 0.98598 | 1.0189 | 1.0193 | 0.81806 | 0.85939 | 0.99708 | 0.99701 | 1.1505 | 0.8848 | 0.99406 | 1.0403 | 1.0933 | 1.002 | 0.98413 | 1.0285 | [6 empty] | 0.96878 | 0.98541 | 1.2091 | 1.1621
|  | RSD |  |  | 2.369% | 5.957% | 14.66% | 9.018% | 6.8% | 3.837% | 1.15% | 1.762% | 8.134% | 2.835% | 7.226% | 0.1711% | 6.948% | 1.223% | 2.454% | 2.481% | 2.15% | 3.362% | 8.398% | 13.73% | 5.205% | 13.23% | 21.49% | 5.993% | 3.093% | 2.166% | 6.782% | 1.823% | 1.399% | [6 empty] | 1.023% | 0.8539% | 0.3795% | 1.383%
|  | % Recovery |  |  | 96.75% | 100.5% | 136.1% | 24.61% | 151.9% | 99.1% | 110.5% | 101.9% | 107.3% | 109.1% | 91.66% | 105.9% | 111.4% | 97.47% | 98.6% | 101.9% | 101.9% | 81.81% | 85.94% | 99.71% | 99.7% | 115.1% | 88.48% | 99.41% | 104% | 109.3% | 100.2% | 98.41% | 102.9% | 64.74% | 48.75% | 57.11% | 58.92% | 48.03% |  | 96.88% | 98.54% | 120.9% | 116.2%
| 233 | CCV 1ppb TRACE |  | 06/17/2023 11:51:11 | 971.2691 | 1012.0527 | 1.2375 | 0.23216 | 1682.2554 | 1039.9667 | 1059.4263 | 1.0437 | 1.0384 | 1.0541 | 0.89143 | 1.083 | 1.0717 | 0.99625 | 1.0157 | 1.0338 | 1.0378 | 0.79909 | 0.86597 | 1.0611 | 0.96681 | 1.4148 | 0.9827 | 0.9507 | 0.98401 | 1.0236 | 1.0158 | 0.97625 | 1.0435 | [6 empty] | 0.95965 | 0.98965 | 1.1992 | 1.1692
|  | RSD |  |  | 0.5947% | 2.679% | 3.768% | 8.819% | 7.594% | 3.916% | 5.848% | 1.114% | 3.019% | 3.512% | 4.177% | 0.9864% | 10.48% | 2.038% | 4.431% | 2.875% | 1.771% | 3.546% | 6.442% | 9.039% | 1.453% | 1.761% | 13.73% | 18.27% | 4.806% | 2.407% | 19.68% | 0.7271% | 1.405% | [6 empty] | 3.373% | 1.048% | 1.823% | 1.287%
|  | % Recovery |  |  | 97.13% | 101.2% | 123.7% | 23.22% | 168.2% | 104% | 105.9% | 104.4% | 103.8% | 105.4% | 89.14% | 108.3% | 107.2% | 99.62% | 101.6% | 103.4% | 103.8% | 79.91% | 86.6% | 106.1% | 96.68% | 141.5% | 98.27% | 95.07% | 98.4% | 102.4% | 101.6% | 97.63% | 104.3% | 64.34% | 48.42% | 56.66% | 58.22% | 48.19% |  | 95.96% | 98.96% | 119.9% | 116.9%
QC STD 4
|  | Sample Id | R | Acquisition Time | Na KED 23/23 | Mg KED 24/24 | Al KED 27/27 | P Oshift 31/47 | S (34) -1 34/34 | K KED 39/39 | Ca KED 43/43 | V KED 51/51 | Cr (52) 52/52 | Cr (53) NH3 53/53 | Fe (54) NH3 54/54 | Mn Oshift 55/71 | Fe (56) KED 56/56 | Co KED 59/59 | Ni KED 60/60 | Cu (63) KED 63/63 | Cu (65) NH3 65/65 | Zn NH3 66/66 | As 75/75 | As KED 75/75 | As Oshift 75/91 | Se (77) 77/77 | Se (77) Oshift 77/93 | Se (78) Oshift 78/94 | Se (80) Oshift 80/96 | Se (82) 82/82 | Se (82) Oshift 82/98 | Mo (95) NH3 95/95 | Mo (95) OShift 95/127 | Rh (IS1) 103/103 (IS) | Rh (IS2) 103/103 (IS) | Rh (IS4) 103/103 (IS) | Rh (IS5) 103/103 (IS) | Rh (IS6) 103/103 (IS) | Rh (IS7) 103/103 (IS) | Cd (111) NH3 111/111 | Cd (114) NH3 114/114 | Pb 208/208 | Pb NH3 208/208
| 21 | LFB |  | 06/16/2023 14:06:07 | 10631.0355 | 10937.2198 | 10.7055 | 11.5311 | 10903.3394 | 10615.1967 | 10398.4198 | 10.7744 | 11.3367 | 9.9786 | 9.5522 | 9.9937 | 10.6854 | 10.9986 | 10.5865 | 10.8235 | 9.914 | 10.0895 | 9.9048 | 11.207 | 9.9616 | 10.3388 | 10.1419 | 10.1303 | 9.9546 | 10.4514 | 10.5175 | 9.7079 | 9.7943 | [6 empty] | 9.669 | 9.5549 | 9.1443 | 9.426
|  | RSD |  |  | 9.294% | 9.795% | 11.63% | 0.5237% | 12.11% | 9.65% | 8.647% | 10.41% | 11.27% | 4.291% | 5.056% | 0.9218% | 12.85% | 10.29% | 10.32% | 10.63% | 3.289% | 4.34% | 1.462% | 14.67% | 1.092% | 2.59% | 1.71% | 2.934% | 2.562% | 2.87% | 3.766% | 5.365% | 1.34% | [6 empty] | 4.806% | 4.649% | 0.5717% | 1.751%
|  | % Recovery |  |  | 106.3% | 109.4% | 107.1% | 115.3% | 109% | 106.2% | 104% | 107.7% | 113.4% | 99.79% | 95.52% | 99.94% | 106.9% | 110% | 105.9% | 108.2% | 99.14% | 100.9% | 99.05% | 112.1% | 99.62% | 103.4% | 101.4% | 101.3% | 99.55% | 104.5% | 105.2% | 97.08% | 97.94% | 89.22% | 83.36% | 85.25% | 86.67% | 82.9% |  | 96.69% | 95.55% | 91.44% | 94.26%
| 55 | LFB |  | 06/16/2023 17:42:12 | 9318.2332 | 9688.4626 | 9.7361 | 11.4387 | 10212.8116 | 9450.1073 | 9627.478 | 9.6244 | 10.0298 | 9.8195 | 9.6886 | 9.8477 | 9.3763 | 9.8275 | 9.545 | 9.8513 | 9.8229 | 9.7185 | 9.5215 | 10.0351 | 9.8384 | 10.3697 | 10.479 | 9.8586 | 9.9861 | 10.1133 | 9.8555 | 9.5679 | 9.8013 | [6 empty] | 9.4964 | 9.4689 | 9.4318 | 9.6216
|  | RSD |  |  | 1.369% | 2.521% | 2.386% | 0.7684% | 3.189% | 1.411% | 1.501% | 1.374% | 0.5744% | 0.582% | 0.7025% | 1.102% | 1.538% | 0.5445% | 0.9103% | 0.6605% | 0.7689% | 0.6492% | 1.751% | 3.019% | 0.9623% | 4.602% | 3.837% | 3.355% | 0.3693% | 4.766% | 3.396% | 0.8523% | 0.4636% | [6 empty] | 1.629% | 1.534% | 0.7322% | 1.589%
|  | % Recovery |  |  | 93.18% | 96.88% | 97.36% | 114.4% | 102.1% | 94.5% | 96.27% | 96.24% | 100.3% | 98.2% | 96.89% | 98.48% | 93.76% | 98.28% | 95.45% | 98.51% | 98.23% | 97.19% | 95.22% | 100.4% | 98.38% | 103.7% | 104.8% | 98.59% | 99.86% | 101.1% | 98.56% | 95.68% | 98.01% | 81.51% | 70.96% | 75.83% | 77.13% | 71.52% |  | 94.96% | 94.69% | 94.32% | 96.22%
| 86 | LFB |  | 06/16/2023 20:49:46 | 9466.6541 | 9904.8327 | 9.4772 | 11.9438 | 10401.5438 | 9535.8912 | 9640.5655 | 9.5913 | 10.0779 | 10.0633 | 9.7932 | 9.9504 | 9.2658 | 9.741 | 9.4367 | 9.7016 | 9.8291 | 9.7731 | 9.8013 | 10.0316 | 9.739 | 10.624 | 9.8981 | 9.9175 | 9.9831 | 10.3206 | 9.9436 | 9.5256 | 9.8087 | [6 empty] | 9.4004 | 9.4604 | 9.7719 | 9.6345
|  | RSD |  |  | 0.1591% | 1.854% | 6.106% | 0.8385% | 2.574% | 3.269% | 1.119% | 0.3248% | 0.5305% | 0.8769% | 2.719% | 1.788% | 0.9701% | 0.8488% | 0.1847% | 1.197% | 1.338% | 1.257% | 4.444% | 1.391% | 1.756% | 4.664% | 4.444% | 3.118% | 2.126% | 5.945% | 3.36% | 1.475% | 1.294% | [6 empty] | 1.642% | 1.409% | 1.869% | 0.3545%
|  | % Recovery |  |  | 94.67% | 99.05% | 94.77% | 119.4% | 104% | 95.36% | 96.41% | 95.91% | 100.8% | 100.6% | 97.93% | 99.5% | 92.66% | 97.41% | 94.37% | 97.02% | 98.29% | 97.73% | 98.01% | 100.3% | 97.39% | 106.2% | 98.98% | 99.17% | 99.83% | 103.2% | 99.44% | 95.26% | 98.09% | 76.21% | 66.1% | 70.78% | 72.1% | 66.42% |  | 94% | 94.6% | 97.72% | 96.35%
| 116 | LFB |  | 06/16/2023 23:51:38 | 8856.5512 | 9233.4125 | 9.0278 | 11.5031 | 10225.7146 | 8971.6175 | 9523.1364 | 9.4157 | 9.8454 | 10.0143 | 9.6841 | 9.7627 | 8.711 | 9.4044 | 9.2996 | 9.5611 | 9.7698 | 9.8099 | 9.0941 | 9.6085 | 9.6807 | 10.0371 | 9.4857 | 9.6289 | 9.9364 | 9.6183 | 9.8764 | 9.6503 | 9.8539 | [6 empty] | 9.5391 | 9.5837 | 9.4655 | 10.0337
|  | RSD |  |  | 1.594% | 0.7547% | 11.11% | 1.733% | 1.505% | 1.78% | 2.146% | 2.31% | 2.211% | 1.168% | 2.72% | 1.044% | 1.817% | 1.172% | 1.045% | 0.9967% | 0.2913% | 0.8767% | 2.868% | 2.997% | 1.596% | 3.222% | 4.437% | 2.106% | 1.531% | 4.891% | 2.991% | 0.3969% | 0.2004% | [6 empty] | 1.158% | 0.354% | 3.092% | 0.7532%
|  | % Recovery |  |  | 88.57% | 92.33% | 90.28% | 115% | 102.3% | 89.72% | 95.23% | 94.16% | 98.45% | 100.1% | 96.84% | 97.63% | 87.11% | 94.04% | 93% | 95.61% | 97.7% | 98.1% | 90.94% | 96.08% | 96.81% | 100.4% | 94.86% | 96.29% | 99.36% | 96.18% | 98.76% | 96.5% | 98.54% | 76.6% | 59.15% | 65.61% | 67.66% | 58.8% |  | 95.39% | 95.84% | 94.66% | 100.3%
| 117 | LFB |  | 06/16/2023 23:57:55 | 8924.9738 | 9503.9294 | 9.1653 | 11.5659 | 10370.3317 | 9199.315 | 9709.2367 | 9.5653 | 10.0675 | 10.037 | 9.7676 | 9.8653 | 8.7551 | 9.6751 | 9.3953 | 9.6844 | 9.9154 | 9.7938 | 9.2828 | 9.5281 | 9.8623 | 10.2711 | 9.9673 | 10.1907 | 10.0648 | 9.8316 | 10.3033 | 9.6507 | 9.8937 | [6 empty] | 9.5445 | 9.6102 | 9.6924 | 9.9435
|  | RSD |  |  | 1.365% | 2.48% | 5.128% | 0.4739% | 1.547% | 2.123% | 1.58% | 2.39% | 1.702% | 0.5856% | 0.8188% | 0.3211% | 1.497% | 0.5277% | 3.956% | 1.208% | 0.5225% | 0.2959% | 0.2409% | 4.096% | 0.6044% | 2.839% | 2.64% | 1.459% | 2.424% | 4.216% | 0.5887% | 1.449% | 0.6604% | [6 empty] | 0.8615% | 0.7795% | 1.529% | 1.334%
|  | % Recovery |  |  | 89.25% | 95.04% | 91.65% | 115.7% | 103.7% | 91.99% | 97.09% | 95.65% | 100.7% | 100.4% | 97.68% | 98.65% | 87.55% | 96.75% | 93.95% | 96.84% | 99.15% | 97.94% | 92.83% | 95.28% | 98.62% | 102.7% | 99.67% | 101.9% | 100.6% | 98.32% | 103% | 96.51% | 98.94% | 75.47% | 59.81% | 66.63% | 68.64% | 60.52% |  | 95.45% | 96.1% | 96.92% | 99.43%
| 148 | LFB |  | 06/17/2023 03:07:39 | 9123.3373 | 9345.9849 | 9.3734 | 11.6738 | 10722.5977 | 9204.8392 | 9932.1496 | 9.6151 | 10.052 | 10.2338 | 9.729 | 9.9726 | 9.1071 | 9.643 | 9.3195 | 9.6476 | 9.8048 | 9.8163 | 9.5112 | 9.7902 | 9.8163 | 10.2543 | 10.7167 | 10.0953 | 10.0937 | 10.3666 | 10.3098 | 9.5825 | 9.8779 | [6 empty] | 9.5213 | 9.5604 | 9.9174 | 10.4426
|  | RSD |  |  | 2.164% | 1.353% | 9.975% | 0.7427% | 3.509% | 1.359% | 3.178% | 1.218% | 3.29% | 1.362% | 0.2341% | 0.3317% | 1.898% | 2.1% | 1.477% | 1.677% | 1.568% | 0.5384% | 1.627% | 1.982% | 1.311% | 3.945% | 6.429% | 1.691% | 3.432% | 4.106% | 3.154% | 0.7472% | 0.7683% | [6 empty] | 1.39% | 0.6683% | 1.19% | 2.759%
|  | % Recovery |  |  | 91.23% | 93.46% | 93.73% | 116.7% | 107.2% | 92.05% | 99.32% | 96.15% | 100.5% | 102.3% | 97.29% | 99.73% | 91.07% | 96.43% | 93.19% | 96.48% | 98.05% | 98.16% | 95.11% | 97.9% | 98.16% | 102.5% | 107.2% | 101% | 100.9% | 103.7% | 103.1% | 95.82% | 98.78% | 71.49% | 55.88% | 62.55% | 61.51% | 56.6% |  | 95.21% | 95.6% | 99.17% | 104.4%
| 149 | LFB |  | 06/17/2023 03:13:57 | 8972.1227 | 9656.0419 | 9.6568 | 11.9956 | 10536.2556 | 9338.0058 | 9596.5639 | 9.7802 | 10.0446 | 10.2039 | 9.831 | 9.9603 | 9.1934 | 9.7123 | 9.1212 | 9.6313 | 9.8862 | 9.8124 | 9.3311 | 9.4447 | 9.7939 | 10.1977 | 9.7979 | 10.1906 | 10.1112 | 10.0706 | 10.0154 | 9.5099 | 9.9398 | [6 empty] | 9.5094 | 9.6022 | 9.6689 | 10.0364
|  | RSD |  |  | 0.2334% | 2.946% | 3.196% | 0.7212% | 2.494% | 1.163% | 2.823% | 2.842% | 1.315% | 1.48% | 0.8276% | 1.551% | 1.226% | 1.584% | 2.099% | 1.827% | 0.6081% | 0.8279% | 1.708% | 3.987% | 0.3363% | 1.376% | 6.936% | 0.4469% | 0.3965% | 5.398% | 4.839% | 1.274% | 0.6709% | [6 empty] | 1.862% | 1.167% | 1.193% | 1.674%
|  | % Recovery |  |  | 89.72% | 96.56% | 96.57% | 120% | 105.4% | 93.38% | 95.97% | 97.8% | 100.4% | 102% | 98.31% | 99.6% | 91.93% | 97.12% | 91.21% | 96.31% | 98.86% | 98.12% | 93.31% | 94.45% | 97.94% | 102% | 97.98% | 101.9% | 101.1% | 100.7% | 100.2% | 95.1% | 99.4% | 72.51% | 56.35% | 62.72% | 63.99% | 56.37% |  | 95.09% | 96.02% | 96.69% | 100.4%
| 182 | LFB |  | 06/17/2023 06:37:00 | 8970.3074 | 9183.3327 | 9.2986 | 11.5407 | 10881.8268 | 9176.0433 | 9770.5853 | 9.6765 | 10.1655 | 10.2248 | 9.854 | 9.8512 | 8.8593 | 9.5357 | 9.1662 | 9.5941 | 9.9232 | 10.0133 | 9.4709 | 9.8337 | 9.8297 | 9.9138 | 9.7686 | 9.9008 | 9.8871 | 10.1287 | 10.0335 | 9.6339 | 9.9816 | [6 empty] | 9.6126 | 9.6824 | 10.1483 | 10.3675
|  | RSD |  |  | 0.3726% | 1.303% | 8.569% | 2.075% | 2.091% | 1.642% | 0.7999% | 1.807% | 1.107% | 2.625% | 1.495% | 0.7119% | 2.885% | 0.5268% | 1.769% | 2.134% | 1.664% | 3.782% | 1.944% | 3.049% | 1.05% | 1.178% | 3.36% | 1.31% | 0.7824% | 2.271% | 2.5% | 1.138% | 0.99% | [6 empty] | 2.027% | 1.94% | 1.144% | 0.7658%
|  | % Recovery |  |  | 89.7% | 91.83% | 92.99% | 115.4% | 108.8% | 91.76% | 97.71% | 96.77% | 101.7% | 102.2% | 98.54% | 98.51% | 88.59% | 95.36% | 91.66% | 95.94% | 99.23% | 100.1% | 94.71% | 98.34% | 98.3% | 99.14% | 97.69% | 99.01% | 98.87% | 101.3% | 100.3% | 96.34% | 99.82% | 67.11% | 50.06% | 57.27% | 59.5% | 51.03% |  | 96.13% | 96.82% | 101.5% | 103.7%
| 183 | LFB |  | 06/17/2023 06:43:18 | 8867.5677 | 9124.608 | 9.0696 | 11.6353 | 10358.6309 | 9136.001 | 9792.1015 | 9.651 | 9.9827 | 9.9947 | 9.5324 | 9.8232 | 8.8429 | 9.5589 | 9.2102 | 9.5548 | 9.6558 | 9.4595 | 9.0273 | 9.728 | 9.669 | 9.5589 | 10.1294 | 9.8792 | 10.122 | 9.8282 | 9.9723 | 9.3787 | 9.762 | [6 empty] | 9.3795 | 9.4313 | 9.8528 | 11.2168
|  | RSD |  |  | 2.955% | 1.206% | 1.339% | 1.127% | 1.746% | 3.002% | 2.003% | 3.285% | 1.444% | 0.1806% | 0.5698% | 0.5927% | 2.184% | 1.29% | 1.443% | 2.79% | 0.5672% | 0.4506% | 1.061% | 4.665% | 1.291% | 2.641% | 5.212% | 3.663% | 0.5591% | 1.544% | 3.427% | 0.4208% | 0.697% | [6 empty] | 0.6314% | 0.4678% | 0.3102% | 5.154%
|  | % Recovery |  |  | 88.68% | 91.25% | 90.7% | 116.4% | 103.6% | 91.36% | 97.92% | 96.51% | 99.83% | 99.95% | 95.32% | 98.23% | 88.43% | 95.59% | 92.1% | 95.55% | 96.56% | 94.6% | 90.27% | 97.28% | 96.69% | 95.59% | 101.3% | 98.79% | 101.2% | 98.28% | 99.72% | 93.79% | 97.62% | 69.45% | 50.89% | 59.35% | 56.14% | 51.01% |  | 93.8% | 94.31% | 98.53% | 112.2%
| 216 | LFB |  | 06/17/2023 10:06:18 | 8604.3312 | 8848.0206 | 9.4861 | 10.8847 | 10437.6528 | 9135.7801 | 9755.5838 | 9.5967 | 9.8197 | 9.7091 | 9.2974 | 9.6264 | 8.5047 | 9.2071 | 8.9238 | 9.2904 | 9.4477 | 9.3339 | 9.072 | 9.2887 | 9.5888 | 9.8521 | 9.3578 | 9.4377 | 9.9006 | 9.5877 | 10.1963 | 9.3895 | 9.8947 | [6 empty] | 9.4288 | 9.514 | 10.5896 | 10.8361
|  | RSD |  |  | 2.546% | 2.686% | 8.492% | 0.3324% | 1.646% | 2.507% | 4.836% | 1.761% | 3.712% | 0.9378% | 1.141% | 0.6715% | 2.837% | 1.542% | 1.777% | 2.433% | 0.3666% | 0.7172% | 1.662% | 4.356% | 0.6179% | 5.393% | 9.387% | 2.174% | 2.604% | 1.503% | 7.005% | 1.206% | 0.6866% | [6 empty] | 0.6128% | 0.1444% | 0.8397% | 0.5003%
|  | % Recovery |  |  | 86.04% | 88.48% | 94.86% | 108.8% | 104.4% | 91.36% | 97.56% | 95.97% | 98.2% | 97.09% | 92.97% | 96.26% | 85.05% | 92.07% | 89.24% | 92.9% | 94.48% | 93.34% | 90.72% | 92.89% | 95.89% | 98.52% | 93.58% | 94.38% | 99.01% | 95.88% | 102% | 93.89% | 98.95% | 61.95% | 44.92% | 53.11% | 54.51% | 45.72% |  | 94.29% | 95.14% | 105.9% | 108.4%
| 217 | LFB |  | 06/17/2023 10:12:36 | 8488.4552 | 8863.8757 | 9.0709 | 11.3655 | 10232.5317 | 8918.4416 | 9541.5482 | 9.341 | 9.5082 | 10.0825 | 9.539 | 9.7991 | 8.522 | 9.1402 | 8.9403 | 9.3112 | 9.6675 | 9.5192 | 9.166 | 9.9727 | 9.709 | 10.0455 | 9.4721 | 9.5825 | 9.8053 | 10.0227 | 10.4228 | 9.6155 | 10.0292 | [6 empty] | 9.588 | 9.728 | 10.5883 | 10.7092
|  | RSD |  |  | 1.53% | 0.5538% | 10.36% | 1.963% | 5.573% | 1.134% | 1.6% | 2.04% | 1.389% | 3.166% | 4.739% | 1.192% | 1.613% | 0.7677% | 1.404% | 1.389% | 3.713% | 3.257% | 0.503% | 3.49% | 0.4911% | 4.786% | 9.76% | 4.339% | 3.261% | 3.503% | 6.01% | 3.565% | 0.4086% | [6 empty] | 4.553% | 4.469% | 2.098% | 0.5975%
|  | % Recovery |  |  | 84.88% | 88.64% | 90.71% | 113.7% | 102.3% | 89.18% | 95.42% | 93.41% | 95.08% | 100.8% | 95.39% | 97.99% | 85.22% | 91.4% | 89.4% | 93.11% | 96.67% | 95.19% | 91.66% | 99.73% | 97.09% | 100.5% | 94.72% | 95.82% | 98.05% | 100.2% | 104.2% | 96.15% | 100.3% | 62.44% | 45.73% | 53.05% | 55.54% | 46.1% |  | 95.88% | 97.28% | 105.9% | 107.1%
QC STD 5
|  | Sample Id | R | Acquisition Time | Na KED 23/23 | Mg KED 24/24 | Al KED 27/27 | P Oshift 31/47 | S (34) -1 34/34 | K KED 39/39 | Ca KED 43/43 | V KED 51/51 | Cr (52) 52/52 | Cr (53) NH3 53/53 | Fe (54) NH3 54/54 | Mn Oshift 55/71 | Fe (56) KED 56/56 | Co KED 59/59 | Ni KED 60/60 | Cu (63) KED 63/63 | Cu (65) NH3 65/65 | Zn NH3 66/66 | As 75/75 | As KED 75/75 | As Oshift 75/91 | Se (77) 77/77 | Se (77) Oshift 77/93 | Se (78) Oshift 78/94 | Se (80) Oshift 80/96 | Se (82) 82/82 | Se (82) Oshift 82/98 | Mo (95) NH3 95/95 | Mo (95) OShift 95/127 | Rh (IS1) 103/103 (IS) | Rh (IS2) 103/103 (IS) | Rh (IS4) 103/103 (IS) | Rh (IS5) 103/103 (IS) | Rh (IS6) 103/103 (IS) | Rh (IS7) 103/103 (IS) | Cd (111) NH3 111/111 | Cd (114) NH3 114/114 | Pb 208/208 | Pb NH3 208/208
| 22 | LLOQ 0.01 |  | 06/16/2023 14:12:05 | 68.7627 | 11.998 | 0.31844 | 1.2213 | 404.9005 | 74.1205 | 49.5998 | 0.011602 | 0.0069612 | 0.013291 | -0.064712 | 0.029726 | 0.26583 | 0.0096876 | 0.02016 | 0.01962 | 0.02022 | -0.0037104 | 0.05627 | 0.0034775 | 0.0077831 | 0.028379 | 0.019141 | 0.023973 | 9.0899e-05 | 0.031049 | 0.001602 | 0.050454 | 0.021506 | [6 empty] | 0.010088 | 0.0088365 | 0.0077222 | 0.013587
|  | RSD |  |  | 1.245% | 5.38% | 200.8% | 1.933% | 26.96% | 8.771% | 10.67% | 46.88% | 51.1% | 12.75% | 49.97% | 11.32% | 6.145% | 11.26% | 34.83% | 27.72% | 9.577% | 262% | 71.46% | 294.6% | 12.76% | 308.8% | 71.42% | 28.17% | 97.05% | 198% | 1910% | 2.543% | 9.055% | [6 empty] | 13.82% | 16.79% | 19.29% | 3.558%
|  | % Recovery |  |  | 687.6% | 120% | 3184% | 12213% | 4049% | 741.2% | 496% | 116% | 69.61% | 132.9% | -647.1% | 297.3% | 2658% | 96.88% | 201.6% | 196.2% | 202.2% | -37.1% | 562.7% | 34.77% | 77.83% | 283.8% | 191.4% | 239.7% | 0.909% | 310.5% | 16.02% | 504.5% | 215.1% | 93.13% | 89.42% | 92.11% | 92.22% | 87.86% |  | 100.9% | 88.37% | 77.22% | 135.9%
QC STD 6
|  | Sample Id | R | Acquisition Time | Na KED 23/23 | Mg KED 24/24 | Al KED 27/27 | P Oshift 31/47 | S (34) -1 34/34 | K KED 39/39 | Ca KED 43/43 | V KED 51/51 | Cr (52) 52/52 | Cr (53) NH3 53/53 | Fe (54) NH3 54/54 | Mn Oshift 55/71 | Fe (56) KED 56/56 | Co KED 59/59 | Ni KED 60/60 | Cu (63) KED 63/63 | Cu (65) NH3 65/65 | Zn NH3 66/66 | As 75/75 | As KED 75/75 | As Oshift 75/91 | Se (77) 77/77 | Se (77) Oshift 77/93 | Se (78) Oshift 78/94 | Se (80) Oshift 80/96 | Se (82) 82/82 | Se (82) Oshift 82/98 | Mo (95) NH3 95/95 | Mo (95) OShift 95/127 | Rh (IS1) 103/103 (IS) | Rh (IS2) 103/103 (IS) | Rh (IS4) 103/103 (IS) | Rh (IS5) 103/103 (IS) | Rh (IS6) 103/103 (IS) | Rh (IS7) 103/103 (IS) | Cd (111) NH3 111/111 | Cd (114) NH3 114/114 | Pb 208/208 | Pb NH3 208/208
| 23 | CCV 10 ppb MAJORS |  | 06/16/2023 14:18:05 | 10122.8889 | 10275.35 | 9.7477 | 9.1705 | 10255.1656 | 9972.5124 | 9698.068 | 10.1342 | 10.5512 | 10.4501 | 10.0215 | 10.4576 | 9.83 | 10.4451 | 10.2808 | 10.3056 | 10.327 | 10.4275 | 9.9744 | 10.2396 | 10.2403 | 10.5057 | 9.9803 | 9.9928 | 10.1894 | 10.2047 | 10.2849 | 10.1979 | 10.2561 | [6 empty] | 9.8422 | 9.8537 | 9.3505 | 9.6711
|  | RSD |  |  | 1.695% | 2.042% | 1.571% | 1.54% | 2.497% | 0.494% | 1.971% | 0.7571% | 0.4387% | 1.003% | 1.004% | 1.13% | 1.838% | 1.492% | 1.427% | 1.356% | 0.3781% | 0.8226% | 1.872% | 5.737% | 1.145% | 2.319% | 6.124% | 2.95% | 0.883% | 3.276% | 2.883% | 1.265% | 1.507% | [6 empty] | 0.2624% | 0.3827% | 0.3781% | 1.141%
|  | % Recovery |  |  | 101.2% | 102.8% | 97.48% | 91.7% | 102.6% | 99.73% | 96.98% | 101.3% | 105.5% | 104.5% | 100.2% | 104.6% | 98.3% | 104.5% | 102.8% | 103.1% | 103.3% | 104.3% | 99.74% | 102.4% | 102.4% | 105.1% | 99.8% | 99.93% | 101.9% | 102% | 102.8% | 102% | 102.6% | 88.31% | 81.2% | 83.3% | 85.97% | 82.26% |  | 98.42% | 98.54% | 93.5% | 96.71%
| 29 | CCV 10 ppb MAJORS |  | 06/16/2023 15:05:12 | 9998.191 | 10263.8003 | 8.857 | 9.0615 | 10259.6438 | 10055.9356 | 9931.5792 | 10.1704 | 10.489 | 10.3977 | 10.0399 | 10.3779 | 9.8225 | 10.2719 | 10.0946 | 10.1712 | 10.319 | 10.349 | 10.0115 | 10.2913 | 10.2486 | 10.1398 | 10.4099 | 10.2028 | 10.292 | 10.3317 | 10.3477 | 10.1689 | 10.3237 | [6 empty] | 9.8926 | 9.9207 | 9.3714 | 9.774
|  | RSD |  |  | 0.4342% | 2.684% | 9.996% | 0.6815% | 2.317% | 1.358% | 1.589% | 0.5377% | 0.7941% | 1.013% | 1.235% | 0.8408% | 1.068% | 1.479% | 1.743% | 1.758% | 1.334% | 2.05% | 1.984% | 1.771% | 0.9794% | 5.42% | 3.02% | 1.871% | 0.3719% | 3.44% | 0.9664% | 0.3917% | 0.6337% | [6 empty] | 0.9161% | 1.446% | 0.8602% | 1.518%
|  | % Recovery |  |  | 99.98% | 102.6% | 88.57% | 90.61% | 102.6% | 100.6% | 99.32% | 101.7% | 104.9% | 104% | 100.4% | 103.8% | 98.23% | 102.7% | 100.9% | 101.7% | 103.2% | 103.5% | 100.1% | 102.9% | 102.5% | 101.4% | 104.1% | 102% | 102.9% | 103.3% | 103.5% | 101.7% | 103.2% | 84.03% | 74.88% | 78.34% | 80.25% | 76.75% |  | 98.93% | 99.21% | 93.71% | 97.74%
| 40 | CCV 10 ppb MAJORS |  | 06/16/2023 16:11:33 | 9961.7153 | 10510.6335 | 9.7507 | 8.8572 | 10582.1499 | 10164.7328 | 9939.21 | 10.1587 | 10.4579 | 10.3163 | 10.2467 | 10.2249 | 9.8969 | 10.2916 | 9.9266 | 10.1153 | 10.131 | 10.3464 | 9.7164 | 10.2678 | 9.9965 | 10.2444 | 9.977 | 10.1643 | 10.0885 | 10.0647 | 10.1743 | 9.9111 | 10.0907 | [6 empty] | 9.8342 | 9.8947 | 9.6594 | 9.8891
|  | RSD |  |  | 1.535% | 2.383% | 3.285% | 1.847% | 4.852% | 4.303% | 1.925% | 2.621% | 1.894% | 0.6639% | 1.571% | 1.647% | 1.602% | 2.015% | 2.75% | 1.97% | 0.3064% | 0.6809% | 1.125% | 4.044% | 0.7572% | 2.395% | 5.466% | 3.089% | 1.737% | 4.882% | 2.827% | 0.6091% | 0.6696% | [6 empty] | 1.235% | 0.6602% | 0.9531% | 0.9137%
|  | % Recovery |  |  | 99.62% | 105.1% | 97.51% | 88.57% | 105.8% | 101.6% | 99.39% | 101.6% | 104.6% | 103.2% | 102.5% | 102.2% | 98.97% | 102.9% | 99.27% | 101.2% | 101.3% | 103.5% | 97.16% | 102.7% | 99.96% | 102.4% | 99.77% | 101.6% | 100.9% | 100.6% | 101.7% | 99.11% | 100.9% | 82.64% | 73.78% | 76.88% | 78.33% | 74.42% |  | 98.34% | 98.95% | 96.59% | 98.89%
| 41 | CCV 10 ppb MAJORS |  | 06/16/2023 16:17:52 | 10132.4338 | 10727.9424 | 9.4693 | 8.9241 | 10627.4135 | 10161.8632 | 9907.9138 | 10.2231 | 10.6301 | 10.4494 | 10.0884 | 10.3631 | 10.0249 | 10.3471 | 9.9207 | 10.238 | 10.3375 | 10.5247 | 9.926 | 9.9607 | 10.1212 | 10.1338 | 9.9371 | 10.052 | 9.8966 | 10.2714 | 10.1326 | 10.0575 | 10.1465 | [6 empty] | 9.7552 | 9.8221 | 9.5135 | 9.8035
|  | RSD |  |  | 1.112% | 1.095% | 5.305% | 2.297% | 4.474% | 1.888% | 0.9133% | 2.681% | 1.103% | 1.539% | 1.238% | 0.9717% | 1.512% | 1.66% | 1.621% | 1.297% | 1.353% | 0.5107% | 1.301% | 2.171% | 1.621% | 0.2758% | 7.471% | 1.714% | 2.408% | 2.509% | 3.996% | 1.615% | 1.114% | [6 empty] | 1.158% | 0.8829% | 1.157% | 2.216%
|  | % Recovery |  |  | 101.3% | 107.3% | 94.69% | 89.24% | 106.3% | 101.6% | 99.08% | 102.2% | 106.3% | 104.5% | 100.9% | 103.6% | 100.2% | 103.5% | 99.21% | 102.4% | 103.4% | 105.2% | 99.26% | 99.61% | 101.2% | 101.3% | 99.37% | 100.5% | 98.97% | 102.7% | 101.3% | 100.6% | 101.5% | 83.34% | 74.42% | 77.21% | 79.09% | 75.88% |  | 97.55% | 98.22% | 95.13% | 98.04%
| 56 | CCV 10 ppb MAJORS |  | 06/16/2023 17:48:11 | 9642.9197 | 10168.3327 | 9.0177 | 8.9859 | 10240.0218 | 9932.3643 | 9801.3678 | 9.9058 | 10.2795 | 10.4564 | 10.2546 | 10.287 | 9.766 | 10.216 | 9.7392 | 10.0267 | 10.2615 | 10.4433 | 9.5325 | 9.8829 | 9.9566 | 10.3203 | 9.9852 | 9.9118 | 10.1161 | 9.7431 | 9.7936 | 10.0474 | 10.2082 | [6 empty] | 9.7785 | 9.8097 | 9.5589 | 9.9094
|  | RSD |  |  | 0.9674% | 0.7618% | 3.171% | 2.339% | 0.5388% | 1.966% | 2.246% | 1.596% | 1.96% | 0.8779% | 1.475% | 0.1265% | 0.8655% | 1.808% | 1.36% | 1.066% | 0.6891% | 0.1179% | 1.27% | 1.529% | 0.6102% | 2.065% | 2.425% | 1.878% | 1.028% | 2.458% | 2.368% | 1.008% | 0.314% | [6 empty] | 0.3958% | 0.655% | 1.588% | 2.124%
|  | % Recovery |  |  | 96.43% | 101.7% | 90.18% | 89.86% | 102.4% | 99.32% | 98.01% | 99.06% | 102.8% | 104.6% | 102.5% | 102.9% | 97.66% | 102.2% | 97.39% | 100.3% | 102.6% | 104.4% | 95.32% | 98.83% | 99.57% | 103.2% | 99.85% | 99.12% | 101.2% | 97.43% | 97.94% | 100.5% | 102.1% | 82.99% | 71.63% | 76.33% | 77.5% | 72.42% |  | 97.78% | 98.1% | 95.59% | 99.09%
| 57 | CCV 10 ppb MAJORS |  | 06/16/2023 17:54:30 | 9842.1164 | 10306.0446 | 9.6058 | 9.0394 | 10768.3658 | 9919.952 | 10003.6006 | 10.227 | 10.4489 | 10.5586 | 10.1865 | 10.2453 | 9.9996 | 10.4423 | 10.1498 | 10.3751 | 10.3184 | 10.385 | 9.8983 | 9.9591 | 9.8456 | 10.486 | 9.6585 | 10.0605 | 9.9058 | 10.3007 | 10.396 | 10.0906 | 10.1703 | [6 empty] | 9.6967 | 9.7855 | 9.6044 | 9.8314
|  | RSD |  |  | 2.143% | 2.367% | 4.505% | 0.8531% | 1.112% | 1.116% | 2.768% | 0.6755% | 1.611% | 1.134% | 1.451% | 2.294% | 1.667% | 1.117% | 2.696% | 0.9469% | 0.4459% | 0.6287% | 2.948% | 0.2651% | 1.583% | 3.728% | 5.637% | 3.164% | 1.791% | 1.494% | 3.231% | 0.893% | 1.623% | [6 empty] | 1.033% | 0.42% | 2.247% | 1.692%
|  | % Recovery |  |  | 98.42% | 103.1% | 96.06% | 90.39% | 107.7% | 99.2% | 100% | 102.3% | 104.5% | 105.6% | 101.9% | 102.5% | 100% | 104.4% | 101.5% | 103.8% | 103.2% | 103.9% | 98.98% | 99.59% | 98.46% | 104.9% | 96.59% | 100.6% | 99.06% | 103% | 104% | 100.9% | 101.7% | 83.53% | 73.2% | 76.77% | 77.85% | 73.62% |  | 96.97% | 97.86% | 96.04% | 98.31%
| 71 | CCV 10 ppb MAJORS |  | 06/16/2023 19:19:08 | 9635.2676 | 10151.3372 | 9.4319 | 8.8801 | 10594.1633 | 9744.4391 | 9907.7371 | 9.8795 | 10.1958 | 10.4225 | 10.0754 | 10.2562 | 9.6368 | 10.0051 | 9.6712 | 10.0261 | 10.1498 | 10.3049 | 9.8162 | 9.9649 | 9.9857 | 10.6706 | 9.934 | 10.0132 | 10.0792 | 10.4874 | 9.8281 | 9.9381 | 10.214 | [6 empty] | 9.6553 | 9.7762 | 10.0501 | 10.1056
|  | RSD |  |  | 0.2673% | 1.216% | 2.956% | 1.233% | 2.981% | 1.18% | 1.701% | 0.8202% | 1.177% | 0.3951% | 0.4945% | 0.7408% | 2.243% | 1.373% | 1.642% | 1.687% | 0.3238% | 0.8813% | 1.695% | 3.015% | 0.5142% | 5.825% | 4.268% | 2.688% | 2.16% | 4.038% | 3.417% | 1.514% | 0.1612% | [6 empty] | 1.223% | 1.082% | 2.325% | 0.8533%
|  | % Recovery |  |  | 96.35% | 101.5% | 94.32% | 88.8% | 105.9% | 97.44% | 99.08% | 98.8% | 102% | 104.2% | 100.8% | 102.6% | 96.37% | 100.1% | 96.71% | 100.3% | 101.5% | 103% | 98.16% | 99.65% | 99.86% | 106.7% | 99.34% | 100.1% | 100.8% | 104.9% | 98.28% | 99.38% | 102.1% | 78.21% | 66.14% | 71.46% | 73.35% | 66.94% |  | 96.55% | 97.76% | 100.5% | 101.1%
| 72 | CCV 10 ppb MAJORS |  | 06/16/2023 19:25:27 | 9689.1906 | 9953.3571 | 9.403 | 8.9625 | 10606.1927 | 9782.6699 | 9988.6639 | 10.0685 | 10.5402 | 10.6579 | 10.2485 | 10.4787 | 9.9744 | 10.2245 | 9.8293 | 10.151 | 10.2338 | 10.3994 | 9.7787 | 10.1096 | 10.0639 | 10.5556 | 10.3708 | 10.2393 | 10.1633 | 10.5873 | 10.6943 | 10.0407 | 10.2687 | [6 empty] | 9.7836 | 9.7555 | 9.7335 | 9.9283
|  | RSD |  |  | 2.963% | 3.026% | 6.918% | 1.823% | 1.906% | 2.18% | 2.537% | 0.6022% | 2.094% | 1.828% | 1.666% | 0.7612% | 2.671% | 1.309% | 1.735% | 0.9542% | 0.3518% | 0.7785% | 1.726% | 1.137% | 0.9207% | 1.092% | 2.749% | 2.122% | 1.33% | 2.723% | 4.202% | 0.5111% | 1.202% | [6 empty] | 1.634% | 0.3777% | 2.443% | 0.9257%
|  | % Recovery |  |  | 96.89% | 99.53% | 94.03% | 89.62% | 106.1% | 97.83% | 99.89% | 100.7% | 105.4% | 106.6% | 102.5% | 104.8% | 99.74% | 102.2% | 98.29% | 101.5% | 102.3% | 104% | 97.79% | 101.1% | 100.6% | 105.6% | 103.7% | 102.4% | 101.6% | 105.9% | 106.9% | 100.4% | 102.7% | 80.29% | 67.91% | 72.92% | 75.02% | 68.57% |  | 97.84% | 97.56% | 97.33% | 99.28%
| 87 | CCV 10 ppb MAJORS |  | 06/16/2023 20:55:46 | 9907.5254 | 10350.9786 | 9.9565 | 9.377 | 10682.0945 | 9945.1814 | 10117.4077 | 10.2223 | 10.5432 | 10.7419 | 10.3071 | 10.4417 | 9.7082 | 10.1628 | 9.7941 | 10.0052 | 10.3036 | 10.5344 | 9.7441 | 10.2478 | 9.9968 | 10.3172 | 10.0712 | 10.1348 | 10.1377 | 10.582 | 10.2023 | 9.9813 | 10.2245 | [6 empty] | 9.7546 | 9.7844 | 9.7097 | 9.9626
|  | RSD |  |  | 1.986% | 2.077% | 5.835% | 0.2603% | 2.208% | 1.176% | 1.583% | 0.6887% | 0.9248% | 1.415% | 0.7486% | 0.7985% | 0.8694% | 0.579% | 0.9823% | 0.5914% | 0.9213% | 0.375% | 1.526% | 1.63% | 0.5155% | 5.455% | 2.69% | 2.557% | 1.252% | 3.559% | 3.112% | 1.465% | 0.8863% | [6 empty] | 0.4793% | 0.3063% | 0.6561% | 2.009%
|  | % Recovery |  |  | 99.08% | 103.5% | 99.57% | 93.77% | 106.8% | 99.45% | 101.2% | 102.2% | 105.4% | 107.4% | 103.1% | 104.4% | 97.08% | 101.6% | 97.94% | 100.1% | 103% | 105.3% | 97.44% | 102.5% | 99.97% | 103.2% | 100.7% | 101.3% | 101.4% | 105.8% | 102% | 99.81% | 102.2% | 79.38% | 66.35% | 70.25% | 72.9% | 67.44% |  | 97.55% | 97.84% | 97.1% | 99.63%
| 101 | CCV 10 ppb MAJORS |  | 06/16/2023 22:20:23 | 9730.104 | 10148.7696 | 9.5046 | 9.0768 | 10884.6744 | 9980.2188 | 9967.1371 | 10.1639 | 10.6508 | 10.6538 | 10.304 | 10.5466 | 9.7354 | 10.2861 | 9.7647 | 10.1525 | 10.228 | 10.429 | 9.6983 | 9.8855 | 10.1663 | 10.4224 | 10.3729 | 10.0838 | 9.9256 | 10.2341 | 10.3914 | 9.995 | 10.3779 | [6 empty] | 9.7695 | 9.8099 | 9.8755 | 10.1874
|  | RSD |  |  | 2.429% | 3.628% | 11.06% | 0.309% | 1.244% | 2.828% | 2.14% | 3.612% | 1.086% | 0.1597% | 2.146% | 1.286% | 3.59% | 2% | 0.978% | 2.29% | 0.3346% | 1.08% | 0.8617% | 5.029% | 0.7% | 0.6303% | 2.068% | 2.901% | 2.523% | 1.705% | 4.471% | 0.6719% | 1.543% | [6 empty] | 0.882% | 0.3755% | 1.594% | 0.8672%
|  | % Recovery |  |  | 97.3% | 101.5% | 95.05% | 90.77% | 108.8% | 99.8% | 99.67% | 101.6% | 106.5% | 106.5% | 103% | 105.5% | 97.35% | 102.9% | 97.65% | 101.5% | 102.3% | 104.3% | 96.98% | 98.86% | 101.7% | 104.2% | 103.7% | 100.8% | 99.26% | 102.3% | 103.9% | 99.95% | 103.8% | 76.72% | 61.69% | 67.72% | 68.64% | 62.94% |  | 97.7% | 98.1% | 98.75% | 101.9%
| 118 | CCV 10 ppb MAJORS |  | 06/17/2023 00:03:55 | 9146.4813 | 9732.3602 | 9.1596 | 9.2258 | 10290.665 | 9419.1562 | 9950.6236 | 9.838 | 10.2238 | 10.3956 | 10.111 | 10.221 | 9.3725 | 9.9119 | 9.504 | 9.8709 | 10.1029 | 10.2163 | 9.8929 | 9.2593 | 9.9075 | 10.9237 | 9.476 | 9.8004 | 10.0228 | 10.2829 | 10.5384 | 9.9149 | 10.2118 | [6 empty] | 9.6348 | 9.7148 | 10.4492 | 10.3041
|  | RSD |  |  | 0.9527% | 0.9423% | 6.881% | 0.8498% | 3.458% | 1.298% | 1.486% | 0.9653% | 1.204% | 0.3609% | 1.36% | 0.3994% | 0.8611% | 1.662% | 1.682% | 2.276% | 0.1869% | 1.132% | 2.46% | 3.53% | 0.8935% | 2.816% | 3.105% | 0.5264% | 2.644% | 3.972% | 3.936% | 0.2169% | 0.9715% | [6 empty] | 0.3733% | 0.7945% | 0.7553% | 1.26%
|  | % Recovery |  |  | 91.46% | 97.32% | 91.6% | 92.26% | 102.9% | 94.19% | 99.51% | 98.38% | 102.2% | 104% | 101.1% | 102.2% | 93.72% | 99.12% | 95.04% | 98.71% | 101% | 102.2% | 98.93% | 92.59% | 99.08% | 109.2% | 94.76% | 98% | 100.2% | 102.8% | 105.4% | 99.15% | 102.1% | 73.91% | 60.71% | 67.92% | 69.11% | 61.08% |  | 96.35% | 97.15% | 104.5% | 103%
| 133 | CCV 10 ppb MAJORS |  | 06/17/2023 01:35:32 | 9198.8992 | 9642.5842 | 8.8068 | 8.8042 | 10567.5571 | 9467.1743 | 10040.5209 | 9.9471 | 10.2946 | 10.3894 | 10.2326 | 10.2764 | 9.1999 | 9.9296 | 9.4652 | 9.813 | 9.9928 | 10.2096 | 9.5101 | 10.0143 | 9.8402 | 10.1058 | 9.9663 | 9.8738 | 9.9824 | 10.0215 | 9.9297 | 9.8692 | 10.3203 | [6 empty] | 9.7451 | 9.8086 | 10.3905 | 10.5557
|  | RSD |  |  | 2.002% | 2.445% | 7.702% | 1.204% | 2.307% | 1.067% | 2.269% | 1.838% | 1.383% | 0.4253% | 1.793% | 1.522% | 1.676% | 0.4678% | 2.687% | 2.348% | 0.4027% | 1.332% | 2.144% | 1.954% | 0.2037% | 2.164% | 4.884% | 4.564% | 1.115% | 0.4638% | 2.866% | 0.3259% | 0.4193% | [6 empty] | 0.6582% | 0.7845% | 2.857% | 0.5708%
|  | % Recovery |  |  | 91.99% | 96.43% | 88.07% | 88.04% | 105.7% | 94.67% | 100.4% | 99.47% | 102.9% | 103.9% | 102.3% | 102.8% | 92% | 99.3% | 94.65% | 98.13% | 99.93% | 102.1% | 95.1% | 100.1% | 98.4% | 101.1% | 99.66% | 98.74% | 99.82% | 100.2% | 99.3% | 98.69% | 103.2% | 71.84% | 55.53% | 63.03% | 64.79% | 55.81% |  | 97.45% | 98.09% | 103.9% | 105.6%
| 134 | CCV 10 ppb MAJORS |  | 06/17/2023 01:41:51 | 9210.9318 | 9666.5118 | 9.5342 | 8.785 | 10531.8412 | 9347.6247 | 10086.3066 | 9.9852 | 10.1448 | 10.6126 | 10.2737 | 10.2175 | 9.2229 | 9.8789 | 9.5724 | 9.6853 | 10.1415 | 10.2698 | 9.4948 | 9.719 | 9.8399 | 10.1636 | 10.0204 | 9.9062 | 9.97 | 10.3088 | 10.1556 | 9.9226 | 10.2306 | [6 empty] | 9.7404 | 9.771 | 10.2194 | 10.4249
|  | RSD |  |  | 1.866% | 0.4532% | 4.581% | 0.8886% | 0.7322% | 0.2673% | 0.9153% | 0.3192% | 1.01% | 3.602% | 4.058% | 0.7115% | 0.669% | 0.8262% | 2.204% | 1.517% | 3.469% | 4.398% | 1.709% | 5.412% | 0.8446% | 2.946% | 4.203% | 3.207% | 1.788% | 3.753% | 5.236% | 3.818% | 1.194% | [6 empty] | 5.009% | 4.018% | 0.774% | 1.25%
|  | % Recovery |  |  | 92.11% | 96.67% | 95.34% | 87.85% | 105.3% | 93.48% | 100.9% | 99.85% | 101.4% | 106.1% | 102.7% | 102.2% | 92.23% | 98.79% | 95.72% | 96.85% | 101.4% | 102.7% | 94.95% | 97.19% | 98.4% | 101.6% | 100.2% | 99.06% | 99.7% | 103.1% | 101.6% | 99.23% | 102.3% | 74.38% | 57.44% | 64.92% | 66.46% | 57.54% |  | 97.4% | 97.71% | 102.2% | 104.2%
| 150 | CCV 10 ppb MAJORS |  | 06/17/2023 03:19:56 | 9501.9728 | 10048.5273 | 9.3208 | 9.5016 | 10864.3158 | 9851.0134 | 10385.0432 | 10.3208 | 10.5674 | 10.6308 | 10.5023 | 10.3868 | 9.8696 | 10.1599 | 9.681 | 10.0892 | 10.2151 | 10.3869 | 9.4912 | 10.348 | 9.9173 | 10.5936 | 9.8965 | 10.0221 | 10.1009 | 10.2473 | 10.755 | 9.902 | 10.3375 | [6 empty] | 9.7112 | 9.7188 | 10.031 | 10.3503
|  | RSD |  |  | 3.333% | 1.876% | 3.872% | 1.544% | 5.465% | 1.088% | 4.909% | 1.963% | 2.921% | 0.7549% | 0.9334% | 0.0879% | 1.681% | 3.266% | 1.874% | 2.25% | 0.5504% | 0.221% | 1.462% | 8.065% | 0.3989% | 1.753% | 9.072% | 2.095% | 2.063% | 2.149% | 1.081% | 0.3528% | 0.7392% | [6 empty] | 0.2161% | 0.9853% | 0.8047% | 1.446%
|  | % Recovery |  |  | 95.02% | 100.5% | 93.21% | 95.02% | 108.6% | 98.51% | 103.9% | 103.2% | 105.7% | 106.3% | 105% | 103.9% | 98.7% | 101.6% | 96.81% | 100.9% | 102.2% | 103.9% | 94.91% | 103.5% | 99.17% | 105.9% | 98.96% | 100.2% | 101% | 102.5% | 107.5% | 99.02% | 103.4% | 73.59% | 56.86% | 63.74% | 65.15% | 57.57% |  | 97.11% | 97.19% | 100.3% | 103.5%
| 151 | CCV 10 ppb MAJORS |  | 06/17/2023 03:26:14 | 9456.6663 | 9965.1394 | 9.5303 | 9.2573 | 10804.9655 | 9722.0743 | 10241.1771 | 10.3568 | 10.5133 | 10.7993 | 10.4783 | 10.5147 | 9.6045 | 10.1885 | 9.6873 | 9.9861 | 10.2508 | 10.472 | 9.3425 | 9.7767 | 9.9689 | 10.1017 | 9.7904 | 9.905 | 10.1544 | 9.6905 | 10.3999 | 10.0468 | 10.3589 | [6 empty] | 9.7324 | 9.8083 | 9.7941 | 10.3318
|  | RSD |  |  | 3.109% | 4.693% | 5.385% | 2.507% | 1.149% | 1.808% | 2.967% | 1.002% | 1.221% | 0.3388% | 1.731% | 1.1% | 1.744% | 3.286% | 2.689% | 2.031% | 0.204% | 1.486% | 3.749% | 1.685% | 1.079% | 6.684% | 5.421% | 2.696% | 3.142% | 6.179% | 0.6847% | 0.3929% | 0.6422% | [6 empty] | 0.6004% | 1.088% | 4.757% | 0.6702%
|  | % Recovery |  |  | 94.57% | 99.65% | 95.3% | 92.57% | 108% | 97.22% | 102.4% | 103.6% | 105.1% | 108% | 104.8% | 105.1% | 96.04% | 101.9% | 96.87% | 99.86% | 102.5% | 104.7% | 93.42% | 97.77% | 99.69% | 101% | 97.9% | 99.05% | 101.5% | 96.9% | 104% | 100.5% | 103.6% | 75.33% | 57.22% | 63.82% | 65.71% | 58.15% |  | 97.32% | 98.08% | 97.94% | 103.3%
| 166 | CCV 10 ppb MAJORS |  | 06/17/2023 04:58:31 | 9082.5039 | 9580.7748 | 8.9554 | 9.0381 | 10710.8149 | 9556.8428 | 10161.4654 | 9.8986 | 10.3285 | 10.4954 | 10.2133 | 10.198 | 8.9676 | 9.7084 | 9.3707 | 9.7398 | 10.0391 | 10.1715 | 9.4579 | 9.7517 | 9.8535 | 10.3911 | 9.875 | 9.7636 | 9.9566 | 10.2759 | 9.9948 | 9.8994 | 10.2608 | [6 empty] | 9.734 | 9.7572 | 10.5532 | 10.797
|  | RSD |  |  | 2.277% | 1.348% | 8.512% | 1.552% | 3.514% | 1.806% | 4.008% | 2.343% | 2.551% | 1.316% | 1.272% | 0.4426% | 1.211% | 0.2009% | 2.314% | 0.5512% | 1.106% | 0.8839% | 0.6866% | 3.939% | 0.3463% | 1.515% | 3.727% | 4.339% | 3.421% | 2.597% | 6.439% | 0.383% | 0.8137% | [6 empty] | 0.6313% | 0.7514% | 1.769% | 0.996%
|  | % Recovery |  |  | 90.83% | 95.81% | 89.55% | 90.38% | 107.1% | 95.57% | 101.6% | 98.99% | 103.3% | 105% | 102.1% | 102% | 89.68% | 97.08% | 93.71% | 97.4% | 100.4% | 101.7% | 94.58% | 97.52% | 98.54% | 103.9% | 98.75% | 97.64% | 99.57% | 102.8% | 99.95% | 98.99% | 102.6% | 68.97% | 52.07% | 59.53% | 60.56% | 52.49% |  | 97.34% | 97.57% | 105.5% | 108%
| 167 | CCV 10 ppb MAJORS |  | 06/17/2023 05:04:50 | 9381.6047 | 9735.7882 | 9.5173 | 8.783 | 10954.4274 | 9589.6856 | 10173.9004 | 10.0587 | 10.435 | 10.4507 | 10.1238 | 10.2891 | 9.4533 | 9.9703 | 9.6684 | 9.7898 | 10.0039 | 10.2206 | 9.5555 | 10.1011 | 9.9143 | 10.3912 | 10.362 | 9.79 | 9.9271 | 10.0439 | 10.2064 | 9.9061 | 10.2198 | [6 empty] | 9.702 | 9.7842 | 10.4492 | 10.6178
|  | RSD |  |  | 1.833% | 3.413% | 1.964% | 2.001% | 2.701% | 2.558% | 1.564% | 2.562% | 1.531% | 1.134% | 0.9041% | 0.7042% | 2.8% | 3.879% | 5.023% | 2.24% | 0.9808% | 1.264% | 1.439% | 4.418% | 1.043% | 4.172% | 8.022% | 0.6407% | 1.412% | 1.656% | 4.412% | 0.9118% | 0.7109% | [6 empty] | 0.9409% | 0.7783% | 3.307% | 0.4813%
|  | % Recovery |  |  | 93.82% | 97.36% | 95.17% | 87.83% | 109.5% | 95.9% | 101.7% | 100.6% | 104.3% | 104.5% | 101.2% | 102.9% | 94.53% | 99.7% | 96.68% | 97.9% | 100% | 102.2% | 95.55% | 101% | 99.14% | 103.9% | 103.6% | 97.9% | 99.27% | 100.4% | 102.1% | 99.06% | 102.2% | 70.01% | 53.02% | 60.58% | 61.85% | 53.53% |  | 97.02% | 97.84% | 104.5% | 106.2%
| 184 | CCV 10 ppb MAJORS |  | 06/17/2023 06:49:17 | 9306.831 | 9713.4271 | 9.1386 | 9.2248 | 10647.9564 | 9569.9926 | 10321.301 | 10.2197 | 10.4124 | 10.5272 | 10.534 | 10.2872 | 9.3711 | 9.8855 | 9.4574 | 9.9306 | 10.1239 | 10.3161 | 9.2763 | 9.8277 | 9.9325 | 10.1599 | 9.835 | 10.2077 | 9.7385 | 10.0393 | 10.3327 | 9.9262 | 10.3355 | [6 empty] | 9.7562 | 9.8718 | 10.3118 | 10.5736
|  | RSD |  |  | 1.247% | 1.262% | 5.364% | 0.677% | 1.978% | 1.487% | 2.696% | 1.147% | 0.4319% | 1.366% | 9.036% | 1.02% | 3.459% | 0.912% | 1.451% | 1.109% | 0.6885% | 0.1334% | 3.512% | 1.165% | 0.5967% | 3.82% | 5.935% | 2.004% | 0.9336% | 4.081% | 3.311% | 0.347% | 0.3777% | [6 empty] | 1.672% | 0.5653% | 2.366% | 1.481%
|  | % Recovery |  |  | 93.07% | 97.13% | 91.39% | 92.25% | 106.5% | 95.7% | 103.2% | 102.2% | 104.1% | 105.3% | 105.3% | 102.9% | 93.71% | 98.85% | 94.57% | 99.31% | 101.2% | 103.2% | 92.76% | 98.28% | 99.32% | 101.6% | 98.35% | 102.1% | 97.38% | 100.4% | 103.3% | 99.26% | 103.4% | 69.18% | 51.46% | 59.22% | 61.5% | 52.07% |  | 97.56% | 98.72% | 103.1% | 105.7%
| 185 | CCV 10 ppb MAJORS |  | 06/17/2023 06:55:37 | 9321.9399 | 9621.4959 | 9.6796 | 9.0552 | 10832.7867 | 9651.5489 | 10077.7649 | 10.1472 | 10.5289 | 10.5721 | 10.2148 | 10.5022 | 9.5828 | 9.9356 | 9.5294 | 9.8754 | 10.1215 | 10.1735 | 9.5486 | 10.0409 | 9.8948 | 10.1073 | 10.3434 | 10.1782 | 9.9266 | 10.2306 | 10.2963 | 9.9969 | 10.3937 | [6 empty] | 9.6756 | 9.772 | 10.4837 | 10.7976
|  | RSD |  |  | 2.606% | 5.241% | 7.575% | 2.58% | 3.388% | 1.252% | 2.365% | 2.42% | 2.161% | 0.1906% | 2.071% | 1.025% | 3.942% | 1.828% | 1.642% | 1.254% | 0.4354% | 1.196% | 2.376% | 8.08% | 0.985% | 1.883% | 4.695% | 2.268% | 0.3783% | 1.494% | 5.263% | 0.5068% | 0.3829% | [6 empty] | 0.7469% | 1.158% | 1.09% | 1.055%
|  | % Recovery |  |  | 93.22% | 96.21% | 96.8% | 90.55% | 108.3% | 96.52% | 100.8% | 101.5% | 105.3% | 105.7% | 102.1% | 105% | 95.83% | 99.36% | 95.29% | 98.75% | 101.2% | 101.7% | 95.49% | 100.4% | 98.95% | 101.1% | 103.4% | 101.8% | 99.27% | 102.3% | 103% | 99.97% | 103.9% | 69.38% | 51.94% | 60.3% | 61.29% | 52.48% |  | 96.76% | 97.72% | 104.8% | 108%
| 200 | CCV 10 ppb MAJORS |  | 06/17/2023 08:27:48 | 9196.0967 | 9579.6142 | 9.6821 | 8.8967 | 11025.4809 | 9642.3472 | 10209.7829 | 9.8544 | 10.2666 | 10.5153 | 9.984 | 10.2443 | 9.1805 | 9.7764 | 9.3451 | 9.8159 | 9.9135 | 10.1591 | 9.4528 | 9.8617 | 9.9315 | 10.4755 | 10.0874 | 10.0847 | 10.1008 | 10.2158 | 10.2704 | 10.0021 | 10.3519 | [6 empty] | 9.8043 | 9.8706 | 10.8231 | 11.1097
|  | RSD |  |  | 0.2681% | 1.064% | 10.09% | 1.69% | 1.779% | 1.603% | 3.494% | 0.8177% | 1.589% | 0.4368% | 1.107% | 0.3435% | 1.668% | 0.5773% | 2.167% | 0.6747% | 0.3925% | 0.1129% | 1.784% | 6.14% | 1.169% | 0.4529% | 5.329% | 1.069% | 2.031% | 1.759% | 0.5253% | 0.3159% | 1.73% | [6 empty] | 0.5169% | 0.3038% | 1.528% | 0.8682%
|  | % Recovery |  |  | 91.96% | 95.8% | 96.82% | 88.97% | 110.3% | 96.42% | 102.1% | 98.54% | 102.7% | 105.2% | 99.84% | 102.4% | 91.81% | 97.76% | 93.45% | 98.16% | 99.13% | 101.6% | 94.53% | 98.62% | 99.31% | 104.8% | 100.9% | 100.8% | 101% | 102.2% | 102.7% | 100% | 103.5% | 64.18% | 47.35% | 55.14% | 56.66% | 47.89% |  | 98.04% | 98.71% | 108.2% | 111.1%
| 201 | CCV 10 ppb MAJORS |  | 06/17/2023 08:34:07 | 9343.0282 | 9610.9488 | 9.238 | 8.8704 | 10996.6849 | 9457.8814 | 10283.2552 | 10.1219 | 10.5213 | 10.4927 | 10.1373 | 10.3406 | 9.3419 | 9.8617 | 9.4333 | 9.8444 | 10.0404 | 10.1828 | 8.909 | 10.0913 | 9.973 | 9.5723 | 10.678 | 9.9592 | 10.05 | 10.1431 | 10.2223 | 9.9624 | 10.4288 | [6 empty] | 9.759 | 9.8639 | 10.1266 | 11.0322
|  | RSD |  |  | 1.339% | 1.933% | 4.406% | 2.834% | 0.9372% | 1.104% | 1.375% | 1.505% | 1.339% | 1.003% | 2.275% | 1.669% | 0.9572% | 2.216% | 3.986% | 0.7517% | 0.9434% | 1.361% | 2.723% | 3.078% | 1.196% | 4.579% | 3.112% | 5.701% | 1.852% | 3.691% | 3.079% | 0.7197% | 1.974% | [6 empty] | 0.3928% | 0.8562% | 4.919% | 1.363%
|  | % Recovery |  |  | 93.43% | 96.11% | 92.38% | 88.7% | 110% | 94.58% | 102.8% | 101.2% | 105.2% | 104.9% | 101.4% | 103.4% | 93.42% | 98.62% | 94.33% | 98.44% | 100.4% | 101.8% | 89.09% | 100.9% | 99.73% | 95.72% | 106.8% | 99.59% | 100.5% | 101.4% | 102.2% | 99.62% | 104.3% | 69.06% | 48.3% | 56.36% | 57.59% | 48.85% |  | 97.59% | 98.64% | 101.3% | 110.3%
| 218 | CCV 10 ppb MAJORS |  | 06/17/2023 10:18:36 | 9088.8882 | 9458.5212 | 9.7384 | 9.1108 | 11032.9994 | 9584.4373 | 10268.8251 | 10.0607 | 10.3768 | 10.457 | 10.1439 | 10.2826 | 9.2199 | 9.7878 | 9.2477 | 9.7816 | 10.0346 | 9.9822 | 9.4649 | 9.426 | 9.8989 | 10.2056 | 10.015 | 10.1258 | 9.9487 | 10.0483 | 10.2024 | 9.9111 | 10.431 | [6 empty] | 9.6437 | 9.7748 | 10.8962 | 11.1018
|  | RSD |  |  | 2.687% | 0.8758% | 0.6871% | 1.208% | 2.751% | 1.488% | 2.113% | 2.2% | 2.158% | 1.167% | 1.262% | 1.536% | 2.854% | 1.104% | 1.042% | 2.547% | 0.6427% | 1.155% | 0.5745% | 4.628% | 1.095% | 0.8475% | 2.878% | 1.346% | 2.301% | 1.847% | 2.357% | 0.6816% | 1.139% | [6 empty] | 1.437% | 1.729% | 1.443% | 1.189%
|  | % Recovery |  |  | 90.89% | 94.59% | 97.38% | 91.11% | 110.3% | 95.84% | 102.7% | 100.6% | 103.8% | 104.6% | 101.4% | 102.8% | 92.2% | 97.88% | 92.48% | 97.82% | 100.3% | 99.82% | 94.65% | 94.26% | 98.99% | 102.1% | 100.2% | 101.3% | 99.49% | 100.5% | 102% | 99.11% | 104.3% | 64.34% | 46.69% | 54.84% | 56.92% | 47.11% |  | 96.44% | 97.75% | 109% | 111%
| 219 | CCV 10 ppb MAJORS |  | 06/17/2023 10:24:55 | 9136.6389 | 9344.4324 | 9.3527 | 8.8583 | 11013.0283 | 9452.938 | 10278.224 | 10.0455 | 10.1677 | 10.581 | 10.0385 | 10.2585 | 9.2147 | 9.7741 | 9.4132 | 9.7052 | 10.0953 | 10.1275 | 9.4815 | 9.6238 | 9.8198 | 10.1195 | 10.0042 | 9.8117 | 9.8013 | 10.4542 | 10.049 | 9.9984 | 10.4653 | [6 empty] | 9.7188 | 9.8832 | 10.7595 | 10.8331
|  | RSD |  |  | 2.446% | 2.865% | 4.48% | 1.508% | 3.839% | 1.97% | 3.176% | 2.503% | 1.561% | 1.41% | 1.197% | 0.3845% | 2.583% | 3.213% | 2.028% | 2.125% | 0.8129% | 0.9016% | 1.077% | 1.68% | 0.974% | 1.822% | 5.334% | 3.203% | 1.157% | 0.8085% | 6.222% | 1.003% | 0.4736% | [6 empty] | 0.9213% | 0.3083% | 1.426% | 1.045%
|  | % Recovery |  |  | 91.37% | 93.44% | 93.53% | 88.58% | 110.1% | 94.53% | 102.8% | 100.5% | 101.7% | 105.8% | 100.4% | 102.6% | 92.15% | 97.74% | 94.13% | 97.05% | 101% | 101.3% | 94.82% | 96.24% | 98.2% | 101.2% | 100% | 98.12% | 98.01% | 104.5% | 100.5% | 99.98% | 104.7% | 64.84% | 47.31% | 54.99% | 58.39% | 48.17% |  | 97.19% | 98.83% | 107.6% | 108.3%
| 234 | CCV 10 ppb MAJORS |  | 06/17/2023 11:57:11 | 9036.4133 | 9290.9252 | 9.1368 | 8.8344 | 10684.1627 | 9550.7685 | 10259.8601 | 10.0001 | 10.1686 | 10.5382 | 10.1855 | 10.3267 | 9.0796 | 9.6226 | 9.3076 | 9.688 | 10.0191 | 10.0356 | 9.2331 | 10.1107 | 9.9015 | 10.1159 | 10.0307 | 9.7954 | 10.1018 | 10.1547 | 9.9832 | 10.0086 | 10.4225 | [6 empty] | 9.6742 | 9.7936 | 10.7344 | 11.149
|  | RSD |  |  | 1.275% | 1.424% | 5.272% | 0.5683% | 4.916% | 1.289% | 0.5192% | 1.958% | 0.1688% | 1.364% | 2.84% | 0.7222% | 0.2356% | 0.2678% | 3.044% | 0.8978% | 1.101% | 0.5258% | 2.597% | 4.674% | 1.637% | 2.382% | 7.061% | 3.876% | 2.81% | 4.309% | 6.96% | 0.3073% | 0.6464% | [6 empty] | 1.944% | 0.8456% | 1.898% | 1.628%
|  | % Recovery |  |  | 90.36% | 92.91% | 91.37% | 88.34% | 106.8% | 95.51% | 102.6% | 100% | 101.7% | 105.4% | 101.9% | 103.3% | 90.8% | 96.23% | 93.08% | 96.88% | 100.2% | 100.4% | 92.33% | 101.1% | 99.02% | 101.2% | 100.3% | 97.95% | 101% | 101.5% | 99.83% | 100.1% | 104.2% | 63.41% | 45.29% | 53.25% | 55.18% | 45.69% |  | 96.74% | 97.94% | 107.3% | 111.5%
QC STD 7
|  | Sample Id | R | Acquisition Time | Na KED 23/23 | Mg KED 24/24 | Al KED 27/27 | P Oshift 31/47 | S (34) -1 34/34 | K KED 39/39 | Ca KED 43/43 | V KED 51/51 | Cr (52) 52/52 | Cr (53) NH3 53/53 | Fe (54) NH3 54/54 | Mn Oshift 55/71 | Fe (56) KED 56/56 | Co KED 59/59 | Ni KED 60/60 | Cu (63) KED 63/63 | Cu (65) NH3 65/65 | Zn NH3 66/66 | As 75/75 | As KED 75/75 | As Oshift 75/91 | Se (77) 77/77 | Se (77) Oshift 77/93 | Se (78) Oshift 78/94 | Se (80) Oshift 80/96 | Se (82) 82/82 | Se (82) Oshift 82/98 | Mo (95) NH3 95/95 | Mo (95) OShift 95/127 | Rh (IS1) 103/103 (IS) | Rh (IS2) 103/103 (IS) | Rh (IS4) 103/103 (IS) | Rh (IS5) 103/103 (IS) | Rh (IS6) 103/103 (IS) | Rh (IS7) 103/103 (IS) | Cd (111) NH3 111/111 | Cd (114) NH3 114/114 | Pb 208/208 | Pb NH3 208/208
| 24 | LLOQ 0.1 |  | 06/16/2023 14:24:05 | 157.5781 | 102.3591 | -0.73809 | -0.51602 | 462.677 | 164.6774 | 129.5481 | 0.10451 | 0.088081 | 0.10195 | -0.2596 | 0.10732 | 0.13125 | 0.09552 | 0.10538 | 0.10931 | 0.10692 | 0.014436 | 0.13764 | 0.094082 | 0.094888 | 0.13829 | 0.096396 | 0.10776 | 0.091667 | 0.056754 | 0.13188 | 0.13578 | 0.10093 | [6 empty] | 0.097918 | 0.09334 | 0.089389 | 0.093548
|  | RSD |  |  | 5.96% | 4.792% | 4.423% | 11.01% | 39.67% | 3.805% | 8.24% | 9.959% | 4.757% | 5.109% | 6.217% | 5.866% | 62.07% | 9.499% | 7.081% | 8.64% | 2.645% | 122.9% | 34.34% | 4.707% | 4.781% | 94.57% | 41.23% | 14.52% | 21.68% | 45.59% | 17.63% | 5.174% | 2.726% | [6 empty] | 4.84% | 3.777% | 2.771% | 1.121%
|  | % Recovery |  |  | 157.6% | 102.4% | -738.1% | -516% | 462.7% | 164.7% | 129.5% | 104.5% | 88.08% | 101.9% | -259.6% | 107.3% | 131.3% | 95.52% | 105.4% | 109.3% | 106.9% | 14.44% | 137.6% | 94.08% | 94.89% | 138.3% | 96.4% | 107.8% | 91.67% | 56.75% | 131.9% | 135.8% | 100.9% | 93.22% | 89.09% | 90.83% | 92.57% | 86.35% |  | 97.92% | 93.34% | 89.39% | 93.55%
QC Duplicates (% Difference)
|  | Sample Id | R | Acquisition Time | Na KED 23/23 | Mg KED 24/24 | Al KED 27/27 | P Oshift 31/47 | S (34) -1 34/34 | K KED 39/39 | Ca KED 43/43 | V KED 51/51 | Cr (52) 52/52 | Cr (53) NH3 53/53 | Fe (54) NH3 54/54 | Mn Oshift 55/71 | Fe (56) KED 56/56 | Co KED 59/59 | Ni KED 60/60 | Cu (63) KED 63/63 | Cu (65) NH3 65/65 | Zn NH3 66/66 | As 75/75 | As KED 75/75 | As Oshift 75/91 | Se (77) 77/77 | Se (77) Oshift 77/93 | Se (78) Oshift 78/94 | Se (80) Oshift 80/96 | Se (82) 82/82 | Se (82) Oshift 82/98 | Mo (95) NH3 95/95 | Mo (95) OShift 95/127 | Rh (IS1) 103/103 (IS) | Rh (IS2) 103/103 (IS) | Rh (IS4) 103/103 (IS) | Rh (IS5) 103/103 (IS) | Rh (IS6) 103/103 (IS) | Rh (IS7) 103/103 (IS) | Cd (111) NH3 111/111 | Cd (114) NH3 114/114 | Pb 208/208 | Pb NH3 208/208
| 65 | 10_1_21_W_11_LD, Duplicate of Row Index 64 |  | 06/16/2023 18:42:31 | 184.3% | 199.9% | 220.7% | 368.3% | 1182% | 160.2% | 197.7% | 201.6% | 403.1% | 236.5% | 214.4% | 200.1% | 266.7% | 548.3% | 272.1% | 207% | 206.7% | 350.4% | 5.445% | 203.2% | 202.1% | 241% | 133.9% | 208.5% | 374.4% | 199.3% | 682.2% | 200.9% | 199% | [6 empty] | 155.4% | 63845% | 222.3% | 209.4%
| 66 | 10_1_21_W_11_LD, Duplicate of Row Index 64 |  | 06/16/2023 18:48:48 | 0.3515% | 2.08% | 4.515% | 22.74% | 3.225% | 1.068% | 2.148% | 5.534% | 0.2558% | 26.65% | 1.891% | 4.177% | 0.1478% | 48.71% | 10.27% | 2.833% | 0.2974% | 2.72% | 3.102% | 19.77% | 7.215% | 7.41% | 20.42% | 13.48% | 203.8% | 201.4% | 612190% | 3.957% | 4.338% | [6 empty] | 22.64% | 17.49% | 3.015% | 2.24%
| 91 | 10_1_21_UF_8_LD, Duplicate of Row Index 90 |  | 06/16/2023 21:19:45 | 1.317% | 1.514% | 91.57% | 51.84% | 3.316% | 0.7549% | 2.544% | 1.994% | 1.131% | 8.051% | 0.2915% | 1.472% | 9.076% | 36.7% | 6.274% | 0.2608% | 1.103% | 0.9851% | 0.7454% | 6.752% | 0.5961% | 29.6% | 26.59% | 16.08% | 3.195% | 3.472% | 19.81% | 0.1404% | 0.4515% | [6 empty] | 7.331% | 6.719% | 4.399% | 1.505%
| 92 | 10_1_21_UF_8_LD, Duplicate of Row Index 90 |  | 06/16/2023 21:26:03 | 0.1363% | 1.954% | 91.22% | 56.52% | 2.286% | 1.455% | 2.499% | 1.547% | 5.73% | 5.634% | 2.624% | 0.6477% | 13.39% | 34.74% | 9.53% | 2.138% | 0.5958% | 1.295% | 0.6048% | 10.01% | 0.8987% | 0.8382% | 36.13% | 12.69% | 1.886% | 3.115% | 36.24% | 0.3659% | 0.9128% | [6 empty] | 1.538% | 32.09% | 1.847% | 2.418%
| 143 | 10_15_21_F_6_LD, Duplicate of Row Index 142 |  | 06/17/2023 02:37:02 | 1.182% | 3.518% | 7.703% | 66.14% | 4.808% | 5.567% | 2.96% | 3.558% | 2.952% | 5.224% | 18.28% | 0.07597% | 18.11% | 14.03% | 1.975% | 1.531% | 0.6098% | 2.201% | 1.155% | 1.859% | 1.736% | 17.59% | 60.06% | 22.78% | 6.327% | 4.949% | 13.19% | 0.3173% | 0.7194% | [6 empty] | 3.088% | 5.095% | 3.321% | 2.858%
| 144 | 10_15_21_F_6_LD, Duplicate of Row Index 142 |  | 06/17/2023 02:43:20 | 0.1404% | 1.804% | 6.728% | 77.97% | 1.435% | 2.532% | 0.2077% | 1.38% | 1.674% | 0.3238% | 16.31% | 0.4136% | 15.79% | 25.57% | 0.9034% | 0.125% | 2.526% | 0.7369% | 3.566% | 5.356% | 1.68% | 21.43% | 59.81% | 11.25% | 28.11% | 13.28% | 27.3% | 2.846% | 0.6525% | [6 empty] | 13.33% | 11.62% | 8.307% | 1.179%
| 176 | 10_15_21_UF_8_LD, Duplicate of Row Index 175 |  | 06/17/2023 06:00:03 | 0.2787% | 0.1971% | 66.83% | 60.82% | 0.4301% | 0.9596% | 0.1069% | 1.3% | 11.15% | 13.26% | 41.83% | 5.491% | 19.89% | 21.98% | 7.616% | 2.948% | 5.807% | 73.52% | 0.4421% | 3.146% | 2.782% | 12.93% | 70.53% | 5.153% | 32.07% | 45.52% | 32.74% | 0.6437% | 4.674% | [6 empty] | 11.22% | 12.29% | 13.53% | 16.78%
| 177 | 10_15_21_UF_8_LD, Duplicate of Row Index 175 |  | 06/17/2023 06:06:21 | 1.858% | 3.759% | 69.71% | 69.36% | 2.017% | 2.259% | 1.207% | 0.2076% | 1.476% | 15.97% | 43.38% | 0.6619% | 13.11% | 30.13% | 9.648% | 7.318% | 6.686% | 74.35% | 1.823% | 0.43% | 1.697% | 23.83% | 35.36% | 28.5% | 9.839% | 50.66% | 46.33% | 0.5131% | 0.3828% | [6 empty] | 8.977% | 3.99% | 7.749% | 15.32%
Samples, Blanks & Calibration Standards
|  | Sample Id | R | Acquisition Time | Na KED 23/23 | Mg KED 24/24 | Al KED 27/27 | P Oshift 31/47 | S (34) -1 34/34 | K KED 39/39 | Ca KED 43/43 | V KED 51/51 | Cr (52) 52/52 | Cr (53) NH3 53/53 | Fe (54) NH3 54/54 | Mn Oshift 55/71 | Fe (56) KED 56/56 | Co KED 59/59 | Ni KED 60/60 | Cu (63) KED 63/63 | Cu (65) NH3 65/65 | Zn NH3 66/66 | As 75/75 | As KED 75/75 | As Oshift 75/91 | Se (77) 77/77 | Se (77) Oshift 77/93 | Se (78) Oshift 78/94 | Se (80) Oshift 80/96 | Se (82) 82/82 | Se (82) Oshift 82/98 | Mo (95) NH3 95/95 | Mo (95) OShift 95/127 | Rh (IS1) 103/103 (IS) | Rh (IS2) 103/103 (IS) | Rh (IS4) 103/103 (IS) | Rh (IS5) 103/103 (IS) | Rh (IS6) 103/103 (IS) | Rh (IS7) 103/103 (IS) | Cd (111) NH3 111/111 | Cd (114) NH3 114/114 | Pb 208/208 | Pb NH3 208/208
| 106 | 5_18_23_DB |  | 06/16/2023 22:50:24 | 43769.9619 | 12318.5837 | 113.1533 | 10.8613 | 1207.6708 | 3716.8541 | 49307.7328 | 1.9543 | 1.4534 | 1.4731 | 42.6719 | 3.5835 | 39.5921 | 1.3396 | 1.6309 | 0.77295 | 0.86221 | 5.8381 | 2.2272 | 2.1675 | 2.3597 | 4.8171 | 2.0437 | 1.8249 | 1.8442 | 1.9192 | 2.0587 | 1.0445 | 1.08 | [6 empty] | 3.267 | 3.2022 | 1.3019 | 1.2691
|  | IS % Recovery | [31 empty] | 70.23% | 54.54% | 61.13% | 64.17% | 54.65% | 0%
|  | Sample RSD |  |  | 0.8144% | 0.7885% | 1.035% | 1.158% | 16.16% | 1.369% | 0.7182% | 2.035% | 6.131% | 0.9507% | 0.832% | 1.584% | 0.7303% | 1.206% | 1.794% | 6.246% | 1.433% | 0.7074% | 2.533% | 10.05% | 1.036% | 5.246% | 8.458% | 6.659% | 5.733% | 12.96% | 3.879% | 1.533% | 0.6418% | [6 empty] | 2.014% | 1.66% | 2.553% | 0.24%
|  | Sample Concentration SD |  |  | 356.4601 | 97.1276 | 1.1715 | 0.12578 | 195.2086 | 50.8952 | 354.1038 | 0.039769 | 0.089115 | 0.014005 | 0.35502 | 0.056748 | 0.28913 | 0.016153 | 0.029259 | 0.048276 | 0.012355 | 0.041297 | 0.05641 | 0.2178 | 0.024439 | 0.25271 | 0.17286 | 0.12152 | 0.10572 | 0.24879 | 0.079849 | 0.016011 | 0.0069317 | [6 empty] | 0.065794 | 0.053157 | 0.033239 | 0.0030464
| 107 | 5_18_23_DB |  | 06/16/2023 22:56:42 | 43362.4959 | 12410.459 | 108.812 | 11.8107 | 961.5035 | 3682.17 | 49526.3751 | 1.8454 | 1.4003 | 1.4197 | 41.8412 | 3.573 | 38.854 | 1.357 | 1.5545 | 0.7449 | 0.85922 | 5.795 | 2.2864 | 2.1086 | 2.3171 | 5.0834 | 1.9291 | 1.8769 | 1.853 | 2.1861 | 1.9397 | 1.0555 | 1.0916 | [6 empty] | 3.2344 | 3.1803 | 1.3582 | 1.2892
|  | IS % Recovery | [31 empty] | 68.53% | 54.89% | 62.38% | 63.77% | 54.65% | 0%
|  | Sample RSD |  |  | 1.388% | 0.6112% | 1.091% | 1.281% | 11.15% | 2.966% | 2.118% | 1.2% | 2.905% | 0.566% | 1.131% | 1.617% | 1.741% | 1.131% | 2.027% | 2.118% | 1.804% | 0.5923% | 5.558% | 4.089% | 1.392% | 2.326% | 16.55% | 8.355% | 1.628% | 4.035% | 0.8809% | 0.5967% | 1.6% | [6 empty] | 1.603% | 0.4922% | 1.905% | 0.7208%
|  | Sample Concentration SD |  |  | 601.7522 | 75.8572 | 1.1873 | 0.15128 | 107.2407 | 109.2108 | 1048.8631 | 0.022143 | 0.04068 | 0.0080348 | 0.47304 | 0.057784 | 0.67637 | 0.015348 | 0.031512 | 0.01578 | 0.0155 | 0.034324 | 0.12709 | 0.086224 | 0.032249 | 0.11823 | 0.31934 | 0.15682 | 0.030171 | 0.088213 | 0.017086 | 0.0062983 | 0.017464 | [6 empty] | 0.05186 | 0.015655 | 0.025875 | 0.0092918
| 108 | 5_18_23_T_239 |  | 06/16/2023 23:02:42 | 43818.6967 | 12401.9544 | 112.5601 | 12.1886 | 1049.0914 | 3786.4273 | 50300.4608 | 1.8823 | 1.4043 | 1.4782 | 42.5428 | 3.6191 | 39.7527 | 1.3633 | 1.5978 | 0.75129 | 0.84652 | 5.9218 | 2.2436 | 2.1529 | 2.3504 | 5.183 | 1.8762 | 1.911 | 1.8645 | 2.1089 | 1.7647 | 1.1012 | 1.0772 | [6 empty] | 3.2441 | 3.2062 | 1.3426 | 1.2786
|  | IS % Recovery | [31 empty] | 69.01% | 54.9% | 61.65% | 64.91% | 54.28% | 0%
|  | Sample RSD |  |  | 1.703% | 1.907% | 1.187% | 1.371% | 14.72% | 0.339% | 0.7527% | 2.13% | 1.178% | 1.061% | 1.007% | 0.8387% | 1.57% | 2.339% | 0.9694% | 2.035% | 2.806% | 1.704% | 1.202% | 5.975% | 0.03144% | 1.402% | 3.753% | 1.617% | 4.667% | 6.882% | 16.25% | 0.8764% | 0.9302% | [6 empty] | 0.8258% | 1.934% | 1.459% | 1.801%
|  | Sample Concentration SD |  |  | 746.1488 | 236.5498 | 1.3361 | 0.16706 | 154.3868 | 12.8352 | 378.6341 | 0.0401 | 0.016545 | 0.015685 | 0.42835 | 0.030354 | 0.62429 | 0.031893 | 0.015489 | 0.015292 | 0.023751 | 0.10092 | 0.026977 | 0.12864 | 0.00073899 | 0.072651 | 0.07042 | 0.030899 | 0.087023 | 0.14513 | 0.2867 | 0.0096512 | 0.010021 | [6 empty] | 0.026791 | 0.06202 | 0.019582 | 0.023027
| 109 | 5_18_23_T_239 |  | 06/16/2023 23:09:00 | 43638.5231 | 12318.0966 | 114.4458 | 11.7129 | 1252.4876 | 3804.2649 | 50684.1641 | 1.886 | 1.4278 | 1.4217 | 41.8146 | 3.451 | 39.7999 | 1.3533 | 1.7091 | 0.74796 | 0.83342 | 5.7601 | 2.2777 | 2.1467 | 2.2225 | 5.0414 | 1.6799 | 1.8186 | 1.7585 | 2.1142 | 1.806 | 1.0965 | 1.0362 | [6 empty] | 3.2232 | 3.183 | 1.3481 | 1.3047
|  | IS % Recovery | [31 empty] | 68.16% | 56.3% | 61.92% | 63.18% | 54.43% | 0%
|  | Sample RSD |  |  | 0.4131% | 1.535% | 0.2866% | 0.6331% | 9.522% | 3.363% | 1.463% | 2.404% | 2.238% | 1.682% | 0.1375% | 0.5156% | 0.7733% | 2.193% | 1.64% | 1.236% | 2.467% | 0.9773% | 3.093% | 11.62% | 1.018% | 3.955% | 10.27% | 8.25% | 5.833% | 1.681% | 2.183% | 2.47% | 0.6734% | [6 empty] | 0.4858% | 0.6729% | 0.6073% | 2.037%
|  | Sample Concentration SD |  |  | 180.2823 | 189.1183 | 0.32802 | 0.074158 | 119.2594 | 127.951 | 741.3139 | 0.045347 | 0.031952 | 0.023908 | 0.057483 | 0.017794 | 0.30779 | 0.029679 | 0.028026 | 0.0092412 | 0.020557 | 0.056292 | 0.070455 | 0.24955 | 0.022631 | 0.19937 | 0.1725 | 0.15003 | 0.10256 | 0.035543 | 0.039426 | 0.027087 | 0.0069785 | [6 empty] | 0.015658 | 0.021419 | 0.0081874 | 0.026573
| 123 | 10_15_21_W_9 |  | 06/17/2023 00:33:56 | 1572.1129 | 1292.7546 | 9.5485 | 1.8094 | 3728.0764 | 403.3684 | 6875.8952 | 0.11606 | 0.030872 | 0.034331 | 9.0725 | 2.3189 | 8.4067 | -0.0014588 | 0.056117 | 0.58915 | 0.62044 | 0.25387 | 0.5951 | 0.54605 | 0.58819 | 0.78538 | 0.019037 | 0.018882 | 0.006106 | 0.16521 | 0.0071705 | 1.4202 | 1.4846 | [6 empty] | 0.0044981 | 0.00048436 | 0.06522 | 0.054579
|  | IS % Recovery | [31 empty] | 64.24% | 59.69% | 67.57% | 69.72% | 60.6% | 0%
|  | Sample RSD |  |  | 1.576% | 1.214% | 1.084% | 4.619% | 3.909% | 3.102% | 2.778% | 19.36% | 16.54% | 7.227% | 0.6629% | 2% | 2.815% | 138.8% | 11.63% | 4.619% | 3.004% | 7.523% | 7.955% | 16.27% | 2.315% | 59.78% | 61.42% | 0.7803% | 128.9% | 101.3% | 367.6% | 1.977% | 2.145% | [6 empty] | 16.52% | 217.5% | 16.49% | 1.638%
|  | Sample Concentration SD |  |  | 24.7754 | 15.6942 | 0.10353 | 0.083582 | 145.7293 | 12.5137 | 190.995 | 0.022468 | 0.0051052 | 0.0024812 | 0.060139 | 0.046369 | 0.23664 | 0.0020243 | 0.006525 | 0.027211 | 0.01864 | 0.019099 | 0.047342 | 0.088835 | 0.013618 | 0.46947 | 0.011692 | 0.00014733 | 0.0078718 | 0.16742 | 0.026355 | 0.028079 | 0.031851 | [6 empty] | 0.00074297 | 0.0010534 | 0.010754 | 0.00089401
| 124 | 10_15_21_W_9 |  | 06/17/2023 00:40:15 | 1590.7698 | 1274.7171 | 8.7663 | 1.8495 | 3698.9078 | 410.1781 | 6898.2629 | 0.1177 | 0.029519 | 0.037816 | 9.099 | 2.28 | 8.5242 | -0.0010056 | 0.036501 | 0.62082 | 0.63553 | 0.26782 | 0.57476 | 0.47825 | 0.59738 | 0.19708 | 0.012288 | 0.012467 | 0.0042995 | 0.11182 | 0.00718 | 1.4431 | 1.4681 | [6 empty] | 0.0033199 | 0.00028658 | 0.051986 | 0.05653
|  | IS % Recovery | [31 empty] | 74.86% | 59.47% | 66.53% | 68.83% | 60.02% | 0%
|  | Sample RSD |  |  | 3.118% | 3.263% | 5.987% | 1.911% | 8.385% | 7.008% | 3.319% | 17.01% | 21.78% | 6.161% | 0.5776% | 2.051% | 1.102% | 77.39% | 10.45% | 1.16% | 0.8699% | 1.174% | 13.59% | 28.89% | 6.769% | 34.5% | 95.16% | 51.7% | 369.9% | 80.51% | 273.9% | 0.8042% | 1.829% | [6 empty] | 34.83% | 302.7% | 1.345% | 2.457%
|  | Sample Concentration SD |  |  | 49.5934 | 41.5892 | 0.52483 | 0.035349 | 310.1454 | 28.7434 | 228.9829 | 0.02002 | 0.00643 | 0.0023299 | 0.052552 | 0.046769 | 0.093961 | 0.00077821 | 0.0038141 | 0.0072 | 0.0055285 | 0.0031454 | 0.078114 | 0.13818 | 0.040436 | 0.067983 | 0.011693 | 0.0064449 | 0.015903 | 0.090024 | 0.019667 | 0.011605 | 0.026851 | [6 empty] | 0.0011563 | 0.00086744 | 0.00069945 | 0.0013889
| 125 | 10_15_21_W_10 |  | 06/17/2023 00:46:15 | 1285.8078 | 1505.7855 | 30.7228 | 2.9233 | 3222.2775 | 403.2486 | 6651.2359 | 0.15196 | 0.033668 | 0.034283 | 25.7102 | 5.1996 | 23.6581 | 0.0072383 | 0.048663 | 0.86917 | 0.88971 | 1.0798 | 0.65321 | 0.6412 | 0.6658 | 0.16078 | 0.032907 | 0.023412 | 0.0083637 | 0.046454 | -0.015696 | 0.63991 | 0.66836 | [6 empty] | 0.0067906 | 0.0042742 | 0.1793 | 0.16979
|  | IS % Recovery | [31 empty] | 74.27% | 59.13% | 66.39% | 68.91% | 59.31% | 0%
|  | Sample RSD |  |  | 2.143% | 2.868% | 0.9412% | 3.248% | 3.239% | 2.739% | 3.762% | 3.979% | 11.65% | 24.21% | 1.136% | 1.511% | 0.7344% | 23.09% | 15.33% | 3.807% | 2.392% | 1.247% | 2.751% | 16.55% | 4.952% | 64.04% | 36.27% | 42.46% | 92.36% | 67.01% | 63.75% | 2.451% | 0.6243% | [6 empty] | 43.44% | 19.17% | 2.126% | 2.221%
|  | Sample Concentration SD |  |  | 27.5533 | 43.1793 | 0.28915 | 0.094948 | 104.3776 | 11.0461 | 250.1996 | 0.0060458 | 0.0039225 | 0.0083002 | 0.29218 | 0.078561 | 0.17374 | 0.0016714 | 0.0074594 | 0.033091 | 0.02128 | 0.013468 | 0.01797 | 0.10615 | 0.032971 | 0.10295 | 0.011934 | 0.0099404 | 0.0077246 | 0.03113 | 0.010007 | 0.015687 | 0.0041724 | [6 empty] | 0.0029496 | 0.00081936 | 0.0038115 | 0.0037719
| 126 | 10_15_21_W_10 |  | 06/17/2023 00:52:33 | 1268.8095 | 1473.1627 | 31.2012 | 2.7824 | 3205.0898 | 367.1392 | 6757.8141 | 0.15041 | 0.025625 | 0.031782 | 25.6163 | 5.2486 | 23.8756 | 0.0071848 | 0.040801 | 0.86183 | 0.88084 | 1.0925 | 0.60909 | 0.48603 | 0.68398 | 0.075173 | 0.032777 | 0.0168 | 0.0092596 | 0.060801 | -0.0043024 | 0.63699 | 0.66042 | [6 empty] | 0.0064709 | 0.0032196 | 0.17381 | 0.17129
|  | IS % Recovery | [31 empty] | 75.35% | 59.5% | 67.14% | 69.75% | 60.2% | 0%
|  | Sample RSD |  |  | 1.625% | 2.843% | 1.779% | 3.298% | 4.221% | 6.841% | 2.782% | 15.92% | 23.93% | 14.77% | 0.2636% | 0.7195% | 0.7277% | 18.01% | 10.81% | 1.488% | 1.085% | 2.229% | 9.9% | 3.803% | 3.839% | 102.1% | 71.98% | 59.06% | 58.19% | 41.4% | 398.6% | 0.7919% | 2.675% | [6 empty] | 22.96% | 13.81% | 2.18% | 2.327%
|  | Sample Concentration SD |  |  | 20.6154 | 41.8881 | 0.55494 | 0.091771 | 135.2973 | 25.1158 | 188.006 | 0.023952 | 0.0061322 | 0.0046951 | 0.067529 | 0.037765 | 0.17375 | 0.0012938 | 0.0044101 | 0.012824 | 0.0095554 | 0.024354 | 0.060299 | 0.018483 | 0.026256 | 0.076743 | 0.023591 | 0.0099228 | 0.0053885 | 0.025172 | 0.01715 | 0.0050447 | 0.017668 | [6 empty] | 0.001486 | 0.0004447 | 0.0037887 | 0.0039863
| 127 | 10_15_21_W_10_DUP |  | 06/17/2023 00:58:33 | 1229.5166 | 1420.7281 | 28.7666 | 2.8389 | 3184.337 | 382.2393 | 6571.0237 | 0.1441 | 0.057515 | 0.063047 | 24.983 | 5.0238 | 22.7382 | 0.0059246 | 0.044269 | 0.85086 | 0.8821 | 0.96952 | 0.59848 | 0.54385 | 0.67152 | 0.24149 | 0.0055366 | 0.010246 | -0.00095537 | 0.071134 | -0.0042822 | 0.62367 | 0.63807 | [6 empty] | 0.007231 | 0.0034866 | 0.16196 | 0.16484
|  | IS % Recovery | [31 empty] | 75.79% | 59.93% | 68.01% | 68.51% | 59.87% | 0%
|  | Sample RSD |  |  | 2.87% | 1.47% | 0.5589% | 1.721% | 7.798% | 8.331% | 1.13% | 6.245% | 1.952% | 6.738% | 3.353% | 0.4858% | 2.8% | 23.05% | 28.81% | 1.332% | 2.474% | 1.431% | 12.74% | 19.14% | 1.558% | 31.48% | 0% | 73.3% | 665.5% | 18.96% | 695.4% | 1.778% | 2.27% | [6 empty] | 4.268% | 26.21% | 0.6943% | 0.6915%
|  | Sample Concentration SD |  |  | 35.2817 | 20.8825 | 0.16079 | 0.048854 | 248.317 | 31.8459 | 74.2709 | 0.0089989 | 0.0011227 | 0.004248 | 0.83766 | 0.024404 | 0.63659 | 0.0013654 | 0.012754 | 0.01133 | 0.021823 | 0.013878 | 0.076253 | 0.1041 | 0.010462 | 0.076016 | 0 | 0.0075106 | 0.006358 | 0.013489 | 0.029777 | 0.01109 | 0.014482 | [6 empty] | 0.0003086 | 0.00091372 | 0.0011245 | 0.0011399
| 128 | 10_15_21_W_10_DUP |  | 06/17/2023 01:04:52 | 1247.1092 | 1458.0539 | 27.6383 | 2.85 | 3208.9612 | 387.2612 | 6409.5866 | 0.14541 | 0.055372 | 0.057049 | 25.2636 | 5.0287 | 22.9964 | 0.0059069 | 0.048256 | 0.86047 | 0.87087 | 0.95387 | 0.62072 | 0.60296 | 0.69031 | 0.16657 | 0.0055366 | 0.019184 | 0.013621 | 0.12264 | -0.0098875 | 0.61297 | 0.65128 | [6 empty] | 0.0075176 | 0.0030841 | 0.16071 | 0.15901
|  | IS % Recovery | [31 empty] | 73.61% | 58.74% | 66.43% | 68.29% | 58.67% | 0%
|  | Sample RSD |  |  | 4.556% | 2.151% | 4.333% | 1.396% | 2.657% | 3.472% | 1.834% | 8.36% | 14.02% | 6.172% | 2.494% | 1.456% | 1.536% | 55.61% | 29.96% | 1.496% | 1.285% | 0.5251% | 6.921% | 16.08% | 3.611% | 52.62% | 0% | 33.93% | 25.63% | 19.16% | 203% | 1.378% | 1.028% | [6 empty] | 8.603% | 46.4% | 2.724% | 4.494%
|  | Sample Concentration SD |  |  | 56.8178 | 31.3694 | 1.1976 | 0.039775 | 85.2591 | 13.4469 | 117.5514 | 0.012157 | 0.0077647 | 0.0035212 | 0.63001 | 0.073218 | 0.35331 | 0.0032846 | 0.014456 | 0.012869 | 0.011189 | 0.0050087 | 0.042959 | 0.096959 | 0.024929 | 0.087648 | 0 | 0.0065085 | 0.0034905 | 0.023495 | 0.020068 | 0.0084454 | 0.0066928 | [6 empty] | 0.00064672 | 0.0014309 | 0.0043773 | 0.0071452
| 135 | 10_15_21_W_11 |  | 06/17/2023 01:47:51 | 1202.0649 | 1500.0398 | 41.9447 | 2.8231 | 3194.1838 | 398.3913 | 6621.6961 | 0.17238 | 0.043371 | 0.046943 | 33.9716 | 5.7171 | 31.3393 | 0.010605 | 0.057631 | 1.0295 | 1.0937 | 1.7004 | 0.62582 | 0.60114 | 0.70033 | -0.049296 | 0.019075 | 0.0059015 | 0.011138 | 0.014905 | -0.0042934 | 0.56267 | 0.55949 | [6 empty] | 0.0061421 | 0.0065104 | 0.18482 | 0.20569
|  | IS % Recovery | [31 empty] | 81.75% | 59.94% | 66.66% | 69.27% | 59.93% | 0%
|  | Sample RSD |  |  | 3.404% | 3.088% | 2.403% | 0.4959% | 5.386% | 7.659% | 0.911% | 3.483% | 31.77% | 4.762% | 1.081% | 2.364% | 1.163% | 11.52% | 30.26% | 5.395% | 1.232% | 1.175% | 5.567% | 27.07% | 4.072% | 143.4% | 122.9% | 108.1% | 84.1% | 226.8% | 400.3% | 3.125% | 0.8717% | [6 empty] | 27.49% | 9.276% | 5.949% | 0.8111%
|  | Sample Concentration SD |  |  | 40.923 | 46.3231 | 1.0078 | 0.014001 | 172.0477 | 30.5111 | 60.3255 | 0.0060047 | 0.013778 | 0.0022355 | 0.36722 | 0.13518 | 0.36436 | 0.0012213 | 0.017438 | 0.055546 | 0.013469 | 0.019972 | 0.034842 | 0.16273 | 0.02852 | 0.070685 | 0.023449 | 0.0063819 | 0.0093665 | 0.03381 | 0.017188 | 0.017581 | 0.0048771 | [6 empty] | 0.0016885 | 0.00060393 | 0.010996 | 0.0016683
| 136 | 10_15_21_W_11 |  | 06/17/2023 01:54:10 | 1185.5387 | 1477.3299 | 40.491 | 3.0542 | 2985.4201 | 371.5053 | 6594.0206 | 0.17399 | 0.040188 | 0.045111 | 34.2345 | 5.6832 | 30.7074 | 0.01021 | 0.055387 | 0.99603 | 1.0887 | 1.6934 | 0.736 | 0.74405 | 0.7189 | 0.099548 | 0.033358 | 0.023794 | 0.0014201 | 0.044907 | -0.0038902 | 0.55994 | 0.57723 | [6 empty] | 0.0093589 | 0.0053224 | 0.21451 | 0.20386
|  | IS % Recovery | [31 empty] | 73.51% | 58.22% | 65.48% | 67.98% | 58.67% | 0%
|  | Sample RSD |  |  | 3.122% | 2.723% | 2.325% | 2.079% | 2.796% | 4.277% | 0.5009% | 3.441% | 6.511% | 4.177% | 1.216% | 1.086% | 0.2539% | 25.68% | 49.51% | 4.019% | 0.3721% | 1.18% | 5.135% | 13.97% | 0.9733% | 91.05% | 72.23% | 57.94% | 334.7% | 69.34% | 452.4% | 1.856% | 0.4357% | [6 empty] | 15.51% | 24.87% | 2.895% | 0.7207%
|  | Sample Concentration SD |  |  | 37.0158 | 40.2245 | 0.94156 | 0.063508 | 83.469 | 15.8892 | 33.031 | 0.0059874 | 0.0026165 | 0.0018842 | 0.41643 | 0.061721 | 0.077976 | 0.0026218 | 0.027422 | 0.040029 | 0.0040505 | 0.01998 | 0.037793 | 0.10394 | 0.0069974 | 0.090633 | 0.024095 | 0.013787 | 0.0047529 | 0.031136 | 0.017599 | 0.010392 | 0.0025148 | [6 empty] | 0.0014511 | 0.0013235 | 0.0062092 | 0.0014692
| 137 | 10_15_21_F_1 |  | 06/17/2023 02:00:09 | 10219.3731 | 6847.8694 | 1.9268 | 6.4141 | 32544.0982 | 2787.0104 | 45317.1903 | 0.19546 | 0.067471 | 0.075495 | 12.227 | 19.9237 | 12.1574 | 0.024947 | 0.39748 | 0.77675 | 0.83846 | 1.2518 | 2.7852 | 2.5831 | 2.8186 | 0.61552 | 0.092999 | 0.068957 | 0.046631 | 0.52158 | 0.046131 | 21.3393 | 21.8333 | [6 empty] | 0.0086021 | 0.0075661 | 3.6005 | 3.1027
|  | IS % Recovery | [31 empty] | 71.71% | 55.53% | 62.58% | 71.65% | 55.64% | 0%
|  | Sample RSD |  |  | 2.344% | 2.966% | 10.62% | 0.5722% | 1.513% | 1.136% | 0.06202% | 10.32% | 23.29% | 7.68% | 0.2374% | 0.4545% | 1.895% | 31.36% | 2.514% | 3.965% | 0.7786% | 0.63% | 1.225% | 1.71% | 1.974% | 7.267% | 24.07% | 9.657% | 6.5% | 18.07% | 100.7% | 0.9987% | 0.5171% | [6 empty] | 41.2% | 16.63% | 0.9568% | 4.771%
|  | Sample Concentration SD |  |  | 239.5537 | 203.0881 | 0.20458 | 0.036699 | 492.5498 | 31.6742 | 28.1051 | 0.020167 | 0.015716 | 0.0057977 | 0.029033 | 0.090551 | 0.23042 | 0.0078233 | 0.0099927 | 0.0308 | 0.0065279 | 0.0078864 | 0.034121 | 0.044168 | 0.055634 | 0.04473 | 0.022388 | 0.0066591 | 0.003031 | 0.094265 | 0.046432 | 0.21311 | 0.1129 | [6 empty] | 0.003544 | 0.0012584 | 0.034449 | 0.14802
| 138 | 10_15_21_F_1 |  | 06/17/2023 02:06:28 | 10386.9913 | 7032.3456 | 1.9451 | 6.3023 | 32560.42 | 2844.9986 | 45525.2282 | 0.20668 | 0.067385 | 0.074497 | 12.4622 | 20.0952 | 12.4585 | 0.024122 | 0.34814 | 0.76734 | 0.8441 | 1.2882 | 2.7229 | 2.5591 | 2.8237 | 0.6889 | 0.082766 | 0.070961 | 0.047669 | 0.45425 | 0.07357 | 21.2365 | 21.9279 | [6 empty] | 0.010252 | 0.0079057 | 3.497 | 3.4967
|  | IS % Recovery | [31 empty] | 73.09% | 57.6% | 62.86% | 64.85% | 58.08% | 0%
|  | Sample RSD |  |  | 1.703% | 1.065% | 12.54% | 2.745% | 0.4802% | 2.536% | 2.415% | 2.998% | 17.07% | 2.383% | 0.5763% | 0.1365% | 0.9274% | 21.17% | 13.05% | 1.181% | 1.596% | 0.8654% | 2.184% | 1.297% | 0.6864% | 11.95% | 38.98% | 29.02% | 13.14% | 6.851% | 142.1% | 0.2832% | 0.8433% | [6 empty] | 19.49% | 31.85% | 1.039% | 0.6199%
|  | Sample Concentration SD |  |  | 176.8619 | 74.898 | 0.2439 | 0.17297 | 156.3516 | 72.1615 | 1099.348 | 0.0061956 | 0.011503 | 0.0017751 | 0.071816 | 0.027433 | 0.11555 | 0.0051058 | 0.045423 | 0.0090642 | 0.013468 | 0.011148 | 0.059477 | 0.03319 | 0.019382 | 0.082301 | 0.032263 | 0.020596 | 0.0062637 | 0.031119 | 0.10456 | 0.060143 | 0.18493 | [6 empty] | 0.0019979 | 0.0025181 | 0.036324 | 0.021675
| 139 | 10_15_21_F_1_DUP |  | 06/17/2023 02:12:28 | 9901.0171 | 6587.3786 | 1.1049 | 5.1146 | 30527.5174 | 2650.462 | 42640.337 | 0.203 | 0.053164 | 0.050694 | 11.6102 | 19.3078 | 11.6988 | 0.023372 | 0.35659 | 0.70677 | 0.77966 | 1.2538 | 2.6899 | 2.4528 | 2.7299 | 0.69758 | 0.081364 | 0.056555 | 0.050794 | 0.5058 | 0.071988 | 20.3201 | 20.8434 | [6 empty] | 0.010912 | 0.006162 | 0.05058 | 0.053585
|  | IS % Recovery | [31 empty] | 71.5% | 58.55% | 63.61% | 64.86% | 58.9% | 0%
|  | Sample RSD |  |  | 0.9813% | 1.758% | 23.35% | 1.354% | 0.8837% | 3.488% | 1.101% | 10.29% | 20.24% | 10.15% | 0.2304% | 1.154% | 1.472% | 8.91% | 1.611% | 4.219% | 1.952% | 1.996% | 1.485% | 7.813% | 0.3801% | 20.08% | 52.35% | 28.96% | 16.85% | 10.98% | 61.31% | 1.104% | 0.6435% | [6 empty] | 3.675% | 23.35% | 3.096% | 1.846%
|  | Sample Concentration SD |  |  | 97.1636 | 115.7805 | 0.25795 | 0.069243 | 269.7868 | 92.4359 | 469.4154 | 0.020879 | 0.01076 | 0.0051443 | 0.02675 | 0.22274 | 0.17216 | 0.0020824 | 0.0057443 | 0.029818 | 0.015217 | 0.025025 | 0.039937 | 0.19163 | 0.010377 | 0.14006 | 0.042596 | 0.016379 | 0.0085565 | 0.05553 | 0.044138 | 0.22429 | 0.13413 | [6 empty] | 0.00040105 | 0.0014386 | 0.0015662 | 0.00098894
| 140 | 10_15_21_F_1_DUP |  | 06/17/2023 02:18:46 | 10138.9519 | 6771.2102 | 1.1569 | 5.1716 | 30874.0792 | 2702.4277 | 43277.5899 | 0.17748 | 0.04635 | 0.050235 | 11.6423 | 19.1878 | 11.8251 | 0.021165 | 0.35103 | 0.72759 | 0.77051 | 1.2365 | 2.6951 | 2.2757 | 2.7049 | 0.66134 | 0.053359 | 0.084417 | 0.039096 | 0.47696 | 0.042306 | 20.1097 | 20.7729 | [6 empty] | 0.0092028 | 0.0066434 | 0.04779 | 0.049862
|  | IS % Recovery | [31 empty] | 71.84% | 59% | 64% | 65.73% | 59.23% | 0%
|  | Sample RSD |  |  | 1.848% | 2.069% | 23.22% | 3.607% | 2.219% | 1.754% | 2.242% | 3.99% | 17.84% | 9.168% | 2.261% | 0.4631% | 3.704% | 18.26% | 8.393% | 3% | 1.213% | 0.3319% | 3.188% | 5.917% | 0.9862% | 29.89% | 43.9% | 14.42% | 18.2% | 21.8% | 25% | 0.792% | 1.511% | [6 empty] | 41.54% | 20.2% | 4.694% | 1.743%
|  | Sample Concentration SD |  |  | 187.3229 | 140.1144 | 0.26865 | 0.18655 | 685.0731 | 47.4051 | 970.3108 | 0.0070822 | 0.0082682 | 0.0046053 | 0.26319 | 0.088867 | 0.43803 | 0.0038648 | 0.029461 | 0.021829 | 0.0093467 | 0.0041043 | 0.085908 | 0.13466 | 0.026677 | 0.19767 | 0.023423 | 0.01217 | 0.0071139 | 0.10397 | 0.010577 | 0.15927 | 0.31379 | [6 empty] | 0.0038226 | 0.0013419 | 0.0022431 | 0.00086898
| 141 | 10_15_21_F_6 |  | 06/17/2023 02:24:45 | 9629.9114 | 7669.1188 | 11.2546 | 2.3267 | 24141.9699 | 2245.8268 | 41752.8675 | 0.5675 | 0.080044 | 0.086269 | 18.7791 | 8.2352 | 18.7771 | 0.015592 | 0.28296 | 2.0075 | 2.1282 | 4.9782 | 2.8731 | 2.6881 | 2.8923 | 0.69087 | 0.13604 | 0.099793 | 0.099692 | 0.50131 | 0.030706 | 11.4307 | 11.707 | [6 empty] | 0.015011 | 0.012246 | 0.091037 | 0.079212
|  | IS % Recovery | [31 empty] | 71.79% | 58.9% | 63.23% | 72.67% | 59.7% | 0%
|  | Sample RSD |  |  | 0.5248% | 2.162% | 5.927% | 3.05% | 3.797% | 2.396% | 1.125% | 10.76% | 10.55% | 4.952% | 1.185% | 0.9176% | 2.258% | 13.82% | 4.624% | 2.985% | 2.253% | 0.9331% | 1.253% | 2.717% | 1.55% | 45.3% | 38.23% | 32.42% | 20.54% | 7.415% | 56.31% | 1.102% | 1.002% | [6 empty] | 23.21% | 5.649% | 2.852% | 9.872%
|  | Sample Concentration SD |  |  | 50.5393 | 165.837 | 0.66708 | 0.070971 | 916.5895 | 53.8018 | 469.7373 | 0.061048 | 0.0084475 | 0.0042718 | 0.22248 | 0.075566 | 0.42407 | 0.0021542 | 0.013085 | 0.059921 | 0.047945 | 0.046449 | 0.036006 | 0.073035 | 0.044829 | 0.31295 | 0.052003 | 0.032355 | 0.020481 | 0.037173 | 0.017289 | 0.12601 | 0.11729 | [6 empty] | 0.003484 | 0.00069176 | 0.002596 | 0.0078194
| 142 | 10_15_21_F_6 |  | 06/17/2023 02:31:03 | 9430.877 | 7611.3164 | 11.0525 | 2.3269 | 24142.7901 | 2294.784 | 41845.4663 | 0.54157 | 0.08211 | 0.089646 | 18.656 | 8.2356 | 18.2018 | 0.020376 | 0.31111 | 2.0248 | 2.083 | 5.0895 | 2.7877 | 2.5437 | 2.964 | 0.53509 | 0.1666 | 0.077532 | 0.068499 | 0.46481 | 0.096984 | 11.4216 | 11.8225 | [6 empty] | 0.015578 | 0.014996 | 0.092431 | 0.09269
|  | IS % Recovery | [31 empty] | 70.77% | 57.71% | 62.34% | 64.11% | 58.35% | 0%
|  | Sample RSD |  |  | 0.4645% | 1.797% | 0.6857% | 1.649% | 0.8095% | 1.321% | 1.162% | 3.03% | 6.772% | 6.497% | 0.5189% | 1.107% | 0.4554% | 11.88% | 8.864% | 2.054% | 1.009% | 2.114% | 1.205% | 8.803% | 0.8962% | 16.1% | 19.03% | 42.67% | 26.32% | 5.948% | 28.12% | 0.32% | 0.6178% | [6 empty] | 14.16% | 8.159% | 3.439% | 2.717%
|  | Sample Concentration SD |  |  | 43.8089 | 136.7678 | 0.075787 | 0.038371 | 195.4411 | 30.3256 | 486.0391 | 0.016407 | 0.0055604 | 0.0058245 | 0.096801 | 0.091154 | 0.082882 | 0.0024198 | 0.027578 | 0.041591 | 0.021027 | 0.10761 | 0.033588 | 0.22393 | 0.026562 | 0.086172 | 0.031711 | 0.033086 | 0.018028 | 0.027646 | 0.027274 | 0.036548 | 0.073039 | [6 empty] | 0.0022055 | 0.0012236 | 0.0031789 | 0.0025183
| 152 | 10_15_21_F_6_DUP |  | 06/17/2023 03:32:14 | 8811.1153 | 6992.3724 | 6.8005 | 2.3788 | 22820.0661 | 2141.8205 | 39844.1552 | 0.50363 | 0.081727 | 0.093434 | 10.2046 | 7.7409 | 10.3714 | 0.024951 | 0.31427 | 1.9886 | 2.0722 | 4.8828 | 2.766 | 2.6448 | 2.8368 | 0.59003 | 0.10713 | 0.066391 | 0.072447 | 0.46184 | 0.052217 | 11.2417 | 11.4655 | [6 empty] | 0.019252 | 0.014072 | 0.10432 | 0.10827
|  | IS % Recovery | [31 empty] | 69.09% | 55.72% | 61.47% | 63.14% | 55.86% | 0%
|  | Sample RSD |  |  | 2.425% | 1.751% | 3.5% | 1.255% | 1.344% | 5.187% | 1.174% | 7.516% | 7.753% | 10.08% | 0.8054% | 0.4112% | 0.4426% | 19.24% | 10.55% | 2.759% | 0.3559% | 0.7344% | 2.1% | 2.865% | 1.15% | 33.95% | 11.88% | 49.11% | 58.25% | 16.99% | 93.78% | 0.391% | 0.8335% | [6 empty] | 1.541% | 10.72% | 1.957% | 1.563%
|  | Sample Concentration SD |  |  | 213.6434 | 122.4161 | 0.23803 | 0.029863 | 306.8089 | 111.1007 | 467.7233 | 0.037855 | 0.0063366 | 0.0094139 | 0.082185 | 0.031832 | 0.045902 | 0.0047996 | 0.033168 | 0.054858 | 0.0073754 | 0.035861 | 0.058085 | 0.075777 | 0.032611 | 0.20029 | 0.01273 | 0.032608 | 0.042203 | 0.078476 | 0.048968 | 0.043951 | 0.095565 | [6 empty] | 0.00029669 | 0.0015091 | 0.0020413 | 0.0016926
| 153 | 10_15_21_F_6_DUP |  | 06/17/2023 03:38:32 | 8910.3035 | 7051.8656 | 7.1826 | 2.3587 | 22781.2926 | 2139.8691 | 40020.173 | 0.50234 | 0.093405 | 0.097135 | 9.9386 | 7.8319 | 10.4932 | 0.026948 | 0.29225 | 1.9905 | 2.0379 | 4.8686 | 3.0144 | 2.65 | 2.864 | 0.68447 | 0.057486 | 0.079187 | 0.094594 | 0.53331 | 0.065473 | 10.9494 | 11.57 | [6 empty] | 0.018388 | 0.015016 | 0.11671 | 0.11496
|  | IS % Recovery | [31 empty] | 64.34% | 54.86% | 61.69% | 60.62% | 56.07% | 0%
|  | Sample RSD |  |  | 2.233% | 0.7876% | 7.093% | 2.521% | 2.016% | 0.5349% | 1.516% | 9.699% | 20.2% | 7.825% | 1.465% | 1.302% | 1.548% | 14.58% | 11.22% | 2.247% | 1.246% | 0.8059% | 3.039% | 8.414% | 0.6545% | 17.93% | 80.57% | 9.891% | 21.77% | 27.2% | 70.46% | 0.4461% | 2.105% | [6 empty] | 7.332% | 4.947% | 3.262% | 4.12%
|  | Sample Concentration SD |  |  | 198.972 | 55.5385 | 0.50949 | 0.059464 | 459.3438 | 11.4469 | 606.636 | 0.048722 | 0.018871 | 0.0076007 | 0.14559 | 0.10195 | 0.1624 | 0.0039278 | 0.032795 | 0.044736 | 0.025382 | 0.039234 | 0.091604 | 0.22296 | 0.018745 | 0.12271 | 0.046315 | 0.0078322 | 0.020592 | 0.14506 | 0.046134 | 0.048847 | 0.24353 | [6 empty] | 0.0013482 | 0.00074289 | 0.0038068 | 0.0047366
| 154 | 10_15_21_F_6_DUP_LSP |  | 06/17/2023 03:44:32 | 17563.3886 | 16273.8317 | 16.946 | 14.3103 | 31736.4955 | 11646.9161 | 46706.9872 | 10.6931 | 10.4294 | 10.5597 | 20.3091 | 17.5141 | 19.4848 | 9.9816 | 9.5785 | 11.2966 | 11.593 | 14.3953 | 12.1378 | 11.9845 | 12.4204 | 10.5632 | 10.0777 | 10.0092 | 10.0159 | 10.5568 | 9.4897 | 20.1646 | 20.9537 | [6 empty] | 9.3567 | 9.4372 | 9.8518 | 10.1426
|  | IS % Recovery | [31 empty] | 69.23% | 53.83% | 59.46% | 60.99% | 55% | 0%
|  | Sample RSD |  |  | 1.099% | 2.515% | 1.038% | 1.354% | 0.2235% | 2.128% | 1.777% | 1.915% | 0.3929% | 0.5381% | 0.3757% | 0.3351% | 1.48% | 2.237% | 2.143% | 0.7129% | 1.089% | 0.6671% | 0.919% | 4.198% | 0.8583% | 2.519% | 0.1631% | 0.3215% | 2.501% | 2.8% | 7.324% | 0.6911% | 0.3114% | [6 empty] | 0.3568% | 0.4268% | 1.302% | 1.237%
|  | Sample Concentration SD |  |  | 193.0272 | 409.2976 | 0.17583 | 0.19375 | 70.9403 | 247.8698 | 830.2035 | 0.20476 | 0.04098 | 0.05682 | 0.076294 | 0.058691 | 0.28841 | 0.22325 | 0.2053 | 0.080529 | 0.12621 | 0.096032 | 0.11155 | 0.50306 | 0.1066 | 0.26608 | 0.016439 | 0.032176 | 0.25048 | 0.29555 | 0.69505 | 0.13935 | 0.065243 | [6 empty] | 0.033386 | 0.040275 | 0.12829 | 0.12542
| 155 | 10_15_21_F_6_DUP_LSP |  | 06/17/2023 03:50:52 | 18749.7413 | 17093.0077 | 18.5224 | 15.1751 | 33563.1567 | 12318.1988 | 49989.2252 | 11.3525 | 11.2101 | 10.702 | 20.5879 | 17.8444 | 20.7693 | 10.6252 | 10.1331 | 11.9968 | 11.7464 | 14.5267 | 12.069 | 12.398 | 12.3779 | 10.5394 | 10.3708 | 10.065 | 9.8342 | 10.5624 | 10.2297 | 20.3677 | 20.9185 | [6 empty] | 9.5494 | 9.5547 | 9.7951 | 10.1567
|  | IS % Recovery | [31 empty] | 69.95% | 54.34% | 59.62% | 61.31% | 55.48% | 0%
|  | Sample RSD |  |  | 4.357% | 3.896% | 4.9% | 0.7528% | 3.685% | 2% | 3.463% | 3.468% | 3.774% | 0.9826% | 1.383% | 0.7914% | 3.558% | 3.352% | 3.884% | 3.386% | 0.7585% | 1.72% | 1.173% | 6.843% | 1.194% | 3.297% | 5.93% | 3.002% | 1.916% | 1.602% | 5.337% | 0.1655% | 0.5199% | [6 empty] | 0.4971% | 1.167% | 1.337% | 1.424%
|  | Sample Concentration SD |  |  | 816.9225 | 665.9221 | 0.90752 | 0.11423 | 1236.6821 | 246.3251 | 1731.2694 | 0.39367 | 0.42311 | 0.10515 | 0.28482 | 0.14122 | 0.73887 | 0.35612 | 0.3936 | 0.40626 | 0.089095 | 0.2499 | 0.14153 | 0.84836 | 0.14779 | 0.34744 | 0.61498 | 0.30218 | 0.18838 | 0.16922 | 0.54599 | 0.033718 | 0.10876 | [6 empty] | 0.047468 | 0.1115 | 0.131 | 0.14461
| 156 | 10_15_21_F_9 |  | 06/17/2023 03:56:52 | 8102.7062 | 6954.4507 | 4.8346 | 2.9816 | 18217.2736 | 1955.6885 | 36413.1928 | 0.57781 | 0.10968 | 0.12751 | 7.9751 | 4.7308 | 8.3261 | 0.012742 | 0.23535 | 2.363 | 2.4114 | 4.1252 | 2.6276 | 2.6289 | 2.7552 | 0.42187 | 0.070854 | 0.082572 | 0.065844 | 0.38652 | 0.052124 | 7.3815 | 7.5648 | [6 empty] | 0.013028 | 0.0081197 | 0.086792 | 0.090247
|  | IS % Recovery | [31 empty] | 69.92% | 55.76% | 60.72% | 62.6% | 56.63% | 0%
|  | Sample RSD |  |  | 1.431% | 1.434% | 5.073% | 1.009% | 2.775% | 2.071% | 2.234% | 4.207% | 7.485% | 5.77% | 0.7352% | 0.9241% | 0.9594% | 10.63% | 9.258% | 2.703% | 1.173% | 1.684% | 4.477% | 10.3% | 1.345% | 49.76% | 53.34% | 38.49% | 21.09% | 24.66% | 61.39% | 1.781% | 0.9163% | [6 empty] | 13.94% | 18.74% | 4.923% | 1.716%
|  | Sample Concentration SD |  |  | 115.928 | 99.7293 | 0.24528 | 0.03007 | 505.5629 | 40.5 | 813.3038 | 0.024306 | 0.0082101 | 0.0073576 | 0.058634 | 0.043715 | 0.079882 | 0.0013547 | 0.021789 | 0.063864 | 0.028281 | 0.069453 | 0.11763 | 0.27076 | 0.037064 | 0.20991 | 0.037793 | 0.031786 | 0.013887 | 0.095306 | 0.031997 | 0.13145 | 0.069316 | [6 empty] | 0.0018167 | 0.0015217 | 0.0042725 | 0.0015482
| 157 | 10_15_21_F_9 |  | 06/17/2023 04:03:12 | 8150.1238 | 6942.8337 | 4.8298 | 2.4293 | 18057.9904 | 1921.0344 | 36289.5407 | 0.55985 | 0.10642 | 0.12468 | 7.9007 | 4.7679 | 8.0582 | 0.0095882 | 0.22446 | 2.3313 | 2.4272 | 4.0293 | 2.6614 | 2.5905 | 2.7507 | 0.62515 | 0.07126 | 0.071393 | 0.047871 | 0.43153 | 0.077226 | 7.333 | 7.547 | [6 empty] | 0.015327 | 0.0085878 | 0.088836 | 0.092073
|  | IS % Recovery | [31 empty] | 69.62% | 55.42% | 60.65% | 62.32% | 56.18% | 0%
|  | Sample RSD |  |  | 0.8581% | 2.562% | 6.301% | 2.207% | 1.907% | 5.371% | 0.8316% | 3.627% | 6.469% | 6.99% | 1.096% | 0.3914% | 2.052% | 20.75% | 13.98% | 3.261% | 1.568% | 1.136% | 1.271% | 8.356% | 0.8498% | 4.403% | 31.14% | 19.89% | 27.96% | 14.97% | 14.42% | 1.017% | 1.706% | [6 empty] | 12.04% | 19.08% | 3.257% | 1.615%
|  | Sample Concentration SD |  |  | 69.9386 | 177.8622 | 0.30432 | 0.053605 | 344.3038 | 103.1726 | 301.7972 | 0.020304 | 0.0068844 | 0.0087154 | 0.086605 | 0.018661 | 0.16539 | 0.0019891 | 0.031382 | 0.076029 | 0.03807 | 0.045759 | 0.033815 | 0.21647 | 0.023375 | 0.027527 | 0.022188 | 0.014197 | 0.013386 | 0.06461 | 0.011138 | 0.074567 | 0.12879 | [6 empty] | 0.0018452 | 0.0016382 | 0.0028936 | 0.0014874
| 158 | 10_15_21_F_10 |  | 06/17/2023 04:09:13 | 6299.0039 | 7642.8526 | 5.376 | 3.9528 | 14518.378 | 1736.9165 | 34055.4244 | 0.63737 | 0.13138 | 0.14164 | 7.1129 | 5.3923 | 7.6122 | 0.011977 | 0.33376 | 1.3945 | 1.4571 | 3.226 | 2.9094 | 2.7114 | 3.1781 | 0.53466 | 0.14439 | 0.080377 | 0.066856 | 0.36566 | 0.017569 | 3.2005 | 3.5127 | [6 empty] | 0.0084272 | 0.0047624 | 0.23764 | 0.22821
|  | IS % Recovery | [31 empty] | 69.08% | 52.96% | 60.85% | 63.27% | 55.44% | 0%
|  | Sample RSD |  |  | 1.298% | 0.06311% | 4.801% | 0.9876% | 2.729% | 1.399% | 1.073% | 4.535% | 11.2% | 5.072% | 0.2581% | 5.352% | 2.786% | 27.69% | 1.137% | 2.369% | 1.544% | 1.737% | 1.683% | 7.367% | 7.03% | 22.99% | 32.72% | 29.37% | 27.01% | 5.791% | 193.1% | 0.9782% | 6.912% | [6 empty] | 34.49% | 33.6% | 4.151% | 2.166%
|  | Sample Concentration SD |  |  | 81.7863 | 4.8234 | 0.25811 | 0.039039 | 396.1748 | 24.2948 | 365.5152 | 0.028907 | 0.014709 | 0.0071839 | 0.018357 | 0.28858 | 0.21211 | 0.0033165 | 0.0037961 | 0.03303 | 0.022497 | 0.056038 | 0.04896 | 0.19974 | 0.22341 | 0.12291 | 0.047241 | 0.02361 | 0.018056 | 0.021176 | 0.033922 | 0.031307 | 0.24279 | [6 empty] | 0.0029068 | 0.0016003 | 0.0098649 | 0.0049427
| 159 | 10_15_21_F_10 |  | 06/17/2023 04:15:32 | 6195.3918 | 7319.3991 | 5.6836 | 3.7716 | 14253.6366 | 1690.3526 | 33381.6621 | 0.60657 | 0.12683 | 0.16107 | 8.5265 | 5.0844 | 7.3288 | 0.0087652 | 0.32234 | 1.3962 | 1.7344 | 3.8172 | 2.9324 | 2.6849 | 2.9969 | 0.4144 | 0.10164 | 0.077081 | 0.074064 | 0.3276 | 0.072154 | 3.777 | 3.3264 | [6 empty] | 0.0078382 | 0.0058589 | 0.23614 | 0.23336
|  | IS % Recovery | [31 empty] | 68.79% | 54.69% | 54.14% | 62.08% | 55.7% | 0%
|  | Sample RSD |  |  | 2.017% | 1.784% | 5.283% | 0.8593% | 1.37% | 5.366% | 2.386% | 3.792% | 12.2% | 14.94% | 18.68% | 1.388% | 3.241% | 60.87% | 8.627% | 4.543% | 19.3% | 17.19% | 2.053% | 2.355% | 2.395% | 24.4% | 45.56% | 18.19% | 16.56% | 3.947% | 68.29% | 18.08% | 1.681% | [6 empty] | 35.62% | 37.08% | 1.977% | 1.637%
|  | Sample Concentration SD |  |  | 124.9445 | 130.5962 | 0.30024 | 0.03241 | 195.2145 | 90.7126 | 796.4457 | 0.023002 | 0.015475 | 0.024072 | 1.5924 | 0.070585 | 0.23751 | 0.0053351 | 0.027808 | 0.063425 | 0.33466 | 0.6562 | 0.060194 | 0.063234 | 0.071766 | 0.10111 | 0.046305 | 0.014023 | 0.012266 | 0.012931 | 0.049275 | 0.68284 | 0.055915 | [6 empty] | 0.0027919 | 0.0021722 | 0.0046686 | 0.0038191
| 160 | 10_15_21_F_10_DUP |  | 06/17/2023 04:21:32 | 6217.952 | 7506.678 | 6.1406 | 3.5061 | 14037.0184 | 1770.1345 | 34553.1907 | 0.64102 | 0.099026 | 0.095613 | 6.8738 | 4.9928 | 7.4375 | 0.012175 | 0.15514 | 1.3159 | 1.3214 | 2.582 | 2.4829 | 2.6549 | 2.6722 | 0.65078 | 0.050031 | 0.058871 | 0.044605 | 0.30788 | 0.029009 | 3.0638 | 3.2789 | [6 empty] | 0.0085942 | 0.0035429 | 0.38288 | 0.35871
|  | IS % Recovery | [31 empty] | 70.18% | 54.25% | 61.67% | 63.43% | 55.29% | 0%
|  | Sample RSD |  |  | 0.544% | 2.954% | 7.45% | 0.4536% | 3.687% | 0.8091% | 1.133% | 6.352% | 15.93% | 4.766% | 0.7851% | 2.147% | 3.415% | 17.81% | 2.779% | 3.94% | 0.8292% | 1.488% | 1.405% | 10.23% | 2.494% | 24.01% | 76.27% | 43.36% | 33.53% | 13.75% | 198.6% | 0.5743% | 1.521% | [6 empty] | 35.17% | 33.26% | 3.26% | 1.381%
|  | Sample Concentration SD |  |  | 33.8235 | 221.749 | 0.45746 | 0.015904 | 517.5306 | 14.3221 | 391.325 | 0.040721 | 0.015776 | 0.0045566 | 0.053969 | 0.10722 | 0.25397 | 0.0021679 | 0.0043113 | 0.051851 | 0.010957 | 0.038421 | 0.034883 | 0.2717 | 0.066647 | 0.15623 | 0.038157 | 0.025524 | 0.014954 | 0.042345 | 0.057606 | 0.017594 | 0.049864 | [6 empty] | 0.0030222 | 0.0011782 | 0.012483 | 0.0049533
| 161 | 10_15_21_F_10_DUP |  | 06/17/2023 04:27:51 | 6073.8187 | 7368.8303 | 5.6486 | 3.4927 | 13772.809 | 1723.6082 | 33934.5645 | 0.61669 | 0.10096 | 0.097561 | 7.0339 | 4.9706 | 7.3089 | 0.0073058 | 0.16144 | 1.2764 | 1.3309 | 2.6238 | 2.421 | 2.3436 | 2.6484 | 0.3797 | 0.050117 | 0.058691 | 0.045758 | 0.38427 | 0.041455 | 3.1715 | 3.2682 | [6 empty] | 0.0086233 | 0.0038039 | 0.36084 | 0.36895
|  | IS % Recovery | [31 empty] | 71.4% | 54.25% | 60.34% | 62.5% | 54.7% | 0%
|  | Sample RSD |  |  | 1.179% | 1.729% | 10.49% | 1.24% | 0.1674% | 1.963% | 2.601% | 3.954% | 25.76% | 5.628% | 1.494% | 0.7548% | 3.059% | 67.37% | 6.946% | 2.249% | 1.701% | 2% | 1.67% | 4.884% | 0.8772% | 29.52% | 77.04% | 27.28% | 36.21% | 14.11% | 94.42% | 1.792% | 0.9033% | [6 empty] | 25.69% | 37.26% | 3.461% | 0.805%
|  | Sample Concentration SD |  |  | 71.6029 | 127.4365 | 0.59274 | 0.043299 | 23.0521 | 33.8425 | 882.6216 | 0.024381 | 0.026005 | 0.0054904 | 0.10508 | 0.037516 | 0.22357 | 0.0049217 | 0.011214 | 0.028705 | 0.02264 | 0.052479 | 0.040437 | 0.11446 | 0.023232 | 0.11208 | 0.038608 | 0.016012 | 0.01657 | 0.05423 | 0.03914 | 0.056827 | 0.029523 | [6 empty] | 0.0022155 | 0.0014175 | 0.012487 | 0.0029699
| 168 | 10_15_21_F_11 |  | 06/17/2023 05:10:50 | 5592.9865 | 7242.6501 | 5.3809 | 2.8956 | 13472.9781 | 1637.2603 | 32214.5355 | 0.59902 | 0.045408 | 0.052248 | 6.4671 | 4.3891 | 6.6122 | 0.010852 | 0.16026 | 1.3174 | 1.3805 | 3.8596 | 2.9805 | 2.7283 | 3.0592 | 0.48999 | 0.10636 | 0.06849 | 0.075806 | 0.48096 | 0.083389 | 2.7031 | 2.8195 | [6 empty] | 0.006384 | 0.0037648 | 0.074611 | 0.075565
|  | IS % Recovery | [31 empty] | 67.48% | 52.18% | 59.74% | 61.58% | 53.35% | 0%
|  | Sample RSD |  |  | 2.264% | 2.241% | 8.472% | 0.9493% | 3.843% | 2.161% | 1.936% | 4.223% | 24.92% | 10.78% | 2.3% | 0.5314% | 3.464% | 25.9% | 19.72% | 3.572% | 0.852% | 1.172% | 2.614% | 9.601% | 1.669% | 25.65% | 45.78% | 37.72% | 21.13% | 18.91% | 68.19% | 1.615% | 0.3512% | [6 empty] | 14.27% | 10.46% | 3.101% | 1.132%
|  | Sample Concentration SD |  |  | 126.6145 | 162.3355 | 0.45586 | 0.027488 | 517.7169 | 35.3741 | 623.5303 | 0.025294 | 0.011317 | 0.0056314 | 0.14872 | 0.023325 | 0.22903 | 0.002811 | 0.031597 | 0.047051 | 0.011762 | 0.04524 | 0.077905 | 0.26194 | 0.051048 | 0.1257 | 0.048691 | 0.025833 | 0.016021 | 0.090971 | 0.056859 | 0.043649 | 0.0099015 | [6 empty] | 0.00091095 | 0.00039389 | 0.0023138 | 0.00085511
| 169 | 10_15_21_F_11 |  | 06/17/2023 05:17:09 | 5709.2532 | 7276.4977 | 6.0489 | 2.9814 | 14036.4157 | 1622.1786 | 32411.0066 | 0.63879 | 0.047943 | 0.056146 | 6.4569 | 4.3605 | 6.5465 | 0.0062589 | 0.16589 | 1.3106 | 1.388 | 3.9062 | 2.9168 | 2.8841 | 3.0314 | 0.42432 | 0.060216 | 0.093855 | 0.066084 | 0.44742 | 0.044583 | 2.7347 | 2.8179 | [6 empty] | 0.0081176 | 0.0047789 | 0.073105 | 0.059691
|  | IS % Recovery | [31 empty] | 67.05% | 51.69% | 59.01% | 71.43% | 52.47% | 0%
|  | Sample RSD |  |  | 1.364% | 0.2476% | 3.732% | 4.178% | 2.286% | 2.602% | 1.869% | 2.379% | 14.05% | 5.327% | 0.8801% | 1.6% | 2.218% | 43.13% | 12.97% | 3.327% | 1.849% | 1.053% | 3.297% | 6.843% | 0.9334% | 28.93% | 21.62% | 44.28% | 36.68% | 8.871% | 24.31% | 1.547% | 0.8387% | [6 empty] | 17.49% | 4.917% | 3.104% | 6.176%
|  | Sample Concentration SD |  |  | 77.8462 | 18.0188 | 0.22574 | 0.12458 | 320.8903 | 42.2011 | 605.6787 | 0.015196 | 0.0067367 | 0.002991 | 0.056826 | 0.069782 | 0.14522 | 0.0026991 | 0.021513 | 0.043596 | 0.025666 | 0.041125 | 0.096177 | 0.19736 | 0.028294 | 0.12274 | 0.01302 | 0.041561 | 0.024237 | 0.039692 | 0.01084 | 0.042305 | 0.023632 | [6 empty] | 0.0014199 | 0.00023496 | 0.0022692 | 0.0036868
| 170 | 10_15_21_UF_1 |  | 06/17/2023 05:23:09 | 10677.3627 | 6508.6069 | -0.091374 | 4.4228 | 30497.0149 | 2696.2837 | 43622.0162 | 0.17706 | 0.14795 | 0.13663 | 0.90798 | 18.9121 | 1.8741 | 0.020279 | 0.32587 | 0.51999 | 0.59483 | 1.7301 | 2.5021 | 2.4878 | 2.5515 | 0.66694 | 0.036776 | 0.046769 | 0.043952 | 0.54831 | 0.0049962 | 20.2131 | 20.74 | [6 empty] | 0.0094326 | 0.0058988 | 0.027038 | 0.032148
|  | IS % Recovery | [31 empty] | 67.15% | 51.76% | 58.55% | 60.23% | 52.18% | 0%
|  | Sample RSD |  |  | 1.468% | 2.191% | 151.8% | 1.702% | 0.8603% | 1.83% | 0.7695% | 6.803% | 6.566% | 5.162% | 4.515% | 0.5548% | 3.472% | 20.68% | 10.35% | 1.178% | 2.676% | 1.314% | 1.926% | 10.19% | 0.9397% | 24.31% | 73.57% | 49.07% | 36.78% | 12.15% | 458.8% | 1.781% | 0.7251% | [6 empty] | 19.55% | 23.58% | 4.942% | 9.292%
|  | Sample Concentration SD |  |  | 156.7639 | 142.6239 | 0.13866 | 0.075258 | 262.3631 | 49.3462 | 335.6534 | 0.012045 | 0.0097142 | 0.0070533 | 0.040998 | 0.10492 | 0.065072 | 0.004194 | 0.033734 | 0.0061276 | 0.015915 | 0.022733 | 0.048188 | 0.25363 | 0.023977 | 0.16211 | 0.027055 | 0.022951 | 0.016164 | 0.066609 | 0.022924 | 0.35998 | 0.15038 | [6 empty] | 0.001844 | 0.0013911 | 0.0013362 | 0.0029872
| 171 | 10_15_21_UF_1 |  | 06/17/2023 05:29:28 | 10912.8776 | 6553.5644 | 0.036679 | 4.675 | 31010.5811 | 2777.1165 | 44445.534 | 0.1867 | 0.15768 | 0.13581 | 0.88232 | 19.2145 | 2.0467 | 0.022778 | 0.3689 | 0.52478 | 0.60167 | 1.7154 | 2.416 | 2.3614 | 2.5725 | 0.66903 | 0.043898 | 0.050948 | 0.037463 | 0.48918 | 0.030998 | 20.2482 | 20.9722 | [6 empty] | 0.0091044 | 0.0062453 | 0.026177 | 0.030681
|  | IS % Recovery | [31 empty] | 69.17% | 52.44% | 58.93% | 60.53% | 53.04% | 0%
|  | Sample RSD |  |  | 1.711% | 0.1893% | 252.8% | 1.995% | 0.256% | 1.55% | 0.8389% | 6.119% | 4.594% | 3.794% | 3.123% | 1.597% | 2.502% | 9.278% | 3.345% | 4.454% | 2.059% | 2.1% | 2.241% | 8% | 2.275% | 13.79% | 79.64% | 13.45% | 15.14% | 8.975% | 159.6% | 1.189% | 1.11% | [6 empty] | 11.64% | 29.17% | 7.853% | 8%
|  | Sample Concentration SD |  |  | 186.7452 | 12.4078 | 0.092709 | 0.093262 | 79.3874 | 43.0461 | 372.8345 | 0.011424 | 0.0072436 | 0.0051523 | 0.027554 | 0.30694 | 0.051203 | 0.0021134 | 0.012341 | 0.023376 | 0.012388 | 0.036017 | 0.054155 | 0.1889 | 0.058523 | 0.092247 | 0.034959 | 0.0068503 | 0.0056707 | 0.043902 | 0.049468 | 0.24082 | 0.23277 | [6 empty] | 0.0010594 | 0.0018218 | 0.0020558 | 0.0024546
| 172 | 10_15_21_UF_6 |  | 06/17/2023 05:35:27 | 9778.1275 | 6907.1205 | 0.0462 | 2.8905 | 22982.8535 | 2177.2132 | 40607.0865 | 0.50769 | 0.069257 | 0.068044 | 1.7509 | 7.575 | 2.737 | 0.014942 | 0.2576 | 1.227 | 1.3191 | 3.8486 | 2.5823 | 2.4143 | 2.6986 | 0.63864 | 0.096684 | 0.081297 | 0.071793 | 0.44489 | 0.055334 | 10.8516 | 11.2565 | [6 empty] | 0.013609 | 0.013152 | 0.044547 | 0.044508
|  | IS % Recovery | [31 empty] | 67.7% | 53.32% | 59.25% | 61.12% | 53.65% | 0%
|  | Sample RSD |  |  | 1.18% | 3.343% | 220% | 2.715% | 3.836% | 2.275% | 2.563% | 0.4401% | 7.59% | 13.67% | 4.688% | 0.4486% | 5.939% | 35.19% | 15.13% | 1.706% | 2.453% | 1.638% | 2.474% | 4.836% | 2.895% | 9.856% | 24.55% | 35.03% | 25.11% | 18.61% | 33.74% | 0.9581% | 0.8412% | [6 empty] | 12.25% | 6.596% | 6.095% | 3.257%
|  | Sample Concentration SD |  |  | 115.3996 | 230.8709 | 0.10165 | 0.078477 | 881.6603 | 49.5359 | 1040.9258 | 0.0022344 | 0.0052566 | 0.0093017 | 0.082083 | 0.033984 | 0.16255 | 0.0052582 | 0.038976 | 0.020935 | 0.032355 | 0.063037 | 0.06389 | 0.11675 | 0.078123 | 0.062944 | 0.023735 | 0.02848 | 0.018026 | 0.082797 | 0.018667 | 0.10397 | 0.094685 | [6 empty] | 0.0016666 | 0.00086753 | 0.0027151 | 0.0014494
| 173 | 10_15_21_UF_6 |  | 06/17/2023 05:41:46 | 9509.7211 | 6878.5393 | 0.13504 | 2.8128 | 22388.3122 | 2150.7365 | 39413.0655 | 0.47552 | 0.071657 | 0.065158 | 1.7019 | 7.5941 | 2.5708 | 0.013491 | 0.23676 | 1.264 | 1.3155 | 3.8334 | 2.6357 | 2.7498 | 2.6374 | 0.62927 | 0.074189 | 0.099048 | 0.072203 | 0.51086 | 0.036481 | 10.7183 | 11.203 | [6 empty] | 0.018228 | 0.012559 | 0.044285 | 0.046527
|  | IS % Recovery | [31 empty] | 66.32% | 53% | 59.97% | 60.89% | 53.89% | 0%
|  | Sample RSD |  |  | 0.9649% | 0.5068% | 94.29% | 5.335% | 2.144% | 4.567% | 0.1558% | 5.623% | 13.46% | 7.026% | 3.218% | 0.1559% | 3.237% | 32.37% | 4.324% | 2.837% | 0.9331% | 0.6934% | 4.719% | 1.657% | 0.7632% | 34.99% | 0.2632% | 43.23% | 14.99% | 11.32% | 139.7% | 1.126% | 0.2292% | [6 empty] | 7.524% | 9.38% | 2.142% | 1.53%
|  | Sample Concentration SD |  |  | 91.7622 | 34.8611 | 0.12733 | 0.15006 | 479.9291 | 98.2176 | 61.4175 | 0.026737 | 0.0096428 | 0.004578 | 0.054769 | 0.011836 | 0.083223 | 0.0043669 | 0.010238 | 0.035857 | 0.012275 | 0.02658 | 0.12439 | 0.045571 | 0.020128 | 0.22018 | 0.00019523 | 0.042815 | 0.010825 | 0.057828 | 0.050965 | 0.12067 | 0.02568 | [6 empty] | 0.0013714 | 0.001178 | 0.00094864 | 0.00071183
| 174 | 10_15_21_UF_8 |  | 06/17/2023 05:47:46 | 10167.1898 | 7133.965 | 1.4165 | 2.6727 | 22209.4721 | 2276.7367 | 40130.2502 | 0.65756 | 0.10186 | 0.079606 | 0.76854 | 7.8594 | 1.8432 | 0.014035 | 0.256 | 1.1536 | 1.2398 | 3.4563 | 3.272 | 2.9149 | 3.2465 | 0.59386 | 0.12179 | 0.080902 | 0.06668 | 0.49962 | 0.070402 | 9.9962 | 10.4041 | [6 empty] | 0.0099515 | 0.0074862 | 0.03899 | 0.04215
|  | IS % Recovery | [31 empty] | 66.64% | 52.13% | 58.84% | 60.16% | 53.04% | 0%
|  | Sample RSD |  |  | 1.707% | 1.284% | 14.59% | 2.033% | 0.6255% | 0.3793% | 0.9042% | 1.534% | 3.423% | 10.93% | 1.639% | 0.3297% | 5.609% | 29.69% | 10.31% | 3.86% | 1.117% | 0.4475% | 2.24% | 17.3% | 1.084% | 28.86% | 32.74% | 36.62% | 31.44% | 12.22% | 64.8% | 0.9742% | 0.4214% | [6 empty] | 12.39% | 6.719% | 2.118% | 4.462%
|  | Sample Concentration SD |  |  | 173.5856 | 91.5751 | 0.20668 | 0.054332 | 138.9287 | 8.6351 | 362.8431 | 0.010087 | 0.0034869 | 0.0087048 | 0.012596 | 0.025911 | 0.10339 | 0.0041667 | 0.026393 | 0.044528 | 0.013843 | 0.015468 | 0.073296 | 0.50433 | 0.035182 | 0.1714 | 0.039877 | 0.029623 | 0.020963 | 0.061055 | 0.04562 | 0.097387 | 0.04384 | [6 empty] | 0.0012326 | 0.00050303 | 0.00082574 | 0.0018807
| 175 | 10_15_21_UF_8 |  | 06/17/2023 05:54:04 | 10308.5276 | 7292.6146 | 1.4987 | 2.7549 | 22366.0507 | 2222.1842 | 40438.484 | 0.69611 | 0.10715 | 0.08496 | 0.74773 | 7.8134 | 1.8393 | 0.014904 | 0.24387 | 1.1724 | 1.2325 | 3.4805 | 3.1527 | 2.9835 | 3.2151 | 0.70476 | 0.091005 | 0.075842 | 0.084091 | 0.35646 | 0.044069 | 10.0139 | 10.3562 | [6 empty] | 0.0093285 | 0.0067423 | 0.041126 | 0.042234
|  | IS % Recovery | [31 empty] | 66.7% | 52.1% | 58.41% | 60.65% | 52.45% | 0%
|  | Sample RSD |  |  | 1.965% | 2.433% | 5.567% | 2.77% | 1.548% | 3.293% | 2.203% | 2.82% | 11.68% | 2.197% | 2.917% | 1.33% | 6.825% | 20.44% | 11.16% | 4.33% | 1.149% | 1.625% | 3.617% | 6.008% | 0.3015% | 32.8% | 39.38% | 48.49% | 15.63% | 5.22% | 24.76% | 0.3212% | 1.39% | [6 empty] | 16.43% | 29.46% | 5.711% | 2.792%
|  | Sample Concentration SD |  |  | 202.5784 | 177.4214 | 0.08343 | 0.0763 | 346.2879 | 73.1733 | 890.806 | 0.019629 | 0.012516 | 0.0018669 | 0.021809 | 0.1039 | 0.12554 | 0.0030461 | 0.027223 | 0.050765 | 0.014157 | 0.056563 | 0.11404 | 0.17925 | 0.0096927 | 0.23119 | 0.035836 | 0.036775 | 0.013144 | 0.018608 | 0.010913 | 0.032168 | 0.14397 | [6 empty] | 0.0015323 | 0.001986 | 0.0023486 | 0.0011792
| 186 | 10_15_21_UF_8_DUP |  | 06/17/2023 07:01:37 | 10384.274 | 7292.7311 | 0.011584 | 2.6336 | 22873.8333 | 2289.1992 | 41227.5619 | 0.69471 | 0.12987 | 0.10074 | 1.5959 | 7.6724 | 2.7848 | 0.013787 | 0.25072 | 1.2456 | 1.2584 | 2.6476 | 3.1381 | 2.9242 | 3.2768 | 0.61622 | 0.078082 | 0.13022 | 0.089055 | 0.50312 | 0.067117 | 10.07 | 10.3259 | [6 empty] | 0.0097243 | 0.0070001 | 0.041976 | 0.046897
|  | IS % Recovery | [31 empty] | 65.07% | 50.15% | 56.87% | 58.89% | 50.96% | 0%
|  | Sample RSD |  |  | 1.855% | 3.992% | 1097% | 3.736% | 3.092% | 2.321% | 3.682% | 3.838% | 15.1% | 1.895% | 0.6143% | 0.5931% | 6.223% | 29.37% | 9.614% | 2.524% | 1.309% | 1.002% | 3.263% | 7.596% | 2.573% | 33.08% | 62.01% | 31.01% | 14.84% | 17.5% | 17.76% | 0.516% | 0.8113% | [6 empty] | 34.11% | 7.455% | 2.164% | 3.003%
|  | Sample Concentration SD |  |  | 192.6561 | 291.1578 | 0.12709 | 0.098384 | 707.2249 | 53.1362 | 1517.8288 | 0.026666 | 0.019604 | 0.0019089 | 0.0098038 | 0.045509 | 0.17329 | 0.0040497 | 0.024105 | 0.031438 | 0.016472 | 0.026533 | 0.10238 | 0.22212 | 0.084316 | 0.20387 | 0.048418 | 0.040379 | 0.013212 | 0.088062 | 0.011917 | 0.051959 | 0.083773 | [6 empty] | 0.0033173 | 0.00052186 | 0.00090835 | 0.0014083
| 187 | 10_15_21_UF_8_DUP |  | 06/17/2023 07:07:56 | 9850.2514 | 6887.3176 | -0.048696 | 2.8155 | 22047.339 | 2217.6119 | 39826.8979 | 0.68307 | 0.11885 | 0.10378 | 1.6106 | 7.6629 | 2.5062 | 0.011616 | 0.24294 | 1.1865 | 1.2277 | 2.6478 | 3.1327 | 2.9767 | 3.1707 | 0.55869 | 0.078707 | 0.079626 | 0.086299 | 0.39406 | 0.067957 | 9.8998 | 10.2526 | [6 empty] | 0.012761 | 0.007404 | 0.043746 | 0.046002
|  | IS % Recovery | [31 empty] | 64.78% | 49.73% | 56.66% | 58.3% | 50.21% | 0%
|  | Sample RSD |  |  | 1.71% | 2.83% | 41.93% | 3.883% | 1.353% | 3.154% | 1.648% | 5.957% | 9.958% | 4.271% | 3.917% | 1.992% | 2.37% | 20.63% | 10.96% | 3.942% | 1.161% | 3.14% | 0.7473% | 5.635% | 1.756% | 32.49% | 0.4778% | 40.27% | 22.67% | 27.73% | 63.59% | 0.4942% | 1.553% | [6 empty] | 10.35% | 29.38% | 4.322% | 3.52%
|  | Sample Concentration SD |  |  | 168.402 | 194.8767 | 0.020419 | 0.10932 | 298.3776 | 69.9498 | 656.2974 | 0.040692 | 0.011835 | 0.0044319 | 0.063082 | 0.15266 | 0.059408 | 0.0023962 | 0.026616 | 0.046766 | 0.014252 | 0.083136 | 0.023411 | 0.16773 | 0.055691 | 0.18153 | 0.00037609 | 0.032067 | 0.019564 | 0.10926 | 0.043212 | 0.048924 | 0.15919 | [6 empty] | 0.001321 | 0.0021752 | 0.0018907 | 0.0016191
| 188 | 10_15_21_UF_8_DUP_LSP |  | 06/17/2023 07:13:55 | 18258.2928 | 15957.0251 | 11.0449 | 15.7559 | 30341.579 | 11662.155 | 46464.476 | 10.7604 | 10.5319 | 10.3998 | 12.0851 | 16.998 | 11.9649 | 9.7962 | 9.4143 | 10.5352 | 10.681 | 11.525 | 12.0351 | 12.7318 | 12.5689 | 10.1763 | 10.1885 | 9.8168 | 9.7248 | 10.3647 | 9.7098 | 18.8764 | 19.5579 | [6 empty] | 9.4169 | 9.424 | 10.2377 | 10.4079
|  | IS % Recovery | [31 empty] | 65.14% | 48.71% | 55.52% | 56.7% | 48.89% | 0%
|  | Sample RSD |  |  | 1.772% | 0.2693% | 6.093% | 5.682% | 0.191% | 1.553% | 1.741% | 0.5517% | 1.673% | 0.758% | 0.2707% | 0.9826% | 1.445% | 1.886% | 1.689% | 0.794% | 0.6089% | 0.497% | 0.5844% | 0.8744% | 0.4299% | 1.254% | 7.383% | 1.783% | 2.239% | 2.232% | 2.221% | 0.8135% | 0.4757% | [6 empty] | 0.3111% | 0.8818% | 1.511% | 0.3767%
|  | Sample Concentration SD |  |  | 323.4588 | 42.9754 | 0.67302 | 0.89519 | 57.9617 | 181.0584 | 808.8497 | 0.059361 | 0.17624 | 0.07883 | 0.03271 | 0.16703 | 0.17287 | 0.18472 | 0.15902 | 0.083654 | 0.065036 | 0.057277 | 0.070338 | 0.11133 | 0.054029 | 0.12761 | 0.75224 | 0.17507 | 0.21772 | 0.23139 | 0.21566 | 0.15356 | 0.093034 | [6 empty] | 0.0293 | 0.083097 | 0.15472 | 0.039209
| 189 | 10_15_21_UF_8_DUP_LSP |  | 06/17/2023 07:20:13 | 18290.0209 | 15456.0348 | 10.8726 | 15.6574 | 29759.8218 | 11437.4433 | 46012.4646 | 10.7682 | 10.3547 | 10.4974 | 11.9713 | 16.9751 | 11.7553 | 9.7419 | 9.3663 | 10.2005 | 10.663 | 11.5272 | 12.353 | 12.1812 | 12.4703 | 10.5643 | 10.1203 | 9.74 | 10.0045 | 10.7222 | 9.9553 | 18.9821 | 19.5368 | [6 empty] | 9.4894 | 9.5252 | 10.3605 | 10.4457
|  | IS % Recovery | [31 empty] | 64.76% | 49.64% | 56.32% | 57.22% | 50.47% | 0%
|  | Sample RSD |  |  | 1.851% | 2.039% | 3.382% | 1.264% | 1.211% | 1.744% | 1.72% | 1.905% | 1.93% | 1.483% | 0.6524% | 0.6074% | 0.5302% | 1.571% | 2.005% | 1.207% | 0.57% | 1.167% | 3.119% | 5.389% | 1.164% | 5.378% | 3.179% | 3.568% | 0.6463% | 4.098% | 3.333% | 1.151% | 0.05412% | [6 empty] | 0.3916% | 0.5218% | 3.734% | 1.137%
|  | Sample Concentration SD |  |  | 338.6357 | 315.0773 | 0.36776 | 0.1979 | 360.3649 | 199.5141 | 791.6006 | 0.20519 | 0.19985 | 0.15569 | 0.078099 | 0.1031 | 0.062324 | 0.15305 | 0.18778 | 0.12307 | 0.060777 | 0.13457 | 0.38526 | 0.65643 | 0.14513 | 0.56812 | 0.32169 | 0.34755 | 0.064663 | 0.43937 | 0.33186 | 0.21844 | 0.010573 | [6 empty] | 0.037158 | 0.049698 | 0.38691 | 0.11877
| 190 | 10_15_21_UF_9 |  | 06/17/2023 07:26:13 | 8681.2319 | 6514.3072 | 12.7994 | 6.7919 | 17805.5469 | 1935.3941 | 35857.3928 | 0.55291 | 0.11357 | 0.11179 | 0.52013 | 4.3694 | 1.4831 | 0.0095453 | 0.19127 | 1.0147 | 1.0846 | 2.8383 | 2.5441 | 2.4415 | 2.596 | 0.69525 | 0.069559 | 0.080838 | 0.061884 | 0.42996 | 0.10037 | 7.0353 | 7.2486 | [6 empty] | 0.011369 | 0.0087918 | 0.02821 | 0.028778
|  | IS % Recovery | [31 empty] | 65.56% | 50.56% | 57.25% | 58.82% | 51.24% | 0%
|  | Sample RSD |  |  | 1.027% | 2.016% | 3.737% | 1.085% | 1.76% | 2.617% | 0.3647% | 2.709% | 5.39% | 1.216% | 2.956% | 0.6749% | 2.665% | 44.81% | 3.403% | 0.7483% | 2.289% | 0.3365% | 0.6632% | 14.89% | 1.228% | 23.15% | 20.39% | 30.13% | 31.19% | 22.91% | 70.6% | 1.288% | 1.045% | [6 empty] | 28.98% | 23.06% | 2.058% | 4.369%
|  | Sample Concentration SD |  |  | 89.1306 | 131.3499 | 0.4783 | 0.073665 | 313.3143 | 50.6551 | 130.7579 | 0.014976 | 0.0061214 | 0.0013595 | 0.015373 | 0.029489 | 0.039526 | 0.0042769 | 0.0065094 | 0.0075935 | 0.024829 | 0.0095499 | 0.016873 | 0.36356 | 0.031874 | 0.16092 | 0.01418 | 0.024358 | 0.019303 | 0.098495 | 0.07086 | 0.090595 | 0.075737 | [6 empty] | 0.0032947 | 0.002027 | 0.00058057 | 0.0012573
| 191 | 10_15_21_UF_9 |  | 06/17/2023 07:32:31 | 8598.2121 | 6573.2885 | 12.3818 | 6.0986 | 18024.4361 | 1882.2505 | 35627.2489 | 0.5333 | 0.11919 | 0.10858 | 0.48428 | 4.4017 | 1.4087 | 0.0092074 | 0.1751 | 0.99183 | 1.0864 | 2.8711 | 2.453 | 2.3753 | 2.632 | 0.50747 | 0.046004 | 0.08197 | 0.050878 | 0.36164 | 0.074505 | 7.0407 | 7.275 | [6 empty] | 0.010457 | 0.0080746 | 0.026571 | 0.030637
|  | IS % Recovery | [31 empty] | 65.38% | 49.86% | 56.62% | 58.36% | 50.2% | 0%
|  | Sample RSD |  |  | 1.402% | 2.225% | 1.422% | 2.478% | 2.998% | 2.425% | 2.313% | 7.057% | 17.78% | 3.422% | 4.916% | 1.768% | 4.578% | 34.95% | 19.26% | 1.71% | 1.606% | 1.819% | 3.619% | 7.09% | 1.948% | 16.31% | 60.57% | 21.36% | 28.7% | 29.98% | 84.2% | 0.3901% | 1.215% | [6 empty] | 11.61% | 8.309% | 12.22% | 4.277%
|  | Sample Concentration SD |  |  | 120.5299 | 146.2613 | 0.1761 | 0.15114 | 540.2964 | 45.6365 | 823.936 | 0.037635 | 0.021191 | 0.0037158 | 0.023807 | 0.07784 | 0.064484 | 0.0032177 | 0.033722 | 0.016956 | 0.017444 | 0.052216 | 0.088783 | 0.16841 | 0.051276 | 0.082768 | 0.027866 | 0.017508 | 0.0146 | 0.10842 | 0.062729 | 0.027464 | 0.088391 | [6 empty] | 0.0012142 | 0.00067092 | 0.0032468 | 0.0013102
| 192 | 10_15_21_UF_10 |  | 06/17/2023 07:38:31 | 6908.1782 | 7277.051 | -0.047034 | 4.3194 | 14283.6831 | 1647.0466 | 33688.9073 | 0.61551 | 0.10345 | 0.086187 | 0.28748 | 4.5774 | 1.2897 | 0.0072488 | 0.17595 | 0.80388 | 0.8599 | 2.1185 | 2.7805 | 2.8519 | 2.8633 | 0.54562 | 0.14189 | 0.076286 | 0.070962 | 0.36588 | 0.094407 | 3.1106 | 3.1919 | [6 empty] | 0.0078488 | 0.0029637 | 0.028569 | 0.033127
|  | IS % Recovery | [31 empty] | 65.22% | 50.19% | 56.81% | 60.84% | 50.58% | 0%
|  | Sample RSD |  |  | 0.9982% | 0.8704% | 286% | 2.224% | 1.604% | 0.2545% | 1.188% | 4.562% | 9.503% | 1.609% | 8.203% | 2.113% | 8.125% | 45.64% | 8.169% | 3.936% | 2.085% | 0.357% | 2.218% | 6.205% | 1.635% | 23.3% | 53.93% | 35.96% | 6.471% | 1.799% | 33.87% | 2.524% | 2.253% | [6 empty] | 17.27% | 9.394% | 3.316% | 25.39%
|  | Sample Concentration SD |  |  | 68.9575 | 63.3428 | 0.13452 | 0.096076 | 229.048 | 4.1915 | 400.1344 | 0.028082 | 0.0098313 | 0.001387 | 0.023581 | 0.096717 | 0.10479 | 0.0033081 | 0.014374 | 0.03164 | 0.017925 | 0.0075632 | 0.061659 | 0.17696 | 0.046811 | 0.12711 | 0.076517 | 0.027431 | 0.0045921 | 0.0065836 | 0.031976 | 0.078506 | 0.071923 | [6 empty] | 0.0013552 | 0.00027841 | 0.00094743 | 0.0084105
| 193 | 10_15_21_UF_10 |  | 06/17/2023 07:44:50 | 6816.0919 | 7000.3872 | -0.18804 | 4.0854 | 14019.0841 | 1695.7911 | 33118.622 | 0.59768 | 0.098621 | 0.087848 | 0.26931 | 4.5917 | 1.1492 | 0.0065144 | 0.13791 | 0.79565 | 0.85683 | 2.1546 | 2.7302 | 2.7956 | 2.8861 | 0.61787 | 0.070913 | 0.098112 | 0.059739 | 0.40034 | 0.026922 | 3.0991 | 3.2061 | [6 empty] | 0.0075895 | 0.0031747 | 0.030267 | 0.027758
|  | IS % Recovery | [31 empty] | 65.38% | 49.5% | 56.53% | 65.53% | 49.67% | 0%
|  | Sample RSD |  |  | 0.9659% | 1.605% | 83.32% | 3.265% | 2.624% | 3.125% | 2.516% | 4.673% | 10.26% | 7.207% | 4.487% | 1.321% | 7.212% | 19.59% | 9.21% | 1.539% | 0.5672% | 1.107% | 1.214% | 5.026% | 3.131% | 13.2% | 20.31% | 19.25% | 39.92% | 44.78% | 89.42% | 1.159% | 0.729% | [6 empty] | 6.817% | 39.82% | 2.622% | 2.291%
|  | Sample Concentration SD |  |  | 65.8354 | 112.3912 | 0.15667 | 0.13339 | 367.8345 | 52.9951 | 833.3566 | 0.027931 | 0.010116 | 0.0063309 | 0.012084 | 0.060648 | 0.08288 | 0.0012762 | 0.012702 | 0.012248 | 0.0048601 | 0.023859 | 0.033141 | 0.14049 | 0.090358 | 0.081552 | 0.014405 | 0.01889 | 0.023846 | 0.17926 | 0.024074 | 0.035933 | 0.023373 | [6 empty] | 0.00051736 | 0.0012641 | 0.0007935 | 0.00063607
| 194 | 10_15_21_UF_11 |  | 06/17/2023 07:50:50 | 6563.6549 | 7219.8262 | -0.040966 | 3.3337 | 13955.4355 | 1598.3116 | 32909.7825 | 0.58637 | 0.073131 | 0.056315 | 1.4423 | 4.0029 | 2.1636 | 0.0069465 | 0.12095 | 0.78428 | 0.85014 | 3.0595 | 2.9087 | 2.7622 | 2.9895 | 0.34193 | 0.046214 | 0.071946 | 0.053482 | 0.38851 | 0.095383 | 2.6352 | 2.7668 | [6 empty] | 0.0054095 | 0.0046742 | 0.032175 | 0.032871
|  | IS % Recovery | [31 empty] | 64.29% | 49.72% | 56.69% | 60.77% | 50.14% | 0%
|  | Sample RSD |  |  | 0.837% | 1.84% | 78.91% | 2.196% | 3.591% | 4.1% | 1.248% | 2.821% | 0.3321% | 6.19% | 5.795% | 1.186% | 4.212% | 40.53% | 10.05% | 0.4831% | 0.3757% | 0.246% | 0.9556% | 2.681% | 1.421% | 29.02% | 110.1% | 27.37% | 13.31% | 18.12% | 87.54% | 0.9595% | 1.414% | [6 empty] | 47.08% | 16.53% | 5.987% | 5.437%
|  | Sample Concentration SD |  |  | 54.9379 | 132.8661 | 0.032326 | 0.073222 | 501.0859 | 65.5231 | 410.8712 | 0.016544 | 0.00024285 | 0.0034857 | 0.083575 | 0.047471 | 0.091142 | 0.0028152 | 0.012161 | 0.0037887 | 0.0031938 | 0.0075271 | 0.027795 | 0.074061 | 0.042482 | 0.099222 | 0.050875 | 0.019692 | 0.0071191 | 0.070403 | 0.083497 | 0.025286 | 0.039118 | [6 empty] | 0.0025465 | 0.00077269 | 0.0019264 | 0.0017871
| 195 | 10_15_21_UF_11 |  | 06/17/2023 07:57:09 | 6694.7831 | 7168.185 | 0.095456 | 3.5003 | 14096.4545 | 1566.5985 | 32446.2584 | 0.58537 | 0.077728 | 0.063078 | 1.4567 | 4.0649 | 2.1564 | 0.0075975 | 0.16379 | 0.80875 | 0.86873 | 3.0912 | 2.8091 | 2.7826 | 3.0339 | 0.52947 | 0.071978 | 0.099996 | 0.067411 | 0.38781 | 0.069767 | 2.6288 | 2.7605 | [6 empty] | 0.0066366 | 0.0043425 | 0.03311 | 0.036126
|  | IS % Recovery | [31 empty] | 64.63% | 48.61% | 55.97% | 57.11% | 48.85% | 0%
|  | Sample RSD |  |  | 1.598% | 3.456% | 41.94% | 1.807% | 2.113% | 2.003% | 2.093% | 8.529% | 4.335% | 18.08% | 0.9818% | 0.3526% | 4.138% | 23.65% | 21.11% | 3.054% | 1.867% | 1.404% | 2.014% | 7.943% | 1.051% | 12.41% | 71.77% | 27.67% | 22.36% | 30.29% | 96.6% | 1.032% | 1.215% | [6 empty] | 17.35% | 19.37% | 2.734% | 1.882%
|  | Sample Concentration SD |  |  | 107.0007 | 247.7106 | 0.040032 | 0.063257 | 297.9244 | 31.3832 | 679.0239 | 0.049926 | 0.0033695 | 0.011404 | 0.014302 | 0.014335 | 0.089221 | 0.001797 | 0.034568 | 0.024695 | 0.01622 | 0.043405 | 0.056584 | 0.22102 | 0.031896 | 0.065706 | 0.051658 | 0.027668 | 0.015074 | 0.11748 | 0.067395 | 0.027131 | 0.033543 | [6 empty] | 0.0011516 | 0.00084104 | 0.00090515 | 0.00067986
| 202 | 10_15_21_FB |  | 06/17/2023 08:40:08 | 309.5635 | 5.9134 | 8.7846 | 0.64889 | 660.5355 | 73.1704 | 50.8432 | 0.00037281 | 0.081431 | 0.073486 | 1.4713 | 0.058921 | 1.571 | -0.0056663 | 0.0097588 | 0.10545 | 0.11742 | 7.5315 | -0.052131 | -0.00678 | 0.0090817 | 0.13273 | 0.0055366 | 0.0051056 | -0.0070843 | -0.021988 | -0.021474 | 0.00096373 | 0.012359 | [6 empty] | 0.0011888 | -0.00094353 | 0.050446 | 0.053852
|  | IS % Recovery | [31 empty] | 67.97% | 46.53% | 59.49% | 61.32% | 51.47% | 0%
|  | Sample RSD |  |  | 9.015% | 31.62% | 7.971% | 9.802% | 12.43% | 14.39% | 12.08% | 344.8% | 6.652% | 6.073% | 1.333% | 11.95% | 2.204% | 16.62% | 49.6% | 10.69% | 7.311% | 0.6995% | 116.6% | 116.8% | 32.56% | 92.79% | 0% | 97% | 74.69% | 418.2% | 0% | 70.9% | 4.849% | [6 empty] | 52.69% | 54.57% | 4.832% | 5.13%
|  | Sample Concentration SD |  |  | 27.9066 | 1.8701 | 0.7002 | 0.063607 | 82.1357 | 10.5292 | 6.1422 | 0.0012853 | 0.0054166 | 0.0044625 | 0.019617 | 0.0070407 | 0.034626 | 0.00094156 | 0.0048399 | 0.011277 | 0.0085847 | 0.052686 | 0.060766 | 0.0079157 | 0.0029574 | 0.12316 | 0 | 0.0049526 | 0.0052911 | 0.091953 | 0 | 0.00068332 | 0.00059931 | [6 empty] | 0.00062641 | 0.00051488 | 0.0024376 | 0.0027626
| 203 | 10_15_21_FB |  | 06/17/2023 08:46:26 | 300.7921 | 6.6407 | 8.8997 | 0.585 | 577.2082 | 80.9715 | 51.9309 | 0.0026006 | 0.079389 | 0.076795 | 1.478 | 0.050685 | 1.4772 | -0.0061818 | 0.014396 | 0.09774 | 0.11869 | 7.5547 | -0.035538 | -0.006632 | 0.0082709 | 0.16972 | 0.013568 | 0.0045383 | -0.0084892 | 0.0132 | -0.021474 | -0.00030517 | 0.010018 | [6 empty] | 0.00078386 | -0.0014934 | 0.047406 | 0.052525
|  | IS % Recovery | [31 empty] | 67.01% | 50.73% | 58.38% | 60.81% | 50.53% | 0%
|  | Sample RSD |  |  | 3.93% | 33.41% | 1.202% | 5.522% | 20.36% | 15.61% | 18.49% | 58.54% | 10.9% | 7.519% | 2.129% | 6.98% | 6.016% | 27.67% | 61.07% | 19.93% | 7.222% | 0.96% | 106.9% | 123.2% | 14.11% | 60.12% | 102.5% | 192.6% | 41.91% | 833.6% | 0% | 189.4% | 20.52% | [6 empty] | 30.89% | 35.37% | 8.546% | 2.847%
|  | Sample Concentration SD |  |  | 11.8215 | 2.2187 | 0.107 | 0.032304 | 117.495 | 12.6393 | 9.6038 | 0.0015223 | 0.0086566 | 0.0057745 | 0.031471 | 0.0035377 | 0.08887 | 0.0017106 | 0.0087913 | 0.019479 | 0.0085721 | 0.072527 | 0.038006 | 0.008172 | 0.0011667 | 0.10203 | 0.013911 | 0.0087423 | 0.0035582 | 0.11004 | 0 | 0.00057788 | 0.0020554 | [6 empty] | 0.00024215 | 0.00052825 | 0.0040514 | 0.0014954
| 204 | 10_15_21_BB |  | 06/17/2023 08:52:26 | 56.1217 | 1.5306 | -0.36638 | -0.72755 | 443.6407 | 59.1756 | 44.6656 | -8.8222e-06 | 0.01434 | 0.019423 | -0.35299 | 0.002982 | -0.23378 | -0.0061632 | 0.011964 | 0.075239 | 0.085513 | -0.06074 | 0.0029923 | 0.0072197 | 0.00084329 | 0.10906 | 0.013477 | 0.0045487 | -0.0073402 | 0.0022803 | -0.021474 | -0.007672 | 0.0025852 | [6 empty] | 0.00064456 | -0.0020845 | 0.013018 | 0.019339
|  | IS % Recovery | [31 empty] | 67.81% | 51.03% | 58.8% | 60.92% | 49.91% | 0%
|  | Sample RSD |  |  | 2.838% | 73.27% | 15.01% | 1.556% | 8.1% | 10.96% | 6.041% | 17768% | 53.37% | 15.49% | 8.012% | 69.62% | 38.22% | 9.594% | 48.51% | 17.03% | 3.569% | 6.071% | 1419% | 445.5% | 75.92% | 160.6% | 102% | 96.3% | 55.69% | 966.3% | 0% | 6.872% | 42.25% | [6 empty] | 37.7% | 13.29% | 11.05% | 8.294%
|  | Sample Concentration SD |  |  | 1.5929 | 1.1215 | 0.055007 | 0.011323 | 35.9348 | 6.4862 | 2.6984 | 0.0015676 | 0.0076528 | 0.0030082 | 0.028282 | 0.002076 | 0.089351 | 0.00059129 | 0.0058036 | 0.012812 | 0.0030517 | 0.0036878 | 0.042466 | 0.032164 | 0.00064023 | 0.17515 | 0.013753 | 0.0043802 | 0.0040879 | 0.022035 | 0 | 0.00052723 | 0.0010922 | [6 empty] | 0.00024299 | 0.00027699 | 0.0014379 | 0.001604
| 205 | 10_15_21_BB |  | 06/17/2023 08:58:45 | 56.1369 | 0.23633 | -0.19821 | -0.72812 | 487.6436 | 56.437 | 46.5609 | -0.00091385 | 0.012905 | 0.020062 | -0.3422 | 0.0043811 | -0.16623 | -0.0076969 | 0.0053235 | 0.08384 | 0.082649 | -0.061945 | 0.014109 | -0.01135 | -0.0010676 | 0.00016846 | 0.0055366 | 0.0071314 | -0.0061376 | -0.022243 | -0.014663 | -0.0064993 | 0.0025805 | [6 empty] | 0.00050427 | -0.0022672 | 0.012778 | 0.019384
|  | IS % Recovery | [31 empty] | 66.93% | 50.83% | 59.3% | 60.3% | 49.92% | 0%
|  | Sample RSD |  |  | 1.432% | 1.931e-06% | 55.07% | 0.3062% | 5.446% | 1.604% | 12.54% | 0% | 50.83% | 13.66% | 0.6141% | 40.87% | 43.05% | 4.311% | 100.3% | 7.164% | 1.898% | 21.83% | 367% | 1.777e-06% | 30.8% | 46048% | 0% | 107.8% | 66.27% | 385.6% | 80.45% | 10.76% | 35.21% | [6 empty] | 0% | 4.76% | 7.674% | 4.335%
|  | Sample Concentration SD |  |  | 0.80386 | 4.5625e-09 | 0.10915 | 0.0022297 | 26.5582 | 0.90497 | 5.838 | 0 | 0.0065599 | 0.0027407 | 0.0021015 | 0.0017904 | 0.071569 | 0.00033179 | 0.005338 | 0.006006 | 0.001569 | 0.013525 | 0.051779 | 2.0164e-10 | 0.00032877 | 0.077574 | 0 | 0.0076902 | 0.0040675 | 0.085765 | 0.011797 | 0.00069942 | 0.00090864 | [6 empty] | 0 | 0.00010793 | 0.00098059 | 0.00084031
| 206 | 10_15_21_LAB_A |  | 06/17/2023 09:04:45 | 56.9618 | 0.23633 | -0.39347 | -0.7624 | 312.4661 | 55.5559 | 41.9002 | -0.00091385 | 0.038638 | 0.043247 | -0.18434 | 0.0070936 | -0.07752 | -0.0076986 | 0.00026748 | 0.088188 | 0.08877 | -0.039874 | -0.018785 | -0.01135 | -0.0012161 | 0.023675 | 0.0055366 | -0.00050913 | -0.010861 | 0.032175 | -0.021474 | -0.0030283 | 0.0064621 | [6 empty] | 0.00078049 | -0.0022502 | 0.015723 | 0.020626
|  | IS % Recovery | [31 empty] | 66.76% | 50.48% | 61.5% | 60.16% | 49.67% | 0%
|  | Sample RSD |  |  | 2.456% | 1.931e-06% | 21.09% | 1.234% | 22.87% | 2.482% | 6.631% | 0% | 26.92% | 7.052% | 9.858% | 16.14% | 147.3% | 4.272% | 2908% | 4.885% | 5.132% | 60.02% | 149.2% | 1.777e-06% | 91.55% | 279.9% | 0% | 0% | 19.12% | 165.9% | 0% | 31.65% | 34.51% | [6 empty] | 61.3% | 17.22% | 11.9% | 7.141%
|  | Sample Concentration SD |  |  | 1.3991 | 4.5625e-09 | 0.082999 | 0.0094115 | 71.4519 | 1.3787 | 2.7786 | 0 | 0.010402 | 0.0030497 | 0.018172 | 0.0011449 | 0.11416 | 0.00032892 | 0.0077776 | 0.0043081 | 0.0045553 | 0.02393 | 0.028035 | 2.0164e-10 | 0.0011134 | 0.066266 | 0 | 0 | 0.0020766 | 0.053373 | 0 | 0.00095851 | 0.0022301 | [6 empty] | 0.00047843 | 0.00038755 | 0.0018711 | 0.0014729
| 207 | 10_15_21_LAB_A |  | 06/17/2023 09:11:03 | 56.3986 | 0.23633 | -0.38434 | -0.75607 | 619.7993 | 55.4549 | 42.808 | 0.00041334 | 0.040328 | 0.046871 | -0.16125 | 0.0061968 | -0.053997 | -0.0073192 | 0.0023391 | 0.097948 | 0.093321 | -0.025752 | -0.019415 | -0.0066474 | -0.0017623 | 0.19994 | 0.0055366 | 0.0020185 | -0.0084854 | -0.023321 | -0.021474 | -0.001479 | 0.0073253 | [6 empty] | 0.00078087 | -0.0022045 | 0.017173 | 0.020528
|  | IS % Recovery | [31 empty] | 66.42% | 50.65% | 58.92% | 59.96% | 49.93% | 0%
|  | Sample RSD |  |  | 3.007% | 1.931e-06% | 19.24% | 2.276% | 21.69% | 12.76% | 3.764% | 321.6% | 10.63% | 1.495% | 2.343% | 12.11% | 102.6% | 7.765% | 278.6% | 12.93% | 7.501% | 42.78% | 188.8% | 122.5% | 45.33% | 24.21% | 0% | 216.9% | 0.2341% | 73.05% | 0% | 70.33% | 10.85% | [6 empty] | 61.35% | 30.04% | 12.96% | 5.551%
|  | Sample Concentration SD |  |  | 1.6959 | 4.5625e-09 | 0.073958 | 0.017205 | 134.4465 | 7.0783 | 1.6114 | 0.0013295 | 0.004288 | 0.00070087 | 0.0037778 | 0.0007504 | 0.05541 | 0.00056832 | 0.0065169 | 0.012665 | 0.0069999 | 0.011016 | 0.036651 | 0.0081453 | 0.00079888 | 0.048397 | 0 | 0.0043779 | 1.9864e-05 | 0.017036 | 0 | 0.0010402 | 0.00079457 | [6 empty] | 0.00047909 | 0.00066219 | 0.0022257 | 0.0011395
| 208 | 10_15_21_LAB_B |  | 06/17/2023 09:17:03 | 55.8835 | 0.89899 | -0.32662 | -0.69175 | 463.7233 | 57.0996 | 41.9431 | -0.00091385 | 0.012211 | 0.016121 | -0.36697 | 0.0028359 | -0.1743 | -0.0073101 | 0.007829 | 0.098107 | 0.14728 | 0.098667 | -0.020084 | -0.0017266 | -0.0027726 | 0.11412 | 0.0055366 | 0.0046528 | -0.0084127 | -0.085032 | -0.014538 | -0.0089123 | 0.0021771 | [6 empty] | 0.00064589 | -0.0025438 | 0.016229 | 0.020504
|  | IS % Recovery | [31 empty] | 65.81% | 49.71% | 57.63% | 59.33% | 48.71% | 0%
|  | Sample RSD |  |  | 2.247% | 127.7% | 39.86% | 1.6% | 8.056% | 11.12% | 6.805% | 0% | 18.54% | 18.62% | 3.651% | 46.41% | 56.93% | 7.86% | 88.31% | 8.795% | 56.26% | 11.67% | 144.5% | 482.7% | 40.86% | 184% | 0% | 192.2% | 43.2% | 59.68% | 82.63% | 8.327% | 33.05% | [6 empty] | 37.98% | 10.75% | 3.896% | 5.554%
|  | Sample Concentration SD |  |  | 1.2557 | 1.1478 | 0.1302 | 0.011067 | 37.3556 | 6.3474 | 2.854 | 0 | 0.0022642 | 0.0030015 | 0.013397 | 0.0013163 | 0.099235 | 0.0005746 | 0.0069142 | 0.0086283 | 0.082857 | 0.011517 | 0.029026 | 0.0083349 | 0.0011329 | 0.20995 | 0 | 0.0089408 | 0.0036341 | 0.050749 | 0.012013 | 0.00074211 | 0.00071945 | [6 empty] | 0.0002453 | 0.00027355 | 0.00063223 | 0.0011388
| 209 | 10_15_21_LAB_B |  | 06/17/2023 09:23:20 | 54.2331 | 0.23633 | -0.18095 | -0.6595 | 232.4913 | 61.6954 | 41.8829 | 0.00043462 | 0.0085678 | 0.015795 | -0.3636 | 0.0021673 | -0.23348 | -0.0071177 | 0.0019464 | 0.11115 | 0.10095 | 0.097623 | 0.029191 | -0.01135 | -0.0025741 | 0.074651 | 0.013686 | -0.00050913 | -0.01206 | -0.081376 | -0.014517 | -0.0086601 | 0.0020907 | [6 empty] | 0.00078281 | -0.002016 | 0.017091 | 0.020987
|  | IS % Recovery | [31 empty] | 65.66% | 49.31% | 58.17% | 59.69% | 48.62% | 0%
|  | Sample RSD |  |  | 1.777% | 1.931e-06% | 53.93% | 4.447% | 46.78% | 3.041% | 11.49% | 4.192% | 26.7% | 13.89% | 4.107% | 26.28% | 18.7% | 9.38% | 242.8% | 23.29% | 7.986% | 13.49% | 112.9% | 1.777e-06% | 12.67% | 16.85% | 103.1% | 0% | 0% | 17.36% | 83% | 1.148% | 32.28% | [6 empty] | 61.63% | 13.61% | 10.17% | 2.5%
|  | Sample Concentration SD |  |  | 0.96366 | 4.5625e-09 | 0.097583 | 0.029329 | 108.7672 | 1.8763 | 4.8139 | 1.8218e-05 | 0.0022879 | 0.002194 | 0.014932 | 0.00056959 | 0.04367 | 0.00066767 | 0.0047266 | 0.025891 | 0.008062 | 0.013168 | 0.032946 | 2.0164e-10 | 0.00032619 | 0.012575 | 0.014116 | 0 | 0 | 0.014131 | 0.01205 | 9.943e-05 | 0.00067491 | [6 empty] | 0.00048244 | 0.00027439 | 0.0017389 | 0.00052477
| 210 | 5_18_23_DB |  | 06/17/2023 09:29:20 | 54.7914 | 0.23633 | -0.5834 | 2.7988 | 400.8324 | 54.8712 | 42.8917 | -0.00091385 | -0.0069955 | -0.0014792 | -0.51844 | 0.00065848 | -0.3335 | -0.0071132 | -0.005392 | 0.07844 | 0.083187 | -0.14136 | -0.02964 | -0.01135 | -0.0020394 | -0.0062654 | 0.0055366 | -0.00050913 | -0.0096315 | -0.045747 | -0.021474 | -0.0094574 | 0.0014597 | [6 empty] | 0.00078973 | -0.0027816 | 0.017861 | 0.022752
|  | IS % Recovery | [31 empty] | 65.93% | 49.46% | 57.44% | 59.09% | 48.35% | 0%
|  | Sample RSD |  |  | 3.524% | 1.931e-06% | 20.04% | 4.707% | 2.453% | 5.179% | 3.874% | 0% | 32.04% | 30.96% | 1.157% | 214.7% | 8.329% | 4.646% | 93.4% | 6.362% | 0.7397% | 0.3606% | 132.1% | 1.777e-06% | 40.75% | 3090% | 0% | 0% | 21.84% | 108.4% | 0% | 4.551% | 28.05% | [6 empty] | 62.61% | 3.721% | 12.41% | 2.518%
|  | Sample Concentration SD |  |  | 1.9307 | 4.5625e-09 | 0.11693 | 0.13174 | 9.8308 | 2.8417 | 1.6615 | 0 | 0.0022411 | 0.00045801 | 0.0059977 | 0.0014136 | 0.027776 | 0.00033049 | 0.0050363 | 0.0049901 | 0.00061533 | 0.00050973 | 0.03915 | 2.0164e-10 | 0.00083112 | 0.19362 | 0 | 0 | 0.0021035 | 0.049579 | 0 | 0.00043043 | 0.00040941 | [6 empty] | 0.00049444 | 0.00010351 | 0.0022163 | 0.00057298
| 211 | 5_18_23_DB |  | 06/17/2023 09:35:39 | 54.5141 | 0.23633 | -0.59238 | 3.9776 | 284.2138 | 60.6181 | 42.9465 | 0.00045495 | -0.0039211 | -0.0013658 | -0.50863 | 0.00031322 | -0.41929 | -0.0073018 | -0.0015406 | 0.088481 | 0.08197 | -0.14344 | 0.024899 | -0.0065295 | -0.0018156 | 0.047586 | 0.0055366 | 0.0020976 | -0.010837 | -0.053863 | -0.007759 | -0.0089791 | 0.0012842 | [6 empty] | 0.00092955 | -0.0022477 | 0.018101 | 0.022976
|  | IS % Recovery | [31 empty] | 65.07% | 49.16% | 57.68% | 59.72% | 48.69% | 0%
|  | Sample RSD |  |  | 1.532% | 1.931e-06% | 28.04% | 2.046% | 30.51% | 1.058% | 10.32% | 3.157% | 70.6% | 67.95% | 0.6656% | 52.24% | 13.34% | 7.979% | 285.8% | 11.21% | 7.027% | 4.341% | 70.49% | 127.9% | 140.1% | 106.2% | 0% | 215.2% | 19.55% | 12.19% | 306.2% | 9.477% | 12.26% | [6 empty] | 45.58% | 11.95% | 7.938% | 1.305%
|  | Sample Concentration SD |  |  | 0.8352 | 4.5625e-09 | 0.16611 | 0.081398 | 86.7163 | 0.64147 | 4.432 | 1.4362e-05 | 0.0027682 | 0.00092812 | 0.0033854 | 0.00016361 | 0.055935 | 0.00058261 | 0.0044023 | 0.0099169 | 0.0057602 | 0.0062262 | 0.017552 | 0.0083495 | 0.0025436 | 0.050552 | 0 | 0.004515 | 0.002119 | 0.0065674 | 0.023755 | 0.00085093 | 0.00015747 | [6 empty] | 0.00042367 | 0.00026861 | 0.0014369 | 0.00029987
| 220 | 5_18_23_DB |  | 06/17/2023 10:30:55 | 56.0874 | 0.23633 | -0.34413 | 3.0299 | 483.6795 | 56.5471 | 42.803 | 0.00039696 | 1.6621e-05 | 0.0065821 | 1.1027 | 0.0076569 | 1.2403 | -0.0060061 | -0.0043181 | 0.081987 | 0.075413 | -0.13935 | -0.0097701 | -0.0066049 | -0.0019681 | 0.12984 | 0.0055366 | 0.0095666 | -0.0073449 | -0.0050493 | -0.0014681 | 0.0034523 | 0.0037905 | [6 empty] | 0.00050427 | -0.0019947 | 0.017713 | 0.021232
|  | IS % Recovery | [31 empty] | 66.85% | 51.14% | 60.21% | 63.36% | 50.31% | 0%
|  | Sample RSD |  |  | 2.453% | 1.931e-06% | 33% | 2.102% | 43.73% | 9.954% | 9.988% | 327.3% | 13980% | 38.12% | 3.773% | 25.67% | 12.42% | 5.933% | 49.9% | 17.21% | 4.498% | 7.245% | 307.6% | 124.4% | 45.97% | 166.2% | 0% | 120.6% | 55.6% | 395.7% | 1360% | 38.51% | 21.8% | [6 empty] | 0% | 14.71% | 15.59% | 4.099%
|  | Sample Concentration SD |  |  | 1.3757 | 4.5625e-09 | 0.11355 | 0.063695 | 211.5192 | 5.6287 | 4.275 | 0.0012991 | 0.0023236 | 0.0025092 | 0.041605 | 0.0019652 | 0.15406 | 0.00035633 | 0.0021549 | 0.014107 | 0.0033924 | 0.010096 | 0.030056 | 0.0082189 | 0.00090472 | 0.21579 | 0 | 0.01154 | 0.0040838 | 0.019978 | 0.019971 | 0.0013293 | 0.00082652 | [6 empty] | 0 | 0.00029348 | 0.0027618 | 0.00087026
| 221 | 5_18_23_DB |  | 06/17/2023 10:37:14 | 55.5641 | 0.87935 | -0.45883 | 4.2547 | 414.7607 | 59.2883 | 40.9256 | -0.00047279 | 0.00071555 | 0.0043316 | 1.0919 | 0.0080144 | 1.1945 | -0.0073199 | -0.0050174 | 0.082157 | 0.080496 | -0.14309 | -0.032794 | -0.0020638 | -0.0012618 | 0.084607 | 0.013394 | 0.0020228 | -0.010879 | -0.064205 | -0.0080586 | -0.00069779 | 0.0019706 | [6 empty] | 0.00064062 | -0.002559 | 0.015932 | 0.022555
|  | IS % Recovery | [31 empty] | 67.86% | 51.12% | 59.47% | 61.18% | 49.74% | 0%
|  | Sample RSD |  |  | 3.367% | 126.7% | 2.638% | 2.162% | 4.033% | 11.81% | 3.856% | 161.6% | 167.9% | 38.31% | 3.294% | 32.15% | 18.38% | 13.45% | 67.57% | 16.19% | 6.523% | 6.406% | 217.9% | 389.7% | 127.8% | 106.1% | 101.6% | 216.8% | 18.81% | 30.4% | 144.2% | 147.5% | 15.26% | [6 empty] | 36.86% | 7.783% | 14.28% | 8.205%
|  | Sample Concentration SD |  |  | 1.8709 | 1.1137 | 0.012105 | 0.091995 | 16.7253 | 7.0039 | 1.5781 | 0.00076393 | 0.0012016 | 0.0016596 | 0.035966 | 0.0025764 | 0.21956 | 0.00098477 | 0.0033905 | 0.013304 | 0.0052507 | 0.0091668 | 0.071471 | 0.0080426 | 0.0016132 | 0.089799 | 0.013609 | 0.0043855 | 0.0020459 | 0.019521 | 0.011618 | 0.0010292 | 0.00030066 | [6 empty] | 0.00023616 | 0.00019916 | 0.0022754 | 0.0018505
| 222 | 5_18_23_DB |  | 06/17/2023 10:43:14 | 55.0292 | 0.23633 | -0.41535 | 4.4563 | 526.0722 | 63.4354 | 46.4853 | -3.8842e-05 | 0.007231 | 0.013075 | -0.38083 | 0.0034809 | -0.10681 | 0.06763 | 0.0058478 | 0.089958 | 0.08793 | -0.09427 | -0.091237 | -0.0019545 | -0.0011934 | 0.43327 | 0.0055366 | 0.004616 | -0.0084655 | 0.013129 | -0.014676 | -0.004036 | 0.0023358 | [6 empty] | 0.00078193 | -0.0023205 | 0.025718 | 0.025116
|  | IS % Recovery | [31 empty] | 60.27% | 50.22% | 58.86% | 60.28% | 49.53% | 0%
|  | Sample RSD |  |  | 1.698% | 1.931e-06% | 20.56% | 2.586% | 12.6% | 5.832% | 15.02% | 3902% | 52.5% | 12.43% | 5.711% | 24.48% | 146% | 9.794% | 168.6% | 15.03% | 5.875% | 4.877% | 59.78% | 416.3% | 71.4% | 40.43% | 0% | 96.17% | 42.16% | 696.6% | 80.22% | 17.35% | 23.01% | [6 empty] | 30.75% | 7.585% | 27.09% | 2.63%
|  | Sample Concentration SD |  |  | 0.93463 | 4.5625e-09 | 0.085393 | 0.11522 | 66.2999 | 3.6994 | 6.9834 | 0.0015156 | 0.0037962 | 0.0016246 | 0.021751 | 0.00085222 | 0.15599 | 0.0066235 | 0.0098582 | 0.013519 | 0.0051656 | 0.0045977 | 0.054542 | 0.008137 | 0.00085214 | 0.17519 | 0 | 0.0044393 | 0.0035687 | 0.091462 | 0.011773 | 0.00070026 | 0.00053754 | [6 empty] | 0.00024046 | 0.000176 | 0.006966 | 0.00066043
| 223 | 5_18_23_DB |  | 06/17/2023 10:49:33 | 55.5987 | 0.23633 | -0.542 | 4.249 | 541.8147 | 54.1982 | 41.9035 | -2.0521e-05 | 0.0043401 | 0.0089419 | -0.35624 | 0.0030895 | -0.10843 | 0.063046 | -0.0055082 | 0.086176 | 0.090599 | -0.10312 | 0.0080738 | -0.01135 | -0.0015096 | 0.18504 | 0.037918 | 0.0020514 | -0.007219 | -0.022544 | -0.021474 | -0.0054082 | 0.0025307 | [6 empty] | 0.00092317 | -0.0027262 | 0.020813 | 0.020228
|  | IS % Recovery | [31 empty] | 64.43% | 49.76% | 58.4% | 67.1% | 48.91% | 0%
|  | Sample RSD |  |  | 1.709% | 1.931e-06% | 13.82% | 0.3475% | 33.4% | 4.782% | 0.06262% | 7540% | 158.6% | 25.65% | 1.494% | 46% | 41.28% | 10.61% | 100.8% | 13.53% | 6.35% | 6.042% | 334.6% | 1.777e-06% | 127.2% | 48.7% | 73.96% | 216.2% | 28.42% | 348.5% | 0% | 18.73% | 16.37% | [6 empty] | 45.35% | 9.918% | 1.077% | 21.03%
|  | Sample Concentration SD |  |  | 0.9499 | 4.5625e-09 | 0.074881 | 0.014764 | 180.9765 | 2.5918 | 0.026242 | 0.0015473 | 0.0068812 | 0.0022938 | 0.0053239 | 0.0014211 | 0.044764 | 0.0066892 | 0.0055517 | 0.01166 | 0.005753 | 0.00623 | 0.027019 | 2.0164e-10 | 0.0019195 | 0.090111 | 0.028045 | 0.004435 | 0.0020517 | 0.078576 | 0 | 0.0010127 | 0.00041418 | [6 empty] | 0.00041861 | 0.00027038 | 0.00022409 | 0.0042546
| 224 | 5_18_23_DB |  | 06/17/2023 10:55:33 | 56.4577 | 0.23633 | -0.49485 | 4.087 | 584.6872 | 57.1195 | 41.9141 | -0.00046435 | -0.0015507 | 0.0015378 | -0.46333 | 0.0056896 | -0.10679 | -0.0071152 | 0.00051067 | 0.065643 | 0.060561 | -0.09366 | 0.0048129 | -0.01135 | -0.0031009 | 0.03863 | 0.0055366 | -0.00050913 | -0.0058545 | -0.035078 | -0.021474 | -0.00518 | 0.0018303 | [6 empty] | 0.0010707 | -0.0022452 | 0.0097389 | 0.014763
|  | IS % Recovery | [31 empty] | 65.9% | 48.82% | 57.57% | 59.14% | 48.25% | 0%
|  | Sample RSD |  |  | 3.078% | 1.931e-06% | 30.16% | 1.857% | 40.65% | 6.171% | 11.61% | 167.7% | 248.6% | 28.74% | 4.787% | 34.51% | 87.43% | 4.635% | 461.8% | 4.192% | 12.91% | 7.008% | 756.9% | 1.777e-06% | 36.6% | 30.88% | 0% | 0% | 132.6% | 334.3% | 0% | 24.64% | 22.88% | [6 empty] | 45.82% | 16.63% | 16.73% | 12.61%
|  | Sample Concentration SD |  |  | 1.7376 | 4.5625e-09 | 0.14923 | 0.0759 | 237.6711 | 3.525 | 4.868 | 0.00077855 | 0.0038556 | 0.00044198 | 0.022181 | 0.0019638 | 0.093365 | 0.00032978 | 0.0023581 | 0.002752 | 0.0078195 | 0.0065634 | 0.036429 | 2.0164e-10 | 0.0011349 | 0.011927 | 0 | 0 | 0.0077649 | 0.11726 | 0 | 0.0012764 | 0.00041886 | [6 empty] | 0.00049057 | 0.00037349 | 0.0016297 | 0.0018622
| 225 | 5_18_23_DB |  | 06/17/2023 11:01:51 | 56.1408 | 0.23633 | -0.58117 | 3.9238 | 489.0645 | 57.4775 | 43.8027 | -0.00091385 | -0.0046186 | 1.0946e-05 | -0.46647 | 0.0044928 | -0.19377 | -0.0072986 | 0.0047816 | 0.05814 | 0.06201 | -0.096595 | -0.031544 | 0.003152 | -0.0042019 | -0.023642 | 0.0055366 | 0.012573 | -0.0059201 | -0.035507 | -0.021474 | -0.0060373 | 0.0017349 | [6 empty] | 0.00064873 | -0.0021798 | 0.0081603 | 0.016227
|  | IS % Recovery | [31 empty] | 65.64% | 49.11% | 57.07% | 59.01% | 47.6% | 0%
|  | Sample RSD |  |  | 2.497% | 1.931e-06% | 26.54% | 3.959% | 28.51% | 7.256% | 3.818% | 0% | 70.24% | 13088% | 4.91% | 12.41% | 23.36% | 14% | 208.9% | 15.84% | 5.116% | 16.66% | 75.94% | 464.6% | 15.04% | 516.4% | 0% | 130% | 95.05% | 168.6% | 0% | 15.62% | 23.84% | [6 empty] | 38.57% | 18.97% | 7.353% | 7.583%
|  | Sample Concentration SD |  |  | 1.4017 | 4.5625e-09 | 0.15424 | 0.15535 | 139.4359 | 4.1703 | 1.6725 | 0 | 0.0032443 | 0.0014326 | 0.022903 | 0.00055771 | 0.045268 | 0.0010216 | 0.009987 | 0.0092118 | 0.0031726 | 0.016089 | 0.023954 | 0.014644 | 0.000632 | 0.12207 | 0 | 0.016339 | 0.0056274 | 0.059852 | 0 | 0.00094273 | 0.00041366 | [6 empty] | 0.00025022 | 0.00041356 | 0.00059999 | 0.0012305
| 226 | 5_18_23_MFB |  | 06/17/2023 11:07:51 | 4502.9532 | 4575.0836 | 4.35 | 9.6796 | 5801.114 | 4707.6221 | 4994.9937 | 4.9413 | 5.2152 | 5.117 | 4.8635 | 5.1413 | 4.775 | 4.8546 | 4.6213 | 5.0005 | 5.0256 | 4.6334 | 4.5372 | 4.8443 | 4.9364 | 5.1613 | 5.2911 | 5.0694 | 4.9715 | 4.8808 | 4.9016 | 4.7381 | 5.1323 | [6 empty] | 4.6559 | 4.7335 | 5.656 | 5.6578
|  | IS % Recovery | [31 empty] | 62.84% | 45.61% | 54.22% | 56.11% | 46.34% | 0%
|  | Sample RSD |  |  | 2.403% | 3.422% | 7.453% | 1.294% | 2.839% | 1.912% | 1.242% | 2.341% | 2.503% | 0.8728% | 0.8931% | 1.97% | 0.8765% | 1.308% | 2.166% | 1.987% | 0.7466% | 0.3956% | 2.789% | 7.619% | 0.8321% | 3.715% | 4.501% | 2.062% | 4.971% | 1.31% | 2.689% | 0.5383% | 2.682% | [6 empty] | 1.229% | 1.478% | 0.826% | 1.732%
|  | Sample Concentration SD |  |  | 108.1949 | 156.5622 | 0.32419 | 0.12528 | 164.7136 | 90.003 | 62.0518 | 0.11565 | 0.13055 | 0.04466 | 0.043436 | 0.10128 | 0.041851 | 0.063477 | 0.10008 | 0.099374 | 0.037522 | 0.018328 | 0.12656 | 0.3691 | 0.041076 | 0.19174 | 0.23816 | 0.10455 | 0.24713 | 0.063961 | 0.13181 | 0.025506 | 0.13766 | [6 empty] | 0.05724 | 0.069952 | 0.046721 | 0.097971
| 227 | 5_18_23_MFB |  | 06/17/2023 11:14:09 | 4463.4639 | 4710.4861 | 4.735 | 9.6626 | 5434.759 | 4631.4346 | 5054.1685 | 4.8814 | 5.0092 | 5.1374 | 4.8842 | 5.0269 | 4.5976 | 4.7959 | 4.5058 | 4.9135 | 4.9987 | 4.6561 | 4.4796 | 4.7587 | 4.7943 | 5.0754 | 5.2604 | 4.7412 | 4.8838 | 5.0667 | 4.7738 | 4.749 | 4.9597 | [6 empty] | 4.7477 | 4.7746 | 5.5546 | 5.638
|  | IS % Recovery | [31 empty] | 64.46% | 47.19% | 54.88% | 57.1% | 47.12% | 0%
|  | Sample RSD |  |  | 2.95% | 1.504% | 7.103% | 1.074% | 5.236% | 2.345% | 0.8781% | 2.283% | 2.002% | 1.299% | 1.024% | 1.635% | 0.5998% | 2.625% | 2.944% | 4.292% | 0.5598% | 0.628% | 1.89% | 6.735% | 0.761% | 5.408% | 14.84% | 0.3806% | 2.308% | 5.898% | 5.173% | 0.1382% | 0.4084% | [6 empty] | 1.245% | 0.3789% | 2.594% | 1.302%
|  | Sample Concentration SD |  |  | 131.6604 | 70.824 | 0.33632 | 0.10381 | 284.5448 | 108.5885 | 44.3796 | 0.11145 | 0.10029 | 0.06675 | 0.050015 | 0.082189 | 0.027577 | 0.12589 | 0.13266 | 0.21087 | 0.027984 | 0.029242 | 0.084642 | 0.3205 | 0.036483 | 0.27449 | 0.7807 | 0.018045 | 0.11273 | 0.29883 | 0.24694 | 0.0065627 | 0.020254 | [6 empty] | 0.059087 | 0.01809 | 0.14407 | 0.07342
| 228 | 5_18_23_MFB |  | 06/17/2023 11:20:08 | 4428.9243 | 4430.2363 | 4.6763 | 9.9353 | 5721.9346 | 4590.4168 | 4902.0196 | 4.8057 | 5.0277 | 4.9683 | 4.5221 | 4.8921 | 4.2767 | 4.7135 | 4.5537 | 4.871 | 4.9143 | 4.5065 | 4.4756 | 4.3624 | 4.6548 | 5.2603 | 5.1506 | 4.8765 | 4.7152 | 4.8929 | 4.8166 | 4.5883 | 4.8284 | [6 empty] | 4.6554 | 4.6797 | 5.5783 | 6.1916
|  | IS % Recovery | [31 empty] | 63.85% | 48.18% | 55.96% | 52.41% | 48.54% | 0%
|  | Sample RSD |  |  | 0.3924% | 1.626% | 9.594% | 0.6716% | 4.235% | 3.34% | 3.461% | 3.996% | 1.88% | 1.667% | 1.328% | 0.9603% | 2.94% | 0.6175% | 0.2325% | 2.052% | 0.9718% | 0.9213% | 1.318% | 6.087% | 2.909% | 6.9% | 5.773% | 1.181% | 2.969% | 6.062% | 7.944% | 0.7474% | 0.4237% | [6 empty] | 1.382% | 1.208% | 0.2196% | 5.484%
|  | Sample Concentration SD |  |  | 17.3773 | 72.0452 | 0.44866 | 0.066721 | 242.329 | 153.327 | 169.659 | 0.19205 | 0.094504 | 0.082819 | 0.060035 | 0.046981 | 0.12575 | 0.029105 | 0.010588 | 0.099955 | 0.04776 | 0.041517 | 0.058992 | 0.26552 | 0.13541 | 0.36294 | 0.29735 | 0.057592 | 0.13997 | 0.29663 | 0.38265 | 0.034294 | 0.020458 | [6 empty] | 0.064348 | 0.056516 | 0.01225 | 0.33953
| 229 | 5_18_23_MFB |  | 06/17/2023 11:26:26 | 4388.0558 | 4647.9886 | 4.4676 | 10.0162 | 5605.7713 | 4604.6739 | 4913.2522 | 4.8383 | 5.1078 | 5.1317 | 4.6075 | 4.9353 | 4.368 | 4.7526 | 4.6859 | 4.9266 | 5.0203 | 4.5944 | 4.1423 | 4.7415 | 4.7736 | 4.6372 | 4.8249 | 4.7784 | 4.6967 | 4.4521 | 4.679 | 4.7558 | 4.9286 | [6 empty] | 4.6784 | 4.704 | 5.0777 | 5.5252
|  | IS % Recovery | [31 empty] | 68.61% | 47.39% | 55.16% | 57.05% | 47.94% | 0%
|  | Sample RSD |  |  | 2.251% | 4.854% | 5.927% | 0.9417% | 2.141% | 2.545% | 1.496% | 2.37% | 1.334% | 1.379% | 1.594% | 0.2227% | 0.3518% | 1.66% | 3.9% | 0.5092% | 0.4503% | 0.2917% | 3.753% | 6.499% | 2.473% | 9.624% | 3.464% | 5.631% | 7.676% | 4.317% | 7.358% | 0.7453% | 1.045% | [6 empty] | 0.6244% | 0.4041% | 3.085% | 1.23%
|  | Sample Concentration SD |  |  | 98.771 | 225.6238 | 0.26479 | 0.094321 | 119.9982 | 117.1749 | 73.5031 | 0.11469 | 0.068133 | 0.070779 | 0.073436 | 0.010993 | 0.015368 | 0.078871 | 0.18277 | 0.025085 | 0.022605 | 0.013403 | 0.15544 | 0.30814 | 0.11804 | 0.4463 | 0.16714 | 0.26907 | 0.36052 | 0.19221 | 0.34427 | 0.035444 | 0.051481 | [6 empty] | 0.029213 | 0.019009 | 0.15663 | 0.06796

## Calibration Curves 15
Calibration Table - Na KED 23/23
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 4.6667 | 49.4872 | 0 |  | 0 | 0
Cal. Std.4 | 0.058833 | 8.4997 | 500 | 509.979 | 9.979 | 1.9958
Cal. Std.5 | 0.11869 | 1.6125 | 1000 | 973.0996 | -26.9004 | -2.69
Cal. Std.6 | 0.61018 | 1.7842 | 5000 | 4775.8538 | -224.1462 | -4.4829
Cal. Std.7 | 1.2174 | 1.0552 | 10000 | 9474.3158 | -525.6842 | -5.2568
Cal. Std.8 | 3.2264 | 0.7985 | 25000 | 25017.6593 | 17.6593 | 0.070637
Cal. Std.9 | 6.5472 | 1.2803 | 50000 | 50711.6345 | 711.6345 | 1.4233
Cal. Std.10 | 14.073 | 1.7009 | 100000 | 108940.0898 | 8940.0898 | 8.9401
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99852 | 0.00012925 | -0.0070799 | 3.1474e-06 | 0.0037048 | 1.1052 | 1.6583
Calibration Table - Mg KED 24/24
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 0 |  | 0 |  | 0 | 0
Cal. Std.4 | 0.027711 | 7.7899 | 500 | 507.6499 | 7.6499 | 1.53
Cal. Std.5 | 0.05312 | 2.6138 | 1000 | 972.9214 | -27.0786 | -2.7079
Cal. Std.6 | 0.26913 | 1.7281 | 5000 | 4928.2533 | -71.7467 | -1.4349
Cal. Std.7 | 0.54058 | 2.135 | 10000 | 9898.8248 | -101.1752 | -1.0118
Cal. Std.8 | 1.3706 | 0.48247 | 25000 | 25097.446 | 97.446 | 0.38978
Cal. Std.9 | 2.6981 | 0.75145 | 50000 | 49405.7836 | -594.2164 | -1.1884
Cal. Std.10 | 5.7027 | 1.3065 | 100000 | 104423.2102 | 4423.2102 | 4.4232
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99964 | 5.4612e-05 | -1.2906e-05 | 6.5217e-07 | 0.00076767 | 0 | 0
Calibration Table - Al KED 27/27
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 50.6668 | 12.0594 | 0 |  | 0 | 0
Cal. Std.4 | 0.006558 | 7.6157 | 0.5 | 0.5591 | 0.059102 | 11.8204
Cal. Std.5 | 0.0075838 | 5.7761 | 1 | 0.76018 | -0.23982 | -23.9818
Cal. Std.6 | 0.029782 | 4.6267 | 5 | 5.1116 | 0.11155 | 2.231
Cal. Std.7 | 0.053162 | 3.9471 | 10 | 9.6944 | -0.30558 | -3.0558
Cal. Std.8 | 0.13324 | 1.9118 | 25 | 25.3921 | 0.39207 | 1.5683
Cal. Std.9 | 0.26479 | 1.2679 | 50 | 51.1779 | 1.1779 | 2.3559
Cal. Std.10 | 0.56008 | 1.3321 | 100 | 109.062 | 9.062 | 9.062
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99149 | 0.0051015 | 0.0037058 | 0.00029956 | 0.00035262 | 0.3034 | 0.1217
Calibration Table - P Oshift 31/47
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 950.0317 | 7.1115 | 0 |  | 0 | 0
Cal. Std.3 | 0.0066461 | 3.3592 | 0.1 | 0.069031 | -0.030969 | -30.9692
Cal. Std.4 | 0.022789 | 4.2342 | 0.5 | 1.3261 | 0.8261 | 165.2198
Cal. Std.5 | 0.016916 | 1.2687 | 1 | 0.86876 | -0.13124 | -13.1241
Cal. Std.6 | 0.060928 | 0.59524 | 5 | 4.296 | -0.70396 | -14.0791
Cal. Std.7 | 0.097289 | 1.1114 | 10 | 7.1276 | -2.8724 | -28.724
Cal. Std.8 | 0.23055 | 15.2884 | 25 | 17.5047 | -7.4953 | -29.981
Cal. Std.9 | 0.4855 | 0.20551 | 50 | 37.3586 | -12.6414 | -25.2829
Cal. Std.10 | 0.9938 | 0.92425 | 100 | 76.9405 | -23.0595 | -23.0595
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.81637 | 0.012842 | 0.0057597 | 0.0037088 | 0.0010235 | 0.2353 | 0.044491
Calibration Table - S (34) -1 34/34
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 2914.298 | 6.2565 | 0 |  | 0 | 0
Cal. Std.5 | 0.030514 | 8.8323 | 1000 | 1002.6864 | 2.6864 | 0.26864
Cal. Std.6 | 0.15812 | 3.6815 | 5000 | 4994.1795 | -5.8205 | -0.11641
Cal. Std.7 | 0.30962 | 2.9509 | 10000 | 9732.8947 | -267.1053 | -2.6711
Cal. Std.8 | 0.79785 | 1.8572 | 25000 | 25003.9579 | 3.9579 | 0.015832
Cal. Std.9 | 1.5907 | 1.9854 | 50000 | 49803.9939 | -196.0061 | -0.39201
Cal. Std.10 | 3.2881 | 0.4312 | 100000 | 102894.999 | 2894.999 | 2.895
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99981 | 3.1971e-05 | -0.0015423 | 3.0902e-07 | 0.00073795 | 2780.0714 | 404.6744
Calibration Table - K KED 39/39
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 44.6667 | 15.7249 | 0 |  | 0 | 0
Cal. Std.4 | 0.035091 | 7.6525 | 500 | 506.8855 | 6.8855 | 1.3771
Cal. Std.5 | 0.072723 | 2.6896 | 1000 | 989.698 | -10.302 | -1.0302
Cal. Std.6 | 0.36189 | 2.856 | 5000 | 4699.6773 | -300.3227 | -6.0065
Cal. Std.7 | 0.7272 | 1.9126 | 10000 | 9386.551 | -613.449 | -6.1345
Cal. Std.8 | 1.9125 | 1.405 | 25000 | 24593.4258 | -406.5742 | -1.6263
Cal. Std.9 | 3.9769 | 1.4247 | 50000 | 51079.6663 | 1079.6663 | 2.1593
Cal. Std.10 | 8.6676 | 2.4727 | 100000 | 111261.013 | 11261.013 | 11.261
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99779 | 7.7943e-05 | -0.0044169 | 2.3215e-06 | 0.0027327 | 17.5152 | 9.0734
Calibration Table - Ca KED 43/43
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 0.66667 | 173.2051 | 0 |  | 0 | 0
Cal. Std.4 | 0.018226 | 8.6076 | 500 | 518.7426 | 18.7426 | 3.7485
Cal. Std.5 | 0.033982 | 7.4649 | 1000 | 932.9317 | -67.0683 | -6.7068
Cal. Std.6 | 0.18298 | 3.0221 | 5000 | 4849.6011 | -150.3989 | -3.008
Cal. Std.7 | 0.3699 | 0.23057 | 10000 | 9763.2671 | -236.7329 | -2.3673
Cal. Std.8 | 0.93303 | 0.09784 | 25000 | 24566.666 | -433.334 | -1.7333
Cal. Std.9 | 1.9325 | 0.23035 | 50000 | 50840.0327 | 840.0327 | 1.6801
Cal. Std.10 | 4.1216 | 2.0123 | 100000 | 108386.8988 | 8386.8988 | 8.3869
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99843 | 3.8041e-05 | -0.0015079 | 9.5534e-07 | 0.0011245 | 0.52916 | 2.7496
Calibration Table - V KED 51/51
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 6 | 66.6667 | 0 |  | 0 | 0
Cal. Std.1 | 0.00067436 | 19.7822 | 0.01 | 0.0098554 | -0.00014457 | -1.4457
Cal. Std.2 | 0.0043245 | 10.9172 | 0.05 | 0.055701 | 0.005701 | 11.402
Cal. Std.3 | 0.0073088 | 7.1328 | 0.1 | 0.093184 | -0.0068157 | -6.8157
Cal. Std.4 | 0.037728 | 2.4188 | 0.5 | 0.47525 | -0.024747 | -4.9494
Cal. Std.5 | 0.075328 | 5.349 | 1 | 0.94751 | -0.052491 | -5.2491
Cal. Std.6 | 0.39091 | 2.2477 | 5 | 4.9112 | -0.088818 | -1.7764
Cal. Std.7 | 0.79198 | 0.55858 | 10 | 9.9486 | -0.051383 | -0.51383
Cal. Std.8 | 2.0488 | 0.27767 | 25 | 25.7346 | 0.73462 | 2.9385
Cal. Std.9 | 4.236 | 1.2584 | 50 | 53.2049 | 3.2049 | 6.4097
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99806 | 0.079618 | -0.00011031 | 0.0018752 | 5.4887e-05 | 0.0022867 | 0.0045358
Calibration Table - Cr (52) 52/52
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 27.3334 | 29.5716 | 0 |  | 0 | 0
Cal. Std.1 | 0.0022045 | 8.1378 | 0.01 | 0.010367 | 0.00036724 | 3.6724
Cal. Std.2 | 0.0054782 | 8.799 | 0.05 | 0.043646 | -0.0063544 | -12.7088
Cal. Std.3 | 0.0098994 | 14.7323 | 0.1 | 0.08859 | -0.01141 | -11.41
Cal. Std.4 | 0.050102 | 5.0268 | 0.5 | 0.49727 | -0.0027323 | -0.54646
Cal. Std.5 | 0.10031 | 1.4674 | 1 | 1.0077 | 0.0077013 | 0.77013
Cal. Std.6 | 0.50964 | 1.4971 | 5 | 5.1687 | 0.16871 | 3.3743
Cal. Std.7 | 1.0138 | 0.98635 | 10 | 10.2935 | 0.2935 | 2.935
Cal. Std.8 | 2.6012 | 1.4531 | 25 | 26.43 | 1.43 | 5.72
Cal. Std.9 | 5.3228 | 1.2244 | 50 | 54.0967 | 4.0967 | 8.1934
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.9971 | 0.098372 | 0.0011847 | 0.0028387 | 8.3089e-05 | 0.0084946 | 0.0077245
Calibration Table - Cr (53) NH3 53/53
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 20 | 10 | 0 |  | 0 | 0
Cal. Std.1 | 0.0003117 | 16.8212 | 0.01 | 0.01005 | 4.9731e-05 | 0.49731
Cal. Std.2 | 0.0011921 | 3.8129 | 0.05 | 0.049976 | -2.4159e-05 | -0.048318
Cal. Std.3 | 0.0021786 | 2.8911 | 0.1 | 0.094718 | -0.0052823 | -5.2823
Cal. Std.4 | 0.011417 | 3.8094 | 0.5 | 0.5137 | 0.013702 | 2.7404
Cal. Std.5 | 0.021955 | 1.7613 | 1 | 0.99163 | -0.0083717 | -0.83717
Cal. Std.6 | 0.1072 | 1.3384 | 5 | 4.8575 | -0.14245 | -2.849
Cal. Std.7 | 0.21491 | 0.2118 | 10 | 9.7424 | -0.25757 | -2.5757
Cal. Std.8 | 0.57185 | 0.94247 | 25 | 25.9305 | 0.93052 | 3.7221
Cal. Std.9 | 1.1536 | 0.29612 | 50 | 52.3163 | 2.3163 | 4.6327
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99941 | 0.02205 | 9.0105e-05 | 0.00028662 | 8.3895e-06 | 0.0022727 | 0.00068283
Calibration Table - Fe (54) NH3 54/54
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 663.2662 | 9.2906 | 0 |  | 0 | 0
Cal. Std.4 | 0.015853 | 2.2425 | 0.5 | 0.47673 | -0.023275 | -4.6549
Cal. Std.5 | 0.025356 | 1.1476 | 1 | 1.1122 | 0.11218 | 11.2184
Cal. Std.6 | 0.077711 | 0.86465 | 5 | 4.6132 | -0.38681 | -7.7362
Cal. Std.7 | 0.15077 | 0.734 | 10 | 9.4985 | -0.50154 | -5.0154
Cal. Std.8 | 0.39123 | 0.67778 | 25 | 25.5779 | 0.57789 | 2.3116
Cal. Std.9 | 0.76323 | 0.18915 | 50 | 50.4536 | 0.45365 | 0.9073
Cal. Std.10 | 1.5486 | 0.37728 | 100 | 102.9693 | 2.9693 | 2.9693
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99741 | 0.014954 | 0.0087239 | 0.00048265 | 0.00056813 | 0.11115 | 0.032223
Calibration Table - Mn Oshift 55/71
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 25.3334 | 18.2321 | 0 |  | 0 | 0
Cal. Std.2 | 0.0021696 | 6.0284 | 0.05 | 0.050249 | 0.00024892 | 0.49784
Cal. Std.3 | 0.004309 | 0.46309 | 0.1 | 0.099868 | -0.00013201 | -0.13201
Cal. Std.4 | 0.021026 | 2.1364 | 0.5 | 0.48759 | -0.012412 | -2.4823
Cal. Std.5 | 0.041587 | 1.2183 | 1 | 0.96445 | -0.035548 | -3.5548
Cal. Std.6 | 0.21156 | 1.117 | 5 | 4.9067 | -0.093283 | -1.8657
Cal. Std.7 | 0.43267 | 0.80165 | 10 | 10.0349 | 0.034896 | 0.34896
Cal. Std.8 | 1.1195 | 1.5838 | 25 | 25.9632 | 0.96322 | 3.8529
Cal. Std.9 | 2.2277 | 0.33099 | 50 | 51.6675 | 1.6675 | 3.335
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99958 | 0.043117 | 3.063e-06 | 0.00050986 | 6.417e-05 | 0.001659 | 0.00090747
Calibration Table - Fe (56) KED 56/56
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 9780.6812 | 1.93 | 0 |  | 0 | 0
Cal. Std.4 | 0.10603 | 2.9731 | 0.5 | 0.47623 | -0.023766 | -4.7532
Cal. Std.5 | 0.1717 | 1.0172 | 1 | 1.1114 | 0.11139 | 11.1389
Cal. Std.6 | 0.54588 | 0.99353 | 5 | 4.73 | -0.26997 | -5.3993
Cal. Std.7 | 1.0321 | 1.7463 | 10 | 9.4327 | -0.56726 | -5.6726
Cal. Std.8 | 2.5803 | 0.84158 | 25 | 24.4054 | -0.59461 | -2.3784
Cal. Std.9 | 5.4289 | 2.0481 | 50 | 51.9542 | 1.9542 | 3.9084
Cal. Std.10 | 10.7233 | 2.0797 | 100 | 103.1562 | 3.1562 | 3.1562
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.9975 | 0.1034 | 0.056783 | 0.0032787 | 0.0038594 | 2.8862 | 0.1433
Calibration Table - Co KED 59/59
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 5.3333 | 94.3729 | 0 |  | 0 | 0
Cal. Std.1 | 0.003177 | 10.5998 | 0.01 | 0.010136 | 0.0001356 | 1.356
Cal. Std.2 | 0.010207 | 4.8432 | 0.05 | 0.048077 | -0.0019229 | -3.8457
Cal. Std.3 | 0.018757 | 13.8458 | 0.1 | 0.094221 | -0.0057792 | -5.7792
Cal. Std.4 | 0.094145 | 1.1965 | 0.5 | 0.50108 | 0.0010828 | 0.21657
Cal. Std.5 | 0.18331 | 1.3495 | 1 | 0.98231 | -0.017689 | -1.7689
Cal. Std.6 | 0.93455 | 0.71648 | 5 | 5.0367 | 0.036671 | 0.73342
Cal. Std.7 | 1.8608 | 0.96876 | 10 | 10.0354 | 0.035356 | 0.35356
Cal. Std.8 | 4.8279 | 0.81211 | 25 | 26.049 | 1.049 | 4.196
Cal. Std.9 | 9.6863 | 1.7005 | 50 | 52.2692 | 2.2692 | 4.5383
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99937 | 0.18529 | 0.0012989 | 0.0024859 | 7.2764e-05 | 0.00087853 | 0.0024667
Calibration Table - Ni KED 60/60
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 26.6667 | 15.6125 | 0 |  | 0 | 0
Cal. Std.1 | 0.00050591 | 15.3403 | 0.01 | 0.0099249 | -7.5073e-05 | -0.75073
Cal. Std.2 | 0.0025096 | 18.8661 | 0.05 | 0.051122 | 0.0011218 | 2.2436
Cal. Std.3 | 0.0050156 | 12.4603 | 0.1 | 0.10265 | 0.0026478 | 2.6478
Cal. Std.4 | 0.024968 | 2.6315 | 0.5 | 0.51289 | 0.012894 | 2.5788
Cal. Std.5 | 0.048217 | 3.9349 | 1 | 0.99091 | -0.0090861 | -0.90861
Cal. Std.6 | 0.24072 | 2.2539 | 5 | 4.9488 | -0.051164 | -1.0233
Cal. Std.7 | 0.47291 | 2.385 | 10 | 9.7228 | -0.27715 | -2.7715
Cal. Std.8 | 1.2138 | 0.48586 | 25 | 24.9564 | -0.043552 | -0.17421
Cal. Std.9 | 2.387 | 1.8527 | 50 | 49.0791 | -0.92089 | -1.8418
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99978 | 0.048636 | 2.3204e-05 | 0.0003861 | 1.1301e-05 | 0.016753 | 0.0084465
Calibration Table - Cu (63) KED 63/63
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 16 | 25 | 0 |  | 0 | 0
Cal. Std.1 | 0.0033893 | 10.0579 | 0.01 | 0.0103 | 0.00030008 | 3.0008
Cal. Std.2 | 0.0070187 | 12.7515 | 0.05 | 0.04384 | -0.0061603 | -12.3205
Cal. Std.3 | 0.012306 | 10.3097 | 0.1 | 0.092701 | -0.0072993 | -7.2993
Cal. Std.4 | 0.060912 | 3.0481 | 0.5 | 0.54186 | 0.041858 | 8.3716
Cal. Std.5 | 0.11249 | 3.1437 | 1 | 1.0185 | 0.018517 | 1.8517
Cal. Std.6 | 0.55809 | 1.2517 | 5 | 5.1363 | 0.13626 | 2.7252
Cal. Std.7 | 1.0912 | 0.77011 | 10 | 10.0629 | 0.062913 | 0.62913
Cal. Std.8 | 2.7829 | 0.42947 | 25 | 25.6951 | 0.69515 | 2.7806
Cal. Std.9 | 5.4271 | 2.1107 | 50 | 50.1304 | 0.13042 | 0.26084
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99793 | 0.10821 | 0.0022746 | 0.0026348 | 7.712e-05 | 0.0045228 | 0.0035921
Calibration Table - Cu (65) NH3 65/65
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 127.3339 | 16.3482 | 0 |  | 0 | 0
Cal. Std.1 | 0.0011897 | 10.4783 | 0.01 | 0.010282 | 0.00028195 | 2.8195
Cal. Std.2 | 0.0026637 | 4.6917 | 0.05 | 0.044167 | -0.0058328 | -11.6657
Cal. Std.3 | 0.0047635 | 4.0911 | 0.1 | 0.092436 | -0.0075636 | -7.5636
Cal. Std.4 | 0.025039 | 2.6738 | 0.5 | 0.55853 | 0.058528 | 11.7056
Cal. Std.5 | 0.045779 | 3.4769 | 1 | 1.0353 | 0.03529 | 3.529
Cal. Std.6 | 0.21919 | 2.0207 | 5 | 5.0217 | 0.021668 | 0.43336
Cal. Std.7 | 0.43077 | 0.41665 | 10 | 9.8854 | -0.11463 | -1.1463
Cal. Std.8 | 1.116 | 1.4833 | 25 | 25.6387 | 0.63867 | 2.5547
Cal. Std.9 | 2.1613 | 0.8869 | 50 | 49.6667 | -0.33326 | -0.66652
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99755 | 0.043501 | 0.00074242 | 0.0011528 | 3.3742e-05 | 0.0073283 | 0.0034525
Calibration Table - Zn NH3 66/66
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 1445.4065 | 4.0919 | 0 |  | 0 | 0
Cal. Std.3 | 0.0022327 | 3.1823 | 0.1 | 0.09614 | -0.0038603 | -3.8603
Cal. Std.4 | 0.013185 | 4.453 | 0.5 | 0.60645 | 0.10645 | 21.2891
Cal. Std.5 | 0.021057 | 3.6694 | 1 | 0.97323 | -0.026765 | -2.6765
Cal. Std.6 | 0.10317 | 1.1063 | 5 | 4.7993 | -0.20073 | -4.0146
Cal. Std.7 | 0.20678 | 0.31086 | 10 | 9.6268 | -0.37325 | -3.7325
Cal. Std.8 | 0.53291 | 0.92719 | 25 | 24.8226 | -0.17735 | -0.70941
Cal. Std.9 | 1.0394 | 0.85229 | 50 | 48.4204 | -1.5796 | -3.1593
Cal. Std.10 | 2.0791 | 0.79309 | 100 | 96.8635 | -3.1365 | -3.1365
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99585 | 0.021462 | 0.00016935 | 0.00080045 | 0.00022089 | 0.16877 | 0.023933
Calibration Table - As 75/75
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | -3944.1929 | 4.0034 | 0 |  | 0 | 0
Cal. Std.2 | 0.0011009 | 59.0919 | 0.05 | 0.060638 | 0.010638 | 21.2768
Cal. Std.3 | 0.0010272 | 46.9426 | 0.1 | 0.057167 | -0.042833 | -42.8334
Cal. Std.4 | 0.010875 | 4.4351 | 0.5 | 0.52099 | 0.020989 | 4.1979
Cal. Std.5 | 0.019565 | 7.7071 | 1 | 0.93023 | -0.069768 | -6.9768
Cal. Std.6 | 0.1086 | 0.98849 | 5 | 5.1236 | 0.12355 | 2.4711
Cal. Std.7 | 0.21938 | 0.78908 | 10 | 10.3408 | 0.34081 | 3.4081
Cal. Std.8 | 0.57722 | 1.2169 | 25 | 27.1945 | 2.1945 | 8.7781
Cal. Std.9 | 1.1642 | 0.36799 | 50 | 54.8391 | 4.8391 | 9.6783
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.97877 | 0.021233 | -0.00018656 | 0.0018153 | 0.00022847 | -0.68827 | 0.084785
Calibration Table - As KED 75/75
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 0.66667 | 173.2051 | 0 |  | 0 | 0
Cal. Std.1 | 0.00013903 | 64.5605 | 0.01 | 0.010398 | 0.00039759 | 3.9759
Cal. Std.2 | 0.000349 | 18.6678 | 0.05 | 0.039227 | -0.010773 | -21.5457
Cal. Std.3 | 0.00080189 | 24.3816 | 0.1 | 0.10141 | 0.0014089 | 1.4089
Cal. Std.4 | 0.0040861 | 12.0898 | 0.5 | 0.55233 | 0.052331 | 10.4662
Cal. Std.5 | 0.0071876 | 16.9527 | 1 | 0.97818 | -0.021823 | -2.1823
Cal. Std.6 | 0.037138 | 2.3571 | 5 | 5.0904 | 0.090425 | 1.8085
Cal. Std.7 | 0.071948 | 1.3439 | 10 | 9.8699 | -0.13006 | -1.3006
Cal. Std.8 | 0.19151 | 1.8695 | 25 | 26.2858 | 1.2858 | 5.143
Cal. Std.9 | 0.37233 | 3.074 | 50 | 51.1131 | 1.1131 | 2.2262
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.9957 | 0.0072832 | 6.3301e-05 | 0.00025615 | 7.4976e-06 | 0.0028437 | 0.014776
Calibration Table - As Oshift 75/91
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 14 | 42.8572 | 0 |  | 0 | 0
Cal. Std.1 | 0.00027799 | 12.3228 | 0.01 | 0.0099849 | -1.507e-05 | -0.1507
Cal. Std.2 | 0.001147 | 7.755 | 0.05 | 0.051436 | 0.001436 | 2.8719
Cal. Std.3 | 0.0020896 | 5.0454 | 0.1 | 0.096394 | -0.0036056 | -3.6056
Cal. Std.4 | 0.010411 | 2.0694 | 0.5 | 0.49332 | -0.0066806 | -1.3361
Cal. Std.5 | 0.02034 | 2.5336 | 1 | 0.9669 | -0.033103 | -3.3103
Cal. Std.6 | 0.10279 | 0.73162 | 5 | 4.8995 | -0.10046 | -2.0092
Cal. Std.7 | 0.20515 | 0.71742 | 10 | 9.782 | -0.21801 | -2.1801
Cal. Std.8 | 0.54918 | 1.0364 | 25 | 26.191 | 1.191 | 4.7642
Cal. Std.9 | 1.1003 | 0.9183 | 50 | 52.4779 | 2.4779 | 4.9559
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99938 | 0.020966 | 6.8652e-05 | 0.0002782 | 8.1429e-06 | 0.0018859 | 0.0024302
Calibration Table - Se (77) 77/77
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 1375.3996 | 4.4046 | 0 |  | 0 | 0
Cal. Std.3 | 0.0005956 | 16.4493 | 0.1 | 0.10184 | 0.0018422 | 1.8422
Cal. Std.4 | 0.0014055 | 12.2511 | 0.5 | 0.4492 | -0.050799 | -10.1597
Cal. Std.5 | 0.0027363 | 12.7746 | 1 | 1.02 | 0.019998 | 1.9998
Cal. Std.6 | 0.011588 | 3.7249 | 5 | 4.8164 | -0.18362 | -3.6724
Cal. Std.7 | 0.024771 | 2.7486 | 10 | 10.4706 | 0.47056 | 4.7056
Cal. Std.8 | 0.060263 | 0.73241 | 25 | 25.693 | 0.69295 | 2.7718
Cal. Std.9 | 0.11987 | 0.79087 | 50 | 51.2563 | 1.2563 | 2.5126
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99849 | 0.0023316 | 0.00035815 | 5.7328e-05 | 1.4798e-05 | 2.1855 | 0.27485
Calibration Table - Se (77) Oshift 77/93
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 0 |  | 0 |  | 0 | 0
Cal. Std.2 | 2.2978e-05 | 43.4143 | 0.05 | 0.054886 | 0.0048859 | 9.7717
Cal. Std.3 | 3.4434e-05 | 65.9207 | 0.1 | 0.079491 | -0.020509 | -20.5094
Cal. Std.4 | 0.00024747 | 9.6411 | 0.5 | 0.53703 | 0.037025 | 7.405
Cal. Std.5 | 0.00043924 | 17.2411 | 1 | 0.94888 | -0.051118 | -5.1118
Cal. Std.6 | 0.002322 | 7.2604 | 5 | 4.9924 | -0.0076135 | -0.15227
Cal. Std.7 | 0.0044771 | 3.7226 | 10 | 9.6209 | -0.37909 | -3.7909
Cal. Std.8 | 0.012355 | 2.7854 | 25 | 26.5403 | 1.5403 | 6.1613
Cal. Std.9 | 0.024728 | 3.2897 | 50 | 53.1132 | 3.1132 | 6.2264
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99412 | 0.00046561 | -2.5779e-06 | 2.0712e-05 | 2.6068e-06 | 0 | 0
Calibration Table - Se (78) Oshift 78/94
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 0.66667 | 173.2051 | 0 |  | 0 | 0
Cal. Std.2 | 7.6512e-05 | 29.6898 | 0.05 | 0.053052 | 0.003052 | 6.104
Cal. Std.3 | 0.00012872 | 9.5199 | 0.1 | 0.088721 | -0.011279 | -11.2787
Cal. Std.4 | 0.0007016 | 7.1807 | 0.5 | 0.48013 | -0.019873 | -3.9746
Cal. Std.5 | 0.0014385 | 6.0873 | 1 | 0.98362 | -0.016381 | -1.6381
Cal. Std.6 | 0.0073472 | 3.5353 | 5 | 5.0205 | 0.020538 | 0.41075
Cal. Std.7 | 0.014455 | 2.3095 | 10 | 9.8766 | -0.12344 | -1.2344
Cal. Std.8 | 0.038482 | 0.7979 | 25 | 26.2925 | 1.2925 | 5.1701
Cal. Std.9 | 0.077895 | 2.0657 | 50 | 53.2205 | 3.2205 | 6.4409
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99783 | 0.0014636 | -1.1372e-06 | 3.9465e-05 | 4.967e-06 | 0.001288 | 0.0066927
Calibration Table - Se (80) Oshift 80/96
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 6 | 66.6667 | 0 |  | 0 | 0
Cal. Std.2 | 0.00018217 | 12.1346 | 0.05 | 0.051786 | 0.0017857 | 3.5713
Cal. Std.3 | 0.00031167 | 12.1472 | 0.1 | 0.093313 | -0.0066871 | -6.6871
Cal. Std.4 | 0.0015607 | 6.7092 | 0.5 | 0.49381 | -0.0061854 | -1.2371
Cal. Std.5 | 0.0030934 | 2.336 | 1 | 0.98529 | -0.014706 | -1.4706
Cal. Std.6 | 0.01503 | 0.12481 | 5 | 4.8128 | -0.18717 | -3.7435
Cal. Std.7 | 0.030518 | 1.8716 | 10 | 9.7794 | -0.22064 | -2.2064
Cal. Std.8 | 0.082784 | 0.20293 | 25 | 26.5387 | 1.5387 | 6.1549
Cal. Std.9 | 0.16471 | 1.4744 | 50 | 52.8092 | 2.8092 | 5.6185
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99871 | 0.0031186 | 2.067e-05 | 6.4657e-05 | 8.1375e-06 | 0.0054288 | 0.010841
Calibration Table - Se (82) 82/82
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 63.8984 | 14.1686 | 0 |  | 0 | 0
Cal. Std.2 | 0.00011641 | 46.6271 | 0.05 | 0.05493 | 0.00493 | 9.86
Cal. Std.3 | 0.00016698 | 9.3694 | 0.1 | 0.080448 | -0.019552 | -19.5521
Cal. Std.4 | 0.0010219 | 36.8085 | 0.5 | 0.51184 | 0.011838 | 2.3675
Cal. Std.5 | 0.0018416 | 9.2051 | 1 | 0.92543 | -0.074566 | -7.4566
Cal. Std.6 | 0.010239 | 4.4059 | 5 | 5.1625 | 0.16251 | 3.2502
Cal. Std.7 | 0.019952 | 1.6389 | 10 | 10.0636 | 0.063598 | 0.63598
Cal. Std.8 | 0.052295 | 0.60706 | 25 | 26.3828 | 1.3828 | 5.5312
Cal. Std.9 | 0.10442 | 0.29834 | 50 | 52.6819 | 2.6819 | 5.3638
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99477 | 0.0019819 | 7.5479e-06 | 8.3072e-05 | 1.0455e-05 | 0.11939 | 0.048642
Calibration Table - Se (82) Oshift 82/98
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 1.3333 | 86.6025 | 0 |  | 0 | 0
Cal. Std.2 | 3.8425e-05 | 53.1062 | 0.05 | 0.05512 | 0.0051195 | 10.239
Cal. Std.3 | 5.2056e-05 | 56.2821 | 0.1 | 0.079866 | -0.020134 | -20.1339
Cal. Std.4 | 0.00028897 | 10.9624 | 0.5 | 0.50996 | 0.0099647 | 1.9929
Cal. Std.5 | 0.00051476 | 13.2821 | 1 | 0.91987 | -0.080127 | -8.0127
Cal. Std.6 | 0.0027905 | 0.82573 | 5 | 5.0514 | 0.051426 | 1.0285
Cal. Std.7 | 0.0054652 | 4.1939 | 10 | 9.9071 | -0.092908 | -0.92908
Cal. Std.8 | 0.014861 | 1.1027 | 25 | 26.9641 | 1.9641 | 7.8562
Cal. Std.9 | 0.029742 | 2.8653 | 50 | 53.9795 | 3.9795 | 7.959
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99397 | 0.00055083 | 8.0637e-06 | 2.4807e-05 | 3.1221e-06 | 0.0068296 | 0.017744
Calibration Table - Mo (95) NH3 95/95
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 231.3352 | 14.0563 | 0 |  | 0 | 0
Cal. Std.2 | 0.0020797 | 1.5145 | 0.05 | 0.052029 | 0.0020287 | 4.0575
Cal. Std.3 | 0.0037627 | 1.3747 | 0.1 | 0.092034 | -0.0079664 | -7.9664
Cal. Std.4 | 0.02106 | 1.1398 | 0.5 | 0.5032 | 0.0032049 | 0.64099
Cal. Std.5 | 0.040926 | 2.7986 | 1 | 0.97544 | -0.02456 | -2.456
Cal. Std.6 | 0.20564 | 1.31 | 5 | 4.8908 | -0.1092 | -2.1839
Cal. Std.7 | 0.41416 | 0.54985 | 10 | 9.8473 | -0.15266 | -1.5266
Cal. Std.8 | 1.0987 | 0.73844 | 25 | 26.1195 | 1.1195 | 4.4781
Cal. Std.9 | 2.2076 | 0.58894 | 50 | 52.4782 | 2.4782 | 4.9564
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99881 | 0.042069 | -0.00010904 | 0.00083827 | 0.0001055 | 0.013787 | 0.0061386
Calibration Table - Mo (95) OShift 95/127
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 4 | 100 | 0 |  | 0 | 0
Cal. Std.2 | 0.0021002 | 3.9338 | 0.05 | 0.05088 | 0.00088019 | 1.7604
Cal. Std.3 | 0.0040577 | 3.0921 | 0.1 | 0.097027 | -0.0029731 | -2.9731
Cal. Std.4 | 0.020825 | 2.256 | 0.5 | 0.4923 | -0.0077028 | -1.5406
Cal. Std.5 | 0.041541 | 2.5315 | 1 | 0.98067 | -0.019333 | -1.9333
Cal. Std.6 | 0.20584 | 1.5474 | 5 | 4.8539 | -0.14612 | -2.9225
Cal. Std.7 | 0.41835 | 0.90267 | 10 | 9.8637 | -0.13631 | -1.3631
Cal. Std.8 | 1.1024 | 0.61802 | 25 | 25.9894 | 0.98943 | 3.9577
Cal. Std.9 | 2.2272 | 0.51534 | 50 | 52.5072 | 2.5072 | 5.0145
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.9994 | 0.042419 | -5.8059e-05 | 0.00059859 | 7.5337e-05 | 0.00026613 | 0.00079759
Calibration Table - Cd (111) NH3 111/111
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 2.6667 | 114.5644 | 0 |  | 0 | 0
Cal. Std.1 | 0.00018655 | 19.4321 | 0.01 | 0.0099596 | -4.0438e-05 | -0.40438
Cal. Std.2 | 0.0010253 | 12.522 | 0.05 | 0.051001 | 0.0010009 | 2.0017
Cal. Std.3 | 0.002021 | 1.8139 | 0.1 | 0.099725 | -0.0002755 | -0.2755
Cal. Std.4 | 0.010458 | 2.4514 | 0.5 | 0.51257 | 0.012573 | 2.5147
Cal. Std.5 | 0.020093 | 1.1194 | 1 | 0.98401 | -0.015993 | -1.5993
Cal. Std.6 | 0.10162 | 2.8102 | 5 | 4.9732 | -0.026823 | -0.53645
Cal. Std.7 | 0.20014 | 1.0285 | 10 | 9.7942 | -0.20576 | -2.0576
Cal. Std.8 | 0.51969 | 1.0194 | 25 | 25.4304 | 0.43041 | 1.7216
Cal. Std.9 | 1.0079 | 0.76133 | 50 | 49.3176 | -0.68244 | -1.3649
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99985 | 0.020436 | -1.6987e-05 | 0.00013547 | 3.9652e-06 | 0.00032868 | 0.0011345
Calibration Table - Cd (114) NH3 114/114
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 29.3334 | 34.9882 | 0 |  | 0 | 0
Cal. Std.1 | 0.0005551 | 14.023 | 0.01 | 0.0099045 | -9.5545e-05 | -0.95545
Cal. Std.2 | 0.0026461 | 4.9365 | 0.05 | 0.052899 | 0.0028992 | 5.7984
Cal. Std.3 | 0.0048288 | 4.7909 | 0.1 | 0.097781 | -0.0022187 | -2.2187
Cal. Std.4 | 0.024781 | 1.0759 | 0.5 | 0.50805 | 0.0080477 | 1.6095
Cal. Std.5 | 0.048151 | 2.5318 | 1 | 0.98856 | -0.011436 | -1.1436
Cal. Std.6 | 0.24148 | 1.4449 | 5 | 4.9639 | -0.036119 | -0.72238
Cal. Std.7 | 0.47714 | 0.16556 | 10 | 9.8096 | -0.19045 | -1.9045
Cal. Std.8 | 1.2361 | 1.001 | 25 | 25.4157 | 0.41573 | 1.6629
Cal. Std.9 | 2.38 | 1.3833 | 50 | 48.9369 | -1.0631 | -2.1263
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99963 | 0.048633 | 7.3419e-05 | 0.00050335 | 1.4733e-05 | 0.0015087 | 0.0015588
Calibration Table - Pb 208/208
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 51.3334 | 14.7504 | 0 |  | 0 | 0
Cal. Std.1 | 0.001943 | 1.533 | 0.01 | 0.0099377 | -6.2297e-05 | -0.62297
Cal. Std.2 | 0.0056527 | 4.1108 | 0.05 | 0.049167 | -0.00083301 | -1.666
Cal. Std.3 | 0.010951 | 5.0145 | 0.1 | 0.10519 | 0.0051924 | 5.1924
Cal. Std.4 | 0.055695 | 0.78446 | 0.5 | 0.57835 | 0.078349 | 15.6698
Cal. Std.5 | 0.10655 | 0.78591 | 1 | 1.1162 | 0.11616 | 11.6162
Cal. Std.6 | 0.51213 | 1.6547 | 5 | 5.4051 | 0.40508 | 8.1015
Cal. Std.7 | 0.95655 | 0.8505 | 10 | 10.1047 | 0.10472 | 1.0472
Cal. Std.8 | 2.1293 | 1.9815 | 25 | 22.5062 | -2.4938 | -9.9753
Cal. Std.9 | 3.3409 | 2.6397 | 50 | 35.3187 | -14.6813 | -29.3627
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.9903 | 0.094565 | 0.0010032 | 0.0050158 | 0.00014681 | 0.002012 | 0.00091365
Calibration Table - Pb NH3 208/208
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 95.3335 | 20.376 | 0 |  | 0 | 0
Cal. Std.1 | 0.0032856 | 4.0933 | 0.01 | 0.010133 | 0.00013326 | 1.3326
Cal. Std.2 | 0.01028 | 0.57481 | 0.05 | 0.04701 | -0.0029903 | -5.9805
Cal. Std.3 | 0.019766 | 0.93948 | 0.1 | 0.097027 | -0.0029726 | -2.9726
Cal. Std.4 | 0.10113 | 0.81346 | 0.5 | 0.52599 | 0.025988 | 5.1975
Cal. Std.5 | 0.20064 | 1.8814 | 1 | 1.0507 | 0.050666 | 5.0666
Cal. Std.6 | 0.98303 | 0.31243 | 5 | 5.1758 | 0.17577 | 3.5154
Cal. Std.7 | 1.9055 | 1.6567 | 10 | 10.0395 | 0.03951 | 0.3951
Cal. Std.8 | 4.7414 | 0.74425 | 25 | 24.9913 | -0.0086563 | -0.034625
Cal. Std.9 | 8.8664 | 1.8078 | 50 | 46.7403 | -3.2597 | -6.5195
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99895 | 0.18967 | 0.0013636 | 0.0032919 | 9.6354e-05 | 0.0015621 | 0.00095171