## Raw Intensities
| Sample Id | R | Acquisition Time | QC Status | Dataset File | Method File | Na KED 23/23 Helium KED_Low (cps) | Mg KED 24/24 Helium KED_Low (cps) | Al KED 27/27 Helium KED_Low (cps) | P Oshift 31/47 Oxygen DRC (cps) | S (34) -1 34/34 Helium KED_Low (cps) | K KED 39/39 Helium KED_Low (cps) | Ca KED 43/43 Helium KED_Low (cps) | V KED 51/51 Helium KED_Low (cps) | Cr (52) 52/52 Helium KED_Low (cps) | Cr (53) NH3 53/53 Ammonia DRC (cps) | Fe (54) NH3 54/54 Ammonia DRC (cps) | Mn Oshift 55/71 Oxygen DRC (cps) | Fe (56) KED 56/56 Helium KED_Low (cps) | Co KED 59/59 Helium KED_Low (cps) | Ni KED 60/60 Helium KED_Low (cps) | Cu (63) KED 63/63 Helium KED_Low (cps) | Cu (65) NH3 65/65 Ammonia DRC (cps) | Zn NH3 66/66 Ammonia DRC (cps) | As 75/75 (cps) | As KED 75/75 Helium KED_Low (cps) | As RH Oshift 75/91 Oxygen DRC (cps) | As Oshift 75/91 Oxygen DRC (cps) | Se (77) 77/77 (cps) | Se (77) Oshift 77/93 Oxygen DRC (cps) | Se (78) Oshift 78/94 Oxygen DRC (cps) | Se (80) Oshift 80/96 Oxygen DRC (cps) | Se (82) 82/82 (cps) | Se (82) Oshift 82/98 Oxygen DRC (cps) | Mo (95) NH3 95/95 Ammonia DRC (cps) | Mo (95) OShift 95/127 Oxygen DRC (cps) | Rh (IS1) 103/103 (IS) (cps) | Rh (IS2) 103/103 (IS) Oxygen DRC (cps) | Rh (IS3) 103/103 (IS) Helium KED_Low (cps) | Rh (IS4) 103/103 (IS) Ammonia DRC (cps) | Rh (IS5) 103/103 (IS) Ammonia DRC (cps) | Rh (IS6) 103/103 (IS) Oxygen DRC (cps) | Rh OShift (IS1) 103/119 (IS) Oxygen DRC (cps) | Cd (111) NH3 111/111 Ammonia DRC (cps) | Cd (114) NH3 114/114 Ammonia DRC (cps) | Pb 208/208 (cps) | Pb NH3 208/208 Ammonia DRC (cps)
1 | CAL BLANK | R | 06/02/2023 15:17:36 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CAL BLANK.004 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.33333 | 0 | 4 | 2280.1821 | 372.3382 | 9.3333 | 0.33333 | 0.66667 | 4.3333 | 18.6667 | 718.1719 | 8.6667 | 952.3651 | 1.6667 | 5.6667 | 7.3333 | 48.0001 | 676.683 | -2539.7451 | 1 | 24.6667 | 24.6667 | 980.0336 | 0 | 2 | 3.3333 | 29.0906 | 0 | 9.3333 | 4.6667 | 136153.5367 | 200107.0943 | 10262.6851 | 220973.2575 | 166481.7999 | 182757.8968 | 3107.6714 | 0 | 5.3333 | 26.6667 | 57.3334
2 | LOW 0.01 | R | 06/02/2023 15:23:43 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LOW 0.01.005 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 12.6667 | 5 | 10.6667 | 3141.6789 | 363.338 | 18 | 1.6667 | 7.6667 | 13.3333 | 94.0003 | 1146.2491 | 146.0008 | 1011.7025 | 21.3334 | 13.3333 | 35.3334 | 290.0029 | 1414.0701 | -2503.587 | 1 | 94.0003 | 94.0003 | 926.0301 | 1.3333 | 8.6667 | 9.3333 | 35.1394 | 4.6667 | 334.6706 | 292.6697 | 135983.4907 | 196439.5642 | 10202.9768 | 216608.7837 | 165740.6331 | 181647.6968 | 3052.3263 | 86.6669 | 195.3347 | 350.6683 | 811.3424
3 | LOW 0.05 | R | 06/02/2023 15:29:46 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LOW 0.05.006 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 50.6668 | 24.3334 | 17.6667 | 4329.3228 | 357.3378 | 42.0001 | 9.6667 | 33.3334 | 63.6668 | 441.3402 | 1427.8301 | 636.6809 | 1075.0405 | 108.0004 | 41.0001 | 100.0004 | 937.3642 | 1284.0578 | -2295.7206 | 4.6667 | 387.3386 | 387.3386 | 932.0304 | 4.6667 | 26.6667 | 58.6668 | 59.7955 | 11.3333 | 775.3544 | 690.0167 | 136570.7617 | 191527.9639 | 9835.0601 | 215509.8351 | 158594.455 | 176415.5452 | 2976.9775 | 391.3387 | 945.3648 | 1615.3692 | 3540.8423
4 | LOW 0.1 | R | 06/02/2023 15:35:48 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LOW 0.1.007 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 120.6672 | 48.0001 | 12.3333 | 3905.2004 | 359.6712 | 63.6668 | 14.6667 | 75.3335 | 109.0004 | 760.687 | 2300.3469 | 1242.0541 | 1273.3901 | 217.6683 | 60.6668 | 169.3343 | 1726.7715 | 2087.4859 | -2052.5688 | 8 | 665.3489 | 665.3489 | 1003.3687 | 12 | 61.3335 | 106.0004 | 91.086 | 18 | 1478.0765 | 1337.3959 | 136777.7305 | 197752.0331 | 10333.4036 | 216000.8763 | 171533.0741 | 178731.7931 | 2868.2881 | 800.6891 | 1925.4632 | 3074.1272 | 6849.3089
5 | LOW 0.5 | R | 06/02/2023 15:41:51 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LOW 0.5.008 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 574.0115 | 223.6684 | 35 | 5767.1643 | 434.6733 | 317.0035 | 91.3336 | 386.0053 | 541.677 | 3724.4855 | 5923.1913 | 6125.981 | 1899.1263 | 1038.0377 | 290.3363 | 874.6934 | 7995.5705 | 5274.3068 | -195.4186 | 47.6667 | 3410.4073 | 3410.4073 | 1244.0543 | 70.6668 | 222.0017 | 497.342 | 265.2106 | 78.6669 | 7376.5716 | 6519.488 | 143511.248 | 210180.4162 | 10266.3611 | 220777.046 | 164559.0387 | 161616.1654 | 3165.6858 | 4009.8963 | 9633.2474 | 15565.9831 | 35089.5009
6 | STD 1 | R | 06/02/2023 15:47:53 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\STD 1.009 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1183.049 | 460.0074 | 48.3334 | 7037.7336 | 549.6772 | 640.6811 | 193.0013 | 790.3552 | 1110.7099 | 7892.1797 | 11571.4983 | 12244.5789 | 2948.6376 | 2123.4912 | 607.0129 | 1690.7668 | 16371.3799 | 10054.871 | 2600.0499 | 83.0002 | 6809.623 | 6809.623 | 1557.4184 | 155.3342 | 460.0074 | 975.3666 | 567.1896 | 186.6679 | 15445.012 | 13595.1327 | 149820.2053 | 204953.0789 | 10386.7758 | 225869.0355 | 141872.1783 | 188348.2289 | 3079.6654 | 8062.2746 | 19110.1088 | 32326.1942 | 75708.9452
7 | STD 5 | R | 06/02/2023 15:53:56 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\STD 5.010 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 5538.7403 | 2246.8435 | 141.0007 | 17491.3937 | 1125.3777 | 2909.963 | 952.3652 | 3946.212 | 5447.7055 | 39216.4428 | 27830.7134 | 59150.877 | 5891.8816 | 10592.5924 | 2904.2952 | 7962.2191 | 76769.102 | 37940.9838 | 24415.6376 | 427.0064 | 32450.816 | 32450.816 | 4018.5654 | 666.6823 | 2244.8431 | 4809.4767 | 2643.3792 | 844.025 | 77589.5016 | 65187.0598 | 155400.1918 | 199674.5214 | 10442.1492 | 229347.7605 | 170792.2626 | 189936.6217 | 2968.3085 | 40445.8505 | 95665.9248 | 156336.7853 | 348479.493
8 | STD 10 | R | 06/02/2023 15:59:58 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\STD 10.011 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 10750.3768 | 4352.6631 | 250.6689 | 26836.5176 | 1912.4615 | 5562.416 | 1895.7925 | 7737.4283 | 10891.1501 | 73603.8138 | 50584.0568 | 113226.9546 | 10299.7118 | 20895.6042 | 5837.5258 | 15355.2485 | 142226.5764 | 69706.9995 | 51124.3607 | 860.6926 | 61168.6876 | 61168.6876 | 6614.866 | 1304.0595 | 4123.9288 | 8675.9673 | 4971.456 | 1566.7527 | 149546.8255 | 124277.5882 | 151074.6348 | 188390.8451 | 10337.4056 | 218989.7884 | 165985.472 | 178791.366 | 2764.9344 | 75712.1123 | 183202.6135 | 290402.6045 | 681302.8474
9 | STD 25 | R | 06/02/2023 16:06:00 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\STD 25.012 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 22699.6957 | 9025.8507 | 567.678 | 52537.1006 | 3666.8039 | 11852.9182 | 4122.5952 | 17829.7895 | 24538.7271 | 159345.8113 | 108010.2446 | 236057.9096 | 21440.0848 | 47273.7953 | 12960.8795 | 34000.4168 | 304439.4699 | 150604.3734 | 122135.4045 | 1957.8008 | 133790.9262 | 133790.9262 | 13779.3154 | 2860.2863 | 9186.9537 | 19471.2627 | 11436.5723 | 3467.0875 | 341755.7156 | 272064.2395 | 139005.7051 | 159585.1454 | 8542.2228 | 191644.2411 | 162940.7854 | 153484.1099 | 2392.8672 | 168839.8912 | 404454.2592 | 601793.4439 | 1503585.1238
10 | STD 50 | R | 06/02/2023 16:12:02 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\STD 50.013 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 38463.045 | 15190.4054 | 898.0283 | 86440.0839 | 6182.6712 | 19912.8691 | 7355.5599 | 31655.0356 | 43099.9216 | 290668.9263 | 195114.2269 | 421123.9033 | 36496.5617 | 82101.2591 | 22204.9113 | 58832.2321 | 546158.4482 | 267722.1408 | 238845.3239 | 3320.7193 | 236021.0894 | 236021.0894 | 25938.8687 | 5225.6223 | 16352.6887 | 34831.0804 | 21725.3667 | 6268.7085 | 631889.3705 | 496538.3563 | 144558.1694 | 141863.5909 | 8202.3541 | 177018.0631 | 136971.5515 | 135937.0912 | 2258.8453 | 308124.952 | 735831.4152 | 886168.7547 | 2795008.3253
11 | STD 100 | R | 06/02/2023 16:18:04 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\STD 100.014 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 77435.9724 | 29557.882 | 1844.119 | 168035.1612 | 12293.2871 | 38706.0314 | 15560.8041 | 66831.63 | 89568.5809 | 592299.3353 | 393314.8201 | 847573.4955 | 75398.4508 | 174944.6985 | 45795.6196 | 117553.0255 | 1060631.718 | 514573.0891 | 485772.1132 | 6724.9163 | 459457.7445 | 459457.7445 | 52848.2485 | 10565.239 | 32603.8341 | 69610.5197 | 45351.0403 | 12495.4625 | 1313992.7064 | 1017143.7007 | 126049.1057 | 139734.0772 | 8070.2796 | 176024.5478 | 137735.472 | 134260.625 | 2268.18 | 604903.6077 | 1447740.6678 | 1279504.4866 | 5620093.6516
12 | Calibration Curves | R | 06/02/2023 16:18:04
13 | LBLANK | R | 06/02/2023 16:24:07 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LBLANK.015 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 3.6667 | 0.33333 | 2.6667 | 1532.749 | 451.0071 | 8 | 1 | 3 | 9 | 34.6667 | 934.0912 | 24.6667 | 1041.7046 | 7 | 7.3333 | 14 | 104.0004 | 557.3442 | -2856.6296 | 0.66667 | 42.6667 | 42.6667 | 1103.376 | 0 | 4 | 8 | 61.0152 | 0.66667 | 154.0008 | 150.0008 | 152331.0678 | 202290.3113 | 11170.3657 | 259075.6478 | 197216.0209 | 180605.1451 | 3391.736 | 9.3333 | 31.3334 | 63.3334 | 98.6668
14 | ICV PE | R | 06/02/2023 16:30:09 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\ICV PE.016 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 11302.4706 | 4784.8012 | 314.0035 | 47710.2043 | 2200.5028 | 5715.4767 | 2245.1765 | 8848.4062 | 12296.9576 | 79638.7034 | 55356.4758 | 119852.6795 | 12252.9191 | 24451.5765 | 6649.2138 | 17478.6884 | 154760.4465 | 75005.7266 | 55260.8218 | 839.358 | 64712.931 | 64712.931 | 5552.4124 | 1403.4023 | 4568.0637 | 9702.6274 | 5485.9912 | 1730.7715 | 159506.2415 | 129256.1184 | 167954.8939 | 198192.0532 | 11340.5016 | 237648.5696 | 184397.8319 | 181738.1293 | 3276.3757 | 83367.8891 | 198590.8734 | 287139.3176 | 713930.0734
15 | CCV 1ppb TRACE | R | 06/02/2023 16:36:12 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.017 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1322.728 | 567.3446 | 48.3334 | 4069.9131 | 752.0198 | 650.6815 | 251.3355 | 1030.3705 | 1416.7369 | 8999.5008 | 7327.1174 | 13896.7564 | 2617.9068 | 2853.285 | 787.3551 | 2095.4872 | 18526.0061 | 9279.0148 | 2615.5884 | 105.0004 | 7138.45 | 7138.45 | 1758.1082 | 160.0009 | 479.3414 | 1042.038 | 614.4017 | 183.3345 | 17773.0512 | 14828.3588 | 167751.1728 | 219473.7251 | 12405.7181 | 253755.3037 | 192724.737 | 178916.1584 | 3781.167 | 9137.5882 | 21992.918 | 31997.9533 | 75283.2706
16 | CCV 1ppb TRACE | R | 06/02/2023 16:42:44 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.018 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1433.4053 | 566.3446 | 45.0001 | 3646.4655 | 814.0232 | 719.6848 | 247.3355 | 1040.7046 | 1437.0723 | 8511.2015 | 6788.0152 | 14105.6273 | 2629.9088 | 2912.2969 | 826.3573 | 2086.4857 | 17905.8828 | 13467.6848 | 2585.5024 | 100.0004 | 7211.8207 | 7211.8207 | 1706.7686 | 160.6676 | 510.0091 | 1062.0395 | 547.0967 | 207.3348 | 17034.1503 | 14291.8121 | 159879.8628 | 210573.2893 | 12601.5564 | 242172.7819 | 184857.6825 | 191460.5238 | 3723.8187 | 8682.6378 | 20465.9856 | 30693.6894 | 71290.4124
17 | LFB 10ppb | R | 06/02/2023 16:48:46 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LFB 10ppb.019 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 13144.3783 | 5193.2773 | 316.6702 | 27980.7097 | 2411.5369 | 6425.1112 | 2415.2042 | 9306.6979 | 12752.023 | 78444.7877 | 57707.4897 | 127041.0701 | 13470.0142 | 26334.585 | 7080.421 | 17844.472 | 150655.5967 | 98031.8797 | 54324.3604 | 931.6971 | 65799.8598 | 65799.8598 | 7028.3952 | 1512.7468 | 4580.7344 | 9732.6482 | 5285.8312 | 1789.4454 | 155809.0683 | 129172.7797 | 154025.2739 | 193170.6946 | 11563.6791 | 222034.8778 | 170737.5701 | 178051.3927 | 3349.7264 | 77340.1522 | 183647.5567 | 257941.3637 | 604965.405
18 | LFB 10ppb | R | 06/02/2023 16:55:16 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LFB 10ppb.020 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 12935.1877 | 5263.6362 | 345.6708 | 28545.1594 | 2397.5347 | 6343.4081 | 2404.2023 | 9312.7011 | 12779.0468 | 76344.1257 | 56590.6509 | 128339.2417 | 13287.8441 | 26206.6829 | 7111.4363 | 17801.0848 | 149028.65 | 96780.0654 | 52920.2814 | 990.701 | 65215.8609 | 65215.8609 | 7123.109 | 1514.7471 | 4564.7293 | 9852.063 | 5157.3201 | 1745.4401 | 152193.159 | 129653.7018 | 152595.3823 | 192459.8529 | 11437.2434 | 218308.8433 | 167749.982 | 175515.7858 | 3287.0449 | 75435.3189 | 179279.9645 | 251213.944 | 592500.643
19 | LLOQ 0.01 | R | 06/02/2023 17:01:19 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LLOQ 0.01.021 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 13.3333 | 5 | 12.3333 | 3315.7186 | 579.0117 | 14.3333 | 4.3333 | 12 | 25 | 112.0004 | 1362.5565 | 162.6676 | 1450.0736 | 32 | 16.6667 | 41.3334 | 358.6712 | 1289.3916 | -2936.0884 | 0.66667 | 108.6671 | 108.6671 | 1062.0396 | 4 | 6 | 14.6667 | 43.7791 | 1.3333 | 402.0057 | 360.6712 | 151811.7147 | 207220.2345 | 12185.5285 | 235343.3661 | 171146.8218 | 185712.7347 | 3680.4742 | 84.0002 | 221.3351 | 338.0016 | 748.0079
20 | CCV 10 ppb MAJORS | R | 06/02/2023 17:07:22 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.022 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 13728.2598 | 5550.0779 | 324.0037 | 26978.1218 | 2440.8752 | 6574.5132 | 2443.209 | 9280.3469 | 12878.47 | 77380.3389 | 53327.1346 | 130354.7732 | 13054.629 | 26469.8344 | 7110.4358 | 17751.0236 | 148270.1433 | 71625.1102 | 52096.6463 | 932.3638 | 64986.1437 | 64986.1437 | 7065.7476 | 1421.4041 | 4530.0515 | 9618.5715 | 5096.2882 | 1702.7683 | 150087.6368 | 128635.8972 | 151695.1769 | 188879.1644 | 11114.9936 | 214186.4725 | 170593.2057 | 173564.123 | 3235.033 | 74383.1489 | 177508.098 | 253768.8841 | 591874.5474
21 | CCV 10 ppb MAJORS | R | 06/02/2023 17:13:54 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.023 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 14060.583 | 5591.7609 | 329.0038 | 26495.215 | 2489.5503 | 6822.2953 | 2525.89 | 9581.212 | 13190.4206 | 78103.5908 | 54064.7188 | 134177.9539 | 13143.0432 | 27047.5846 | 7284.5235 | 18098.792 | 149908.4416 | 72293.1314 | 52547.3023 | 944.3646 | 65340.4282 | 65340.4282 | 7053.7415 | 1458.0746 | 4591.4046 | 9782.0158 | 5250.7766 | 1717.4366 | 151204.0119 | 129831.3473 | 151280.8643 | 182241.17 | 11431.239 | 215573.8082 | 164422.9991 | 181938.5843 | 3227.6988 | 74609.6677 | 177866.4416 | 255066.2407 | 594661.2494
22 | LLOQ 0.1 | R | 06/02/2023 17:19:57 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LLOQ 0.1.024 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 153.3342 | 51.3334 | 17 | 2948.9712 | 624.0137 | 85.3336 | 29 | 96.667 | 148.0008 | 840.6915 | 2618.4043 | 1522.0811 | 1695.7674 | 279.3361 | 92.3336 | 219.6684 | 1917.4622 | 2223.5065 | -2573.3224 | 10 | 729.352 | 729.352 | 1147.3794 | 17.3333 | 50.0001 | 110.6671 | 99.6932 | 14.6667 | 1650.7622 | 1456.7413 | 163654.5928 | 206499.229 | 12089.4481 | 244179.7047 | 181189.4878 | 184214.2113 | 3683.8085 | 792.6887 | 1954.1337 | 2902.7807 | 6785.9629
23 | LBLANK | R | 06/02/2023 17:25:59 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LBLANK.025 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.6667 | 0 | 4.6667 | 1642.7611 | 629.6806 | 11.3333 | 0 | 1.3333 | 6.6667 | 29.3334 | 928.4781 | 29.3334 | 1397.4018 | 1 | 6.6667 | 11.6667 | 106.0004 | 638.681 | -2823.5466 | 0.66667 | 15.3333 | 15.3333 | 1068.04 | 1.3333 | 2 | 4.6667 | 25.0906 | 2.6667 | 28.6667 | 15.3333 | 150831.4389 | 204707.5949 | 12272.2697 | 230293.733 | 177548.5211 | 186473.8448 | 3747.1582 | 6.6667 | 14.6667 | 35.3334 | 76.6667
24 | LFB 6.25ppb & 2500ppb | R | 06/02/2023 17:32:01 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LFB 6.25ppb & 2500ppb.026 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 33645.5743 | 13306.1955 | 608.013 | 22135.8217 | 5165.934 | 16496.854 | 5996.9252 | 5781.503 | 7927.8659 | 47575.7598 | 36105.9891 | 80355.5534 | 9233.3163 | 16453.1361 | 4351.9964 | 10550.562 | 88230.2984 | 44945.9427 | 30795.7056 | 559.011 | 39883.6103 | 39883.6103 | 4632.7516 | 909.3624 | 2804.9421 | 5913.2236 | 3227.704 | 1073.3737 | 91614.2181 | 77215.5444 | 138233.6768 | 181501.9314 | 10377.4352 | 196261.6271 | 152534.7005 | 159055.9435 | 3273.045 | 44733.2799 | 105439.7482 | 153882.0882 | 347451.9292
25 | MT 1 | R | 06/02/2023 17:38:02 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 1.027 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 4675.7651 | 134.334 | 11052.2742 | 5026533.6134 | 945.3646 | 2085.1522 | 143.3341 | 401.339 | 693.6835 | 4021.2327 | 439869.3728 | 10235.6674 | 101541.2699 | 173.3344 | 254.0023 | 7283.523 | 60088.1271 | 89677.9495 | 5284.7171 | 131.6673 | 10534.5506 | 10534.5506 | 1397.4017 | 128.0006 | 382.0051 | 824.6906 | 505.071 | 157.3342 | 1162.0473 | 948.6982 | 144844.6375 | 197024.032 | 11442.9149 | 219361.3746 | 171321.164 | 160237.9986 | 3575.115 | 542.677 | 1308.06 | 4576.953 | 10280.1005
26 | MT 1 LD | R | 06/02/2023 17:44:04 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 1 LD.028 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 4917.5129 | 137.0007 | 11614.3863 | 5214157.8844 | 962.3658 | 2218.839 | 150.3341 | 394.6721 | 687.3499 | 4103.9228 | 429123.4942 | 10487.1816 | 97812.3897 | 184.0012 | 247.6688 | 7150.7895 | 60166.4394 | 92492.4653 | 5585.6316 | 131.6673 | 10880.1431 | 10880.1431 | 1380.0667 | 141.334 | 398.6722 | 906.6955 | 525.8086 | 167.3343 | 1156.0469 | 962.6991 | 134023.9811 | 194272.0411 | 11274.4472 | 206447.5151 | 167556.2411 | 161413.4497 | 3406.4065 | 583.3453 | 1324.0614 | 4500.2771 | 10496.8406
27 | LBLANK | R | 06/02/2023 18:05:56 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LBLANK.029 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 3.3333 | 0.33333 | 4 | 1856.1207 | 434.3407 | 9.3333 | 0 | 0.66667 | 10.6667 | 29.3334 | 950.7696 | 28 | 1167.7156 | 1 | 8.6667 | 13.6667 | 99.3337 | 517.3427 | -2748.3509 | 0 | 24 | 24 | 1016.0363 | 0.66667 | 3.3333 | 6 | 29.741 | 0 | 21.3334 | 30 | 142124.1783 | 197414.6228 | 11533.0139 | 215951.5156 | 168258.4464 | 175336.8371 | 3644.4651 | 0 | 19.3333 | 28.6667 | 58.6667
28 | STD 1 | R | 06/02/2023 18:11:58 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\STD 1.030 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1563.4189 | 611.6798 | 60.0001 | 6284.0487 | 795.0221 | 769.6874 | 263.0024 | 978.7002 | 1421.7374 | 7989.5677 | 12138.8373 | 14232.4216 | 3755.4938 | 2804.2752 | 745.6861 | 2013.4752 | 16323.3222 | 9871.4098 | 2792.0441 | 87.0003 | 6644.8786 | 6644.8786 | 1556.085 | 170.001 | 485.3416 | 1036.7043 | 514.2143 | 206.6682 | 14796.9938 | 12976.5578 | 139194.8733 | 188660.9318 | 11125.9981 | 208790.0466 | 164122.7696 | 171734.3065 | 3353.0603 | 7464.6169 | 17612.851 | 26630.3868 | 62263.864
29 | STD 0.5 | R | 06/02/2023 18:18:06 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\STD 0.5.031 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 816.3567 | 291.003 | 52.3334 | 6074.6247 | 729.352 | 378.3384 | 122.0005 | 487.0083 | 705.3508 | 4001.894 | 6368.0946 | 7101.0987 | 2610.2385 | 1369.7323 | 392.0054 | 1004.3687 | 8461.1717 | 5443.0368 | -98.1369 | 50.3334 | 3565.1117 | 3565.1117 | 1207.3844 | 73.3335 | 255.3356 | 499.3421 | 245.2424 | 94.0003 | 7538.6552 | 6557.5048 | 144190.7011 | 185612.7502 | 11343.5028 | 206054.1387 | 162058.7423 | 167877.3475 | 3285.7115 | 3679.141 | 8866.7511 | 12876.2601 | 30805.0542
30 | ICV PE | R | 06/02/2023 18:24:08 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\ICV PE.032 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 14582.1059 | 6051.6155 | 355.0044 | 56314.1213 | 2525.89 | 7070.4182 | 2555.8954 | 9310.0336 | 12757.6941 | 74928.0005 | 53118.9567 | 131411.6516 | 12896.8193 | 25790.9285 | 6936.6839 | 16864.9492 | 138473.8983 | 66923.7349 | 49737.9416 | 767.354 | 61747.8511 | 61747.8511 | 5038.2218 | 1365.3986 | 4390.0078 | 9348.3911 | 5030.6929 | 1640.0945 | 132853.5698 | 116423.8163 | 138961.264 | 170011.6725 | 10382.7717 | 195913.7733 | 154115.5183 | 155238.2888 | 3058.9943 | 67662.5343 | 161460.7174 | 232797.8725 | 559702.5405
31 | ICV PE | R | 06/02/2023 18:30:39 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\ICV PE.033 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 14626.8178 | 6020.6019 | 352.671 | 56267.9343 | 2480.5487 | 6905.002 | 2603.5707 | 9438.7842 | 12836.0978 | 74964.1862 | 53821.2539 | 134355.0141 | 12944.195 | 26282.4882 | 7040.0677 | 17180.9925 | 139456.0791 | 67466.2697 | 49225.8947 | 795.6889 | 62560.0173 | 62560.0173 | 5018.2146 | 1461.4081 | 4450.6934 | 9469.8044 | 4947.143 | 1704.7685 | 134547.3187 | 119694.0507 | 140173.7185 | 176288.4629 | 10594.2605 | 195773.885 | 155613.7196 | 167457.2363 | 3167.6846 | 68888.3672 | 163797.7668 | 231363.2596 | 558093.2284
32 | CCV 1ppb TRACE | R | 06/02/2023 18:36:41 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.034 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1560.752 | 636.3475 | 37.6667 | 4737.4531 | 770.3541 | 773.6876 | 268.3359 | 1010.3691 | 1400.402 | 8095.6266 | 6571.453 | 14460.6502 | 2611.9054 | 2864.2871 | 798.0223 | 1943.4655 | 16004.9606 | 7860.829 | 2647.3667 | 98.667 | 6797.6172 | 6797.6172 | 1592.7556 | 162.6676 | 466.6744 | 1051.3721 | 513.2726 | 171.3344 | 15113.991 | 13314.2017 | 140177.1374 | 182573.1799 | 11606.049 | 211623.3522 | 163291.273 | 172642.3142 | 3330.3889 | 7491.9642 | 17880.5186 | 26094.033 | 61596.788
33 | CCV 1ppb TRACE | R | 06/02/2023 18:43:13 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.035 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1617.0915 | 659.3486 | 53.0001 | 3844.5173 | 848.6919 | 749.6863 | 289.6696 | 1022.3699 | 1457.7411 | 8188.3462 | 6402.4757 | 14542.7321 | 2714.9247 | 2934.6348 | 807.3561 | 2004.8074 | 15928.2089 | 7748.7678 | 2401.179 | 84.6669 | 6774.2725 | 6774.2725 | 1541.4165 | 145.3341 | 460.6741 | 1036.7043 | 524.8168 | 182.6678 | 14720.2475 | 12961.211 | 139511.253 | 188081.5374 | 11315.1461 | 206739.3439 | 161535.5529 | 168004.1199 | 3365.7301 | 7464.6166 | 17498.0436 | 26150.0039 | 61004.3702
34 | CCV 10 ppb MAJORS | R | 06/02/2023 18:49:16 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.036 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 15494.3988 | 6277.7128 | 362.3379 | 28401.5463 | 2720.259 | 7182.1387 | 2700.5886 | 9771.6749 | 13414.9624 | 76131.6557 | 52826.1225 | 135986.1649 | 13359.9109 | 27578.9296 | 7454.6115 | 17832.1232 | 142233.246 | 68521.9421 | 50067.7542 | 895.0281 | 63052.8417 | 63052.8417 | 6761.5999 | 1416.7369 | 4538.7209 | 9370.4059 | 4957.0312 | 1650.0954 | 139804.0917 | 122609.2982 | 139313.3907 | 176965.3537 | 10864.1297 | 194370.8024 | 154444.0407 | 160979.4707 | 3164.3507 | 68254.6844 | 161996.7433 | 231495.9532 | 548816.1667
35 | CCV 10 ppb MAJORS | R | 06/02/2023 18:55:47 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.037 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 15000.205 | 5903.553 | 342.6708 | 27376.2085 | 2603.9041 | 7002.3825 | 2541.5595 | 9655.2618 | 13205.4344 | 76428.5887 | 52476.0976 | 132944.426 | 13102.0059 | 26926.0213 | 7265.5139 | 17724.9903 | 141308.1955 | 69028.3946 | 50055.364 | 915.3627 | 63257.7531 | 63257.7531 | 6743.5915 | 1452.7406 | 4432.6882 | 9260.3342 | 4904.4475 | 1649.4286 | 140566.9364 | 123339.4893 | 137711.9193 | 174914.3109 | 10709.6798 | 195365.2362 | 147201.2731 | 156556.6226 | 3156.3487 | 68201.7797 | 161068.2705 | 229102.1888 | 544823.3175
36 | MT 1 | R | 06/02/2023 19:01:50 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 1.038 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 4993.8728 | 139.6673 | 12076.1021 | 5310204.5811 | 994.368 | 2181.8333 | 165.001 | 404.6724 | 696.6837 | 3925.873 | 431657.257 | 10459.8303 | 102817.3517 | 189.3346 | 244.3354 | 7043.7361 | 56348.2609 | 85981.9875 | 5120.8704 | 125.3339 | 10398.4508 | 10398.4508 | 1346.0635 | 126.6672 | 388.672 | 844.025 | 475.7795 | 138.6673 | 1127.3778 | 969.3662 | 136255.3949 | 167334.0119 | 11002.9024 | 203483.6597 | 158164.14 | 165982.7923 | 2971.6451 | 530.0099 | 1267.3896 | 4081.564 | 9293.84
37 | MT 1 LD | R | 06/02/2023 19:07:53 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 1 LD.039 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 5094.9085 | 142.6674 | 12463.1021 | 5389603.7971 | 964.6993 | 2272.5143 | 170.001 | 385.0052 | 673.6826 | 3877.8596 | 417948.5011 | 10473.8384 | 99934.7056 | 185.0012 | 268.3359 | 7109.4355 | 57717.6968 | 88991.658 | 5450.1042 | 123.6672 | 10450.4892 | 10450.4892 | 1304.7263 | 130.6673 | 404.0058 | 872.0266 | 507.1788 | 163.3343 | 1046.7051 | 901.3618 | 133188.6655 | 180148.2406 | 10453.4981 | 199046.5042 | 157483.2956 | 160897.8818 | 3304.3822 | 568.0113 | 1274.0568 | 4071.5595 | 9478.5701
38 | MT 2 | R | 06/02/2023 19:13:55 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 2.040 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 5684.4641 | 136.6673 | 13708.5761 | 5753199.7599 | 1020.7031 | 2429.8734 | 169.3343 | 414.006 | 727.3519 | 4125.2623 | 425274.0469 | 10028.1853 | 100430.4823 | 190.0013 | 252.0022 | 7587.6814 | 61426.4569 | 91681.2553 | 5787.5016 | 147.0008 | 10995.5652 | 10995.5652 | 1332.7289 | 119.3338 | 434.6733 | 926.0301 | 515.9321 | 158.6675 | 1430.0717 | 1249.388 | 132124.199 | 177085.044 | 10714.0171 | 194790.4249 | 155296.9091 | 159628.3091 | 3157.0175 | 536.0101 | 1271.3899 | 4070.8938 | 9541.914
39 | MT 2 LSP | R | 06/02/2023 19:19:58 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 2 LSP.041 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 40349.5763 | 13847.041 | 13016.2608 | 6021485.1167 | 5581.09 | 18940.2155 | 6040.2768 | 6035.2746 | 8350.44 | 48879.4798 | 449698.185 | 87886.8656 | 101842.7325 | 15426.659 | 4320.6536 | 16556.5903 | 135266.7699 | 127881.9018 | 36168.5142 | 661.3487 | 47008.5579 | 47008.5579 | 4639.4202 | 960.0323 | 3012.3177 | 6303.3909 | 3366.6288 | 1181.3822 | 83059.437 | 71194.2998 | 124900.3051 | 160710.2982 | 9356.0679 | 178301.8517 | 138288.8266 | 143964.3133 | 2800.2744 | 39820.7593 | 94343.8422 | 136419.2451 | 316552.8324
40 | MT 2 LSP | R | 06/02/2023 19:26:09 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 2 LSP.042 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 39802.0482 | 13557.7652 | 12958.5421 | 5975735.763 | 5630.7762 | 18514.9923 | 6011.9315 | 6069.9561 | 8237.7081 | 49880.9355 | 456266.2417 | 87664.8448 | 102132.138 | 15742.0028 | 4409.0136 | 16908.6689 | 139061.6063 | 130451.5859 | 37282.2835 | 677.016 | 47637.963 | 47637.963 | 4709.4433 | 984.0339 | 2984.3119 | 6413.4394 | 3429.7608 | 1156.0468 | 85559.4705 | 73564.9464 | 128148.2593 | 160164.9291 | 9883.4199 | 183852.8584 | 143600.1409 | 148532.8516 | 2858.9529 | 40672.4903 | 96518.2885 | 140844.9954 | 326785.5061
41 | MT 3 | R | 06/02/2023 19:32:11 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 3 .043 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 4686.1018 | 124.6672 | 11234.4171 | 4868747.3062 | 1005.3687 | 2058.1483 | 130.0006 | 365.6714 | 821.6903 | 4689.4372 | 377515.8504 | 8981.4902 | 86232.1625 | 172.3344 | 281.3361 | 6590.854 | 54561.3458 | 81245.7088 | 5028.8481 | 141.6674 | 10019.5136 | 10019.5136 | 1368.7324 | 141.334 | 389.3386 | 818.0234 | 503.8993 | 150.6675 | 1154.0467 | 934.6972 | 135927.6343 | 188286.7732 | 11285.1229 | 206055.5583 | 157327.6248 | 166069.064 | 3387.0684 | 484.6749 | 1208.718 | 4114.8944 | 9326.5267
42 | MT 4 | R | 06/02/2023 19:38:14 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 4.044 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 5825.5215 | 138.334 | 14235.7564 | 6066386.6517 | 1133.045 | 2586.5676 | 170.6677 | 455.3406 | 771.6875 | 4357.3316 | 584217.5787 | 12377.3599 | 138085.5582 | 215.335 | 266.6692 | 8004.2419 | 64878.3252 | 98458.1364 | 6694.9147 | 155.3342 | 11999.7047 | 11999.7047 | 1373.3994 | 144.6674 | 444.6736 | 990.701 | 559.1542 | 165.3343 | 1274.0568 | 1060.7061 | 133692.7936 | 179527.8694 | 10609.9386 | 199132.4504 | 156175.7518 | 161023.3211 | 3259.7052 | 619.3468 | 1459.4079 | 4426.9349 | 10161.4206
43 | LBLANK | R | 06/02/2023 19:44:17 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LBLANK.045 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.66667 | 0.33333 | 2 | 3183.6881 | 601.3461 | 15.3333 | 0 | 1 | 7.3333 | 22.6667 | 896.344 | 26 | 1283.0576 | 1.3333 | 9 | 9.6667 | 104.0004 | 537.3434 | -2741.4745 | 0.33333 | 26.6667 | 26.6667 | 969.3662 | 0.66667 | 3.3333 | 7.3333 | 42.5323 | 1.3333 | 34 | 14 | 135236.4674 | 180037.4239 | 10735.6995 | 195111.13 | 155773.6461 | 174898.4378 | 3292.3794 | 0.66667 | 24.6667 | 22 | 56
44 | CCV 1ppb TRACE | R | 06/02/2023 19:50:20 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.046 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1596.0892 | 641.3477 | 49.0001 | 3973.8862 | 827.024 | 773.3543 | 282.3361 | 1039.3712 | 1376.733 | 8096.2946 | 6461.9842 | 13860.721 | 2615.2395 | 2844.95 | 798.0223 | 1912.4613 | 15157.3705 | 7598.6872 | 2136.337 | 92.667 | 6301.3898 | 6301.3898 | 1543.4168 | 140.0007 | 440.6735 | 971.3664 | 450.2057 | 170.6677 | 13931.459 | 12245.2468 | 128252.7365 | 172602.4461 | 10644.9647 | 193088.909 | 151309.8731 | 159650.519 | 3160.3496 | 6933.0156 | 16524.218 | 23735.0488 | 56276.9253
45 | CCV 1ppb TRACE | R | 06/02/2023 19:56:52 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.047 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1602.4233 | 662.6821 | 53.0001 | 3588.4508 | 848.6919 | 739.0191 | 269.0025 | 1007.0355 | 1395.7349 | 7958.217 | 6252.6422 | 13771.3025 | 2590.2348 | 2801.9415 | 782.0214 | 1950.4665 | 15340.8996 | 7391.9134 | 2489.4752 | 92.667 | 6256.7033 | 6256.7033 | 1480.7434 | 138.0007 | 445.3403 | 962.6991 | 480.9224 | 192.0013 | 13861.3884 | 12054.4183 | 128243.058 | 175163.3795 | 10753.0455 | 193599.0231 | 152742.968 | 159227.1506 | 3173.0192 | 6810.9569 | 16334.6671 | 23153.9859 | 56139.1503
46 | CCV 10 ppb MAJORS | R | 06/02/2023 20:02:55 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.048 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 15103.9804 | 6122.9791 | 354.0044 | 27945.3136 | 2670.2497 | 7021.3921 | 2535.8919 | 9462.7999 | 12841.1021 | 73341.7905 | 51431.5042 | 127452.6953 | 12738.6776 | 26433.0997 | 7044.0698 | 16728.1226 | 133230.4097 | 65237.9607 | 47110.114 | 858.3592 | 58590.5854 | 58590.5854 | 6388.0949 | 1332.0623 | 4061.9108 | 8644.6148 | 4618.0729 | 1540.7498 | 129440.514 | 112787.6335 | 128424.0928 | 161333.3506 | 9884.0852 | 180318.8873 | 141062.5391 | 145793.7914 | 2900.2944 | 63027.4037 | 148736.9842 | 203372.6412 | 491472.1308
47 | CCV 10 ppb MAJORS | R | 06/02/2023 20:09:27 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.049 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 15374.2696 | 6142.654 | 362.0046 | 27995.4056 | 2594.9023 | 7112.7703 | 2581.2334 | 9668.9385 | 13006.5864 | 74274.6082 | 51464.5186 | 130413.2621 | 12755.3597 | 26991.1444 | 7179.1368 | 17123.5961 | 136304.5586 | 65230.5958 | 46895.6316 | 885.0274 | 59532.453 | 59532.453 | 6537.496 | 1338.0627 | 4114.5926 | 8963.478 | 4647.9767 | 1574.0869 | 131765.6116 | 115371.3371 | 129573.0782 | 162414.0373 | 10226.9935 | 179926.8098 | 135480.1754 | 149038.7763 | 2944.3034 | 63320.0495 | 149921.9292 | 206793.2157 | 501746.7127
48 | MT 5 | R | 06/02/2023 20:15:30 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 5.050 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 5897.884 | 157.0009 | 13792.3219 | 6042779.1702 | 1084.0412 | 2475.5478 | 181.0011 | 431.3398 | 736.6857 | 4119.2606 | 468399.0749 | 11227.7443 | 113220.5736 | 216.0017 | 267.0025 | 7574.3409 | 59759.405 | 92709.8897 | 5921.6857 | 134.6673 | 10756.7151 | 10756.7151 | 1384.067 | 133.334 | 415.3394 | 871.3601 | 521.8462 | 142.6674 | 1136.0452 | 938.0308 | 130798.729 | 178717.776 | 10548.2274 | 193578.1334 | 151337.4828 | 154341.3133 | 3337.7232 | 565.3445 | 1319.3943 | 3887.5388 | 9243.8611
49 | MT 6 | R | 06/02/2023 20:21:32 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 6.051 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 13607.8119 | 330.0038 | 30886.0312 | 13839215.654 | 1744.1065 | 5781.8367 | 408.3392 | 955.3653 | 1458.7412 | 8091.6244 | 990303.8762 | 23914.6709 | 237382.8618 | 430.3398 | 560.3443 | 17131.266 | 134124.7765 | 206890.1015 | 16744.2341 | 347.6709 | 24875.6398 | 24875.6398 | 1820.7827 | 286.0029 | 974.7 | 1952.8001 | 1093.8826 | 368.0047 | 2477.5485 | 1949.4664 | 127129.159 | 168286.7897 | 10156.2764 | 186880.4102 | 143389.9955 | 154815.7027 | 3046.9916 | 1283.3911 | 2926.9665 | 7086.6841 | 16674.4652
50 | MT 7 | R | 06/02/2023 20:27:35 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 7.052 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 13462.3415 | 346.6709 | 32226.6462 | 13519870.5745 | 1741.4395 | 5856.8671 | 364.338 | 929.6969 | 1464.075 | 8410.4782 | 963050.5235 | 23176.1185 | 231784.6803 | 444.0069 | 612.6798 | 17046.4979 | 133811.9896 | 205479.3881 | 17145.7094 | 333.6706 | 24853.6022 | 24853.6022 | 1899.46 | 274.6693 | 972.6998 | 1970.1359 | 1195.913 | 338.6707 | 2644.912 | 2030.811 | 129252.1457 | 167696.549 | 10167.2835 | 188551.2353 | 147612.9305 | 150956.3899 | 3067.6633 | 1186.0493 | 2906.9626 | 8529.6589 | 19479.259
51 | MT 8 | R | 06/02/2023 20:33:38 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 8.053 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 12927.8499 | 324.6704 | 30540.2837 | 12806690.6619 | 1701.768 | 5591.0941 | 380.0051 | 895.028 | 1483.077 | 8218.3654 | 897376.003 | 22548.449 | 215614.6779 | 421.3396 | 584.6786 | 16193.1739 | 127149.3987 | 193165.2434 | 15582.166 | 317.0035 | 23379.1161 | 23379.1161 | 1806.1143 | 268.0025 | 897.3615 | 1870.1224 | 1045.2764 | 338.004 | 2383.5322 | 1968.8024 | 128172.4594 | 169297.3043 | 10030.871 | 189685.047 | 148808.5385 | 151524.0481 | 2979.6443 | 1182.0491 | 2700.2553 | 7014.6625 | 16528.4029
52 | MT 9 | R | 06/02/2023 20:39:40 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 9.054 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 12612.2323 | 303.0032 | 29940.343 | 12810864.0335 | 1670.7644 | 5304.3182 | 361.3379 | 900.3618 | 1378.7332 | 7775.4489 | 894097.7286 | 21923.4791 | 212982.6554 | 426.0064 | 542.0103 | 18575.0686 | 146106.7379 | 194593.8088 | 15901.0861 | 316.0035 | 23173.4499 | 23173.4499 | 1808.1144 | 296.6698 | 873.3601 | 1862.7884 | 1095.1089 | 345.3375 | 2396.201 | 1860.1212 | 120959.5698 | 166886.0691 | 10207.646 | 183249.8391 | 146179.4366 | 147215.3498 | 3008.3167 | 1146.7127 | 2748.9314 | 6777.2978 | 15741.3924
53 | MT 10 | R | 06/02/2023 20:45:42 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 10.055 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 12846.1095 | 340.6707 | 31316.2942 | 13548939.5864 | 1681.0989 | 5586.4254 | 368.3381 | 918.3629 | 1628.4261 | 9046.8641 | 982339.3404 | 23648.8926 | 232297.6414 | 427.3397 | 580.6785 | 18561.0516 | 145889.9369 | 206825.7853 | 16565.7492 | 316.0035 | 24107.6576 | 24107.6576 | 1786.7784 | 273.336 | 905.362 | 1965.4688 | 1075.1983 | 332.6706 | 2479.5485 | 2064.1492 | 123539.8879 | 161519.2994 | 9978.1514 | 183500.4057 | 144693.1356 | 147368.2118 | 2880.2904 | 1196.0501 | 2731.5945 | 6841.9752 | 15902.1144
54 | MT 11 | R | 06/02/2023 20:51:44 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 11 .056 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 24851.5997 | 626.0137 | 59059.5018 | 26430304.0247 | 2711.5907 | 10716.6879 | 783.6882 | 1816.7823 | 2687.5861 | 15664.584 | 1845775.0498 | 46521.628 | 441256.5968 | 837.0245 | 1108.043 | 30594.3923 | 241610.0079 | 392105.0939 | 33970.2561 | 608.6797 | 46333.0196 | 46333.0196 | 2635.5766 | 570.6781 | 1748.1071 | 3807.1744 | 2204.0241 | 689.35 | 4987.5372 | 3724.4856 | 121253.7655 | 159139.6013 | 9787.6852 | 177527.5964 | 139817.6084 | 144517.4363 | 2816.9444 | 2319.5216 | 5343.6662 | 11733.869 | 27965.5072
55 | MT 11 LD | R | 06/02/2023 20:57:47 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 11 LD.057 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 24996.8523 | 595.3458 | 58398.4839 | 26010903.1829 | 2652.2462 | 10394.1133 | 792.3554 | 1748.7737 | 2731.2612 | 15798.0643 | 1819009.6493 | 45638.1349 | 429413.3384 | 2030.811 | 1073.0403 | 30316.8047 | 241585.7664 | 389335.404 | 34104.7151 | 621.3468 | 45564.5544 | 45564.5544 | 2562.8968 | 548.0106 | 1748.107 | 3791.8366 | 2125.292 | 680.0162 | 4732.7839 | 3646.4655 | 120725.3317 | 161174.7755 | 9701.9611 | 178193.9609 | 140444.9968 | 149055.6547 | 2799.6078 | 2243.5096 | 5248.2973 | 12021.9834 | 27721.2104
56 | MT 11 LD | R | 06/02/2023 21:04:18 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 11 LD.058 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 25056.9597 | 629.3472 | 59782.8945 | 26905260.0004 | 2694.5875 | 10803.0843 | 802.0225 | 1791.4458 | 2834.6146 | 15922.869 | 1887954.9982 | 46667.437 | 443108.991 | 2054.8146 | 1114.3769 | 31182.6646 | 245920.7078 | 397203.0604 | 34436.848 | 632.3473 | 46606.5694 | 46606.5694 | 2536.2253 | 535.3434 | 1774.1103 | 3730.487 | 2151.9793 | 650.0148 | 4652.7576 | 3729.8205 | 123525.8371 | 160919.9653 | 9337.7394 | 183494.2989 | 146405.128 | 144846.0727 | 2879.6237 | 2268.1801 | 5276.9746 | 11873.2686 | 27946.719
57 | MT 12 | R | 06/02/2023 21:10:20 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 12.059 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 25683.0664 | 658.6819 | 59309.5314 | 27282499.9307 | 2712.9246 | 10923.8424 | 762.687 | 1791.4457 | 2729.5941 | 15756.7022 | 1841166.1531 | 45502.3551 | 430761.005 | 881.3605 | 1045.3716 | 31197.3627 | 250224.1358 | 395691.968 | 35170.3768 | 630.0139 | 47433.9567 | 47433.9567 | 2645.5783 | 612.0131 | 1896.1265 | 3940.5434 | 2186.0212 | 712.0178 | 4671.432 | 3703.1466 | 120311.7469 | 159872.3432 | 9710.637 | 174475.8863 | 140914.4996 | 143179.5855 | 2937.6353 | 2367.5297 | 5357.672 | 12050.6602 | 28437.7432
58 | MT 12 LSP | R | 06/02/2023 21:16:23 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 12 LSP.060 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 61755.5887 | 14616.1438 | 56331.5179 | 27533500.2998 | 7519.3121 | 28055.1911 | 6803.9533 | 7525.6485 | 10185.9639 | 60847.9806 | 1828794.1271 | 120170.0425 | 412829.0333 | 16318.3149 | 5172.9365 | 38511.841 | 309415.6243 | 418296.4102 | 65888.0607 | 1122.3775 | 81147.1591 | 81147.1591 | 6194.0092 | 1394.0681 | 4375.3367 | 9117.5755 | 5110.7251 | 1674.7649 | 83221.0553 | 71369.8545 | 115192.398 | 148272.1867 | 8982.8234 | 164505.223 | 130282.0417 | 139379.3903 | 2731.5948 | 38584.0386 | 91393.4338 | 131919.9202 | 309858.0952
59 | LBLANK | R | 06/02/2023 21:22:26 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LBLANK.061 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 5.6667 | 1.3333 | 6 | 2210.8383 | 629.6806 | 11.6667 | 0 | 2.6667 | 9 | 22.6667 | 966.171 | 28 | 1370.0658 | 1.3333 | 14 | 17.3333 | 126.0006 | 736.0194 | -2725.7127 | 0.66667 | 16 | 16 | 1004.0353 | 0.66667 | 4.6667 | 6 | 51.2261 | 0.66667 | 178.6678 | 18.6667 | 134720.4606 | 170805.145 | 10731.363 | 192305.7812 | 155442.5215 | 153069.7151 | 3167.0177 | 4.6667 | 20.6667 | 29.3333 | 54.6667
60 | CCV 1ppb TRACE | R | 06/02/2023 21:28:29 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.062 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1704.435 | 678.3494 | 40.0001 | 4174.6099 | 872.36 | 789.3551 | 279.3361 | 1037.0377 | 1376.3997 | 7958.2161 | 6241.1615 | 13608.48 | 2626.9082 | 2835.6148 | 783.6882 | 1970.8029 | 15324.8824 | 7558.6659 | 2440.9301 | 87.6669 | 6221.3546 | 6221.3546 | 1502.0791 | 119.3338 | 438.6734 | 923.3632 | 508.3272 | 156.6675 | 13537.0782 | 11582.6943 | 127328.3794 | 168464.2953 | 10240.3378 | 184730.1891 | 146846.3703 | 151761.863 | 3091.668 | 6681.5624 | 15809.4108 | 22023.9836 | 52222.769
61 | CCV 1ppb TRACE | R | 06/02/2023 21:35:00 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.063 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1754.4411 | 723.0183 | 47.6667 | 3920.5385 | 894.028 | 801.3558 | 289.0029 | 1039.3711 | 1405.0691 | 7810.1353 | 6279.133 | 13538.4123 | 2656.2469 | 2883.9578 | 771.3542 | 1950.1331 | 15231.4517 | 7322.5433 | 2185.8313 | 84.6669 | 6154.6595 | 6154.6595 | 1454.7408 | 130.0006 | 455.3406 | 914.696 | 446.1192 | 160.6676 | 13306.8622 | 11543.3299 | 123219.8364 | 163585.5071 | 10398.4512 | 178047.8048 | 144301.8385 | 147804.7071 | 3052.9931 | 6493.4755 | 15366.9304 | 21328.8853 | 51312.5635
62 | LFB 10ppb | R | 06/02/2023 21:41:03 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LFB 10ppb.064 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 16161.1367 | 6322.0653 | 349.3376 | 30491.8447 | 2658.2475 | 7182.1387 | 2524.8898 | 9241.9894 | 12427.0701 | 73006.7522 | 52425.3739 | 118385.9688 | 12785.7198 | 25112.3862 | 6832.9672 | 16252.2398 | 129764.7086 | 83793.7066 | 44005.3802 | 821.6903 | 54509.8885 | 54509.8885 | 6148.6575 | 1227.3861 | 3934.5422 | 8134.9839 | 4175.9187 | 1512.08 | 121809.7838 | 104079.133 | 124459.7536 | 165483.5535 | 9448.458 | 167732.4454 | 134285.6268 | 138930.3167 | 3067.6669 | 58127.3617 | 137617.7218 | 182172.4178 | 437859.1032
63 | LFB 10ppb | R | 06/02/2023 21:47:33 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LFB 10ppb.065 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 15938.8877 | 6278.0461 | 369.0048 | 32197.579 | 2697.5882 | 7092.7604 | 2519.5557 | 9284.6834 | 12536.4992 | 72633.516 | 52447.7337 | 122844.6766 | 12984.5652 | 25716.1317 | 6971.3674 | 16497.8572 | 128767.8047 | 82743.606 | 44913.6089 | 823.0237 | 56716.3638 | 56716.3638 | 6139.9871 | 1286.0579 | 3868.5243 | 8158.3289 | 4451.716 | 1482.0769 | 121665.959 | 107108.0878 | 118736.0864 | 150460.2357 | 9509.4977 | 170571.7415 | 135062.8438 | 137357.2477 | 2718.2587 | 58431.2782 | 138768.1007 | 181170.2563 | 438914.0168
64 | CCV 10 ppb MAJORS | R | 06/02/2023 21:53:35 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.066 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 16315.312 | 6496.4771 | 366.3381 | 30022.8561 | 2787.2721 | 7283.19 | 2560.2296 | 9423.7739 | 12821.0847 | 73057.8097 | 50444.4815 | 126172.863 | 12592.5493 | 26111.1762 | 7121.4413 | 16711.4358 | 129926.3462 | 64133.6806 | 46924.952 | 884.6941 | 57284.6382 | 57284.6382 | 6252.0355 | 1224.7193 | 3932.5418 | 8108.3008 | 4528.9257 | 1495.4117 | 122632.2107 | 108653.6324 | 107144.1744 | 153110.1867 | 9617.5698 | 175211.1692 | 134292.2827 | 136320.006 | 2799.6076 | 58333.6442 | 137931.1635 | 188460.7683 | 448617.1479
65 | CCV 10 ppb MAJORS | R | 06/02/2023 22:00:06 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.067 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 16557.2593 | 6439.1186 | 356.6711 | 27916.5942 | 2718.2587 | 7392.5798 | 2621.574 | 9616.2356 | 13092.6652 | 74634.0217 | 50864.0322 | 124811.525 | 12758.3628 | 26509.9228 | 7050.4077 | 16895.6568 | 132726.1251 | 63900.6885 | 45334.9204 | 868.0264 | 56941.9252 | 56941.9252 | 6176.0023 | 1270.7233 | 3909.8684 | 8439.1597 | 4418.6766 | 1455.4075 | 124863.1863 | 107934.9302 | 119054.0769 | 151671.5693 | 9761.3517 | 163865.8021 | 133803.7286 | 151602.3969 | 2822.9456 | 59352.4262 | 140113.6134 | 184198.9377 | 447246.0769
66 | MT 13 | R | 06/02/2023 22:06:09 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 13.068 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 28628.9942 | 732.3522 | 65578.5341 | 29909647.2039 | 2986.3124 | 11911.9654 | 822.6904 | 1983.8044 | 2947.6374 | 16710.4345 | 1973345.879 | 49230.6915 | 459942.1833 | 812.6898 | 1101.7092 | 33137.7232 | 259496.9837 | 410777.5989 | 37291.4193 | 649.6814 | 49618.0247 | 49618.0247 | 2684.919 | 605.3462 | 1831.4508 | 3934.5419 | 2175.9187 | 660.682 | 5252.9657 | 3731.8207 | 112314.8915 | 154320.4279 | 9272.0083 | 171008.927 | 136888.0986 | 133641.6149 | 2789.606 | 2418.2047 | 5395.0187 | 11803.2576 | 27687.8016
67 | MT 14 | R | 06/02/2023 22:12:11 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 14.069 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 28680.4337 | 714.6846 | 63836.9918 | 28784315.6288 | 2914.6307 | 11848.9119 | 762.0204 | 1931.1306 | 2878.957 | 17831.8838 | 2167450.1378 | 52756.5698 | 454874.4568 | 873.36 | 1144.0458 | 32633.5692 | 256300.0629 | 401349.1201 | 36714.4897 | 652.0149 | 48069.4128 | 48069.4128 | 2610.9057 | 597.3458 | 1826.1167 | 3925.8727 | 2176.7344 | 709.351 | 5395.6863 | 3768.4977 | 116802.4074 | 152026.7135 | 9298.3586 | 171074.9254 | 134229.0125 | 134756.6352 | 2774.2696 | 2342.8588 | 5315.656 | 11970.6401 | 28034.0998
68 | MT 15 | R | 06/02/2023 22:18:14 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 15.070 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 27926.9464 | 706.3508 | 62444.8949 | 29242735.8518 | 2902.295 | 11742.1583 | 752.6865 | 1894.4591 | 2930.3005 | 17229.0499 | 1909251.5614 | 87297.2884 | 444973.2673 | 827.3573 | 1400.402 | 32041.8963 | 250479.5996 | 401460.994 | 36098.7824 | 610.0131 | 48223.9259 | 48223.9259 | 2627.575 | 586.6788 | 1868.1222 | 3933.8749 | 2168.0344 | 705.3509 | 5466.3792 | 3918.5375 | 114543.3948 | 151612.3087 | 9167.6077 | 169976.581 | 134488.0865 | 136857.7794 | 2768.9352 | 2376.8646 | 5456.3754 | 11678.5339 | 27459.6386
69 | MT 16 | R | 06/02/2023 22:24:17 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 16.071 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 8864.0832 | 502.6755 | 570.0114 | 24556124.9974 | 2252.5111 | 4865.8286 | 91.3336 | 188.3346 | 2922.6323 | 16578.2819 | 4741571.1142 | 76596.1292 | 986871.0723 | 453.6739 | 1197.0502 | 34196.8812 | 267522.5386 | 360879.3803 | 24128.3963 | 426.3397 | 33406.3523 | 33406.3523 | 2495.5516 | 549.3439 | 1766.1092 | 3793.1709 | 1960.0831 | 648.0147 | 25926.1754 | 22210.5873 | 106564.0197 | 157611.3892 | 9741.6538 | 174871.8643 | 140233.603 | 144563.1083 | 2915.631 | 73352.5303 | 166934.5344 | 4622.9569 | 10437.4957
70 | MT 17 | R | 06/02/2023 22:30:20 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 17.072 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 19383.4764 | 1047.3718 | 1110.3765 | 53653680.0959 | 4132.2644 | 10670.317 | 190.3346 | 401.339 | 5985.9208 | 34213.5916 | 10094299.4024 | 168287.2583 | 2206030.6165 | 967.6994 | 2511.221 | 71203.3537 | 554868.3686 | 756214.9309 | 53867.2678 | 941.031 | 73917.4706 | 73917.4706 | 4384.6733 | 1258.7221 | 3995.8921 | 8335.0985 | 4147.2099 | 1466.742 | 53752.9383 | 48078.787 | 111905.9737 | 144462.9091 | 9068.2106 | 169229.2388 | 133963.8739 | 141494.3302 | 2590.9025 | 155114.97 | 356887.0773 | 8613.0052 | 20477.04
71 | MT 18 | R | 06/02/2023 22:36:32 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 18.073 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 16849.5984 | 883.3607 | 969.3662 | 46427339.3155 | 3627.1273 | 9261.6685 | 167.001 | 354.3377 | 4577.0666 | 27145.1043 | 8805742.5833 | 139708.6038 | 1743368.1506 | 832.3576 | 2227.8404 | 59934.141 | 482886.1121 | 648974.7556 | 46633.0738 | 786.355 | 61897.8083 | 61897.8083 | 3829.8481 | 1038.0378 | 3311.0505 | 6974.3691 | 3577.6663 | 1267.3896 | 47030.6281 | 41083.6616 | 110759.7225 | 151578.6203 | 8981.1566 | 167778.1924 | 132772.1802 | 134960.9842 | 2699.5884 | 133208.1977 | 303086.9316 | 6820.6402 | 15887.444
72 | MT 19 | R | 06/02/2023 22:42:43 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 19.074 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 8464.8407 | 472.3411 | 517.6761 | 23394074.3245 | 2080.4848 | 4605.4089 | 85.3336 | 174.3344 | 2235.5107 | 12601.5674 | 4418387.2418 | 69623.9522 | 871515.4105 | 433.6733 | 1745.7733 | 35036.9158 | 280233.9576 | 327205.785 | 21698.3476 | 392.3387 | 30808.5234 | 30808.5234 | 2358.8617 | 484.0082 | 1637.4272 | 3435.747 | 1761.4581 | 598.0125 | 22501.7232 | 19838.4362 | 114056.8542 | 150063.4284 | 9047.1989 | 168498.4729 | 137121.638 | 135145.5984 | 2828.2799 | 65917.0956 | 150208.7847 | 4464.9443 | 10332.8056
73 | MT 20 | R | 06/02/2023 22:48:45 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 20.075 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 17715.3138 | 896.3615 | 987.0342 | 47844261.4059 | 3689.143 | 9428.1109 | 161.3342 | 334.3372 | 4524.3848 | 26976.1309 | 8873001.5934 | 143699.1145 | 1757315.1692 | 802.0226 | 1958.801 | 61213.2235 | 493586.5023 | 662839.0803 | 46048.2761 | 799.3557 | 62958.4268 | 62958.4268 | 3875.1927 | 1056.0391 | 3448.4162 | 7069.7504 | 3608.3203 | 1264.0559 | 45369.2679 | 40187.1163 | 112145.1364 | 150600.4209 | 8730.6674 | 163576.5208 | 132230.5354 | 132943.8871 | 2700.2554 | 135449.7914 | 307667.9178 | 7024.6748 | 16294.9231
74 | MT 21 | R | 06/02/2023 22:54:57 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 21.076 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 9357.7311 | 541.0103 | 552.6774 | 26353908.3218 | 2132.8259 | 5030.8863 | 95.0003 | 175.6677 | 2734.9285 | 16482.86 | 4912775.3024 | 78775.2671 | 957056.6835 | 443.3402 | 1232.3865 | 32663.9711 | 266835.0167 | 372769.5875 | 23743.2822 | 412.3393 | 33227.9351 | 33227.9351 | 2496.2181 | 561.3444 | 1764.109 | 3697.1454 | 1989.5322 | 625.347 | 24516.3619 | 21680.4396 | 117200.7257 | 154500.8557 | 8900.105 | 166878.5796 | 140809.8097 | 125937.3574 | 2838.9491 | 72253.5996 | 167193.5756 | 5140.3636 | 11729.8841
75 | MT 21 LD | R | 06/02/2023 23:00:59 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 21 LD.077 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 9.6667 | 0 | 0.33333 | 2743.598 | 45.3334 | 10.3333 | 0 | 0.66667 | 11.3333 | 27.3334 | 998.853 | 15.3333 | 1094.4006 | 1.3333 | 6.3333 | 46.6668 | 90.0003 | 349.3376 | 438.4111 | 0.66667 | 12 | 12 | 148.6674 | 0.66667 | 3.3333 | 4.6667 | 4.7398 | 0 | 196.668 | 8.6667 | 133608.6325 | 175751.2744 | 10064.2137 | 204223.7019 | 159276.7328 | 158094.7596 | 3002.3155 | 4 | 15.3333 | 9.3333 | 23.3333
76 | LBLANK | R | 06/02/2023 23:07:00 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LBLANK.078 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 3.3333 | 0 | 4.6667 | 2648.9125 | 542.677 | 13.3333 | 0.33333 | 1.3333 | 9.3333 | 42.6667 | 1050.2981 | 21.3333 | 1235.7201 | 1 | 9.6667 | 13.6667 | 104.6671 | 443.3402 | -2501.4321 | 0 | 16 | 16 | 875.3602 | 1.3333 | 2 | 6 | 17.2475 | 0 | 44.0001 | 10.6667 | 113078.4538 | 155329.5126 | 9327.3939 | 171589.7764 | 137289.9279 | 136539.5429 | 2794.2733 | 3.3333 | 14.6667 | 19.3333 | 50
77 | CCV 1ppb TRACE | R | 06/02/2023 23:13:03 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.079 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1823.4497 | 712.3511 | 41.3334 | 5020.2156 | 784.0215 | 784.6882 | 259.669 | 950.3649 | 1218.0519 | 7449.9421 | 5875.6596 | 12863.7899 | 2246.1766 | 2409.5366 | 652.3482 | 1686.0995 | 13385.602 | 6656.8847 | 2397.7188 | 76.0002 | 5588.4263 | 5588.4263 | 1306.0597 | 120.0005 | 412.6726 | 818.0234 | 479.0248 | 153.3342 | 11981.6898 | 10487.1814 | 114126.3262 | 148660.8798 | 8806.7137 | 173714.0749 | 135349.0451 | 125986.4378 | 2653.5802 | 5806.5132 | 13831.3623 | 19285.7751 | 47259.993
78 | CCV 1ppb TRACE | R | 06/02/2023 23:19:35 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.080 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1841.1186 | 721.0182 | 50.6668 | 4821.4839 | 841.3581 | 813.3565 | 247.6688 | 929.6969 | 1268.723 | 7750.1019 | 6065.9722 | 12688.9665 | 2356.5281 | 2539.5591 | 671.6825 | 1693.7672 | 13964.8224 | 6836.9691 | 1920.9486 | 84.3336 | 5595.0955 | 5595.0955 | 1338.7294 | 128.6673 | 408.0059 | 808.6896 | 434.2644 | 149.3341 | 12208.5482 | 10493.8559 | 116004.5935 | 147128.4808 | 8828.0607 | 166992.3963 | 137485.5039 | 130908.414 | 2659.5809 | 5969.914 | 14014.875 | 19676.0198 | 47179.4053
79 | CCV 10 ppb MAJORS | R | 06/02/2023 23:25:38 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.081 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 17604.5074 | 6566.5092 | 351.3377 | 29595.297 | 2626.9083 | 7500.6354 | 2474.8811 | 8537.8842 | 11668.7639 | 70210.1236 | 49041.3942 | 119324.3741 | 11233.0814 | 23169.1068 | 6299.7221 | 14817.0139 | 121768.1562 | 59712.572 | 42706.4613 | 774.6877 | 52731.8291 | 52731.8291 | 5654.453 | 1169.3812 | 3775.8324 | 7603.3562 | 4194.0834 | 1338.0627 | 113231.6525 | 99485.921 | 111569.9995 | 139171.9787 | 8299.0766 | 157317.5594 | 124034.8126 | 124546.6164 | 2550.8945 | 54365.2716 | 129321.4567 | 167382.4962 | 409135.2208
80 | CCV 10 ppb MAJORS | R | 06/02/2023 23:32:09 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.082 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 17906.5514 | 6731.5862 | 378.3384 | 29598.64 | 2675.9173 | 7511.9755 | 2456.2113 | 8768.6905 | 11830.5639 | 71246.5794 | 48959.5037 | 121902.8297 | 11429.9045 | 23715.3366 | 6365.0846 | 15268.8229 | 122540.6827 | 59831.031 | 43479.3152 | 785.0216 | 53141.997 | 53141.997 | 5739.8198 | 1141.3791 | 3635.796 | 7782.1192 | 4428.0079 | 1394.0682 | 115790.0402 | 99806.7855 | 112880.8874 | 139179.4175 | 8439.1594 | 156632.6845 | 127820.0602 | 129222.6311 | 2542.2262 | 54492.4057 | 131012.0249 | 169528.8008 | 409001.743
81 | MT 22 | R | 06/02/2023 23:38:12 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 22.083 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 6667.8892 | 398.6723 | 421.0062 | 17956913.808 | 1751.7741 | 3581.449 | 62.3335 | 150.3341 | 1547.7506 | 9491.8255 | 3425176.5055 | 52953.992 | 650822.1233 | 306.67 | 742.686 | 22463.3179 | 181977.1346 | 243181.2385 | 15267.1577 | 277.0027 | 21912.1258 | 21912.1258 | 1963.4687 | 371.3382 | 1201.3838 | 2472.2139 | 1315.3594 | 420.6729 | 16562.934 | 13688.5551 | 115211.3981 | 148398.208 | 8894.4354 | 167616.2884 | 135681.6873 | 131491.9249 | 2607.5715 | 48347.0116 | 109544.3974 | 4312.9236 | 9922.677
82 | MT 22 LSP | R | 06/02/2023 23:44:15 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 22 LSP.084 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 48816.6097 | 16174.819 | 598.3459 | 2550.8945 | 6418.1082 | 21376.9872 | 6070.9572 | 5590.0936 | 9039.1989 | 53121.9187 | 3285512.8361 | 122757.8538 | 594643.4521 | 14343.5327 | 4420.6841 | 28777.9569 | 235133.2746 | 262135.9483 | 42302.3407 | 703.684 | 52649.5302 | 52649.5302 | 4824.149 | 1073.3738 | 3451.7505 | 7068.4153 | 3896.9101 | 1229.3862 | 86535.9748 | 73427.5664 | 105105.2427 | 132192.1374 | 7873.1691 | 152325.9554 | 120102.868 | 52527.0623 | 2404.8691 | 77004.3204 | 180051.5757 | 104928.9182 | 253753.2838
83 | MT 22 LSP | R | 06/02/2023 23:50:45 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 22 LSP.085 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 49172.8382 | 15860.8037 | 550.0106 | 18783004.7819 | 6345.7441 | 21646.0583 | 5988.9224 | 5441.3698 | 8671.9669 | 53643.2003 | 3301531.1897 | 125438.3137 | 589551.2018 | 14268.4602 | 4372.336 | 28762.9366 | 235660.0776 | 262906.5456 | 42207.8761 | 716.3513 | 53252.404 | 53252.404 | 4815.4783 | 1115.3769 | 3481.0909 | 7173.134 | 3845.2837 | 1238.7204 | 87470.3528 | 74560.7437 | 104725.8797 | 134185.2674 | 8820.4072 | 153323.1957 | 123120.9683 | 121277.3737 | 2414.8708 | 78373.0957 | 180962.9831 | 106059.0367 | 251667.148
84 | MT 23 | R | 06/02/2023 23:56:47 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 23.086 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 3851.8527 | 219.0017 | 269.0025 | 10703436.254 | 1289.7249 | 2128.1585 | 40.0001 | 74.6669 | 1092.3751 | 6270.0428 | 1838621.2647 | 32131.4399 | 403789.5334 | 195.0013 | 481.0081 | 13276.5012 | 108012.1658 | 149612.7739 | 8269.3573 | 169.6677 | 13314.2021 | 13314.2021 | 1589.4217 | 238.6687 | 705.3508 | 1444.073 | 822.728 | 276.0027 | 10874.805 | 9148.9291 | 111197.7891 | 153920.4423 | 8925.1205 | 167676.4403 | 151266.3111 | 134601.3612 | 2847.6179 | 29013.4333 | 65919.1453 | 2399.4129 | 5293.7201
85 | MT 24 | R | 06/03/2023 00:02:49 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 24.087 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 3709.4823 | 195.0013 | 228.3352 | 9801350.4838 | 1220.3856 | 1956.8008 | 32.6667 | 73.6669 | 1070.7068 | 6505.4834 | 1682882.0344 | 29385.5265 | 363878.9041 | 177.3344 | 470.6745 | 12702.9798 | 103546.5852 | 139363.8414 | 7752.8811 | 151.3342 | 12357.3425 | 12357.3425 | 1464.0752 | 236.002 | 679.3495 | 1322.7279 | 776.0585 | 256.0023 | 9489.8177 | 8147.6566 | 115136.194 | 150083.787 | 9041.8609 | 171106.6127 | 140053.8466 | 141044.2998 | 2728.9273 | 27288.0413 | 61936.6452 | 1828.0457 | 4323.5869
86 | MT 25 | R | 06/03/2023 00:08:52 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 25.088 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 4280.3078 | 242.3354 | 291.003 | 11559207.4579 | 1264.7227 | 2316.8545 | 49.6668 | 88.0003 | 771.6875 | 4969.5313 | 1911814.8256 | 33646.2491 | 409196.1067 | 197.668 | 350.0043 | 14549.4066 | 121715.6741 | 160958.516 | 8870.8722 | 175.3344 | 14212.3997 | 14212.3997 | 1650.0953 | 260.6691 | 718.6848 | 1614.7581 | 856.5798 | 256.669 | 9861.4039 | 8485.8534 | 106805.7954 | 148912.8855 | 8928.1234 | 168182.9478 | 137312.21 | 133182.643 | 2673.5836 | 30813.1963 | 70271.7475 | 2404.7452 | 5649.7713
87 | MT 26 | R | 06/03/2023 00:14:55 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 26.089 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 3974.553 | 218.0017 | 312.6701 | 10663398.1757 | 1256.722 | 2063.149 | 35 | 81.3336 | 1173.0485 | 6858.3126 | 1760888.2929 | 31058.3936 | 378209.7073 | 182.0012 | 503.3422 | 13475.6864 | 110748.4338 | 145891.177 | 8281.7255 | 167.001 | 12988.5699 | 12988.5699 | 1542.0834 | 223.3351 | 668.6824 | 1434.7387 | 821.3947 | 246.6688 | 9956.8032 | 8655.9552 | 96650.1443 | 148183.4215 | 8965.1459 | 166650.472 | 161035.9928 | 132369.2567 | 2612.2388 | 28066.893 | 64662.019 | 2834.1107 | 5839.1296
88 | MT 27 | R | 06/03/2023 00:21:03 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 27.090 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 10852.121 | 646.6813 | 615.68 | 28812497.3884 | 2414.5374 | 5611.7692 | 119.0005 | 201.3348 | 2579.233 | 14925.129 | 5334125.3683 | 85019.6038 | 988286.9466 | 465.3409 | 1156.3802 | 33616.5065 | 276690.4905 | 374403.5394 | 25021.34 | 419.6728 | 34634.6006 | 34634.6006 | 2536.2255 | 609.3463 | 1778.7776 | 3843.1838 | 2058.0335 | 664.0155 | 25598.9166 | 22118.4431 | 115912.501 | 148144.1742 | 8726.9986 | 166952.6005 | 136788.5712 | 132802.4223 | 2651.5798 | 73924.7993 | 169074.7247 | 4426.9363 | 10332.7915
89 | MT 28 | R | 06/03/2023 00:27:06 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 28.091 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 11621.3922 | 614.0132 | 640.3477 | 31050285.523 | 2543.2265 | 5944.5701 | 107.0004 | 206.0015 | 2346.8595 | 14843.0413 | 5715719.2936 | 87983.6105 | 1062004.4809 | 492.6752 | 1115.7102 | 36684.0407 | 296937.7428 | 406670.5171 | 26893.6294 | 484.0082 | 36840.8037 | 36840.8037 | 2705.5897 | 636.681 | 1992.1389 | 4173.9437 | 2187.4294 | 746.6862 | 26899.9703 | 22810.8708 | 105673.9642 | 150187.1507 | 8372.7894 | 163741.2064 | 133237.8398 | 131865.2749 | 2632.2459 | 79537.5023 | 184674.7121 | 4286.9219 | 9805.9724
90 | MT 29 | R | 06/03/2023 00:33:08 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 29.092 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 21703.8169 | 1183.049 | 1144.0458 | 59285523.4678 | 4157.9384 | 11409.2215 | 212.6683 | 387.672 | 4921.181 | 30777.1296 | 10762392.8061 | 170125.0746 | 1941326.5243 | 905.362 | 2327.1896 | 66936.1206 | 551721.9471 | 754964.4406 | 54146.5652 | 890.6944 | 70994.6573 | 70994.6573 | 4353.3301 | 1268.7231 | 3783.8348 | 8056.9382 | 4252.7104 | 1432.7385 | 52246.0308 | 46309.6103 | 107013.376 | 144635.2106 | 8323.7583 | 157815.3436 | 141360.3244 | 135902.1665 | 2519.5559 | 150501.313 | 344182.4411 | 7517.4457 | 17037.3849
91 | LBLANK | R | 06/03/2023 00:39:20 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LBLANK.093 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 3 | 0 | 3.6667 | 4920.8478 | 588.0121 | 14.6667 | 0.66667 | 2 | 7.3333 | 34.6667 | 1131.5519 | 28 | 1242.7207 | 2 | 8.6667 | 14.3333 | 102.6671 | 500.6755 | -2609.0949 | 1.3333 | 13.3333 | 13.3333 | 967.3661 | 0 | 4.6667 | 3.3333 | 29.3883 | 0.66667 | 116.6671 | 11.3333 | 113464.8341 | 147087.3837 | 8695.9794 | 167910.1203 | 142143.6684 | 134062.1236 | 2688.2529 | 8 | 16.6667 | 23.3333 | 51.3334
92 | CCV 1ppb TRACE | R | 06/03/2023 00:45:23 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.094 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1905.4606 | 732.6855 | 47.6668 | 5176.2712 | 781.3547 | 806.3561 | 257.0023 | 918.0295 | 1218.7187 | 7662.0554 | 6202.4602 | 13154.7222 | 2217.1721 | 2409.5368 | 647.0147 | 1681.0989 | 13824.6895 | 6741.5906 | 2032.5806 | 80.3336 | 5741.1537 | 5741.1537 | 1424.7378 | 136.6673 | 394.0054 | 808.0229 | 433.6785 | 142.6674 | 12070.4315 | 10768.0568 | 113894.9341 | 141164.5371 | 8679.6363 | 161936.3549 | 133478.7491 | 129732.4821 | 2590.2364 | 5807.1801 | 13876.7377 | 18480.6983 | 46015.1128
93 | CCV 1ppb TRACE | R | 06/03/2023 00:51:55 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.095 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1979.1371 | 750.0197 | 50.3334 | 3951.2132 | 834.6911 | 822.0237 | 261.0024 | 914.696 | 1254.0552 | 7553.9968 | 6029.0429 | 12987.2365 | 2303.5191 | 2477.5482 | 680.6829 | 1675.4317 | 13746.6121 | 6711.5771 | 1934.0656 | 74.3335 | 5516.3986 | 5516.3986 | 1380.7336 | 120.6672 | 389.3387 | 824.6905 | 424.9575 | 121.3339 | 11798.8715 | 10377.7683 | 113278.8358 | 144330.1355 | 8624.6029 | 167898.3355 | 132823.9645 | 130807.533 | 2578.8996 | 5747.1559 | 13736.6015 | 18535.3915 | 45025.9875
94 | LFB 10ppb | R | 06/03/2023 00:57:57 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LFB 10ppb.096 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 17566.7956 | 6662.8869 | 361.0046 | 33062.2189 | 2568.2314 | 7292.1941 | 2447.543 | 8449.1646 | 11644.4107 | 68389.3754 | 50957.6356 | 115195.346 | 11588.3679 | 22707.0321 | 6005.596 | 14596.7871 | 117324.495 | 76206.0562 | 42183.6537 | 749.353 | 51499.394 | 51499.394 | 5871.8738 | 1139.3788 | 3631.7952 | 7531.319 | 4302.5377 | 1393.4013 | 108851.8094 | 95608.3682 | 111714.6027 | 145491.3593 | 7693.7426 | 151509.1968 | 121769.4629 | 121544.9096 | 2584.9032 | 52236.6655 | 122850.6508 | 157441.1499 | 381112.2958
95 | LFB 10ppb | R | 06/03/2023 01:04:27 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LFB 10ppb.097 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 17936.2541 | 6658.8854 | 349.671 | 32949.2894 | 2613.239 | 7629.3704 | 2480.8822 | 8440.16 | 11523.9808 | 69335.8575 | 50538.1516 | 116589.2113 | 11489.6187 | 22738.0821 | 6096.9676 | 14666.5258 | 118431.5577 | 77885.0695 | 42480.7063 | 724.6851 | 51935.5792 | 51935.5792 | 5739.8198 | 1126.7112 | 3604.4551 | 7532.6522 | 4227.7685 | 1340.7297 | 110651.5528 | 96151.1789 | 110317.0089 | 137368.6584 | 8059.6074 | 157916.7456 | 122369.9012 | 124286.3629 | 2338.1914 | 52155.0663 | 123843.8564 | 158224.2304 | 379855.1462
96 | CCV 10 ppb MAJORS | R | 06/03/2023 01:10:29 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.098 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 18283.0281 | 6943.6878 | 337.004 | 31978.4439 | 2618.5734 | 7572.6751 | 2500.2189 | 8743.0086 | 11917.6358 | 72222.783 | 49645.9894 | 120711.1869 | 10931.1808 | 22852.2643 | 6195.3434 | 14712.5739 | 122738.4458 | 60506.5671 | 42486.1708 | 733.6855 | 52482.2316 | 52482.2316 | 5791.8409 | 1221.3856 | 3630.4616 | 7505.9716 | 4192.5668 | 1395.4015 | 113979.5796 | 98482.3238 | 112283.5417 | 134529.8103 | 8043.5976 | 152503.2308 | 123634.0608 | 124774.0374 | 2454.2113 | 53862.0229 | 127503.1956 | 162552.9089 | 394274.1336
97 | CCV 10 ppb MAJORS | R | 06/03/2023 01:17:01 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.099 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 18215.9406 | 6791.6151 | 345.3375 | 31021.6499 | 2638.5771 | 7743.0982 | 2458.2115 | 8762.3547 | 11784.8603 | 72131.6784 | 49587.5709 | 121719.6858 | 11099.311 | 23169.4406 | 6304.0573 | 15246.465 | 121778.1834 | 59337.6586 | 42895.035 | 743.3527 | 53449.142 | 53449.142 | 5963.2447 | 1170.7147 | 3627.7942 | 7574.0078 | 4081.273 | 1344.73 | 111741.3805 | 98774.9749 | 105776.0356 | 138432.1965 | 8312.7513 | 156152.9508 | 123959.4911 | 124070.475 | 2394.8676 | 53465.8655 | 127361.9861 | 164235.8784 | 394424.8262
98 | MT 30 | R | 06/03/2023 01:23:04 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 30.100 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 33 | 98.3337 | 14309.8312 | 4904658.0185 | 919.3629 | 92.667 | 326.0037 | 609.3463 | 575.3449 | 3365.0633 | 2516941.9028 | 1067040.7104 | 448231.1652 | 1025.3702 | 491.6751 | 94631.1805 | 804104.9071 | 723285.7901 | 3735.7777 | 104.3337 | 7543.9924 | 7543.9924 | 1025.3702 | 43.3334 | 136.0006 | 326.0038 | 234.0256 | 57.3335 | 3671.8055 | 3017.6525 | 119644.9382 | 148429.9092 | 8605.9254 | 151186.7287 | 131847.6435 | 136521.7218 | 2630.9091 | 9748.6601 | 22212.589 | 56489.8092 | 134723.6058
99 | MT 31 | R | 06/03/2023 01:29:15 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 31.101 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 187.6679 | 173.001 | 10869.4681 | 4604398.7203 | 791.3553 | 420.3395 | 184.6679 | 445.3403 | 441.3402 | 2642.2444 | 1578898.5221 | 787129.8597 | 335803.5701 | 1022.7033 | 320.6703 | 70481.1185 | 573832.9581 | 839633.5779 | 3055.8604 | 90.0003 | 6778.2746 | 6778.2746 | 992.0346 | 26 | 122.0005 | 224.0018 | 175.3091 | 33.3334 | 2152.8289 | 1954.8007 | 116814.3859 | 151600.5842 | 8676.6345 | 169458.9726 | 132909.5109 | 129850.9095 | 2691.5874 | 10685.9955 | 24438.2191 | 40474.7897 | 95883.3122
100 | MT 31 LD | R | 06/03/2023 01:35:29 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 31 LD.102 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 192.668 | 159.6676 | 10974.2141 | 4722639.8671 | 805.6895 | 428.0064 | 186.3346 | 462.0075 | 427.3398 | 2652.2462 | 1552886.479 | 768445.8899 | 333413.7981 | 993.0345 | 311.6701 | 70102.5981 | 557589.6187 | 839503.1197 | 3152.3793 | 84.3336 | 6652.2152 | 6652.2152 | 991.3677 | 28 | 100.0004 | 219.335 | 190.6319 | 44.0001 | 2140.8271 | 1920.7959 | 115322.3341 | 144721.3635 | 8467.5111 | 161432.3987 | 131203.7645 | 129686.0012 | 2583.567 | 10459.8284 | 24092.2984 | 40071.3604 | 94435.5275
101 | MT 31 LD | R | 06/03/2023 01:42:00 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 31 LD.103 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 198.3347 | 155.0008 | 10935.5183 | 4637496.5145 | 806.0228 | 408.3392 | 181.0012 | 444.6736 | 459.0074 | 2634.9097 | 1572928.8965 | 772233.4134 | 330626.2507 | 1046.705 | 303.3366 | 68848.5566 | 566054.2834 | 841054.2453 | 3242.9613 | 75.0002 | 6652.2166 | 6652.2166 | 979.367 | 27.3334 | 110.0004 | 216.0016 | 162.6854 | 32 | 2132.8259 | 1819.4492 | 116583.199 | 145060.1186 | 8706.3217 | 164067.4556 | 122121.7896 | 127676.0124 | 2600.9037 | 10595.9284 | 24217.8443 | 39698.4401 | 93701.1729
102 | MT 32 | R | 06/03/2023 01:48:02 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 32.104 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 3.6667 | 0.66667 | 45.6667 | 25435.2915 | 588.0121 | 17.6667 | 1.6667 | 0.33333 | 20 | 77.3335 | 1279.2227 | 76.0002 | 1250.7215 | 2 | 14 | 77.0002 | 640.681 | 662.682 | -2472.174 | 0 | 17.3333 | 17.3333 | 866.693 | 1.3333 | 1.3333 | 2.6667 | 37.383 | 0.66667 | 161.3342 | 152.6675 | 111292.5176 | 144762.6031 | 8541.5531 | 164784.467 | 124049.1002 | 123670.4919 | 2557.5623 | 2 | 20 | 712.0069 | 1497.3651
103 | MT 32 LSP | R | 06/03/2023 01:54:03 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 32 LSP.105 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 45165.3041 | 16112.7484 | 290.003 | 44882.4026 | 5448.7057 | 18757.6406 | 6066.955 | 5439.3688 | 7158.1268 | 44245.7487 | 36758.3489 | 77240.2543 | 8434.8227 | 13848.3765 | 5365.0076 | 8816.3864 | 72220.1057 | 135043.959 | 24745.8891 | 444.6736 | 31582.2081 | 31582.2081 | 3725.1524 | 672.0159 | 2211.5047 | 4699.4397 | 2616.7579 | 816.69 | 66505.1116 | 58566.4738 | 100484.1853 | 126873.6279 | 7525.6485 | 141266.3904 | 113790.8631 | 106976.9713 | 2278.1817 | 30922.0968 | 73124.0071 | 94395.2874 | 225824.4544
104 | MT 32 LSP | R | 06/03/2023 02:00:34 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 32 LSP.106 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 44451.7464 | 15898.8427 | 302.6699 | 40251.2986 | 5385.3483 | 18242.9793 | 5915.5579 | 5381.0134 | 7177.8029 | 44344.7305 | 36829.3186 | 75435.4104 | 8514.5367 | 13997.856 | 5416.36 | 9068.8781 | 73585.7893 | 137979.8675 | 24789.3148 | 466.6743 | 31822.0728 | 31822.0728 | 3724.4859 | 714.0179 | 2231.5076 | 4588.0704 | 2515.9784 | 804.0226 | 67892.3423 | 58077.8214 | 105605.9135 | 128938.7378 | 7598.6873 | 143298.1828 | 115624.0624 | 114485.6788 | 2227.5072 | 31345.0195 | 74737.0314 | 96363.5041 | 226590.4549
105 | MT 33 | R | 06/03/2023 02:06:37 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 33.107 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 5.6667 | 0.66667 | 58.0001 | 23806.4913 | 638.3476 | 12.3333 | 2.6667 | 1 | 19.6667 | 86.6669 | 1323.9125 | 68.0002 | 1320.0611 | 6 | 17 | 81.0002 | 766.6873 | 5512.4357 | -2635.4749 | 0 | 16.6667 | 16.6667 | 964.0325 | 0 | 2.6667 | 2.6667 | 23.2584 | 1.3333 | 170.6677 | 31.3334 | 115843.1289 | 148915.4702 | 8934.7936 | 165698.852 | 115217.5325 | 131456.7634 | 2730.9277 | 3.3333 | 20.6667 | 574.0047 | 1250.0215
106 | MT 34 | R | 06/03/2023 02:12:39 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 34.108 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 6 | 0 | 63.6668 | 26628.7965 | 656.6818 | 12 | 1.6667 | 1 | 29.3334 | 158.6676 | 16170.03 | 227.3351 | 4421.351 | 3 | 32.6667 | 35.6667 | 264.0024 | 609.3464 | -2709.0354 | 0.33333 | 15.3333 | 15.3333 | 950.6983 | 0.66667 | 4 | 2.6667 | 37.9143 | 0.66667 | 64.6668 | 17.3333 | 112212.5078 | 148035.0531 | 8771.0261 | 168553.2336 | 135223.2567 | 132593.2492 | 2617.5732 | 2.6667 | 22.6667 | 486.6699 | 1216.6868
107 | MT 35 | R | 06/03/2023 02:18:42 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 35.109 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.6667 | 0.33333 | 50.0001 | 26025.018 | 675.3493 | 12.3333 | 2.6667 | 0.66667 | 20.3333 | 82.6669 | 3360.0479 | 63.3335 | 1704.1016 | 3 | 14.3333 | 25.3334 | 248.0022 | 749.3532 | -2854.6778 | 0.33333 | 18 | 18 | 942.0311 | 2 | 2 | 4 | 36.5918 | 0 | 52.0001 | 16.6667 | 111357.9354 | 142431.9955 | 8819.3885 | 173481.1996 | 135276.9941 | 130614.6629 | 2499.5523 | 5.3333 | 19.3333 | 532.0041 | 1175.3526
108 | LBLANK | R | 06/03/2023 02:24:45 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LBLANK.110 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.3333 | 0 | 5.3333 | 4086.5852 | 674.3493 | 16.6667 | 0.66667 | 1 | 10.3333 | 22.6667 | 943.4653 | 24 | 1225.7193 | 1 | 13.6667 | 13.3333 | 78.6669 | 450.0071 | -2547.5539 | 0 | 10 | 10 | 938.0308 | 0 | 2.6667 | 0.66667 | 30.6621 | 0 | 44.0001 | 7.3333 | 114932.4903 | 145070.9449 | 8665.6276 | 162453.8264 | 124689.2028 | 130952.9097 | 2567.5643 | 0.66667 | 12.6667 | 22 | 44.6667
109 | CCV 1ppb TRACE | R | 06/03/2023 02:30:48 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.111 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1949.7997 | 722.3516 | 46.0001 | 4576.7332 | 854.0255 | 844.3583 | 270.3359 | 950.6983 | 1259.0555 | 7535.3222 | 6140.5763 | 13421.6366 | 2346.8595 | 2453.544 | 688.6833 | 1765.4425 | 13535.7494 | 6788.9471 | 1967.6202 | 81.3336 | 5655.786 | 5655.786 | 1296.0588 | 134.0006 | 360.6712 | 790.6886 | 446.3188 | 126.6672 | 11521.9791 | 10431.1409 | 113112.0466 | 141159.0472 | 8620.9347 | 156769.8035 | 135380.5013 | 127717.8207 | 2572.8984 | 5545.7434 | 13326.2157 | 18137.2011 | 41594.3077
110 | CCV 1ppb TRACE | R | 06/03/2023 02:37:20 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.112 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2001.807 | 749.353 | 53.3334 | 4093.2534 | 892.6946 | 842.3582 | 269.0025 | 980.7004 | 1306.0597 | 7581.345 | 6006.9421 | 13198.0951 | 2327.8564 | 2511.2207 | 691.6834 | 1745.44 | 13484.3612 | 6583.5169 | 2355.2671 | 86.6669 | 5413.6923 | 5413.6923 | 1378.7332 | 128.6672 | 388.672 | 796.0222 | 493.0146 | 150.6675 | 11282.4544 | 10060.2081 | 113799.0502 | 142998.1984 | 8411.1436 | 154916.2631 | 127859.4334 | 130729.0522 | 2542.2262 | 5523.0678 | 13252.8113 | 18076.5163 | 42649.7334
111 | CCV 10 ppb MAJORS | R | 06/03/2023 02:43:23 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.113 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 18802.3691 | 7026.061 | 359.3379 | 31495.3482 | 2724.5931 | 7914.8596 | 2647.2453 | 8886.7635 | 12016.0516 | 70701.8689 | 47824.0363 | 122056.6422 | 11174.3705 | 23687.6243 | 6345.7425 | 15170.0517 | 118541.1298 | 57520.2443 | 42857.4787 | 747.6863 | 51954.314 | 51954.314 | 5983.2534 | 1164.0474 | 3561.1105 | 7508.6403 | 4112.3014 | 1357.3978 | 108437.338 | 97268.6867 | 112414.6413 | 137071.5047 | 7791.7923 | 149687.471 | 119628.1918 | 122170.862 | 2497.552 | 50738.6189 | 120691.0252 | 156009.3163 | 379692.5481
112 | CCV 10 ppb MAJORS | R | 06/03/2023 02:49:56 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.114 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 19047.023 | 7230.4967 | 365.0047 | 31345.6872 | 2684.5856 | 7862.4968 | 2591.2351 | 9002.8364 | 12152.1668 | 72630.1761 | 49208.9289 | 123144.5859 | 11393.8757 | 24084.2852 | 6381.7585 | 15327.8856 | 121965.8109 | 58746.5532 | 42362.681 | 743.3527 | 53156.7164 | 53156.7164 | 5705.8062 | 1178.7153 | 3633.1287 | 7632.7064 | 4066.3626 | 1335.3959 | 111955.009 | 98583.6541 | 118179.1992 | 137302.0397 | 8287.7389 | 147867.6147 | 121829.2941 | 123835.1305 | 2432.2071 | 52720.4717 | 124757.7282 | 156773.7661 | 384318.4996
113 | MT 36 | R | 06/03/2023 02:55:59 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 36.115 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 9 | 1 | 51.3334 | 22220.6045 | 625.6804 | 13.3333 | 4.3333 | 2 | 37.3334 | 172.001 | 2432.0827 | 89.3336 | 1526.7483 | 3.6667 | 14.3333 | 36.6667 | 265.3358 | 766.6873 | -2558.1779 | 0.33333 | 12.6667 | 12.6667 | 937.3642 | 0 | 1.3333 | 0.66667 | 45.3073 | 0 | 86.0003 | 33.3334 | 116851.3061 | 144317.3146 | 8875.0917 | 160959.7375 | 132710.3417 | 132676.6704 | 2543.5598 | 2 | 22 | 520.6705 | 1204.0201
114 | MT 37 | R | 06/03/2023 03:02:01 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 37.116 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 7.3333 | 1.6667 | 49.0001 | 30195.8796 | 657.0151 | 20.6667 | 4.6667 | 0.33333 | 18.6667 | 80.0002 | 1384.9611 | 120.0005 | 1350.7305 | 1 | 19.6667 | 150.6675 | 1206.0509 | 1304.0595 | -2508.0692 | 0.33333 | 16 | 16 | 907.3622 | 0.66667 | 0.66667 | 5.3333 | 9.9358 | 0 | 35.3334 | 20 | 118805.4331 | 146892.8744 | 8379.1276 | 164762.08 | 134050.6899 | 134252.0717 | 2651.5795 | 3.3333 | 25.3334 | 550.6708 | 1300.69
115 | MT 38 | R | 06/03/2023 03:08:04 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 38.117 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 4 | 0.66667 | 42.3334 | 25981.6391 | 670.3491 | 15.6667 | 1.6667 | 1.3333 | 21.6667 | 96.667 | 1428.6325 | 62.6668 | 1330.0619 | 2 | 10.6667 | 48.0001 | 389.3387 | 684.6831 | -2587.8264 | 0.33333 | 15.3333 | 15.3333 | 964.0326 | 0 | 4.6667 | 1.3333 | 48.0606 | 0 | 37.3334 | 18 | 113606.6253 | 144764.5414 | 8798.0426 | 161297.5206 | 114891.818 | 137215.2503 | 2584.234 | 2 | 12 | 521.3371 | 1219.3534
116 | MT 39 | R | 06/03/2023 03:14:07 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 39.118 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 12.3333 | 0.66667 | 52.3334 | 26683.5727 | 632.6807 | 10.6667 | 3.3333 | 1.3333 | 17.3333 | 109.3338 | 4784.7471 | 92.667 | 2025.8103 | 2.3333 | 9.3333 | 47.3334 | 378.005 | 1068.04 | -2556.777 | 0.33333 | 12 | 12 | 879.3604 | 0.66667 | 3.3333 | 5.3333 | 22.5321 | 2 | 29.3334 | 19.3333 | 105375.9139 | 143915.977 | 8647.9515 | 159820.3609 | 128839.7196 | 128766.473 | 2602.2372 | 1.3333 | 17.3333 | 560.0045 | 1192.0195
117 | MT 40 | R | 06/03/2023 03:20:10 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 40.119 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 16.6667 | 6 | 53.3334 | 26271.4783 | 628.3472 | 16.6667 | 14.3333 | 0.66667 | 13.6667 | 62.6668 | 1610.6327 | 358.6712 | 1323.728 | 1.3333 | 12 | 71.0002 | 498.6754 | 606.0129 | -2376.6528 | 0.66667 | 62.0001 | 62.0001 | 893.3613 | 0 | 2.6667 | 3.3333 | 28.7217 | 0.66667 | 107.3337 | 98.0003 | 123590.2816 | 142050.2008 | 8844.4046 | 155237.1126 | 127982.0705 | 125369.7924 | 2598.903 | 1.3333 | 13.3333 | 499.3367 | 1149.3512
118 | MT 31 | R | 06/03/2023 03:26:12 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 31 .120 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 192.668 | 180.6678 | 11407.5541 | 4701792.7259 | 868.6931 | 407.0058 | 192.0013 | 466.341 | 429.6731 | 2370.1972 | 1496907.3947 | 758684.7375 | 335432.3902 | 1013.3694 | 308.3367 | 69961.5734 | 533780.6761 | 796713.3146 | 3113.4655 | 80.6669 | 6576.847 | 6576.847 | 982.7006 | 30 | 106.6671 | 215.335 | 161.9647 | 32 | 1942.7994 | 1791.4458 | 117943.6701 | 141716.9074 | 8287.0697 | 160814.8585 | 124708.6568 | 124744.9812 | 2495.5513 | 9946.1322 | 22690.6853 | 39024.951 | 89571.7216
119 | MT 31 LD | R | 06/03/2023 03:32:23 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 31 LD.121 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 189.3346 | 156.0009 | 11070.2915 | 4797740.839 | 803.0226 | 399.6723 | 188.0012 | 428.0064 | 413.6727 | 2565.5638 | 1498707.3072 | 771730.4494 | 330448.7851 | 993.0345 | 323.337 | 68103.3323 | 534460.5057 | 799022.6098 | 2701.8631 | 84.0002 | 6594.1888 | 6594.1888 | 958.6989 | 32 | 104.6671 | 211.3349 | 133.3138 | 46.0001 | 2061.4822 | 1837.4515 | 114630.7847 | 137246.4134 | 8555.9069 | 155203.8268 | 142754.4613 | 122987.1447 | 2470.8804 | 9873.4122 | 23032.5587 | 38777.3782 | 88267.6829
120 | MT 31 LD | R | 06/03/2023 03:38:54 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 31 LD.122 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 179.6678 | 149.3341 | 10687.6648 | 4742639.1896 | 802.3559 | 413.3393 | 182.3345 | 437.0067 | 428.0064 | 2367.5301 | 1405022.0498 | 750966.9835 | 332775.9945 | 993.7012 | 295.0031 | 68560.4818 | 501565.2226 | 753487.9015 | 2677.5816 | 84.3336 | 6323.3996 | 6323.3996 | 838.0246 | 34.6667 | 90.667 | 196.0014 | 108.69 | 40.0001 | 1894.7923 | 1765.4425 | 108281.5072 | 139303.9149 | 8315.0863 | 166720.6075 | 123065.958 | 124537.8303 | 2503.5528 | 9346.3944 | 21570.2828 | 32766.2115 | 82829.1755
121 | MT 32 | R | 06/03/2023 03:44:56 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 32.123 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 7 | 0 | 41.6667 | 24276.6223 | 513.0092 | 9 | 1.3333 | 0.33333 | 21 | 78.6669 | 1262.5757 | 76.0002 | 1197.7169 | 1.6667 | 13.6667 | 79.3336 | 656.0151 | 558.6776 | -2361.095 | 0 | 16 | 16 | 864.0261 | 0.66667 | 2 | 2.6667 | 42.8463 | 0 | 138.6674 | 118.0005 | 108716.1944 | 138074.0731 | 8192.0153 | 152216.7577 | 125208.325 | 125646.8112 | 2438.875 | 2 | 18.6667 | 666.6728 | 1435.3615
122 | MT 32 LSP | R | 06/03/2023 03:50:58 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 32 LSP.124 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 44044.1694 | 15963.5813 | 303.3366 | 40413.1093 | 5460.0439 | 18225.9535 | 6136.3176 | 5344.9998 | 7205.8171 | 42533.895 | 35470.4414 | 74283.9803 | 8348.7722 | 13633.1697 | 5374.0106 | 8818.0542 | 69948.8547 | 130802.093 | 24564.7212 | 427.3398 | 30594.7328 | 30594.7328 | 3719.8176 | 686.0166 | 2139.4935 | 4420.684 | 2610.2699 | 778.688 | 64370.0367 | 56110.6494 | 100580.1661 | 121543.6968 | 7423.2617 | 136585.8893 | 99201.9141 | 149819.6452 | 2236.1752 | 30084.9825 | 70447.9463 | 93063.9544 | 220528.466
123 | MT 32 LSP | R | 06/03/2023 03:57:29 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 32 LSP.125 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 43317.5778 | 15942.2254 | 290.0029 | 41127.1291 | 5504.7272 | 18509.653 | 5912.5567 | 5347.6675 | 7290.1936 | 41131.1318 | 34926.4445 | 74769.8314 | 8460.171 | 13851.3792 | 5379.0126 | 8959.4759 | 69114.1179 | 130099.7367 | 22757.0938 | 446.6737 | 30841.9276 | 30841.9276 | 3513.766 | 681.3496 | 2096.8207 | 4448.0261 | 2310.8314 | 777.3545 | 63448.6025 | 57189.5833 | 104905.1045 | 125993.2018 | 7581.678 | 145551.9066 | 102666.3737 | 112465.0014 | 2258.8456 | 29384.1906 | 70414.4583 | 86769.6597 | 214214.1624
124 | LBLANK | R | 06/03/2023 04:03:32 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LBLANK.126 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 4 | 0.33333 | 4.3333 | 2292.8515 | 633.3474 | 9.3333 | 0.66667 | 0.66667 | 12.6667 | 43.3334 | 896.6532 | 26.6667 | 1275.0569 | 2.3333 | 12.6667 | 12.3333 | 100.667 | 420.0062 | -2602.2297 | 0.66667 | 20.6667 | 20.6667 | 918.6962 | 0 | 1.3333 | 5.3333 | 22.5917 | 0 | 100.0004 | 16 | 108541.5402 | 145195.574 | 9040.8602 | 164160.5262 | 133585.7036 | 129373.8503 | 2626.9086 | 1.3333 | 16 | 18.6667 | 46.6667
125 | CCV 1ppb TRACE | R | 06/03/2023 04:09:35 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.127 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1983.1377 | 768.0206 | 55.3334 | 4440.6902 | 886.0275 | 813.0232 | 266.0025 | 940.3643 | 1305.7264 | 7742.7649 | 6044.1623 | 13744.6096 | 2318.5215 | 2596.2361 | 702.0173 | 1724.4375 | 13929.4564 | 6760.9334 | 2167.9855 | 82.6669 | 5671.7925 | 5671.7925 | 1416.0703 | 122.0005 | 380.6718 | 827.3573 | 456.8996 | 109.3338 | 11812.2225 | 10641.9631 | 117299.6396 | 147169.5718 | 8555.2288 | 161020.3929 | 130402.5142 | 135110.6101 | 2706.9231 | 5513.731 | 13350.9033 | 18115.1436 | 43017.5567
126 | CCV 1ppb TRACE | R | 06/03/2023 04:16:07 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.128 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1992.1389 | 746.3528 | 55.3334 | 3972.5524 | 878.3604 | 853.6922 | 269.3359 | 1006.3688 | 1354.0642 | 7675.3951 | 6064.5096 | 13216.7791 | 2380.1983 | 2555.8955 | 686.0165 | 1735.7722 | 13469.0154 | 6599.524 | 2144.5716 | 80.3336 | 5525.07 | 5525.07 | 1394.0682 | 104.6671 | 368.6714 | 816.0235 | 429.6999 | 153.3342 | 11281.1202 | 10276.3634 | 114620.884 | 142520.7534 | 8779.7153 | 151688.9276 | 123995.8343 | 127335.1107 | 2596.2359 | 5345.6674 | 13247.4734 | 17682.9966 | 41665.9308
127 | LFB 10ppb | R | 06/03/2023 04:22:09 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LFB 10ppb.129 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 17661.9121 | 6534.4946 | 354.6711 | 32321.191 | 2600.9034 | 7294.8628 | 2446.8762 | 8637.9441 | 11632.068 | 66787.7614 | 48841.73 | 114744.9979 | 11469.2698 | 22412.2366 | 6096.6346 | 14403.5911 | 113529.3572 | 74377.8086 | 42809.3401 | 707.3509 | 50181.3204 | 50181.3204 | 5702.4713 | 1128.7113 | 3432.4127 | 7293.8615 | 4186.6636 | 1270.7232 | 104445.8038 | 92368.3658 | 110426.5119 | 132264.148 | 7942.8742 | 145594.8476 | 114791.6429 | 117840.0348 | 2403.5355 | 48989.1946 | 117048.93 | 148776.8752 | 358260.9863
128 | LFB 10ppb | R | 06/03/2023 04:28:39 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LFB 10ppb.130 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 17564.7942 | 6495.1429 | 379.3384 | 34122.0359 | 2622.2407 | 7257.51 | 2415.8712 | 8667.6287 | 11489.2858 | 69112.7867 | 49919.453 | 117174.606 | 11432.573 | 22811.5321 | 6184.3389 | 14623.4822 | 114582.4016 | 74407.9524 | 42209.8127 | 756.6867 | 50975.4574 | 50975.4574 | 5774.5016 | 1053.3722 | 3504.4299 | 7497.3011 | 4176.1447 | 1290.725 | 106755.441 | 94570.0977 | 109086.3127 | 130165.8418 | 8095.9602 | 144878.5144 | 116165.7947 | 120676.2441 | 2327.523 | 49918.4026 | 117342.6734 | 150730.2111 | 356453.9195
129 | CCV 10 ppb MAJORS | R | 06/03/2023 04:34:42 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.131 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 17849.484 | 6719.9152 | 376.005 | 31792.0038 | 2690.9203 | 7562.3346 | 2457.2115 | 8854.0773 | 11622.395 | 71777.2767 | 48034.7543 | 125351.8113 | 10984.2241 | 22852.6192 | 6248.7008 | 14803.6688 | 118631.2851 | 57512.2238 | 42100.9419 | 736.6857 | 54133.091 | 54133.091 | 5805.1794 | 1201.3839 | 3687.1426 | 7640.7097 | 4140.4641 | 1392.7346 | 109148.7256 | 100377.4837 | 110081.9172 | 102892.7334 | 8550.9311 | 145162.4859 | 117233.8115 | 119906.4461 | 1911.463 | 50440.9006 | 120834.2573 | 151317.0224 | 368503.2375
130 | CCV 10 ppb MAJORS | R | 06/03/2023 04:41:13 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.132 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 17782.3954 | 6661.554 | 355.3378 | 30522.5741 | 2632.5759 | 7604.6903 | 2488.8837 | 8875.0902 | 11738.4885 | 71792.6177 | 49828.8212 | 119106.574 | 11143.6787 | 23178.791 | 6279.38 | 14971.8415 | 120519.0024 | 57258.5382 | 43406.4482 | 743.3527 | 50935.9941 | 50935.9941 | 5885.2123 | 1163.3807 | 3419.0761 | 7453.2774 | 4187.4605 | 1306.0597 | 110490.3184 | 95446.4991 | 112791.543 | 132452.5288 | 8118.3064 | 146822.2091 | 113600.7444 | 118709.9045 | 2410.2033 | 50853.0345 | 122634.1487 | 151158.1461 | 374463.8748
131 | MT 33 | R | 06/03/2023 04:47:16 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 33.133 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 4.3333 | 0.33333 | 54.6668 | 22067.7015 | 653.6816 | 11.3333 | 2.3333 | 0.33333 | 19.6667 | 81.3336 | 1201.592 | 60.0001 | 1325.7283 | 5.6667 | 21.3334 | 89.6669 | 767.3541 | 469.341 | -2718.2761 | 0.66667 | 12.6667 | 12.6667 | 956.6988 | 0 | 4 | 2.6667 | 39.4588 | 0 | 161.3342 | 42.0001 | 115843.1871 | 145028.4809 | 8904.1077 | 159531.1079 | 133381.1205 | 127813.2727 | 2576.8991 | 2 | 19.3333 | 550.0043 | 1219.3538
132 | MT 34 | R | 06/03/2023 04:53:19 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 34.134 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 8.3333 | 0.66667 | 46.0001 | 25747.1832 | 627.6805 | 13 | 2.3333 | 1.3333 | 28.6667 | 139.334 | 15126.9072 | 252.669 | 4395.0096 | 3.3333 | 25.3334 | 33.3334 | 220.6684 | 582.0119 | -2559.5549 | 0.33333 | 10.6667 | 10.6667 | 948.6983 | 1.3333 | 1.3333 | 2.6667 | 37.3234 | 0 | 66.6668 | 14.6667 | 115490.7872 | 145323.4411 | 9023.5196 | 163037.4842 | 129838.704 | 127702.9364 | 2676.2507 | 1.3333 | 22 | 502.67 | 1117.3506
133 | MT 35 | R | 06/03/2023 04:59:21 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 35.135 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.3333 | 0 | 52.0001 | 25980.2721 | 633.014 | 12.6667 | 1.6667 | 0.33333 | 21.3333 | 92.667 | 3113.3651 | 62.0001 | 1778.1107 | 1 | 14 | 34.3334 | 231.3352 | 688.6833 | -2382.0817 | 1 | 14 | 14 | 892.6946 | 0 | 1.3333 | 6 | 40.1092 | 0.66667 | 50.6668 | 16 | 113440.0482 | 140648.3455 | 8522.5418 | 150450.5642 | 126431.7139 | 127141.3377 | 2561.563 | 2 | 18 | 505.3368 | 1085.3496
134 | MT 36 | R | 06/03/2023 05:05:25 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 36.136 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 5 | 0.33333 | 46.3334 | 25355.1645 | 632.6807 | 13 | 2 | 0.33333 | 29.6667 | 164.6676 | 2394.1385 | 79.3336 | 2007.1526 | 2.3333 | 12 | 31 | 267.3358 | 697.3504 | -2250.7045 | 0.66667 | 15.3333 | 15.3333 | 860.0259 | 0.66667 | 2.6667 | 2.6667 | 20.5267 | 1.3333 | 48.0001 | 22 | 105765.0366 | 140803.8994 | 8580.9104 | 153688.7898 | 127503.1423 | 127402.9563 | 2612.9056 | 1.3333 | 15.3333 | 471.3365 | 1159.352
135 | MT 37 | R | 06/03/2023 05:11:28 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 37.137 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 5.6667 | 0 | 52.6668 | 30831.2405 | 624.347 | 15 | 4.6667 | 0.66667 | 16.6667 | 76.0002 | 1419.2412 | 116.0005 | 1336.7293 | 0.33333 | 15.3333 | 155.0008 | 1158.7137 | 1302.7261 | -2416.4633 | 0.33333 | 11.3333 | 11.3333 | 883.3607 | 2 | 0.66667 | 2 | 25.3721 | 0 | 36.6667 | 18 | 106779.4494 | 135159.9698 | 8501.5297 | 177351.56 | 127372.0955 | 125632.7287 | 2519.5557 | 0.66667 | 24 | 573.3377 | 1209.3529
136 | MT 38 | R | 06/03/2023 05:17:31 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 38.138 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.6667 | 1 | 51.6668 | 27112.3727 | 680.6829 | 13.6667 | 1 | 0.66667 | 23.6667 | 88.0003 | 1347.5735 | 65.3335 | 1376.0663 | 4 | 15 | 42.3334 | 346.0042 | 640.6811 | -2728.5474 | 0.66667 | 8 | 8 | 890.6945 | 0 | 2 | 2 | 13.3233 | 0.66667 | 32 | 17.3333 | 110181.9876 | 142945.6135 | 8703.3171 | 151220.8616 | 124320.036 | 126385.9625 | 2529.5573 | 2.6667 | 17.3333 | 484.0034 | 1088.6827
137 | MT 39 | R | 06/03/2023 05:23:33 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 39.139 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 12 | 0.66667 | 53.6668 | 26171.2899 | 637.3476 | 13.3333 | 2.3333 | 0.66667 | 21.6667 | 107.3337 | 4825.885 | 86.0003 | 2019.8095 | 3.3333 | 18.6667 | 49.3334 | 364.0047 | 1008.7023 | -2510.7419 | 0.33333 | 14 | 14 | 886.0276 | 0.66667 | 2 | 4.6667 | 62.1852 | 0 | 30.6667 | 14 | 110909.6427 | 125513.1679 | 8285.0685 | 148834.307 | 123948.0834 | 126617.5194 | 2293.5187 | 2.6667 | 14.6667 | 500.67 | 1102.683
138 | MT 40 | R | 06/03/2023 05:29:36 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 40.140 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 14 | 6.3333 | 51.3334 | 26274.8099 | 637.3476 | 12 | 11 | 0.66667 | 22.6667 | 85.3336 | 1640.3897 | 342.6708 | 1389.4009 | 2 | 11.3333 | 63.6668 | 495.342 | 582.0119 | -2450.974 | 0.66667 | 58.0001 | 58.0001 | 842.0248 | 0.66667 | 0 | 2.6667 | 28.7108 | 1.3333 | 114.0005 | 88.0003 | 109001.6313 | 141436.8137 | 8087.6293 | 151037.001 | 120807.4171 | 123997.8204 | 2527.5569 | 2 | 13.3333 | 488.0031 | 1081.3499
139 | MT 32 B | R | 06/03/2023 05:35:38 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 32 B.141 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 3.3333 | 0.33333 | 36.3334 | 23022.5383 | 652.0149 | 13.3333 | 0.33333 | 0.66667 | 22 | 90.0003 | 1323.4525 | 51.3334 | 1324.7281 | 1.6667 | 22.3334 | 75.0002 | 556.6775 | 540.0103 | -2395.3819 | 0.33333 | 14 | 14 | 868.0264 | 0 | 0 | 4 | 36.7271 | 0 | 52.6668 | 29.3334 | 109735.1973 | 134380.7531 | 8458.1717 | 147382.1668 | 124485.3752 | 125243.9642 | 2466.2131 | 2 | 19.3333 | 625.3386 | 1461.363
140 | MT 32 B LD | R | 06/03/2023 05:41:40 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 32 B LD.142 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 4.3333 | 0.33333 | 34.3334 | 23074.6207 | 630.6806 | 17.6667 | 2 | 0.33333 | 18 | 83.3336 | 1056.3411 | 35.3334 | 1207.051 | 1 | 11.3333 | 96.3337 | 654.6817 | 878.027 | -2299.1944 | 0.33333 | 12 | 12 | 842.0248 | 0 | 0.66667 | 2 | 32.0009 | 0.66667 | 35.3334 | 34.6667 | 110384.8308 | 135764.8934 | 8176.3394 | 148159.7448 | 120643.3072 | 112508.637 | 2491.5507 | 4.6667 | 16 | 734.0076 | 1547.3658
141 | LBLANK | R | 06/03/2023 05:47:43 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LBLANK.143 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.3333 | 0 | 2.6667 | 3635.1292 | 639.3477 | 11.6667 | 0.33333 | 1 | 12.3333 | 35.3334 | 866.3797 | 30 | 1193.3832 | 0.66667 | 10.6667 | 7.6667 | 77.3335 | 388.672 | -2385.3457 | 0 | 8 | 8 | 861.3594 | 1.3333 | 1.3333 | 0.66667 | 42.7434 | 0.66667 | 17.3333 | 10.6667 | 118892.6303 | 135669.2228 | 8135.9833 | 156052.6038 | 124628.191 | 123071.2276 | 2460.2119 | 2 | 10.6667 | 14 | 41.3334
142 | CCV 1ppb TRACE | R | 06/03/2023 05:53:46 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.144 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1908.1275 | 731.3521 | 48.0001 | 4800.8067 | 848.0253 | 768.6874 | 289.0029 | 926.6967 | 1287.058 | 7277.1864 | 5838.8933 | 12929.1819 | 2253.8445 | 2466.5463 | 686.0165 | 1674.7649 | 12777.0473 | 6194.6768 | 2247.0572 | 68.0002 | 5441.7035 | 5441.7035 | 1368.7324 | 114.6671 | 357.3378 | 773.3544 | 413.1512 | 136.0007 | 10587.2564 | 9689.9524 | 108962.6699 | 135163.7506 | 8107.9679 | 144719.4209 | 128114.0546 | 121753.3017 | 2481.5489 | 5114.9156 | 12166.5131 | 17310.8025 | 39777.1911
143 | CCV 1ppb TRACE | R | 06/03/2023 06:00:18 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.145 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1909.4611 | 730.0187 | 52.3334 | 4069.9137 | 872.6934 | 828.3574 | 266.0025 | 943.0312 | 1265.7228 | 7077.753 | 5781.2589 | 12731.6732 | 2300.5186 | 2430.8735 | 695.3503 | 1673.4314 | 12700.9771 | 6149.3234 | 2112.4589 | 78.6669 | 5403.0222 | 5403.0222 | 1353.3975 | 116.0005 | 365.338 | 746.6862 | 421.0868 | 132.6673 | 10658.6417 | 9864.071 | 109308.0459 | 132202.04 | 8255.7182 | 145354.418 | 118235.9765 | 120817.503 | 2398.2014 | 5050.2259 | 11958.3365 | 16674.5041 | 39084.1967
144 | CCV 10 ppb MAJORS | R | 06/03/2023 06:06:21 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.146 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 17629.8718 | 6766.9367 | 355.0044 | 30209.2513 | 2583.9004 | 7440.2706 | 2457.5447 | 8587.9143 | 11613.7196 | 66897.608 | 45182.1446 | 114164.3745 | 10710.0133 | 22806.8598 | 6106.3054 | 14279.4665 | 110456.0537 | 54091.5492 | 41637.8485 | 725.0184 | 49455.465 | 49455.465 | 5566.418 | 1058.0393 | 3391.7361 | 7155.1254 | 4064.2377 | 1219.3855 | 100711.084 | 91080.1385 | 104410.9496 | 126236.6498 | 7482.2926 | 138433.5256 | 111639.2755 | 112936.7634 | 2317.5214 | 47166.4088 | 112262.0415 | 147758.1941 | 351868.3101
145 | CCV 10 ppb MAJORS | R | 06/03/2023 06:12:53 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.147 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 17443.9812 | 6573.8457 | 371.3382 | 30695.6259 | 2701.5889 | 7382.9077 | 2457.2113 | 8565.234 | 11650.0819 | 67726.8733 | 45046.4037 | 118268.311 | 10977.8832 | 22568.4808 | 6101.9698 | 14620.4778 | 113641.5483 | 55031.7985 | 41131.4353 | 719.0181 | 49740.4557 | 49740.4557 | 5635.1117 | 1132.045 | 3501.0958 | 7169.7991 | 3875.225 | 1246.0544 | 102853.614 | 92144.8909 | 108184.0893 | 128484.638 | 7761.1079 | 140634.1918 | 112430.1349 | 119538.1061 | 2263.5127 | 48546.3603 | 113838.4769 | 148657.8676 | 352888.2948
146 | MT 33 B | R | 06/03/2023 06:18:55 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 33 B.148 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.6667 | 1 | 40.3334 | 19696.9057 | 609.3464 | 9 | 2 | 1.6667 | 20 | 82.0002 | 930.5229 | 51.3334 | 1273.7235 | 4 | 17.3333 | 67.3335 | 612.6798 | 490.6751 | -2324.211 | 0.66667 | 8.6667 | 8.6667 | 872.0266 | 0.66667 | 2 | 0 | 31.3883 | 0 | 65.3335 | 45.3334 | 113407.894 | 140808.6413 | 8418.1479 | 154598.0178 | 123486.8257 | 112713.6922 | 2576.2325 | 4.6667 | 17.3333 | 513.3368 | 1175.352
147 | MT 33 B LSP | R | 06/03/2023 06:24:56 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 33 B LSP.149 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 44741.9625 | 16415.7597 | 876.6936 | 91424.9503 | 5667.4573 | 18807.7084 | 6409.1042 | 21932.4919 | 28856.1143 | 167982.7671 | 114043.5202 | 289307.0243 | 25461.0038 | 56512.2743 | 14868.7339 | 35071.3371 | 265835.72 | 125753.355 | 107168.0632 | 1711.1025 | 117863.5789 | 117863.5789 | 12617.5726 | 2571.5648 | 7874.8368 | 16594.9671 | 9755.8177 | 2918.2982 | 255894.5443 | 234568.2003 | 110645.5313 | 125336.1184 | 7438.2704 | 151595.3223 | 109527.0168 | 111565.954 | 2262.846 | 112328.7417 | 266504.5636 | 331997.8917 | 830119.6016
148 | MT 34 B | R | 06/03/2023 06:30:59 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 34 B.150 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 4.3333 | 1 | 30.3334 | 20717.6821 | 605.3462 | 18 | 2 | 2.3333 | 22.6667 | 118.0005 | 1342.8616 | 34.6667 | 1341.063 | 3.3333 | 34 | 44.0001 | 406.6725 | 614.6799 | -2433.3745 | 1.6667 | 31.3334 | 31.3334 | 948.6982 | 1.3333 | 2.6667 | 3.3333 | 34.6622 | 0 | 378.6717 | 216.6683 | 111560.6606 | 140747.333 | 8704.6518 | 156238.4703 | 113549.1462 | 121911.3732 | 2562.2298 | 4 | 28.6667 | 525.3372 | 1124.0168
149 | MT 35 B | R | 06/03/2023 06:37:01 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 35 B.151 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 6 | 0.33333 | 56.0001 | 22707.3684 | 634.3474 | 12.3333 | 2.6667 | 1.3333 | 25.6667 | 90.0003 | 1008.3678 | 34.6667 | 1310.0601 | 1.6667 | 15.6667 | 30.6667 | 296.6697 | 448.6737 | -2601.8883 | 0.33333 | 16 | 16 | 903.362 | 0.66667 | 2.6667 | 4 | 38.7488 | 0 | 78.6669 | 24.6667 | 113701.5025 | 143031.8332 | 8913.7866 | 157711.7943 | 127678.7591 | 128777.9023 | 2584.2341 | 3.3333 | 14.6667 | 488.0033 | 1104.017
150 | MT 36 B | R | 06/03/2023 06:43:03 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 36 B.152 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 5 | 0 | 51.3334 | 21619.683 | 637.0142 | 13.6667 | 0.66667 | 1.3333 | 35.6667 | 170.6677 | 2176.3418 | 60.6668 | 1507.0795 | 2.6667 | 15.6667 | 40.0001 | 286.0029 | 21520.2019 | -2459.6605 | 0.66667 | 12.6667 | 12.6667 | 846.6918 | 0 | 0 | 2.6667 | 46.7489 | 0.66667 | 637.3476 | 548.6772 | 107517.2534 | 140482.7286 | 8423.8178 | 155251.8342 | 120076.0843 | 129946.5825 | 2604.2374 | 2 | 16 | 474.0034 | 1088.016
151 | MT 37 B | R | 06/03/2023 06:49:06 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 37 B.153 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 16.3333 | 0.33333 | 37.6667 | 22927.7184 | 623.0136 | 10.3333 | 2 | 1 | 18.6667 | 97.3337 | 926.2298 | 47.3334 | 1196.7168 | 2 | 15.6667 | 44.6667 | 328.0038 | 489.3417 | -2436.4522 | 0.33333 | 6.6667 | 6.6667 | 862.6927 | 0.66667 | 3.3333 | 0.66667 | 22.1091 | 0.66667 | 40.0001 | 23.3334 | 107769.8696 | 139907.8237 | 8521.8747 | 145882.8466 | 123289.7411 | 123509.6225 | 2501.5524 | 3.3333 | 13.3333 | 475.3363 | 1104.0173
152 | MT 38 B | R | 06/03/2023 06:55:09 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 38 B.154 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 28.3334 | 0.33333 | 37.6667 | 22187.8853 | 608.3463 | 14 | 1.6667 | 1 | 21.3333 | 98.0003 | 1058.0466 | 30 | 1208.3844 | 3 | 15.3333 | 29.3334 | 262.0024 | 546.6772 | -2391.2994 | 0 | 6.6667 | 6.6667 | 867.3597 | 0.66667 | 1.3333 | 0.66667 | 32.6676 | 1.3333 | 83.3336 | 62.0001 | 111174.9387 | 137161.4801 | 8562.2338 | 159501.2907 | 127672.7171 | 123520.5176 | 2592.902 | 2.6667 | 17.3333 | 520.0037 | 1125.3508
153 | MT 39 B | R | 06/03/2023 07:01:12 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 39 B.155 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 4 | 0.66667 | 34.3334 | 22717.3859 | 625.6804 | 13.6667 | 2.6667 | 1.6667 | 24.3334 | 102.0004 | 1355.7651 | 34.6667 | 1278.0573 | 0.66667 | 16 | 35 | 268.6692 | 593.3457 | -2385.1716 | 0.66667 | 6 | 6 | 877.3604 | 0.66667 | 2 | 2 | 24.7108 | 0.66667 | 168.6677 | 134.6673 | 109188.3921 | 140477.3161 | 8279.3992 | 166545.4759 | 122897.0454 | 128990.5374 | 2613.5724 | 4 | 14 | 510.0035 | 1070.0164
154 | MT 40 B | R | 06/03/2023 07:07:15 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 40 B.156 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.6667 | 0.66667 | 40.0001 | 21251.1334 | 620.3468 | 10 | 2 | 1 | 20.6667 | 76.6669 | 1109.8153 | 31.3334 | 1208.3844 | 6 | 14.3333 | 72.3335 | 506.6757 | 544.6771 | -2384.1054 | 0.33333 | 15.3333 | 15.3333 | 844.025 | 0.66667 | 2 | 4 | 19.3666 | 0.66667 | 30.6667 | 26 | 109975.7523 | 135427.5603 | 8220.0356 | 153551.8656 | 123180.851 | 122051.8341 | 2490.8839 | 2.6667 | 13.3333 | 530.6704 | 1026.6816
155 | MT 32 C | R | 06/03/2023 07:13:17 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 32 C.157 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2 | 0.33333 | 36 | 21448.7586 | 598.3459 | 11.6667 | 1 | 1.3333 | 20.3333 | 86.6669 | 946.4412 | 48.0001 | 1178.3819 | 1.3333 | 14.3333 | 71.6668 | 565.3445 | 660.6819 | -2428.5088 | 0.66667 | 10 | 10 | 819.3569 | 0 | 1.3333 | 2.6667 | 32.055 | 0 | 43.3334 | 24 | 108777.3017 | 139746.3902 | 8298.4096 | 145988.16 | 121473.0822 | 120652.241 | 2636.2433 | 3.3333 | 16 | 628.0055 | 1377.3593
156 | LBLANK | R | 06/03/2023 07:19:20 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LBLANK.158 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1 | 0 | 4 | 3460.4192 | 604.0128 | 14.3333 | 0 | 1.3333 | 9.6667 | 30.6667 | 760.336 | 20.6667 | 1131.0448 | 1 | 11 | 10.3333 | 78.0002 | 383.3385 | -2482.8591 | 0.33333 | 6.6667 | 6.6667 | 823.3571 | 0 | 2 | 2 | 15.4099 | 1.3333 | 26 | 4 | 108185.5663 | 133665.4787 | 8389.4652 | 147146.6733 | 119876.9546 | 122861.4512 | 2468.8805 | 0.66667 | 13.3333 | 16.6667 | 36
157 | CCV 1ppb TRACE | R | 06/03/2023 07:25:23 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.159 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1921.796 | 754.0199 | 46.6667 | 4518.7147 | 816.0234 | 795.6889 | 265.3358 | 932.0304 | 1257.0553 | 7249.1721 | 5786.0683 | 12680.9594 | 2261.5123 | 2459.5451 | 658.0152 | 1652.7623 | 12827.0895 | 6237.3615 | 2032.1648 | 75.6669 | 5359.0051 | 5359.0051 | 1346.0635 | 113.3338 | 376.005 | 754.6866 | 417.1622 | 132.0006 | 10599.2644 | 9750.6609 | 107282.1269 | 135732.5163 | 8100.9661 | 145758.489 | 118264.9241 | 118534.4169 | 2425.5392 | 5027.5512 | 12231.2352 | 17066.6379 | 39673.6302
158 | CCV 1ppb TRACE | R | 06/03/2023 07:31:56 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.160 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1935.1311 | 716.3513 | 54.3334 | 3983.8888 | 873.0267 | 796.6889 | 267.3358 | 932.0304 | 1228.7196 | 7195.1459 | 5657.3091 | 12311.305 | 2225.5068 | 2428.2064 | 703.3506 | 1721.7706 | 12670.2859 | 6161.3291 | 2142.7554 | 78.3336 | 5176.2731 | 5176.2731 | 1240.7206 | 116.6672 | 345.3375 | 748.0198 | 429.1843 | 137.334 | 10559.235 | 9770.0093 | 106965.6327 | 133417.6497 | 8035.5929 | 130688.3426 | 114753.0851 | 119092.3798 | 2506.8881 | 5110.2483 | 12225.9012 | 16612.4495 | 38141.1927
159 | LFB 10ppb | R | 06/03/2023 07:37:58 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LFB 10ppb.161 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 16782.1899 | 6332.07 | 368.6714 | 31727.8689 | 2524.5566 | 6834.6349 | 2363.5289 | 8305.7473 | 11150.6836 | 65304.2801 | 46590.1414 | 112647.5659 | 10982.2204 | 21797.9511 | 5945.9038 | 13871.3983 | 107889.2259 | 70658.3161 | 40506.4975 | 695.0169 | 47311.0038 | 47311.0038 | 5545.7431 | 1008.0358 | 3357.7283 | 6945.0218 | 3832.2584 | 1282.7243 | 99012.0018 | 87801.4522 | 105510.8761 | 134505.7236 | 7576.3419 | 138475.9196 | 109901.8529 | 112617.5662 | 2588.2394 | 46686.1649 | 109293.1504 | 140724.7301 | 337616.396
160 | LFB 10ppb | R | 06/03/2023 07:44:29 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LFB 10ppb.162 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 15885.4974 | 6112.9758 | 324.0037 | 31153.2843 | 2475.5484 | 6763.9353 | 2264.5128 | 8032.9262 | 10863.1334 | 3309.6052 | 3056.9029 | 106703.7973 | 10760.3852 | 20771.4278 | 5646.1157 | 13258.4852 | 6545.7222 | 3775.7941 | 378.2931 | 684.3497 | 46911.5723 | 46911.5723 | 148.6674 | 998.7017 | 3341.7242 | 6963.0304 | -12.491 | 1184.0491 | 5763.4801 | 85448.8807 | 107236.4077 | 121388.3141 | 8105.6386 | 157401.7136 | 135358.385 | 115099.1941 | 2286.1831 | 2442.517 | 5655.4768 | 53.3334 | 87.3334
161 | CCV 10 ppb MAJORS | R | 06/03/2023 07:50:31 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.163 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 17029.8152 | 6415.7739 | 341.6708 | 31018.3065 | 2562.8966 | 7139.1171 | 2336.8578 | 8435.8237 | 11386.5364 | 67942.5566 | 45680.2594 | 115974.8616 | 10616.2792 | 22197.8998 | 5860.5359 | 14149.0037 | 111637.9584 | 54377.9652 | 42208.6043 | 688.6833 | 49657.5126 | 49657.5126 | 5577.0901 | 1097.3755 | 3364.3963 | 7107.768 | 4021.763 | 1211.3847 | 102243.9661 | 91930.2224 | 92513.7479 | 126923.3733 | 7568.6714 | 137014.649 | 102452.1228 | 112436.745 | 2377.5316 | 47340.9859 | 112751.2473 | 147303.0634 | 352748.1035
162 | CCV 10 ppb MAJORS | R | 06/03/2023 07:57:03 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.164 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 17227.3843 | 6487.8083 | 330.3372 | 30116.395 | 2568.8978 | 7345.5555 | 2437.8749 | 8668.6302 | 11350.5093 | 69044.4484 | 46459.096 | 113182.6617 | 10726.3593 | 22389.8677 | 6012.9333 | 14441.631 | 113558.941 | 54907.9883 | 41372.7551 | 750.6864 | 48882.8492 | 48882.8492 | 5545.7435 | 1088.0415 | 3337.0565 | 6987.0416 | 3972.6261 | 1260.7223 | 104042.1623 | 91093.5586 | 112585.4202 | 143947.2876 | 7557.3336 | 141370.0921 | 101985.7573 | 114759.1125 | 2632.2431 | 48151.6819 | 113445.3968 | 142896.6112 | 346517.1897
163 | MT 32 C LD | R | 06/03/2023 08:03:06 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 32 C LD.165 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 6.6667 | 0.33333 | 32.6667 | 19532.0107 | 600.346 | 16.6667 | 1.3333 | 1.3333 | 21 | 81.3336 | 1037.3909 | 35.3334 | 1252.7216 | 1.6667 | 15.3333 | 78.0002 | 578.6786 | 718.6847 | -2273.5784 | 0.33333 | 16 | 16 | 872.6933 | 0 | 2.6667 | 2.6667 | 28.7541 | 0.66667 | 93.3336 | 36.6667 | 113745.265 | 126424.5215 | 8069.9471 | 151879.135 | 122677.8627 | 123253.4766 | 2297.5186 | 2.6667 | 23.3334 | 650.0055 | 1411.3601
164 | MT 33 C | R | 06/03/2023 08:09:08 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 33 C.166 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.6667 | 0 | 35.3334 | 20338.468 | 619.0135 | 10.3333 | 1.3333 | 0.66667 | 20 | 78.0002 | 1039.4982 | 48.6668 | 1249.388 | 2 | 20 | 72.6669 | 562.0111 | 594.6791 | -2521.8711 | 0 | 6 | 6 | 846.6919 | 0.66667 | 0.66667 | 2 | 20.0387 | 0 | 50.6668 | 29.3334 | 102830.0991 | 136268.1935 | 8059.6092 | 146983.639 | 123638.9099 | 121637.7582 | 2522.8895 | 4 | 15.3333 | 523.3371 | 1107.3503
165 | MT 33 C LSP | R | 06/03/2023 08:15:09 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 33 C LSP.167 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 43686.0327 | 15932.2195 | 891.0278 | 92216.6957 | 5403.0217 | 17946.6008 | 6209.016 | 21764.233 | 28537.4762 | 77993.6716 | 49435.8226 | 285115.3328 | 24811.5279 | 54801.5816 | 14598.789 | 34033.1577 | 121399.6896 | 65483.0818 | 113272.9672 | 1712.1026 | 114881.4787 | 114881.4787 | 13276.8336 | 2573.5655 | 7827.4773 | 16153.1287 | 10073.8764 | 2862.9536 | 283.3362 | 230702.2215 | 103961.4826 | 126128.0384 | 7388.2451 | 81125.7755 | 48003.1913 | 109728.4891 | 2293.518 | 75256.3952 | 181684.0907 | 257886.7888 | 638792.9594
166 | MT 33 C LSP | R | 06/03/2023 08:21:40 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 33 C LSP.168 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 43439.303 | 16114.0852 | 860.3592 | 89465.507 | 5457.0422 | 18363.4636 | 6160.9955 | 21928.486 | 28647.3663 | 171721.3086 | 117252.5411 | 283216.0283 | 25114.3968 | 56082.8917 | 14775.6406 | 34743.2036 | 274424.1441 | 128039.9073 | 106898.4534 | 1768.4428 | 114199.9758 | 114199.9758 | 12672.9567 | 2515.5549 | 7732.759 | 16427.4406 | 9775.2254 | 2796.9405 | 261152.0872 | 228967.6898 | 107237.2125 | 122716.8485 | 7409.5892 | 136288.4253 | 108838.3988 | 110484.9563 | 2284.1828 | 113213.5216 | 270836.6739 | 330616.9044 | 826415.7634
167 | MT 34 C | R | 06/03/2023 08:27:42 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 34 C.169 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 4.6667 | 0.66667 | 41.3334 | 20236.9903 | 616.0133 | 11.6667 | 0.66667 | 1.3333 | 24.6667 | 120.6672 | 1370.2676 | 45.3334 | 1345.7301 | 5.6667 | 40.6667 | 30.6667 | 247.3355 | 475.3412 | -2468.0703 | 1 | 26 | 26 | 871.3599 | 0.66667 | 2 | 1.3333 | 21.3937 | 0 | 371.3382 | 226.0018 | 98199.9639 | 137868.6005 | 8358.4446 | 154848.0424 | 126369.1749 | 123458.5553 | 2480.2154 | 1.3333 | 13.3333 | 470.0032 | 1011.3471
168 | MT 35 C | R | 06/03/2023 08:33:45 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 35 C.170 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 5.3333 | 0.66667 | 35.6667 | 20433.2864 | 608.3463 | 13.6667 | 1.3333 | 1 | 15 | 84.6669 | 1003.6141 | 34 | 1260.3889 | 2 | 16.6667 | 31.6667 | 289.3363 | 536.6768 | -2520.3217 | 0 | 11.3333 | 11.3333 | 894.6947 | 0 | 2.6667 | 0 | 28.7812 | 0.66667 | 100.0004 | 20.6667 | 108502.6917 | 144322.8135 | 8376.4573 | 154816.2195 | 123058.7341 | 120223.9515 | 2600.237 | 2.6667 | 17.3333 | 451.3361 | 1054.6817

## Raw Intensities RSDs
| Sample Id | R | Acquisition Time | QC Status | Dataset File | Method File | Na KED 23/23 Helium KED_Low | Mg KED 24/24 Helium KED_Low | Al KED 27/27 Helium KED_Low | P Oshift 31/47 Oxygen DRC | S (34) -1 34/34 Helium KED_Low | K KED 39/39 Helium KED_Low | Ca KED 43/43 Helium KED_Low | V KED 51/51 Helium KED_Low | Cr (52) 52/52 Helium KED_Low | Cr (53) NH3 53/53 Ammonia DRC | Fe (54) NH3 54/54 Ammonia DRC | Mn Oshift 55/71 Oxygen DRC | Fe (56) KED 56/56 Helium KED_Low | Co KED 59/59 Helium KED_Low | Ni KED 60/60 Helium KED_Low | Cu (63) KED 63/63 Helium KED_Low | Cu (65) NH3 65/65 Ammonia DRC | Zn NH3 66/66 Ammonia DRC | As 75/75 | As KED 75/75 Helium KED_Low | As RH Oshift 75/91 Oxygen DRC | As Oshift 75/91 Oxygen DRC | Se (77) 77/77 | Se (77) Oshift 77/93 Oxygen DRC | Se (78) Oshift 78/94 Oxygen DRC | Se (80) Oshift 80/96 Oxygen DRC | Se (82) 82/82 | Se (82) Oshift 82/98 Oxygen DRC | Mo (95) NH3 95/95 Ammonia DRC | Mo (95) OShift 95/127 Oxygen DRC | Rh (IS1) 103/103 (IS) | Rh (IS2) 103/103 (IS) Oxygen DRC | Rh (IS3) 103/103 (IS) Helium KED_Low | Rh (IS4) 103/103 (IS) Ammonia DRC | Rh (IS5) 103/103 (IS) Ammonia DRC | Rh (IS6) 103/103 (IS) Oxygen DRC | Rh OShift (IS1) 103/119 (IS) Oxygen DRC | Cd (111) NH3 111/111 Ammonia DRC | Cd (114) NH3 114/114 Ammonia DRC | Pb 208/208 | Pb NH3 208/208 Ammonia DRC
1 | CAL BLANK | R | 6/2/2023 3:17:36 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CAL BLANK.004 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 173.2% |  | 50% | 3.249% | 5.07% | 26.96% | 173.2% | 86.6% | 35.25% | 16.37% | 2.584% | 35.25% | 3.096% | 34.64% | 71.32% | 69.98% | 8.333% | 16.99% | 7.525% | 173.2% | 20.4% | 20.4% | 1.594% |  | 100% | 69.28% | 65.6% |  | 44.61% | 65.47% | 3.277% | 0.3095% | 0.7573% | 0.4141% | 0.5558% | 2.017% | 1.979% |  | 21.65% | 30.31% | 2.014%
2 | LOW 0.01 | R | 6/2/2023 3:23:43 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LOW 0.01.005 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 36.46% | 40% | 19.52% | 2.835% | 10.05% | 20.03% | 124.9% | 30.12% | 30.31% | 12.77% | 0.3086% | 11.86% | 2.332% | 35.18% | 37% | 31.05% | 3.006% | 4.33% | 4.698% | 100% | 22.42% | 22.42% | 4.982% | 173.2% | 81.04% | 53.93% | 81.04% | 65.47% | 3.008% | 1.72% | 0.6899% | 1.464% | 2.25% | 1.661% | 1% | 1.345% | 3.114% | 13.12% | 9.945% | 5.54% | 4.757%
3 | LOW 0.05 | R | 6/2/2023 3:29:46 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LOW 0.05.006 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 17.02% | 14.43% | 14.24% | 2.411% | 3.555% | 10.38% | 5.973% | 6.245% | 11.89% | 9.079% | 3.395% | 2.539% | 1.98% | 5.632% | 4.225% | 19.52% | 5.766% | 3.663% | 12.89% | 44.61% | 9.242% | 9.242% | 1.136% | 98.97% | 28.39% | 17.49% | 28.99% | 26.96% | 0.6492% | 4.377% | 3.037% | 3.901% | 6.34% | 2.055% | 2.358% | 3.639% | 5.768% | 7.306% | 8.714% | 5.326% | 1.816%
4 | LOW 0.1 | R | 6/2/2023 3:35:48 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LOW 0.1.007 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 15.04% | 15.73% | 12.39% | 1.405% | 2.84% | 10.1% | 20.83% | 15.5% | 9.666% | 6.095% | 2.419% | 4.839% | 1.728% | 2.989% | 10.6% | 8.605% | 7.911% | 3.22% | 5.971% | 12.5% | 5.157% | 5.157% | 7.325% | 16.67% | 11.45% | 4.992% | 53.31% | 11.11% | 2.371% | 1.102% | 1.229% | 0.8802% | 2.017% | 1.21% | 0.9493% | 0.0713% | 2.965% | 2.182% | 3.855% | 0.7916% | 0.9116%
5 | LOW 0.5 | R | 6/2/2023 3:41:51 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LOW 0.5.008 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.836% | 6.36% | 4.949% | 2.218% | 6.179% | 7.962% | 9.566% | 10.21% | 7.332% | 1.43% | 1.935% | 3.574% | 2.219% | 2.086% | 4.036% | 1.062% | 1.223% | 0.09546% | 57.43% | 11.93% | 2.889% | 2.889% | 5.175% | 15.59% | 6.242% | 3.071% | 18.25% | 24.69% | 2.763% | 2.696% | 0.4323% | 8.264% | 5.283% | 1.121% | 0.4188% | 2.735% | 8.717% | 2.671% | 1.475% | 0.651% | 1.501%
6 | STD 1 | R | 6/2/2023 3:47:53 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\STD 1.009 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.4473% | 4.277% | 27.94% | 1.972% | 2.755% | 4.261% | 2.374% | 0.8977% | 3.62% | 1.433% | 2.184% | 0.8919% | 0.9839% | 3.161% | 0.8237% | 2.975% | 2.671% | 1.214% | 7.843% | 8.688% | 1.984% | 1.984% | 5.628% | 10.49% | 5.977% | 2.946% | 17.53% | 9.722% | 1.063% | 0.5592% | 0.6733% | 0.3953% | 2.17% | 1.01% | 9.13% | 0.9068% | 2.259% | 1.309% | 1.593% | 0.5761% | 2.412%
7 | STD 5 | R | 6/2/2023 3:53:56 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\STD 5.010 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.254% | 4.076% | 2.128% | 5.966% | 2.133% | 0.6055% | 6.316% | 3.38% | 2.121% | 2.501% | 1.111% | 0.7405% | 1.458% | 0.4575% | 0.96% | 2.405% | 1.897% | 0.4943% | 2.704% | 3.063% | 0.6049% | 0.6049% | 2.76% | 5.804% | 2.293% | 3.24% | 7.939% | 6.074% | 1.569% | 0.6612% | 1.937% | 1.007% | 1.888% | 0.7846% | 0.8302% | 1.916% | 2.833% | 1.576% | 0.6211% | 0.6847% | 1.242%
8 | STD 10 | R | 6/2/2023 3:59:58 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\STD 10.011 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.1747% | 1.176% | 2.438% | 0.7705% | 4.8% | 0.21% | 2.221% | 0.9789% | 0.3491% | 1.298% | 1.841% | 0.683% | 0.8106% | 0.3918% | 0.6821% | 0.9175% | 1.386% | 1.064% | 1.862% | 2.755% | 1.13% | 1.13% | 3.834% | 2.812% | 2.919% | 1.047% | 1.421% | 4.716% | 1.722% | 0.8192% | 0.4851% | 1.248% | 0.6937% | 0.8391% | 1.152% | 1.515% | 3.184% | 0.9192% | 0.8274% | 0.9596% | 1.71%
9 | STD 25 | R | 6/2/2023 4:06:00 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\STD 25.012 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.687% | 1.14% | 3.255% | 0.9798% | 1.071% | 3.02% | 3.322% | 2.099% | 1.543% | 0.945% | 0.8676% | 1.378% | 2.769% | 2.757% | 2.656% | 1.209% | 1.414% | 0.4512% | 2.159% | 0.7249% | 0.5514% | 0.5514% | 3.822% | 1.329% | 1.811% | 1.56% | 0.8396% | 2.435% | 0.7318% | 0.2724% | 0.7074% | 1.095% | 4.204% | 0.3375% | 1.562% | 0.3715% | 3.595% | 0.4057% | 0.5119% | 0.4125% | 0.6391%
10 | STD 50 | R | 6/2/2023 4:12:02 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\STD 50.013 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.7812% | 0.5663% | 4.028% | 1.161% | 1.854% | 0.7505% | 0.7678% | 1.121% | 0.8834% | 1.208% | 0.4596% | 0.8835% | 0.6368% | 0.9289% | 0.9361% | 0.7582% | 0.4882% | 0.9022% | 1.26% | 1.765% | 0.4139% | 0.4139% | 2.244% | 0.6948% | 1.452% | 1.073% | 2.146% | 0.6544% | 0.9626% | 0.9559% | 3.075% | 0.651% | 0.2639% | 1.039% | 0.9888% | 1.18% | 2.279% | 0.8009% | 1.266% | 3.733% | 0.5847%
11 | STD 100 | R | 6/2/2023 4:18:04 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\STD 100.014 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.3222% | 0.951% | 1.705% | 0.5434% | 0.2721% | 0.2372% | 0.9574% | 0.4122% | 0.411% | 0.2799% | 0.7363% | 0.6018% | 0.3959% | 0.5327% | 0.2606% | 0.7711% | 0.7858% | 1.23% | 0.4142% | 2.218% | 0.5258% | 0.5258% | 1.212% | 0.9026% | 1.344% | 0.3857% | 0.592% | 0.4646% | 1.807% | 0.2448% | 1.824% | 0.3891% | 2.22% | 0.9116% | 0.2953% | 0.2254% | 0.3844% | 1.442% | 1.211% | 3.837% | 0.04454%
13 | LBLANK | R | 6/2/2023 4:24:07 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LBLANK.015 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 41.66% | 173.2% | 21.65% | 4.866% | 4.656% | 54.49% | 100% | 33.33% | 50.92% | 18.55% | 3.921% | 30.7% | 1.137% | 65.47% | 15.75% | 55.79% | 5.088% | 2.307% | 13.13% | 173.2% | 21.65% | 21.65% | 4.676% |  | 50% | 25% | 33.3% | 173.2% | 10.63% | 16.38% | 17.26% | 1.191% | 0.8535% | 1.196% | 0.8029% | 0.3994% | 1.61% | 24.74% | 7.37% | 55.63% | 9.14%
14 | ICV PE | R | 6/2/2023 4:30:09 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\ICV PE.016 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.087% | 0.8433% | 3.978% | 0.4769% | 1.658% | 1.861% | 2.005% | 0.6399% | 1.4% | 0.6292% | 0.8771% | 0.7045% | 0.2028% | 1.168% | 0.8993% | 1.777% | 0.5739% | 0.5399% | 3.337% | 3.681% | 1.567% | 1.567% | 1.928% | 1.86% | 1.726% | 1.151% | 1.817% | 1.765% | 0.8283% | 0.2208% | 0.5684% | 0.7782% | 2.672% | 1.203% | 0.7208% | 1.13% | 0.4998% | 0.6286% | 0.8142% | 0.1791% | 0.7105%
15 | CCV 1ppb TRACE | R | 6/2/2023 4:36:12 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.017 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 4.3% | 6.117% | 10.41% | 1.889% | 3.2% | 4.386% | 3.216% | 1.865% | 1.871% | 1.261% | 1.023% | 1.018% | 4.335% | 2.647% | 5.017% | 3.846% | 1.322% | 3.356% | 8.83% | 17.14% | 1.396% | 1.396% | 3.128% | 5.728% | 2.549% | 2.327% | 3.717% | 2.745% | 1.758% | 0.7447% | 0.3011% | 0.2138% | 1.258% | 0.8222% | 1.164% | 1.04% | 0.2663% | 0.8851% | 1.324% | 1.781% | 0.6384%
16 | CCV 1ppb TRACE | R | 6/2/2023 4:42:44 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.018 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 3.109% | 4.999% | 12.37% | 2.438% | 3.692% | 2.696% | 2.599% | 0.5293% | 3.684% | 1.098% | 0.1848% | 0.5507% | 2.043% | 1.8% | 5.211% | 0.246% | 1.366% | 3.846% | 9.349% | 10.15% | 2.499% | 2.499% | 1.231% | 8.743% | 7.472% | 1.305% | 13.56% | 8.753% | 0.9852% | 0.3226% | 1.782% | 0.6125% | 1.518% | 0.4966% | 0.2556% | 0.8769% | 1.613% | 0.3101% | 1.697% | 0.6929% | 0.8235%
17 | LFB 10ppb | R | 6/2/2023 4:48:46 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LFB 10ppb.019 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.216% | 1.562% | 4.834% | 0.7181% | 0.7259% | 0.5747% | 1.539% | 1.877% | 1.399% | 0.3726% | 1.294% | 0.8019% | 0.4502% | 0.8497% | 1.187% | 0.9849% | 0.7759% | 0.4483% | 2.504% | 2.099% | 0.6722% | 0.6722% | 0.6581% | 1.577% | 1.174% | 1.698% | 2.859% | 2.452% | 0.08195% | 1.43% | 1.052% | 1.135% | 1.665% | 0.5565% | 0.1487% | 0.4907% | 3.728% | 1.51% | 0.2299% | 0.3137% | 1.938%
18 | LFB 10ppb | R | 6/2/2023 4:55:16 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LFB 10ppb.020 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.461% | 0.7818% | 1.926% | 1.088% | 3.679% | 1.038% | 1.766% | 0.3726% | 1.029% | 0.7598% | 0.9601% | 0.2638% | 1.086% | 0.6225% | 0.8301% | 1.11% | 0.5667% | 0.8437% | 1.802% | 1.545% | 0.8873% | 0.8873% | 1.509% | 4.899% | 1.569% | 1.177% | 5.292% | 4.7% | 1.31% | 0.4158% | 1.041% | 1.084% | 0.8078% | 0.9704% | 1.585% | 1.534% | 2.448% | 0.8557% | 0.676% | 0.2894% | 1.643%
19 | LLOQ 0.01 | R | 6/2/2023 5:01:19 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LLOQ 0.01.021 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 18.87% | 40% | 28.47% | 4.536% | 3.95% | 24.5% | 35.25% | 38.19% | 6.928% | 7.143% | 5.355% | 15.86% | 1.562% | 24.8% | 3.464% | 15.74% | 8.752% | 4.235% | 7.36% | 86.6% | 14.88% | 14.88% | 6.243% | 50% | 88.19% | 43.83% | 32.71% | 86.6% | 3.483% | 5.929% | 0.3595% | 1.23% | 0.9884% | 0.27% | 2.44% | 0.9701% | 2.43% | 8.248% | 12.03% | 13.08% | 4.418%
20 | CCV 10 ppb MAJORS | R | 6/2/2023 5:07:22 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.022 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.1951% | 0.3698% | 4.376% | 1.805% | 0.8696% | 2.612% | 2.971% | 1.034% | 1.596% | 0.332% | 0.9776% | 1.383% | 0.8238% | 0.7791% | 0.712% | 1.656% | 0.3926% | 0.5568% | 1.851% | 2.781% | 0.2208% | 0.2208% | 2.477% | 3.578% | 0.6761% | 2.314% | 1.545% | 4.421% | 0.2399% | 0.4561% | 0.7594% | 0.8657% | 4.005% | 0.9703% | 0.4442% | 1.381% | 1.44% | 0.3882% | 1.104% | 1.188% | 2.296%
21 | CCV 10 ppb MAJORS | R | 6/2/2023 5:13:54 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.023 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.7775% | 0.9444% | 0.304% | 0.8525% | 1.405% | 0.1139% | 1.901% | 0.6684% | 1.152% | 0.131% | 1.141% | 1.572% | 0.3769% | 1.503% | 0.631% | 1.206% | 0.3915% | 0.7023% | 2.183% | 3.223% | 0.89% | 0.89% | 2.132% | 5.802% | 2.273% | 2.339% | 5.745% | 2.875% | 0.7741% | 1.133% | 1.208% | 5.284% | 1.377% | 1.12% | 1.755% | 0.5646% | 5.952% | 0.9129% | 0.4638% | 1.334% | 2.598%
22 | LLOQ 0.1 | R | 6/2/2023 5:19:57 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LLOQ 0.1.024 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 12.62% | 9.206% | 11.76% | 3.269% | 4.874% | 15.11% | 18.25% | 14.64% | 8.22% | 6.206% | 3.496% | 2.971% | 4.022% | 0.5469% | 13.97% | 2.507% | 4.524% | 2.921% | 4.479% | 17.32% | 4.269% | 4.269% | 1.021% | 74.18% | 12% | 2.761% | 21.51% | 20.83% | 4.405% | 8.797% | 9.201% | 1.188% | 2.087% | 9.1% | 2.183% | 1.131% | 3.106% | 5.124% | 3.133% | 0.8725% | 1.488%
23 | LBLANK | R | 6/2/2023 5:25:59 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LBLANK.025 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 34.64% |  | 44.61% | 2.009% | 5.649% | 28.36% |  | 86.6% | 52.68% | 17.16% | 12.6% | 19.68% | 5.016% | 100% | 22.91% | 51.67% | 4.992% | 5.878% | 6.328% | 173.2% | 19.92% | 19.92% | 6.571% | 86.6% | 100% | 24.74% | 115.4% | 114.6% | 26.41% | 45.81% | 0.7912% | 0.3726% | 1.346% | 6.44% | 0.4287% | 0.6385% | 1.826% | 17.32% | 28.39% | 40.16% | 11.76%
24 | LFB 6.25ppb & 2500ppb | R | 6/2/2023 5:32:01 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LFB 6.25ppb & 2500ppb.026 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.1626% | 1.885% | 4.019% | 4.024% | 1.381% | 1.277% | 0.8771% | 0.3248% | 0.6361% | 0.8657% | 1.615% | 3.582% | 0.4957% | 0.4702% | 2.596% | 1.723% | 0.7458% | 1.825% | 0.8187% | 4.353% | 1.899% | 1.899% | 3.366% | 6.323% | 1.874% | 0.244% | 1.37% | 1.681% | 1.802% | 2.625% | 1.492% | 8.57% | 1.654% | 0.2609% | 0.5971% | 0.9686% | 11.66% | 2.005% | 1.592% | 0.5574% | 1.01%
25 | MT 1 | R | 6/2/2023 5:38:02 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 1.027 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.167% | 10.11% | 1.33% | 1.159% | 2.262% | 2.421% | 2.131% | 6.641% | 0.4404% | 1.87% | 1.05% | 2.733% | 0.6197% | 9.59% | 1.804% | 0.4602% | 1.617% | 1.092% | 7.907% | 7.905% | 2.134% | 2.134% | 1.628% | 21.71% | 7.198% | 7.52% | 9.231% | 12.73% | 4.754% | 1.525% | 0.9372% | 0.5176% | 1.204% | 0.4976% | 0.2489% | 2.413% | 5.928% | 0.9275% | 5.595% | 1.882% | 1.232%
26 | MT 1 LD | R | 6/2/2023 5:44:04 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 1 LD.028 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.4317% | 11.95% | 0.9381% | 4.023% | 3.087% | 2.412% | 15.35% | 3.761% | 1.807% | 0.9959% | 0.1825% | 1.11% | 0.7134% | 8.98% | 6.904% | 1.543% | 0.8213% | 0.7641% | 7.808% | 9.061% | 2.272% | 2.272% | 1.015% | 7.262% | 3.698% | 4.66% | 13.51% | 10.17% | 6.083% | 2.836% | 4.661% | 0.2501% | 0.2011% | 4.336% | 0.3626% | 12.94% | 3.78% | 8.844% | 0.7994% | 1.677% | 2.655%
27 | LBLANK | R | 6/2/2023 6:05:56 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LBLANK.029 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 91.65% | 173.2% | 66.14% | 4.442% | 42.17% | 40.56% |  | 173.2% | 35.49% | 14.19% | 1.454% | 39.77% | 19.57% | 0% | 13.32% | 18.41% | 21.15% | 4.276% | 4.195% |  | 28.87% | 28.87% | 8.722% | 173.2% | 91.65% | 57.74% | 43.59% |  | 76.36% | 30.55% | 0.8819% | 3.788% | 9.354% | 1.37% | 0.708% | 0.9342% | 2.821% |  | 23.89% | 10.66% | 18.78%
28 | STD 1 | R | 6/2/2023 6:11:58 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\STD 1.030 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.409% | 2.623% | 13.64% | 1.443% | 2.849% | 3.657% | 6.497% | 0.413% | 1.758% | 2.314% | 2.225% | 2.113% | 2.929% | 1.431% | 3.524% | 1.617% | 1.655% | 1.337% | 6.502% | 8.678% | 1.537% | 1.537% | 6.255% | 1.176% | 2.981% | 2.081% | 5.351% | 12.89% | 1.502% | 0.5033% | 0.8004% | 0.4601% | 1.337% | 0.4157% | 0.9085% | 1.846% | 2.49% | 2.045% | 0.6403% | 1.535% | 1.208%
29 | STD 0.5 | R | 6/2/2023 6:18:06 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\STD 0.5.031 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.662% | 3.314% | 17.75% | 0.8926% | 5.884% | 5.799% | 7.285% | 3.025% | 6.47% | 2.581% | 4.277% | 2.757% | 1.55% | 0.4278% | 4.676% | 5.991% | 0.8088% | 1.116% | 191% | 5% | 2.538% | 2.538% | 4.324% | 28.91% | 6.283% | 6.895% | 9.528% | 4.255% | 0.1362% | 1.099% | 5.292% | 0.7411% | 1.748% | 0.8597% | 0.9448% | 0.8309% | 3.994% | 4.555% | 1.144% | 1.612% | 0.1692%
30 | ICV PE | R | 6/2/2023 6:24:08 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\ICV PE.032 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.212% | 2.716% | 4.337% | 1.253% | 0.6976% | 4.417% | 2.169% | 2.11% | 0.5089% | 1.153% | 1.218% | 0.279% | 1.049% | 1.094% | 1.318% | 0.4108% | 0.4646% | 0.9981% | 1.973% | 3.055% | 1.66% | 1.66% | 2% | 1.539% | 1.417% | 0.4347% | 2.108% | 7.561% | 0.4917% | 0.1372% | 0.4835% | 0.5159% | 0.3361% | 1.553% | 0.3284% | 0.9607% | 2.663% | 0.8972% | 1.132% | 1.155% | 1.333%
31 | ICV PE | R | 6/2/2023 6:30:39 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\ICV PE.033 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.7599% | 1.56% | 1.661% | 0.8995% | 1.798% | 1.594% | 2.467% | 1.348% | 0.8186% | 1.033% | 0.4821% | 2.063% | 0.2424% | 0.3142% | 1.035% | 0.8251% | 0.9304% | 0.4609% | 1.346% | 5.215% | 0.4426% | 0.4426% | 0.7589% | 1.656% | 2.233% | 0.6206% | 0.2271% | 3.481% | 0.6467% | 1.136% | 0.9306% | 1.079% | 0.9531% | 1.59% | 1.321% | 1.55% | 2.529% | 0.589% | 1.139% | 0.4655% | 1.477%
32 | CCV 1ppb TRACE | R | 6/2/2023 6:36:41 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.034 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 3.273% | 3.43% | 13.36% | 4.447% | 2.636% | 2.467% | 3.733% | 0.756% | 4.025% | 0.7708% | 2.856% | 1.896% | 0.02211% | 0.8813% | 1.149% | 1.456% | 0.2317% | 1.085% | 8.361% | 9.739% | 1.735% | 1.735% | 4.542% | 6.772% | 11.58% | 3.51% | 2.328% | 16.89% | 0.451% | 0.7625% | 1.422% | 4.075% | 3.134% | 0.9271% | 0.1638% | 0.9257% | 5.339% | 1.015% | 1.757% | 1.162% | 0.8806%
33 | CCV 1ppb TRACE | R | 6/2/2023 6:43:13 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.035 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.454% | 4.591% | 6.536% | 1.38% | 0.7667% | 1.111% | 6.415% | 3.514% | 1.952% | 0.8592% | 0.8278% | 0.8822% | 1.872% | 1.793% | 1.324% | 1.556% | 1.151% | 1.24% | 7.359% | 6.505% | 0.4802% | 0.4802% | 2.11% | 11.7% | 5.515% | 1.729% | 5.925% | 9.312% | 1.181% | 1.23% | 0.4997% | 1.523% | 0.2567% | 1.24% | 1.106% | 0.315% | 3.524% | 1.449% | 0.7954% | 0.07068% | 0.2527%
34 | CCV 10 ppb MAJORS | R | 6/2/2023 6:49:16 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.036 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.102% | 2.255% | 6.204% | 2.106% | 0.4947% | 1.916% | 1.905% | 1.777% | 0.363% | 0.1913% | 0.1377% | 0.2425% | 0.4109% | 0.7486% | 1.587% | 0.987% | 0.7521% | 0.3459% | 1.118% | 3.596% | 0.03619% | 0.03619% | 1.158% | 2.028% | 0.5523% | 1.42% | 0.6755% | 4.755% | 0.4199% | 1.28% | 1.213% | 0.4581% | 0.9478% | 1.095% | 2.5% | 1.903% | 3.479% | 1.305% | 1.191% | 0.943% | 2.306%
35 | CCV 10 ppb MAJORS | R | 6/2/2023 6:55:47 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.037 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.024% | 1.026% | 6.082% | 1.225% | 2.75% | 0.9856% | 2.036% | 0.483% | 1.02% | 1.245% | 1.042% | 0.8101% | 1.058% | 1.364% | 1.064% | 1.142% | 1.385% | 1.485% | 0.7191% | 3.285% | 0.9561% | 0.9561% | 1.277% | 2.755% | 3.68% | 1.508% | 1.007% | 1.308% | 1.394% | 0.2687% | 0.4744% | 0.1205% | 0.8444% | 2.254% | 1.46% | 1.43% | 1.266% | 1.729% | 0.6166% | 0.6313% | 1.329%
36 | MT 1 | R | 6/2/2023 7:01:50 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 1.038 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.162% | 2.302% | 0.4937% | 1.979% | 4.325% | 1.5% | 14.25% | 3.373% | 8.41% | 2.853% | 1.15% | 3.098% | 0.7154% | 5.648% | 6.144% | 0.7539% | 1.512% | 0.8478% | 9.087% | 8.531% | 1.982% | 1.982% | 3.649% | 8.696% | 5.18% | 8.026% | 7.191% | 11.2% | 3.195% | 3.007% | 0.6626% | 10.46% | 0.5676% | 0.6905% | 1.108% | 1.269% | 11.49% | 12.61% | 2.011% | 2.793% | 2.258%
37 | MT 1 LD | R | 6/2/2023 7:07:53 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 1 LD.039 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.412% | 5.354% | 1.911% | 1.674% | 4.553% | 3.977% | 16.01% | 1.446% | 6.342% | 0.7727% | 0.6128% | 0.9959% | 1.435% | 8.931% | 1.309% | 1.457% | 0.4913% | 0.3522% | 4.686% | 6.281% | 2.414% | 2.414% | 4.348% | 7.233% | 16.3% | 4.285% | 2.466% | 16.44% | 5.486% | 2.647% | 0.4137% | 1.447% | 5.661% | 0.9494% | 0.5818% | 1.856% | 1.371% | 4.578% | 2.041% | 1.54% | 0.6984%
38 | MT 2 | R | 6/2/2023 7:13:55 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 2.040 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.22% | 8.791% | 2.07% | 0.3736% | 0.8333% | 2.688% | 13.29% | 3.268% | 5.097% | 1.328% | 1.053% | 0.7886% | 1.178% | 5.263% | 8.703% | 1.132% | 1.135% | 0.5299% | 6.783% | 1.178% | 2.608% | 2.608% | 4.111% | 22.32% | 7.894% | 5.727% | 5.545% | 2.624% | 5.28% | 4.297% | 0.7604% | 6.356% | 1.884% | 0.8882% | 0.8442% | 1.233% | 9.465% | 3.896% | 3.036% | 1.548% | 1.732%
39 | MT 2 LSP | R | 6/2/2023 7:19:58 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 2 LSP.041 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.35% | 0.02324% | 0.9613% | 0.7448% | 0.3777% | 1.377% | 1.106% | 0.7969% | 0.9416% | 0.2968% | 0.4817% | 0.9981% | 0.4053% | 1.21% | 2.54% | 1.605% | 1.15% | 1.363% | 2.391% | 6.719% | 1.287% | 1.287% | 2.415% | 6.459% | 2.43% | 2.567% | 1.703% | 1.873% | 0.7818% | 0.6796% | 0.5755% | 0.9017% | 5.028% | 0.8139% | 1.077% | 5.053% | 0.6228% | 1.04% | 0.6056% | 0.5092% | 1.734%
40 | MT 2 LSP | R | 6/2/2023 7:26:09 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 2 LSP.042 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.776% | 1.879% | 1.36% | 1.006% | 0.5247% | 1.536% | 0.739% | 1.518% | 1.529% | 0.7808% | 0.5511% | 1.1% | 0.9339% | 1.254% | 0.6239% | 1.503% | 1.061% | 0.4995% | 1.599% | 0.6439% | 0.2785% | 0.2785% | 3.064% | 3.327% | 2.986% | 0.8424% | 3.516% | 2.911% | 1.104% | 1.258% | 1.223% | 1.209% | 3.109% | 0.5526% | 0.5652% | 0.4011% | 3.085% | 1.367% | 0.1723% | 0.935% | 1.788%
41 | MT 3 | R | 6/2/2023 7:32:11 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 3 .043 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.6938% | 6.074% | 2.193% | 2.259% | 3.146% | 2.577% | 4.615% | 5.848% | 3.142% | 4.404% | 1.275% | 2.41% | 1.064% | 7.348% | 2.313% | 2.448% | 1.62% | 0.146% | 0.8242% | 6.559% | 2.208% | 2.208% | 5.472% | 3.561% | 4.369% | 2.241% | 10.47% | 2.763% | 4.039% | 1.457% | 0.7661% | 0.9923% | 1.467% | 1.192% | 3.551% | 0.8537% | 3.301% | 6.607% | 7.442% | 1.001% | 1.167%
42 | MT 4 | R | 6/2/2023 7:38:14 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 4.044 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.9% | 4.357% | 0.7725% | 1.895% | 3.496% | 2.586% | 9.984% | 5.345% | 2.323% | 3.549% | 0.6815% | 0.8358% | 1.245% | 6.719% | 9.988% | 0.8599% | 0.6202% | 0.7156% | 4.364% | 6.762% | 1.108% | 1.108% | 2.21% | 10% | 9.221% | 1.81% | 5.779% | 7.683% | 1.546% | 2.642% | 1.747% | 1.595% | 0.5574% | 1.393% | 0.9051% | 2.019% | 1.509% | 2.291% | 2.876% | 3.597% | 0.3128%
43 | LBLANK | R | 6/2/2023 7:44:17 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LBLANK.045 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 173.2% | 173.2% | 50% | 0.7496% | 10.45% | 16.41% |  | 100% | 15.75% | 22.21% | 4.932% | 26.65% | 2.195% | 43.3% | 69.39% | 5.973% | 16.77% | 2.614% | 5.199% | 173.2% | 37% | 37% | 3.316% | 173.2% | 34.64% | 15.75% | 35.94% | 86.6% | 29.41% | 14.29% | 0.9507% | 0.8704% | 1.308% | 6.585% | 2.3% | 11.33% | 1.138% | 173.2% | 20.4% | 15.75% | 18.56%
44 | CCV 1ppb TRACE | R | 6/2/2023 7:50:20 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.046 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.9848% | 3.516% | 21.5% | 2.326% | 5.785% | 2.407% | 2.305% | 5.097% | 2.047% | 2.558% | 2.424% | 0.4143% | 2.467% | 1.715% | 0.9788% | 0.8598% | 0.7276% | 1.083% | 5.513% | 12.96% | 2.039% | 2.039% | 4.564% | 5.151% | 11.61% | 3.297% | 5.467% | 10.89% | 2.44% | 1.617% | 8.804% | 0.1298% | 0.7926% | 1.001% | 1.473% | 0.9488% | 1.764% | 1.833% | 0.502% | 0.5261% | 1.738%
45 | CCV 1ppb TRACE | R | 6/2/2023 7:56:52 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.047 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 3.773% | 5.824% | 11.48% | 2.232% | 2.384% | 4.166% | 2.903% | 0.8129% | 2.839% | 2.543% | 1.82% | 1.735% | 1.081% | 2.28% | 2.974% | 1.793% | 0.7607% | 3.384% | 5.526% | 6.594% | 1.033% | 1.033% | 1.978% | 15.68% | 2.26% | 3.063% | 5.252% | 25.79% | 0.6213% | 1.985% | 0.5208% | 0.8058% | 0.5806% | 0.6088% | 1.214% | 0.795% | 2.535% | 1.792% | 0.7123% | 0.4754% | 0.8187%
46 | CCV 10 ppb MAJORS | R | 6/2/2023 8:02:55 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.048 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.2864% | 2.403% | 2.825% | 2.013% | 3.374% | 1.705% | 3.154% | 0.969% | 0.5035% | 0.7434% | 2.069% | 0.4292% | 1.245% | 0.7302% | 1.472% | 1.203% | 0.8744% | 0.8428% | 1.679% | 6.14% | 1.358% | 1.358% | 1.712% | 7.593% | 1.51% | 0.6019% | 1.381% | 4.015% | 1.37% | 2.242% | 1.593% | 1.513% | 1.337% | 0.1982% | 1.956% | 2.399% | 0.9019% | 0.8855% | 0.8094% | 0.8007% | 1.588%
47 | CCV 10 ppb MAJORS | R | 6/2/2023 8:09:27 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.049 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.466% | 2.06% | 2.532% | 0.8802% | 0.8935% | 0.6592% | 3.845% | 2.015% | 1.886% | 1.363% | 1.684% | 0.9253% | 1.734% | 1.441% | 0.4581% | 3.033% | 1.082% | 0.953% | 2.127% | 3.437% | 0.2013% | 0.2013% | 2.14% | 4.014% | 1.807% | 0.8954% | 2.981% | 4.961% | 1.208% | 0.5196% | 1.685% | 0.492% | 1.812% | 1.45% | 4.316% | 0.8608% | 1.801% | 1.791% | 0.6496% | 1.474% | 1.31%
48 | MT 5 | R | 6/2/2023 8:15:30 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 5.050 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.9138% | 7.234% | 1.072% | 0.7289% | 2.815% | 1.08% | 5.976% | 1.508% | 4.876% | 1.551% | 1.812% | 1.328% | 0.7437% | 13.83% | 4.541% | 0.7699% | 0.9641% | 1.295% | 5.65% | 2.608% | 0.7862% | 0.7862% | 1.042% | 13.94% | 8.262% | 9.457% | 9.358% | 11.92% | 3.788% | 3.699% | 0.09245% | 1.465% | 2.185% | 1.019% | 1.38% | 0.9517% | 1.248% | 2.559% | 3.152% | 1.87% | 2.088%
49 | MT 6 | R | 6/2/2023 8:21:32 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 6.051 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.238% | 2.48% | 2.369% | 1.282% | 3.243% | 1.54% | 3.082% | 2.858% | 2.081% | 0.9946% | 1.018% | 1.738% | 1.167% | 1.76% | 1.486% | 0.7356% | 1.454% | 0.916% | 2.22% | 4.496% | 0.7282% | 0.7282% | 0.9907% | 3.7% | 4.326% | 0.4265% | 3.433% | 0.9413% | 4.952% | 2.852% | 0.1585% | 0.8318% | 1.771% | 0.2038% | 3.655% | 0.5005% | 0.8938% | 6.223% | 1.068% | 0.3181% | 2.888%
50 | MT 7 | R | 6/2/2023 8:27:35 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 7.052 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.763% | 12.62% | 1.334% | 0.1295% | 1.511% | 0.8438% | 4.913% | 1.274% | 0.6516% | 4.349% | 3.596% | 0.7876% | 1.045% | 2.655% | 3.038% | 0.566% | 4.621% | 2.959% | 1.975% | 2.105% | 0.9985% | 0.9985% | 6.954% | 4.267% | 5.138% | 1.366% | 5.493% | 2.913% | 5.535% | 1.885% | 1.298% | 0.5251% | 0.5266% | 13.62% | 2.073% | 2.553% | 5.283% | 4.03% | 3.513% | 2.195% | 1.631%
51 | MT 8 | R | 6/2/2023 8:33:38 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 8.053 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.783% | 5.266% | 1.803% | 1.298% | 0.9555% | 1.728% | 10.4% | 0.4469% | 2.236% | 3.634% | 0.9715% | 1.017% | 0.7479% | 4.749% | 3.983% | 1.541% | 1.288% | 0.8724% | 1.539% | 4.267% | 1.006% | 1.006% | 4.138% | 5.829% | 1.731% | 1.132% | 4.023% | 4.843% | 1.99% | 2.759% | 1.041% | 0.941% | 8.569% | 0.6776% | 1.889% | 0.8354% | 3.476% | 8.501% | 2.795% | 2.476% | 0.8628%
52 | MT 9 | R | 6/2/2023 8:39:40 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 9.054 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.339% | 5.453% | 0.6914% | 1.404% | 1.23% | 2.035% | 10.87% | 5.555% | 1.719% | 0.8033% | 1.222% | 1.639% | 0.8781% | 6.381% | 5.166% | 0.635% | 1.055% | 1.1% | 2.1% | 8.497% | 1.599% | 1.599% | 2.031% | 5.734% | 3.895% | 5.895% | 4.231% | 9.8% | 2.183% | 4.257% | 2.996% | 2.092% | 1.329% | 0.3894% | 0.5578% | 0.8843% | 1.274% | 3.627% | 3.748% | 2.567% | 2.302%
53 | MT 10 | R | 6/2/2023 8:45:42 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 10.055 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.786% | 7.463% | 1.743% | 0.7721% | 1.597% | 0.7483% | 7.457% | 1.164% | 1.161% | 1.511% | 0.9675% | 0.8951% | 1.265% | 1.371% | 2.345% | 1.371% | 1.561% | 0.4569% | 2.955% | 9.51% | 0.3031% | 0.3031% | 1.297% | 16.64% | 0.8364% | 5.089% | 4.152% | 13.19% | 1.497% | 2.815% | 0.6311% | 0.1809% | 2.343% | 0.8768% | 1.155% | 0.2671% | 1.945% | 3.954% | 1.818% | 2.124% | 1.596%
54 | MT 11 | R | 6/2/2023 8:51:44 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 11 .056 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.292% | 2.417% | 0.8696% | 0.9949% | 0.4612% | 3.385% | 5.519% | 3.226% | 0.9179% | 0.9078% | 0.3734% | 0.6228% | 1.057% | 0.6208% | 4.548% | 0.1355% | 0.112% | 0.6646% | 1.476% | 4.952% | 1.009% | 1.009% | 3.45% | 6.69% | 4.047% | 3.63% | 0.1123% | 0.4432% | 0.6668% | 2.002% | 0.771% | 1.447% | 0.4026% | 0.4229% | 1.155% | 1.645% | 1.41% | 0.2635% | 2.26% | 2.641% | 2.196%
55 | MT 11 LD | R | 6/2/2023 8:57:47 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 11 LD.057 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.14% | 3.81% | 1.968% | 1.627% | 1.568% | 0.3415% | 5.773% | 1.212% | 2.27% | 0.9686% | 0.808% | 1.954% | 0.3364% | 1.429% | 3.638% | 1.25% | 1.115% | 0.2526% | 1.739% | 1.588% | 1.076% | 1.076% | 3.748% | 8.25% | 3.136% | 2.041% | 4.684% | 2.297% | 0.4454% | 2.087% | 0.9997% | 0.233% | 2.233% | 1.473% | 0.3604% | 1.211% | 2.723% | 2.6% | 0.8459% | 1.886% | 0.3323%
56 | MT 11 LD | R | 6/2/2023 9:04:18 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 11 LD.058 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.684% | 3.004% | 2.819% | 1.164% | 1.984% | 2.043% | 2.275% | 4.195% | 2.192% | 0.5197% | 1.008% | 0.7989% | 2.551% | 4.504% | 5.72% | 1.056% | 0.7455% | 0.6484% | 1.976% | 2.939% | 0.3084% | 0.3084% | 3.08% | 3.02% | 4.225% | 0.2338% | 7.92% | 4.438% | 1.14% | 3.332% | 1.546% | 0.8353% | 10.39% | 0.6253% | 1.273% | 1.641% | 2.827% | 1.244% | 1.294% | 2.245% | 1.939%
57 | MT 12 | R | 6/2/2023 9:10:20 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 12.059 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.5496% | 3.623% | 0.6157% | 2.164% | 4.525% | 1.617% | 0.8531% | 0.1706% | 1.797% | 5.49% | 5.594% | 0.9221% | 1.305% | 2.874% | 2.923% | 0.9037% | 5.124% | 4.817% | 2.124% | 4.733% | 0.9316% | 0.9316% | 2.042% | 3.683% | 8.742% | 0.8311% | 1.013% | 6.479% | 5.767% | 0.3908% | 4.179% | 1.193% | 4.493% | 18.38% | 2.352% | 2.59% | 0.208% | 3.545% | 3.513% | 0.891% | 0.9519%
58 | MT 12 LSP | R | 6/2/2023 9:16:23 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 12 LSP.060 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.647% | 2.793% | 0.9333% | 0.7445% | 1.651% | 1.324% | 0.9603% | 0.7212% | 1.397% | 0.5592% | 0.2072% | 1.381% | 0.2312% | 0.5382% | 1.347% | 0.2411% | 0.3454% | 0.4843% | 2.345% | 4.82% | 0.7397% | 0.7397% | 0.5504% | 4.603% | 1.447% | 1.015% | 2.711% | 4.089% | 1.247% | 1.327% | 6.188% | 0.7726% | 0.7554% | 1.518% | 0.2408% | 1.442% | 4.271% | 0.9825% | 1.003% | 1.335% | 1.149%
59 | LBLANK | R | 6/2/2023 9:22:26 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LBLANK.061 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 26.96% | 86.6% | 28.87% | 7.363% | 3.813% | 19.79% |  | 57.28% | 29.4% | 20.38% | 1.432% | 7.143% | 5.454% | 86.6% | 21.43% | 3.331% | 8.838% | 18.26% | 6.736% | 86.6% | 33.07% | 33.07% | 0.7183% | 173.2% | 89.21% | 120.2% | 82.35% | 173.2% | 21.43% | 12.37% | 9.207% | 1.278% | 1.399% | 0.9841% | 1.208% | 0.6588% | 1.331% | 107.9% | 27.94% | 37.55% | 20.15%
60 | CCV 1ppb TRACE | R | 6/2/2023 9:28:29 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.062 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.086% | 2.351% | 6.614% | 1.37% | 3.489% | 1.231% | 9.158% | 3.296% | 4.699% | 0.6164% | 0.5457% | 1.79% | 1.164% | 2.099% | 1.742% | 5.464% | 0.7171% | 0.823% | 5.082% | 11.43% | 1.572% | 1.572% | 5.386% | 9.82% | 5.862% | 0.3309% | 7.976% | 5.159% | 1.29% | 1.469% | 1.834% | 1.485% | 2.968% | 1.363% | 1.804% | 0.514% | 2.523% | 1.493% | 1.409% | 0.6665% | 1.124%
61 | CCV 1ppb TRACE | R | 6/2/2023 9:35:00 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.063 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.97% | 4.918% | 24.59% | 4.086% | 1.467% | 2.775% | 4.152% | 1.395% | 0.7926% | 2.582% | 2.648% | 0.7989% | 1.203% | 1.4% | 1.968% | 1.689% | 2.27% | 1.763% | 12.05% | 10.72% | 2.747% | 2.747% | 2.581% | 16.06% | 1.663% | 1.281% | 11.7% | 8.473% | 1.441% | 2.123% | 1.099% | 1.942% | 2.372% | 8.414% | 0.7711% | 0.6928% | 3.014% | 0.5757% | 2.388% | 0.6592% | 1.535%
62 | LFB 10ppb | R | 6/2/2023 9:41:03 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LFB 10ppb.064 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.9369% | 0.9903% | 3.725% | 1.479% | 3.527% | 1.946% | 1.499% | 2.039% | 1.384% | 0.9628% | 2.776% | 4.955% | 1.373% | 0.329% | 0.9317% | 0.7043% | 0.5675% | 0.7868% | 3.417% | 2.023% | 3.615% | 3.615% | 3.722% | 1.889% | 3.42% | 3.275% | 3.02% | 1.75% | 0.3814% | 5.314% | 4.147% | 13.4% | 2.311% | 1.59% | 1.582% | 0.8024% | 13.85% | 1.175% | 1.038% | 1.744% | 2.502%
63 | LFB 10ppb | R | 6/2/2023 9:47:33 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LFB 10ppb.065 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.262% | 1.631% | 6.138% | 1.024% | 3.576% | 0.7914% | 3.463% | 1.679% | 1.532% | 0.5644% | 1.993% | 1.632% | 0.8215% | 2.036% | 1.295% | 2.213% | 1.338% | 0.4819% | 1.334% | 4.325% | 0.307% | 0.307% | 3.729% | 2.765% | 4.094% | 0.6392% | 2.424% | 1.686% | 1.377% | 1.536% | 0.8385% | 0.9454% | 1.375% | 1.43% | 1.04% | 1.192% | 2.932% | 1.763% | 1.002% | 0.2653% | 1.404%
64 | CCV 10 ppb MAJORS | R | 6/2/2023 9:53:35 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.066 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.031% | 1.631% | 10.72% | 1.94% | 3.295% | 2.185% | 3.64% | 0.2399% | 1.333% | 3.394% | 3.37% | 1.861% | 2.157% | 1.075% | 0.7992% | 0.7479% | 2.216% | 2.137% | 3.783% | 6.106% | 1.379% | 1.379% | 3.379% | 6.683% | 4.085% | 1.667% | 2.139% | 3.688% | 4.726% | 0.2234% | 9.805% | 1.096% | 0.7243% | 13.02% | 0.8641% | 0.6459% | 0.7573% | 3.928% | 3.338% | 1.034% | 1.575%
65 | CCV 10 ppb MAJORS | R | 6/2/2023 10:00:06 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.067 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.246% | 3.295% | 5.018% | 2.433% | 2.492% | 2.579% | 3.021% | 0.7931% | 2.011% | 4.296% | 5.274% | 0.2734% | 2.036% | 3.041% | 3.791% | 2.704% | 3.173% | 3.117% | 3.56% | 5.879% | 0.3358% | 0.3358% | 3.194% | 5.75% | 1.935% | 2.573% | 1.642% | 1.185% | 3.909% | 1.362% | 1.235% | 0.3058% | 8.985% | 13.17% | 0.8413% | 9.628% | 0.7577% | 2.561% | 3.392% | 1.869% | 1.846%
66 | MT 13 | R | 6/2/2023 10:06:09 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 13.068 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.228% | 6.583% | 1.627% | 0.3389% | 3.697% | 1.835% | 3.57% | 1.261% | 1.29% | 0.4885% | 1.381% | 1.285% | 0.1778% | 1.35% | 2.361% | 0.4697% | 0.2513% | 0.5945% | 2.829% | 0.4948% | 1.018% | 1.018% | 1.891% | 3.654% | 2.496% | 2.125% | 2.632% | 7.268% | 1.386% | 0.8732% | 12.36% | 0.9991% | 1.253% | 1.454% | 1.079% | 1.472% | 4.465% | 0.7211% | 1.563% | 3.23% | 2.499%
67 | MT 14 | R | 6/2/2023 10:12:11 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 14.069 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.775% | 3.998% | 1.149% | 0.9266% | 2.162% | 0.3626% | 7.685% | 2.125% | 3.403% | 11.32% | 6.946% | 0.5611% | 0.3844% | 3.378% | 1.98% | 1.442% | 0.4511% | 0.3596% | 0.4109% | 2.534% | 1.261% | 1.261% | 5.242% | 2.535% | 1.22% | 0.8127% | 3.723% | 9.225% | 3.557% | 4.501% | 2.437% | 1.154% | 0.7913% | 0.5921% | 0.6574% | 0.9968% | 3.947% | 1.744% | 2.833% | 1.515% | 0.8928%
68 | MT 15 | R | 6/2/2023 10:18:14 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 15.070 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.278% | 5.699% | 2.308% | 3.342% | 3.511% | 1.91% | 2.688% | 4.049% | 1.478% | 0.3862% | 0.7107% | 1.488% | 0.6125% | 1.995% | 0.7899% | 1.058% | 0.9004% | 1.528% | 0.5373% | 6.653% | 0.6426% | 0.6426% | 1.297% | 10.38% | 2.09% | 1.184% | 3.225% | 11.36% | 1.706% | 2.343% | 0.8478% | 0.8153% | 1.37% | 0.3174% | 0.3711% | 0.6713% | 3.628% | 3.998% | 1.742% | 3.249% | 0.6927%
69 | MT 16 | R | 6/2/2023 10:24:17 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 16.071 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.021% | 1.131% | 5.744% | 0.5417% | 3.703% | 1.161% | 4.937% | 4.818% | 1.602% | 1.497% | 0.5193% | 0.2732% | 0.4785% | 2.357% | 5.133% | 0.2562% | 1.157% | 0.4057% | 3.519% | 5.437% | 1.41% | 1.41% | 4.306% | 6.052% | 2.565% | 4.694% | 7.579% | 4.792% | 1.648% | 1.133% | 13.4% | 0.8907% | 0.5334% | 0.9951% | 0.3905% | 0.4574% | 3.255% | 1.419% | 1.695% | 1.298% | 0.6995%
70 | MT 17 | R | 6/2/2023 10:30:20 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 17.072 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.424% | 3.392% | 1.467% | 0.6855% | 2.072% | 0.8474% | 3.418% | 7.857% | 1.924% | 1.21% | 0.6984% | 2.797% | 0.5974% | 0.993% | 4.17% | 1.107% | 0.9058% | 0.6696% | 0.4783% | 5.049% | 2.261% | 2.261% | 3.368% | 2.826% | 0.6379% | 2.46% | 2.289% | 1.727% | 0.02917% | 1.828% | 1.278% | 7.663% | 0.6128% | 1.382% | 0.6488% | 2.926% | 7.486% | 1.05% | 0.7942% | 2.857% | 0.1036%
71 | MT 18 | R | 6/2/2023 10:36:32 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 18.073 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.083% | 4.365% | 1.562% | 1.807% | 2.603% | 1.646% | 12.88% | 3.574% | 1.94% | 1.792% | 1.164% | 1.594% | 0.8649% | 3.671% | 0.8355% | 1.296% | 1.494% | 1.17% | 1.14% | 5.886% | 0.3901% | 0.3901% | 6.464% | 4.984% | 2.394% | 1.894% | 1.453% | 4.182% | 1.1% | 1.059% | 0.6094% | 0.8108% | 2.343% | 0.8281% | 0.9725% | 3.367% | 0.6683% | 0.7815% | 0.9507% | 2.664% | 1.264%
72 | MT 19 | R | 6/2/2023 10:42:43 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 19.074 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.229% | 2.481% | 4.741% | 0.7684% | 0.6087% | 0.9787% | 16.42% | 9.501% | 14.55% | 5.659% | 3.956% | 1.781% | 1.043% | 9.041% | 1.152% | 1.017% | 2.03% | 2.617% | 0.7152% | 6.726% | 1.52% | 1.52% | 4.895% | 9.006% | 1.547% | 4.686% | 2.081% | 1.862% | 3.651% | 2.171% | 0.7626% | 0.952% | 2.992% | 16.28% | 0.8677% | 0.7387% | 0.3675% | 1.866% | 3.27% | 1.146% | 2.398%
73 | MT 20 | R | 6/2/2023 10:48:45 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 20.075 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.08% | 6.328% | 5.619% | 1.651% | 1.748% | 1.991% | 6.27% | 3.692% | 6.079% | 3.293% | 1.199% | 0.145% | 0.6739% | 7.428% | 2.905% | 1.643% | 1.366% | 1% | 0.9836% | 5.955% | 0.102% | 0.102% | 3.924% | 3.414% | 0.2529% | 3.583% | 4.04% | 1.606% | 1.104% | 1.019% | 0.6466% | 1.382% | 1.375% | 0.476% | 1.106% | 1.98% | 3.274% | 0.7216% | 0.7798% | 2.055% | 0.7813%
74 | MT 21 | R | 6/2/2023 10:54:57 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 21.076 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.653% | 4.516% | 4.629% | 0.9688% | 2.488% | 3.099% | 2.785% | 3.615% | 1.684% | 6.059% | 1.044% | 0.5037% | 0.6803% | 3.063% | 1.221% | 1.331% | 0.5132% | 0.4261% | 1.031% | 2.726% | 1.141% | 1.141% | 1.001% | 1.145% | 3.142% | 3.44% | 5.69% | 2.972% | 2.646% | 0.7253% | 4.339% | 3.162% | 0.9335% | 1.814% | 9.275% | 0.7241% | 4.421% | 0.7485% | 1.86% | 2.098% | 0.6448%
75 | MT 21 LD | R | 6/2/2023 11:00:59 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 21 LD.077 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 119.5% |  | 173.2% | 8.395% | 18.5% | 75.17% |  | 173.2% | 66.81% | 15.23% | 2.438% | 7.531% | 95.18% | 43.3% | 18.23% | 126.8% | 27.03% | 4.767% | 3.357% | 86.6% | 28.87% | 28.87% | 4.325% | 173.2% | 69.28% | 107.9% | 481.9% |  | 6.214% | 58.08% | 4.263% | 1.299% | 3.683% | 1.791% | 2.635% | 5.031% | 1.434% | 86.6% | 83.86% | 68.88% | 9.897%
76 | LBLANK | R | 6/2/2023 11:07:00 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LBLANK.078 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 62.45% |  | 44.61% | 3.776% | 10.32% | 31.23% | 173.2% | 114.6% | 40.56% | 27.06% | 2.691% | 21.65% | 2.783% | 100% | 5.973% | 16.9% | 17.34% | 4.969% | 6.11% |  | 21.65% | 21.65% | 4.966% | 86.6% | 100% | 88.19% | 23.57% |  | 16.39% | 28.64% | 0.7277% | 2.293% | 9.042% | 1.143% | 1.085% | 1.805% | 1.055% | 34.64% | 7.873% | 56.97% | 22.27%
77 | CCV 1ppb TRACE | R | 6/2/2023 11:13:03 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.079 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.65% | 1.832% | 15.74% | 2.482% | 0.7966% | 4.562% | 7.157% | 1.154% | 2.56% | 0.3366% | 1.435% | 1.431% | 1.743% | 2.982% | 0.8983% | 0.4277% | 1.069% | 2.691% | 13.9% | 11.47% | 1.433% | 1.433% | 3.051% | 9.28% | 5.68% | 3.651% | 6.317% | 4.193% | 1.342% | 0.6711% | 3.006% | 0.9754% | 0.5034% | 6.526% | 2.304% | 0.2884% | 5.331% | 1.016% | 2.643% | 0.9651% | 2.877%
78 | CCV 1ppb TRACE | R | 6/2/2023 11:19:35 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.080 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.535% | 2.846% | 7.977% | 8.343% | 3.806% | 1.303% | 3.434% | 3.683% | 1.954% | 1.294% | 2.338% | 0.9727% | 5.539% | 2.163% | 3.156% | 3.701% | 0.5377% | 1.196% | 9.878% | 14.49% | 0.8333% | 0.8333% | 3.192% | 23.07% | 8.547% | 3.501% | 8.925% | 6.039% | 1.097% | 3.482% | 2.323% | 1.079% | 1.707% | 0.8071% | 3.131% | 0.7261% | 0.7534% | 1.438% | 2.873% | 1.143% | 0.9476%
79 | CCV 10 ppb MAJORS | R | 6/2/2023 11:25:38 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.081 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.6368% | 1.958% | 3.604% | 1.74% | 2.404% | 1.275% | 2.292% | 1.263% | 0.9574% | 1.16% | 0.8045% | 1.797% | 0.5471% | 0.5958% | 0.6587% | 1.03% | 1.361% | 2.103% | 0.8304% | 0.9067% | 1.723% | 1.723% | 3.342% | 4.198% | 2.283% | 0.4636% | 3.231% | 1.943% | 0.5216% | 1.413% | 0.8125% | 0.7045% | 0.6476% | 1.288% | 0.5617% | 1.271% | 3.113% | 1.82% | 1.495% | 0.99% | 1.157%
80 | CCV 10 ppb MAJORS | R | 6/2/2023 11:32:09 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.082 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.855% | 2.414% | 8.688% | 2.097% | 2.246% | 2.793% | 3.706% | 1.056% | 1.293% | 1.47% | 0.7395% | 2.768% | 1.274% | 1.109% | 1.653% | 1.203% | 0.8698% | 0.4232% | 1.642% | 0.995% | 0.8432% | 0.8432% | 1.943% | 8.379% | 1.652% | 0.9556% | 4.478% | 5.86% | 0.5934% | 0.8603% | 0.2871% | 1.117% | 2.235% | 0.5438% | 1.721% | 1.967% | 0.5508% | 1.016% | 0.741% | 1.056% | 2.005%
81 | MT 22 | R | 6/2/2023 11:38:12 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 22.083 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.9454% | 7.501% | 4.851% | 0.4148% | 2.63% | 2.258% | 6.679% | 6.696% | 2.913% | 5.585% | 1.246% | 2.118% | 0.87% | 6.662% | 0.4729% | 1.855% | 1.078% | 1.338% | 1.722% | 1.876% | 0.597% | 0.597% | 7.143% | 6.493% | 0.5847% | 1.699% | 3.795% | 7.73% | 2.689% | 0.2164% | 0.7876% | 1.307% | 1.936% | 0.6126% | 3.071% | 2.294% | 3.163% | 1.081% | 0.2986% | 0.7397% | 3.612%
82 | MT 22 LSP | R | 6/2/2023 11:44:15 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 22 LSP.084 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.443% | 1.066% | 5.354% | 2.276% | 1.193% | 2.168% | 3.033% | 0.8238% | 5.942% | 0.3541% | 1.005% | 0.5724% | 0.9818% | 2.125% | 2.319% | 0.1522% | 0.1134% | 0.85% | 0.9268% | 2.548% | 1.89% | 1.89% | 4.775% | 8.386% | 2.586% | 1.85% | 0.9216% | 2.256% | 0.4113% | 0.9205% | 0.6553% | 0.3218% | 0.9114% | 1.679% | 1.875% | 0.8518% | 1.923% | 0.6942% | 0.6496% | 0.5314% | 2.749%
83 | MT 22 LSP | R | 6/2/2023 11:50:45 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 22 LSP.085 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.951% | 2.546% | 3.037% | 0.6838% | 4.181% | 2.074% | 2.849% | 2.43% | 3.579% | 0.9988% | 1.352% | 0.3306% | 3.651% | 3.178% | 2.906% | 2.284% | 1.349% | 0.6927% | 1.995% | 6.003% | 0.5055% | 0.5055% | 1.561% | 2.601% | 1.988% | 1.466% | 4.563% | 2.204% | 1.674% | 0.7786% | 1.837% | 0.5864% | 10.06% | 1.614% | 0.776% | 1.682% | 2.474% | 1.572% | 0.6415% | 0.6471% | 2.439%
84 | MT 23 | R | 6/2/2023 11:56:47 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 23.086 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.932% | 7.133% | 7.407% | 3.469% | 1.635% | 1.652% | 2.5% | 16.15% | 4.088% | 2.34% | 1.418% | 2.247% | 1.11% | 5.427% | 2.555% | 1.888% | 1.193% | 1.265% | 3.276% | 5.904% | 1.233% | 1.233% | 0.1922% | 6.978% | 4.765% | 2.77% | 3.48% | 12.1% | 1.649% | 1.441% | 0.974% | 2.535% | 0.398% | 1.489% | 18.35% | 2.303% | 6.36% | 0.6666% | 2.735% | 2.001% | 5.521%
85 | MT 24 | R | 6/3/2023 12:02:49 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 24.087 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 4.894% | 9.608% | 9.201% | 1.466% | 5.643% | 3.692% | 9.352% | 10.97% | 3.799% | 5.022% | 1.067% | 0.8426% | 1.39% | 1.419% | 8.61% | 1.788% | 0.5432% | 0.7906% | 3.268% | 19.17% | 0.7133% | 0.7133% | 5.429% | 10.62% | 3.797% | 3.429% | 4.756% | 13.55% | 0.8131% | 1.764% | 1.66% | 1.933% | 1.351% | 0.139% | 0.1458% | 1.582% | 1.749% | 1.494% | 0.6537% | 1.397% | 0.4372%
86 | MT 25 | R | 6/3/2023 12:08:52 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 25.088 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.3857% | 4.693% | 2.996% | 2.573% | 2.332% | 1.307% | 8.137% | 10.23% | 1.839% | 2.764% | 0.4519% | 1.581% | 1.403% | 6.603% | 5.214% | 1.681% | 0.8668% | 1.649% | 3.656% | 7.531% | 0.4649% | 0.4649% | 1.194% | 9.655% | 7.898% | 4.691% | 4.309% | 10.92% | 2.543% | 1.698% | 5.11% | 1.511% | 2.406% | 1.124% | 1.667% | 0.9929% | 2.834% | 0.7511% | 0.5926% | 4.236% | 1.332%
87 | MT 26 | R | 6/3/2023 12:14:55 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 26.089 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.965% | 11.01% | 3.39% | 2.461% | 2.865% | 2.75% | 18.74% | 12.62% | 9.686% | 0.8779% | 0.6354% | 0.8809% | 0.9513% | 3.603% | 5.711% | 0.726% | 2.242% | 0.5215% | 5.406% | 9.803% | 1.862% | 1.862% | 5.759% | 11.69% | 6.623% | 2.031% | 9.337% | 15.73% | 2.332% | 1.132% | 21.85% | 1.883% | 1.35% | 1.088% | 4.918% | 2.001% | 1.104% | 2.904% | 0.9232% | 4.443% | 3.169%
88 | MT 27 | R | 6/3/2023 12:21:03 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 27.090 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.571% | 2.598% | 4.347% | 2.441% | 1.332% | 2.67% | 5.042% | 2.736% | 3.185% | 2.456% | 0.4672% | 1.326% | 0.5612% | 2.73% | 2.711% | 0.4689% | 0.1997% | 0.5711% | 2.324% | 3.932% | 0.3184% | 0.3184% | 4.889% | 4.371% | 5.207% | 2.567% | 6.231% | 6.537% | 0.9857% | 0.3836% | 2.21% | 0.2049% | 1.669% | 0.7935% | 1.777% | 0.1771% | 4.994% | 1.326% | 1.181% | 2.98% | 1.676%
89 | MT 28 | R | 6/3/2023 12:27:06 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 28.091 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.19% | 3.051% | 2.661% | 2.093% | 3.092% | 1.511% | 7.418% | 8.108% | 1.577% | 1.303% | 0.6496% | 3.081% | 1.352% | 5.22% | 1.846% | 0.5413% | 0.811% | 0.4759% | 2.7% | 6.831% | 2.976% | 2.976% | 2.798% | 13.11% | 0.9889% | 4.01% | 5.346% | 2.246% | 0.839% | 2.395% | 4.661% | 13.76% | 4.393% | 1.624% | 1.204% | 1.882% | 14.63% | 1.496% | 0.7597% | 2.916% | 1.632%
90 | MT 29 | R | 6/3/2023 12:33:08 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 29.092 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.878% | 3.326% | 2.877% | 1.634% | 0.8688% | 1.507% | 12.46% | 9.158% | 1.779% | 2.353% | 0.496% | 1.11% | 1.607% | 2.544% | 1.198% | 0.6894% | 0.9033% | 1.165% | 1.278% | 2.038% | 1.29% | 1.29% | 2.084% | 3.61% | 3.396% | 1.074% | 1.501% | 2.744% | 0.3725% | 0.9863% | 1.462% | 1.066% | 2.094% | 2.753% | 16.85% | 1.757% | 5.311% | 0.7318% | 1.07% | 1.384% | 5.278%
91 | LBLANK | R | 6/3/2023 12:39:20 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LBLANK.093 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 57.74% |  | 31.49% | 2.882% | 5.914% | 17.16% | 86.6% | 50% | 51.63% | 43.3% | 5.339% | 14.29% | 1.208% | 86.6% | 24.02% | 10.66% | 28.19% | 5.09% | 9.576% | 43.3% | 17.32% | 17.32% | 4.249% |  | 24.74% | 91.65% | 41.72% | 173.2% | 9.442% | 44.41% | 0.7584% | 0.9775% | 1.084% | 2.596% | 3.789% | 0.669% | 0.5209% | 25% | 56.71% | 19.79% | 14.75%
92 | CCV 1ppb TRACE | R | 6/3/2023 12:45:23 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.094 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 4.358% | 4.343% | 30.13% | 1.894% | 3.776% | 2.752% | 2.552% | 2.929% | 0.5586% | 2.984% | 1.565% | 2.057% | 1.758% | 4.38% | 3.091% | 1.158% | 2.802% | 1.667% | 9.832% | 8.473% | 1.905% | 1.905% | 3.485% | 8.942% | 5.765% | 3.465% | 13.1% | 7.194% | 1.744% | 0.4928% | 1.268% | 8.323% | 1.44% | 4.163% | 4.235% | 1.407% | 10.07% | 0.7788% | 1.675% | 0.9769% | 1.394%
93 | CCV 1ppb TRACE | R | 6/3/2023 12:51:55 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.095 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.855% | 3.712% | 12.93% | 2.405% | 1.918% | 2.032% | 4.801% | 3.015% | 5.942% | 0.9272% | 2.625% | 2.499% | 1.715% | 0.8374% | 0.9446% | 4.658% | 1.76% | 2.907% | 14.9% | 8.22% | 2.228% | 2.228% | 6.584% | 5.328% | 6.172% | 3.346% | 2.89% | 7.613% | 1.757% | 0.9311% | 2.273% | 0.9341% | 1.173% | 1.371% | 0.5272% | 1.209% | 3.074% | 0.9572% | 0.9699% | 1.628% | 0.7695%
94 | LFB 10ppb | R | 6/3/2023 12:57:57 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LFB 10ppb.096 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.466% | 0.9427% | 4.901% | 1.289% | 5.912% | 0.9468% | 0.8787% | 0.7379% | 0.7799% | 2.497% | 0.8443% | 1.767% | 3.072% | 0.3561% | 2.533% | 0.7161% | 0.8378% | 0.6616% | 1.444% | 4.354% | 3.318% | 3.318% | 2.726% | 4.484% | 3.259% | 2.67% | 5.658% | 1.445% | 0.6517% | 2.823% | 2.397% | 15.87% | 5.813% | 0.4013% | 0.8652% | 1.072% | 13.21% | 0.6887% | 0.673% | 1.482% | 1.186%
95 | LFB 10ppb | R | 6/3/2023 1:04:27 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LFB 10ppb.097 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.357% | 2.094% | 5.119% | 0.8729% | 1.725% | 1.669% | 3.541% | 2.184% | 2.01% | 0.5793% | 1.42% | 1.324% | 0.6145% | 0.7937% | 1.438% | 1.366% | 0.4411% | 0.1702% | 0.4875% | 1.801% | 1.353% | 1.353% | 2.042% | 5.859% | 3.706% | 0.974% | 0.9475% | 5.949% | 0.5176% | 0.9856% | 1.132% | 0.8321% | 2.777% | 0.9989% | 0.9717% | 1.255% | 1.482% | 2.337% | 0.9856% | 0.9081% | 1.602%
96 | CCV 10 ppb MAJORS | R | 6/3/2023 1:10:29 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.098 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.907% | 2.483% | 4.371% | 3.274% | 1.769% | 3.78% | 3.493% | 1.92% | 0.6449% | 1.001% | 3.375% | 0.4725% | 0.9523% | 0.9139% | 1.898% | 1.862% | 1.718% | 2.02% | 2.714% | 1.715% | 0.8762% | 0.8762% | 2.263% | 3.623% | 3.34% | 1.426% | 2.614% | 0.8151% | 0.8685% | 1.214% | 0.3343% | 0.3827% | 1.653% | 2.011% | 1.12% | 2.201% | 5.925% | 1.105% | 1.262% | 1.286% | 1.406%
97 | CCV 10 ppb MAJORS | R | 6/3/2023 1:17:01 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.099 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.103% | 3.199% | 5.223% | 1.411% | 2.024% | 1.54% | 1.164% | 2.984% | 2.063% | 1.647% | 1.108% | 0.3123% | 1.691% | 0.5304% | 0.6834% | 0.4405% | 0.403% | 1.303% | 2.616% | 0.2055% | 1.025% | 1.025% | 1.998% | 4.384% | 3.227% | 2.016% | 2.602% | 4.104% | 1.54% | 0.5724% | 6.855% | 0.9804% | 0.7228% | 1.684% | 0.4096% | 0.9582% | 4.193% | 0.4693% | 1.647% | 1.385% | 1.51%
98 | MT 30 | R | 6/3/2023 1:23:04 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 30.100 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 15.75% | 7.143% | 1.553% | 1.364% | 1.275% | 6.323% | 5.427% | 1.093% | 4.029% | 3.597% | 3.798% | 1.897% | 1.113% | 4.135% | 2.175% | 2.422% | 1.174% | 0.6956% | 1.847% | 6.236% | 2.736% | 2.736% | 3.219% | 14.84% | 2.547% | 12.85% | 7.129% | 23.74% | 3.139% | 4.74% | 0.4139% | 1.587% | 1.933% | 4.209% | 0.3621% | 5.189% | 3.401% | 2.72% | 0.3652% | 1.159% | 0.8444%
99 | MT 31 | R | 6/3/2023 1:29:15 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 31.101 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 8.431% | 4.515% | 2.127% | 4.185% | 5.121% | 3.603% | 6.506% | 1.87% | 8.588% | 1.438% | 1.117% | 0.3743% | 0.1964% | 2.74% | 4.774% | 1.107% | 1.128% | 1.008% | 4.088% | 8.678% | 1.367% | 1.367% | 7.662% | 7.692% | 5.679% | 8.794% | 17.93% | 22.72% | 1.817% | 5.616% | 1.469% | 2.868% | 1.476% | 5.268% | 1.575% | 1.621% | 5.433% | 1.043% | 0.7053% | 0.1532% | 0.4688%
100 | MT 31 LD | R | 6/3/2023 1:35:29 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 31 LD.102 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 8.046% | 2.96% | 1.346% | 2.028% | 6.559% | 3.861% | 5.101% | 2.552% | 8.324% | 0.8566% | 0.9851% | 0.2042% | 1.518% | 3.166% | 2.803% | 1.185% | 1.126% | 0.8467% | 10.71% | 13.74% | 0.7759% | 0.7759% | 2.863% | 1.475e-06% | 7.211% | 7.862% | 15.48% | 20.83% | 1.418% | 4.132% | 1.459% | 0.4163% | 3.757% | 1.122% | 0.5867% | 0.6212% | 1.875% | 1.403% | 0.3656% | 0.475% | 0.519%
101 | MT 31 LD | R | 6/3/2023 1:42:00 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 31 LD.103 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 11.46% | 1.117% | 1.991% | 0.4366% | 5.972% | 1.98% | 12.44% | 3.91% | 2.426% | 1.636% | 1.75% | 0.7087% | 0.3929% | 3.025% | 8.037% | 2.157% | 1.112% | 0.9989% | 9.263% | 13.13% | 3.719% | 3.719% | 4.826% | 15.23% | 1.818% | 6.991% | 26.63% | 27.24% | 2.693% | 0.5195% | 1.708% | 0.6476% | 4.019% | 0.7374% | 0.9438% | 0.5114% | 3.925% | 1.094% | 0.728% | 0.4363% | 2.798%
102 | MT 32 | R | 6/3/2023 1:48:02 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 32.104 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 56.77% | 86.6% | 18.23% | 1.019% | 1.191% | 34.59% | 34.64% | 173.2% | 18.03% | 5.384% | 5.605% | 2.632% | 3.441% | 100% | 12.37% | 2.597% | 3.425% | 2.44% | 9.206% |  | 24.02% | 24.02% | 5.264% | 86.6% | 173.2% | 43.3% | 53.74% | 173.2% | 9.138% | 8.723% | 0.8806% | 1.86% | 1.304% | 3.723% | 7.835% | 2.144% | 1.085% | 173.2% | 26.46% | 1.46% | 5.403%
103 | MT 32 LSP | R | 6/3/2023 1:54:03 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 32 LSP.105 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.061% | 0.5095% | 8.362% | 1.319% | 1.135% | 0.8971% | 1.971% | 1.515% | 2.217% | 0.7054% | 2.859% | 0.6178% | 0.2706% | 1.421% | 2.478% | 0.4149% | 1.111% | 0.3306% | 2.339% | 0.566% | 1.155% | 1.155% | 1.411% | 7.755% | 3.778% | 1.869% | 4.413% | 3.196% | 0.4247% | 0.4481% | 1.062% | 1.271% | 0.8434% | 0.7635% | 2.126% | 4.725% | 2.145% | 0.2829% | 0.4331% | 0.837% | 1.831%
104 | MT 32 LSP | R | 6/3/2023 2:00:34 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 32 LSP.106 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.499% | 0.9865% | 4.131% | 1.252% | 1.249% | 2.596% | 0.7449% | 1.376% | 1.036% | 1.57% | 1.592% | 2.764% | 0.6726% | 1.796% | 0.8746% | 1.597% | 2.665% | 1.149% | 4.516% | 2.146% | 1.123% | 1.123% | 3.614% | 4.86% | 1.533% | 2.831% | 7.679% | 3.949% | 2.437% | 0.7757% | 10.81% | 1.89% | 1.409% | 6.059% | 1.08% | 1.811% | 4.381% | 1.069% | 1.856% | 2.414% | 0.9622%
105 | MT 33 | R | 6/3/2023 2:06:37 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 33.107 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 83.4% | 86.6% | 5.973% | 1.928% | 6.345% | 16.88% | 43.3% | 100% | 33.86% | 5.808% | 8.477% | 7.782% | 4.823% | 44.1% | 25.64% | 4.277% | 3.444% | 23.48% | 4.64% |  | 34.64% | 34.64% | 1.245% |  | 43.3% | 114.6% | 97.99% | 173.2% | 1.353% | 16.06% | 0.4069% | 0.386% | 1.525% | 18.74% | 1.845% | 0.9518% | 0.1119% | 91.65% | 11.17% | 5.736% | 3.115%
106 | MT 34 | R | 6/3/2023 2:12:39 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 34.108 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 33.33% |  | 8.06% | 0.9049% | 0.9908% | 22.05% | 69.28% | 100% | 13.78% | 15.56% | 0.3859% | 5.375% | 2.083% | 66.67% | 18.7% | 24.17% | 7.576% | 10.61% | 2.199% | 173.2% | 52.71% | 52.71% | 3.535% | 173.2% | 50% | 43.3% | 13.34% | 173.2% | 11.71% | 35.25% | 12.01% | 0.2628% | 2.413% | 0.8015% | 2.602% | 1.5% | 2.47% | 114.6% | 5.094% | 5.105% | 6.686%
107 | MT 35 | R | 6/3/2023 2:18:42 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 35.109 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 43.3% | 173.2% | 12.49% | 0.6393% | 6.307% | 28.47% | 43.3% | 86.6% | 27.09% | 20.86% | 0.8518% | 26.29% | 1.298% | 33.33% | 31.46% | 8.217% | 4.839% | 13.4% | 4.331% | 173.2% | 0% | 0% | 3.417% | 100% | 100% | 86.6% | 32.39% |  | 23.4% | 24.98% | 7.479% | 4.168% | 1.385% | 7.165% | 2.089% | 2.639% | 4.96% | 21.65% | 46.65% | 5.538% | 4.667%
108 | LBLANK | R | 6/3/2023 2:24:45 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LBLANK.110 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 137.8% |  | 28.64% | 4.414% | 1.667% | 9.165% | 86.6% | 100% | 22.35% | 22.21% | 5.356% | 57.74% | 3.188% | 0% | 27.7% | 70.89% | 10.27% | 1.778% | 2.309% |  | 34.64% | 34.64% | 2.962% |  | 114.6% | 173.2% | 37.12% |  | 24.05% | 15.75% | 0.5983% | 1.21% | 0.8168% | 0.6004% | 7.613% | 1.681% | 4.429% | 173.2% | 32.87% | 32.78% | 13.68%
109 | CCV 1ppb TRACE | R | 6/3/2023 2:30:48 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.111 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.012% | 3.936% | 20.74% | 1.966% | 3.298% | 3.166% | 6.944% | 2.586% | 2.582% | 3.967% | 1.347% | 1.751% | 1.959% | 0.3005% | 2.116% | 2.938% | 3.905% | 3.08% | 9.005% | 19.45% | 0.3344% | 0.3344% | 3.675% | 2.585% | 4.482% | 2.596% | 3.941% | 10.15% | 1.967% | 1.563% | 0.601% | 2.968% | 2.31% | 8.74% | 5.73% | 1.739% | 0.9405% | 2.557% | 2.681% | 0.9218% | 1.06%
110 | CCV 1ppb TRACE | R | 6/3/2023 2:37:20 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.112 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.698% | 2.222% | 13.3% | 3.023% | 6.182% | 5.194% | 1.932% | 7.042% | 2.207% | 2.039% | 2.761% | 1.777% | 2.151% | 0.9949% | 3.055% | 3% | 0.6905% | 1.554% | 21.04% | 7.05% | 0.468% | 0.468% | 1.108% | 5.885% | 0.5942% | 4.044% | 4.699% | 1.533% | 1.705% | 1.306% | 2.934% | 1.181% | 3.025% | 1.702% | 3.526% | 23.81% | 1.665% | 2.146% | 0.6962% | 1.841% | 1.173%
111 | CCV 10 ppb MAJORS | R | 6/3/2023 2:43:23 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.113 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.113% | 1.508% | 9.147% | 0.7343% | 0.7535% | 2.582% | 1.211% | 1.154% | 0.8448% | 0.977% | 1.219% | 1.733% | 2.545% | 1.226% | 1.042% | 1.633% | 0.5156% | 1.146% | 1.042% | 7.099% | 1.366% | 1.366% | 2.896% | 1.19% | 1.125% | 2.571% | 2.148% | 1.77% | 0.5359% | 0.2768% | 1.635% | 2.953% | 3.071% | 0.2421% | 1.407% | 0.7519% | 5.327% | 1.134% | 0.6707% | 1.021% | 1.093%
112 | CCV 10 ppb MAJORS | R | 6/3/2023 2:49:56 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.114 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.7547% | 2.318% | 0.274% | 0.9924% | 1.328% | 1.643% | 2.073% | 1.586% | 0.9606% | 1.079% | 1.103% | 1.718% | 1.288% | 0.5664% | 0.5662% | 0.6975% | 1.115% | 1.149% | 3.016% | 5.141% | 0.6463% | 0.6463% | 1.809% | 3.483% | 1.873% | 2.996% | 3.009% | 6.981% | 1.083% | 0.2403% | 8.621% | 1.044% | 3.708% | 2.385% | 0.6965% | 1.192% | 1.563% | 2.39% | 0.4106% | 2.606% | 1.254%
113 | MT 36 | R | 6/3/2023 2:55:59 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 36.115 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 48.43% | 100% | 14.36% | 1.661% | 1.679% | 17.32% | 81.04% | 50% | 13.48% | 3.488% | 2.195% | 8.476% | 2.009% | 68.63% | 29.05% | 27.32% | 4.904% | 5.412% | 3.215% | 173.2% | 9.116% | 9.116% | 6.42% |  | 86.6% | 173.2% | 70.99% |  | 20.67% | 22.72% | 0.8226% | 0.5796% | 3.575% | 0.7782% | 1.6% | 1.374% | 2.5% | 100% | 15.75% | 8.348% | 6.266%
114 | MT 37 | R | 6/3/2023 3:02:01 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 37.116 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 15.75% | 34.64% | 20.71% | 0.5193% | 3.949% | 20.15% | 12.37% | 173.2% | 29.5% | 37.83% | 2.931% | 15.9% | 1.242% | 100% | 16.35% | 14.3% | 3.313% | 1.958% | 3.985% | 173.2% | 64.95% | 64.95% | 3.089% | 173.2% | 173.2% | 57.28% | 183% |  | 17.29% | 20% | 4.949% | 1.962% | 5.157% | 1.024% | 0.7732% | 2.17% | 2.238% | 34.64% | 25.38% | 6.511% | 7.832%
115 | MT 38 | R | 6/3/2023 3:08:04 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 38.117 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 25% | 86.6% | 16.08% | 4.651% | 3.448% | 3.685% | 91.65% | 43.3% | 16.21% | 17.6% | 3.046% | 27.14% | 2.152% | 50% | 27.06% | 9.547% | 6.822% | 6.281% | 0.8647% | 173.2% | 19.92% | 19.92% | 4.503% |  | 65.47% | 86.6% | 68.75% |  | 20.28% | 29.4% | 0.8482% | 1.332% | 2.295% | 0.8309% | 5.724% | 4.168% | 4.579% | 173.2% | 16.67% | 7.501% | 1.901%
116 | MT 39 | R | 6/3/2023 3:14:07 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 39.118 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 4.681% | 86.6% | 3.978% | 2.315% | 3.025% | 10.83% | 45.83% | 86.6% | 31.77% | 7.393% | 2.683% | 13.71% | 1.727% | 98.97% | 24.74% | 9.984% | 5.963% | 3.646% | 4.618% | 173.2% | 16.67% | 16.67% | 2.389% | 173.2% | 91.65% | 86.6% | 165.2% | 100% | 20.83% | 5.973% | 19.11% | 3.826% | 2.82% | 0.4397% | 0.3909% | 1.465% | 3.923% | 86.6% | 13.32% | 13.27% | 6.809%
117 | MT 40 | R | 6/3/2023 3:20:10 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 40.119 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 15.1% | 50% | 9.249% | 2.551% | 4.993% | 17.32% | 17.56% | 173.2% | 37.55% | 28.78% | 7.112% | 3.713% | 1.101% | 43.3% | 8.333% | 10.16% | 6.087% | 7.147% | 3.248% | 86.6% | 24.35% | 24.35% | 6.165% |  | 173.2% | 69.28% | 24.55% | 173.2% | 6.544% | 11.36% | 10.63% | 1.623% | 2.306% | 1.998% | 0.6156% | 1.489% | 0.5008% | 173.2% | 8.66% | 4.502% | 4.579%
118 | MT 31 | R | 6/3/2023 3:26:12 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 31 .120 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 5.903% | 11.13% | 2.098% | 1.154% | 3.109% | 8.153% | 4.541% | 3.838% | 4.1% | 6.658% | 4.633% | 0.5928% | 0.5083% | 5.719% | 3.249% | 1.086% | 3.845% | 2.755% | 3.826% | 1.431% | 0.9953% | 0.9953% | 8.107% | 29.06% | 13.17% | 16.02% | 11.23% | 22.53% | 8.435% | 3.81% | 2.015% | 1.575% | 0.7747% | 12.16% | 0.8676% | 1.007% | 0.9116% | 4.324% | 3.242% | 0.7248% | 0.3427%
119 | MT 31 LD | R | 6/3/2023 3:32:23 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 31 LD.121 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2% | 19.11% | 3.666% | 1.661% | 3.471% | 6.872% | 9.851% | 3.055% | 5.681% | 2.164% | 2.565% | 0.6015% | 3.457% | 1.258% | 7.527% | 2.215% | 2.644% | 1.95% | 3.643% | 7.435% | 2.354% | 2.354% | 2.833% | 37.5% | 17.34% | 11.76% | 31.05% | 13.04% | 3.137% | 1.344% | 1.263% | 2.565% | 8.462% | 5.954% | 6.666% | 0.5211% | 2.805% | 2.483% | 2.249% | 0.5733% | 3.67%
120 | MT 31 LD | R | 6/3/2023 3:38:54 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 31 LD.122 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 9.219% | 8.532% | 2.505% | 1.999% | 5.442% | 4.276% | 11.64% | 8.574% | 2.229% | 6.783% | 2.085% | 2.374% | 0.474% | 1.278% | 9.666% | 1.466% | 2.538% | 2.871% | 5.423% | 15.01% | 2.01% | 2.01% | 5.136% | 31.77% | 8.915% | 10.2% | 4.62% | 21.79% | 1.357% | 2.696% | 0.8365% | 0.5805% | 1.543% | 7.989% | 6.517% | 0.6336% | 2.476% | 4.771% | 3.098% | 2.216% | 2.69%
121 | MT 32 | R | 6/3/2023 3:44:56 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 32.123 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 14.29% |  | 7.332% | 3.004% | 5.572% | 19.25% | 43.3% | 173.2% | 25.2% | 1.468% | 3.782% | 16.01% | 1.622% | 91.65% | 29.57% | 12.37% | 3.183% | 1.802% | 4.105% |  | 57.28% | 57.28% | 1.62% | 173.2% | 173.2% | 43.3% | 22.13% |  | 16.72% | 2.936% | 1.818% | 0.7765% | 1.688% | 1.42% | 0.739% | 0.6329% | 3.134% | 100% | 59.01% | 7.19% | 3.487%
122 | MT 32 LSP | R | 6/3/2023 3:50:58 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 32 LSP.124 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 3.143% | 0.7537% | 7.374% | 2.687% | 3.145% | 1.189% | 0.4574% | 1.317% | 1.46% | 0.8197% | 1.196% | 1.046% | 0.1799% | 1.482% | 0.6842% | 1.017% | 1.437% | 0.101% | 1.255% | 8.38% | 1.736% | 1.736% | 1.227% | 11.84% | 0.7017% | 1.476% | 5.444% | 8.099% | 0.8401% | 0.7859% | 0.2819% | 2.234% | 1.325% | 1.038% | 7.669% | 12.56% | 4.573% | 1.341% | 0.6529% | 0.1447% | 0.7617%
123 | MT 32 LSP | R | 6/3/2023 3:57:29 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 32 LSP.125 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.7985% | 1.672% | 5.08% | 1.87% | 1.384% | 1.872% | 1.207% | 1.228% | 2.189% | 1.126% | 0.8648% | 0.7358% | 0.2487% | 1.219% | 1.201% | 1.785% | 0.3553% | 0.6002% | 2.779% | 6.77% | 1.349% | 1.349% | 4.575% | 7.796% | 3.961% | 2.898% | 1.81% | 5.109% | 1.237% | 0.889% | 0.3078% | 1.06% | 0.3866% | 1.909% | 7.714% | 1.227% | 5.496% | 0.8947% | 1.132% | 0.7343% | 3.155%
124 | LBLANK | R | 6/3/2023 4:03:32 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LBLANK.126 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 66.14% | 173.2% | 48.04% | 8.07% | 1.621% | 50.63% | 86.6% | 173.2% | 22.79% | 11.62% | 4.779% | 22.91% | 2.242% | 24.74% | 18.23% | 20.4% | 9.176% | 3.434% | 9.496% | 173.2% | 14.78% | 14.78% | 3.748% |  | 86.6% | 86.6% | 54.95% |  | 21.07% | 25% | 5.275% | 1.506% | 1.192% | 0.8449% | 0.8595% | 0.4695% | 4.999% | 86.6% | 45.07% | 52.85% | 25.83%
125 | CCV 1ppb TRACE | R | 6/3/2023 4:09:35 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.127 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.446% | 1.242% | 9.953% | 1.375% | 2.305% | 3.63% | 4.887% | 3.672% | 2.669% | 2.016% | 0.3491% | 1.346% | 1.683% | 3.236% | 2.434% | 2.589% | 2.107% | 2.265% | 10.08% | 4.248% | 1.178% | 1.178% | 4.237% | 11.48% | 7.547% | 5.107% | 2.032% | 10.07% | 4.853% | 1.73% | 1.071% | 1.119% | 2.577% | 0.4788% | 1.127% | 0.8431% | 1.445% | 2.697% | 1.524% | 1.328% | 1.294%
126 | CCV 1ppb TRACE | R | 6/3/2023 4:16:07 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.128 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.159% | 3.337% | 1.043% | 1.788% | 4.852% | 6.432% | 8.983% | 4.771% | 3.424% | 1.684% | 4.48% | 1.842% | 1.232% | 3.914% | 0.9559% | 3.376% | 2.279% | 0.3726% | 6.808% | 16.58% | 4.955% | 4.955% | 5.755% | 7.234% | 0.8287% | 10.52% | 6.752% | 5.43% | 1.93% | 2.787% | 5.431% | 0.7013% | 10.07% | 3.118% | 1.005% | 1.87% | 1.716% | 3.296% | 1.159% | 0.8109% | 2.432%
127 | LFB 10ppb | R | 6/3/2023 4:22:09 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LFB 10ppb.129 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.086% | 2.202% | 5.794% | 1.033% | 1.089% | 2.507% | 1.129% | 0.6586% | 1.697% | 0.6488% | 0.7012% | 0.7424% | 1.766% | 1.585% | 2.802% | 0.6383% | 1.015% | 1.328% | 0.5702% | 6.173% | 1% | 1% | 1.082% | 5.554% | 3.905% | 0.09892% | 1.779% | 1.192% | 1.392% | 1.501% | 1.054% | 0.6958% | 0.4979% | 0.5974% | 7.309% | 0.6011% | 1.086% | 1.128% | 1.086% | 1.273% | 2.31%
128 | LFB 10ppb | R | 6/3/2023 4:28:39 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LFB 10ppb.130 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.882% | 0.5866% | 5.79% | 0.2923% | 2.017% | 1.431% | 4.762% | 0.5492% | 1.643% | 0.9452% | 0.8259% | 0.8455% | 0.8893% | 0.6218% | 2.484% | 1.661% | 1.297% | 0.884% | 2.734% | 3.49% | 0.9406% | 0.9406% | 4.336% | 4.887% | 1.715% | 2.623% | 2.143% | 2.958% | 1.241% | 2.284% | 1.593% | 2.194% | 0.86% | 4.347% | 1.464% | 0.8585% | 2.151% | 1.077% | 1.283% | 0.9141% | 2.105%
129 | CCV 10 ppb MAJORS | R | 6/3/2023 4:34:42 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.131 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.949% | 4.003% | 7.196% | 0.6007% | 3.381% | 1.2% | 4.126% | 2.375% | 2.812% | 2.523% | 0.7503% | 2.659% | 3.162% | 4.252% | 3.535% | 2.233% | 1.636% | 1.683% | 3.209% | 4.261% | 2.834% | 2.834% | 1.061% | 1.547% | 2.237% | 1.332% | 3.235% | 1.087% | 0.5569% | 1.745% | 1.699% | 12% | 15.21% | 0.6787% | 1.577% | 0.1978% | 14.5% | 1.132% | 1.143% | 1.435% | 1.361%
130 | CCV 10 ppb MAJORS | R | 6/3/2023 4:41:13 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.132 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.418% | 3.434% | 5.372% | 0.9435% | 2.09% | 0.8556% | 3.843% | 2.179% | 1.873% | 0.761% | 2.388% | 1.852% | 1.569% | 1.349% | 1.419% | 0.498% | 1.746% | 1.573% | 1.184% | 4.702% | 1.614% | 1.614% | 2.02% | 1.502% | 3.383% | 1.527% | 4.361% | 0% | 0.7019% | 1.427% | 0.9461% | 0.08694% | 1.411% | 2.23% | 13.62% | 1.327% | 1.117% | 1.711% | 0.9092% | 0.3443% | 4.837%
131 | MT 33 | R | 6/3/2023 4:47:16 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 33.133 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 35.25% | 173.2% | 10.72% | 1.769% | 0.9348% | 22.21% | 49.49% | 173.2% | 41.41% | 33.02% | 1.984% | 6.667% | 5.273% | 44.41% | 53.31% | 7.083% | 9.33% | 1.922% | 7.16% | 173.2% | 32.87% | 32.87% | 6.763% |  | 50% | 43.3% | 53.15% |  | 7.97% | 4.762% | 1.426% | 0.847% | 1.316% | 4.714% | 0.5451% | 1.164% | 0.8964% | 0% | 60.62% | 4.194% | 1.209%
132 | MT 34 | R | 6/3/2023 4:53:19 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 34.134 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 24.98% | 86.6% | 20.74% | 1.126% | 1.867% | 46.79% | 65.47% | 43.3% | 26.64% | 3.612% | 1.712% | 30.87% | 2.595% | 96.44% | 12.06% | 10.54% | 11.79% | 3.621% | 5.106% | 173.2% | 28.64% | 28.64% | 6.326% | 173.2% | 86.6% | 43.3% | 56.44% |  | 12.12% | 20.83% | 4.029% | 0.6439% | 4.583% | 4.222% | 0.4155% | 0.8662% | 2.321% | 173.2% | 62.98% | 1.001% | 3.128%
133 | MT 35 | R | 6/3/2023 4:59:21 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 35.135 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 65.47% |  | 19.99% | 1.339% | 3.026% | 19.87% | 34.64% | 173.2% | 11.8% | 23.77% | 3.229% | 8.535% | 2.541% | 100% | 12.37% | 18.95% | 10.24% | 3.943% | 9.783% | 100% | 37.8% | 37.8% | 3.444% |  | 86.6% | 33.33% | 49.32% | 173.2% | 2.279% | 45.07% | 1.775% | 0.8806% | 1.538% | 4.688% | 1.014% | 1.391% | 2.195% | 100% | 38.49% | 2.863% | 6.787%
134 | MT 36 | R | 6/3/2023 5:05:25 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 36.136 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0% | 173.2% | 3.297% | 3.353% | 2.42% | 48.04% | 86.6% | 173.2% | 24.85% | 20.12% | 3.006% | 20.53% | 35.14% | 49.49% | 30.05% | 11.17% | 4.12% | 5.534% | 7.764% | 86.6% | 32.83% | 32.83% | 5.809% | 173.2% | 173.2% | 43.3% | 24.72% | 86.6% | 4.167% | 39.63% | 10.71% | 1.222% | 1.98% | 5.821% | 0.4248% | 1.06% | 1.019% | 86.6% | 52.71% | 14.11% | 6.354%
135 | MT 37 | R | 6/3/2023 5:11:28 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 37.137 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 61.97% |  | 5.801% | 1.719% | 5.927% | 26.67% | 24.74% | 173.2% | 29.6% | 9.488% | 8.37% | 17.5% | 3.612% | 173.2% | 13.58% | 5.734% | 2.515% | 4.04% | 4.838% | 173.2% | 26.96% | 26.96% | 5.91% | 100% | 173.2% | 173.2% | 23.78% |  | 26.91% | 57.74% | 4.519% | 2.484% | 2.159% | 18.06% | 1.801% | 1.231% | 3.304% | 173.2% | 43.3% | 1.756% | 1.657%
136 | MT 38 | R | 6/3/2023 5:17:31 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 38.138 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 21.65% | 173.2% | 14.78% | 1.239% | 4.72% | 34.58% | 100% | 173.2% | 2.44% | 6.013% | 5.222% | 16.86% | 2.964% | 25% | 20% | 13.43% | 8.093% | 8.01% | 4.544% | 86.6% | 50% | 50% | 6.816% |  | 0% | 100% | 54.64% | 173.2% | 43.75% | 29.04% | 1.347% | 0.2564% | 0.7302% | 0.8453% | 1.715% | 0.7716% | 1.678% | 43.3% | 29.04% | 5.59% | 4.699%
137 | MT 39 | R | 6/3/2023 5:23:33 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 39.139 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 8.333% | 86.6% | 14.47% | 1.899% | 5.312% | 62.45% | 65.47% | 86.6% | 17.47% | 8.402% | 1.4% | 9.302% | 2.815% | 45.83% | 31.39% | 14.94% | 14.32% | 2.641% | 5.585% | 173.2% | 24.74% | 24.74% | 7.862% | 173.2% | 100% | 24.74% | 56.21% |  | 26.36% | 0% | 1.948% | 9.875% | 0.3066% | 2.624% | 0.9186% | 2.377% | 10.02% | 43.3% | 31.49% | 2.2% | 4.429%
138 | MT 40 | R | 6/3/2023 5:29:36 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 40.140 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 7.143% | 9.116% | 5.951% | 1.276% | 4.298% | 14.43% | 36.36% | 173.2% | 17.83% | 37.67% | 1.659% | 8.505% | 1.288% | 132.3% | 18.37% | 15.26% | 9.592% | 2.996% | 4.122% | 86.6% | 13.79% | 13.79% | 2.138% | 173.2% |  | 43.3% | 95.34% | 86.6% | 0% | 14.19% | 1.063% | 1.543% | 6.92% | 1.427% | 1.69% | 0.3667% | 0.7927% | 173.2% | 43.3% | 3.502% | 6.44%
139 | MT 32 B | R | 6/3/2023 5:35:38 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 32 B.141 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 45.83% | 173.2% | 16.82% | 1.263% | 1.601% | 31.23% | 173.2% | 86.6% | 7.873% | 8.012% | 4.812% | 13.68% | 2.24% | 69.28% | 33.01% | 9.615% | 3.396% | 11.92% | 6.754% | 173.2% | 0% | 0% | 0.8308% |  |  | 50% | 52.93% |  | 10.96% | 10.41% | 0.6845% | 7.508% | 3.367% | 6.031% | 0.6136% | 0.6488% | 4.381% | 100% | 26.03% | 2.177% | 1.868%
140 | MT 32 B LD | R | 6/3/2023 5:41:40 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 32 B LD.142 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 96.08% | 173.2% | 8.898% | 0.6004% | 0.7981% | 29.05% | 50% | 173.2% | 11.11% | 9.086% | 2.867% | 22.88% | 2.124% | 100% | 18.37% | 11.94% | 6.173% | 1.494% | 6.747% | 173.2% | 50% | 50% | 3.889% |  | 173.2% | 100% | 157.3% | 173.2% | 22.88% | 12.01% | 0.3621% | 1.906% | 1.226% | 1.253% | 0.9606% | 0.4011% | 3.324% | 24.74% | 25% | 8.124% | 0.5829%
141 | LBLANK | R | 6/3/2023 5:47:43 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LBLANK.143 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 43.3% |  | 21.65% | 1.394% | 3.599% | 13.09% | 173.2% | 100% | 18.72% | 14.24% | 2.261% | 13.33% | 0.7965% | 86.6% | 28.64% | 39.85% | 12.76% | 5.502% | 4.109% |  | 50% | 50% | 5.699% | 86.6% | 86.6% | 173.2% | 19.06% | 173.2% | 24.02% | 28.64% | 6.251% | 1.114% | 1.838% | 7.988% | 2.69% | 1.065% | 2.821% | 100% | 43.3% | 42.86% | 11.17%
142 | CCV 1ppb TRACE | R | 6/3/2023 5:53:46 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.144 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.12% | 5.375% | 21.14% | 1.793% | 9.86% | 2.167% | 5.097% | 2.37% | 2.596% | 0.7837% | 3.582% | 1.082% | 2.441% | 0.6838% | 1.051% | 3.05% | 2.669% | 2.819% | 7.853% | 20.75% | 2.87% | 2.87% | 5.294% | 11.87% | 7.843% | 8.04% | 8.541% | 11.49% | 2.325% | 1.572% | 1.078% | 0.3289% | 2.55% | 3.953% | 4.668% | 0.1957% | 2.008% | 1.36% | 1.809% | 2.398% | 1.053%
143 | CCV 1ppb TRACE | R | 6/3/2023 6:00:18 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.145 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 4.716% | 4.146% | 4.413% | 4.164% | 4.253% | 3.512% | 8.142% | 4.548% | 2.03% | 1.104% | 2.848% | 2.51% | 2.73% | 2.223% | 5.545% | 2.236% | 0.8144% | 1.64% | 13.7% | 8.271% | 3.006% | 3.006% | 4.054% | 7.901% | 2.073% | 4.187% | 5.755% | 10.04% | 0.9848% | 0.4287% | 8.766% | 3.463% | 0.7633% | 0.6012% | 1.045% | 1.7% | 2.406% | 0.4726% | 0.711% | 0.3746% | 0.5614%
144 | CCV 10 ppb MAJORS | R | 6/3/2023 6:06:21 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.146 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.3007% | 3.095% | 3.684% | 2.253% | 1.407% | 1.51% | 1.499% | 1.562% | 1.575% | 0.7196% | 0.3315% | 0.7606% | 0.8012% | 1.376% | 2.595% | 0.4565% | 0.6703% | 0.6633% | 1.088% | 1.44% | 1.215% | 1.215% | 2.41% | 5.86% | 2.684% | 2.439% | 4.11% | 5.03% | 0.3154% | 1.594% | 2.134% | 0.7508% | 1.599% | 0.7586% | 1.771% | 1.899% | 2.694% | 1.164% | 0.3903% | 2.549% | 2.043%
145 | CCV 10 ppb MAJORS | R | 6/3/2023 6:12:53 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.147 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.428% | 1.304% | 4.983% | 2.778% | 3.249% | 1.787% | 1.323% | 1.505% | 0.6033% | 2.029% | 1.472% | 1.793% | 0.8408% | 1.091% | 1.554% | 0.4463% | 0.2316% | 0.8268% | 2.294% | 1.327% | 1.461% | 1.461% | 2.585% | 6.448% | 2.706% | 1.767% | 0.2591% | 1.787% | 0.7873% | 0.5407% | 0.2414% | 1.628% | 1.252% | 0.6619% | 2.003% | 1.375% | 2.372% | 1.597% | 1.132% | 0.9404% | 2.316%
146 | MT 33 B | R | 6/3/2023 6:18:55 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 33 B.148 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 43.3% | 100% | 23.04% | 1.778% | 6.091% | 22.22% | 132.3% | 124.9% | 10% | 23.27% | 8.107% | 25.35% | 3.522% | 0% | 35.25% | 10.95% | 4.456% | 7.759% | 6.415% | 86.6% | 53.29% | 53.29% | 4.077% | 173.2% | 173.2% |  | 32.09% |  | 28.44% | 13.48% | 2.552% | 1.458% | 3.217% | 0.1093% | 1.454% | 2.021% | 3.766% | 89.21% | 13.32% | 10.38% | 4.892%
147 | MT 33 B LSP | R | 6/3/2023 6:24:56 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 33 B LSP.149 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.273% | 0.4472% | 2.832% | 0.3082% | 0.5442% | 1.784% | 1.452% | 1.311% | 0.1276% | 3.735% | 3.646% | 1.202% | 1.003% | 2.624% | 0.4916% | 1.514% | 2.641% | 2.84% | 2.515% | 1.436% | 0.9559% | 0.9559% | 2.836% | 2.311% | 1.385% | 0.9544% | 1.832% | 2.7% | 2.616% | 0.9094% | 1.071% | 0.9567% | 2.891% | 13.69% | 1.78% | 0.411% | 2.661% | 2.51% | 2.498% | 2.745% | 1.177%
148 | MT 34 B | R | 6/3/2023 6:30:59 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 34 B.150 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 53.29% | 0% | 13.73% | 1.889% | 2.901% | 14.7% | 50% | 24.74% | 14.18% | 13.45% | 8.988% | 28.46% | 4.073% | 69.28% | 10.6% | 21.92% | 2.797% | 5.842% | 5.37% | 69.28% | 51.2% | 51.2% | 3.985% | 173.2% | 114.6% | 34.64% | 38.18% |  | 5.395% | 14.95% | 1.729% | 1.099% | 1.847% | 1.225% | 20.37% | 1.395% | 2.369% | 0% | 10.66% | 6.072% | 2.272%
149 | MT 35 B | R | 6/3/2023 6:37:01 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 35 B.151 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0% | 173.2% | 14.17% | 1.454% | 1.341% | 20.4% | 57.28% | 114.6% | 9.805% | 18.19% | 3.232% | 38.41% | 4.273% | 124.9% | 20.52% | 10.48% | 1.403% | 2.612% | 6.892% | 173.2% | 43.3% | 43.3% | 5.779% | 173.2% | 114.6% | 100% | 59.52% |  | 18.04% | 20.4% | 2.23% | 0.8355% | 5.945% | 0.5605% | 1.321% | 1.406% | 4.894% | 34.64% | 28.39% | 3.253% | 4.397%
150 | MT 36 B | R | 6/3/2023 6:43:03 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 36 B.152 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 52.92% |  | 4.499% | 1.564% | 0.566% | 42.87% | 173.2% | 43.3% | 13.83% | 5.284% | 3.216% | 13.73% | 0.9123% | 78.06% | 16.06% | 18.03% | 12.95% | 2.134% | 2.214% | 86.6% | 81.03% | 81.03% | 2.613% |  |  | 114.6% | 30.44% | 173.2% | 3.924% | 2.008% | 2.537% | 0.8917% | 3.069% | 1.455% | 2.594% | 2.508% | 2.122% | 100% | 33.07% | 8.017% | 5.897%
151 | MT 37 B | R | 6/3/2023 6:49:06 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 37 B.153 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 24.74% | 173.2% | 9.324% | 0.9173% | 8.352% | 24.35% | 132.3% | 100% | 26.96% | 17.23% | 9.96% | 23.27% | 2.759% | 86.6% | 13.29% | 14.91% | 1.056% | 7.315% | 7.335% | 173.2% | 96.44% | 96.44% | 1.318% | 173.2% | 173.2% | 173.2% | 23.19% | 173.2% | 8.66% | 38.65% | 7.634% | 1.203% | 1.467% | 4.744% | 0.639% | 0.4963% | 2.766% | 34.64% | 22.91% | 2.802% | 7.441%
152 | MT 38 B | R | 6/3/2023 6:55:09 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 38 B.154 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 5.391% | 173.2% | 17.68% | 1.108% | 6.619% | 43.45% | 91.65% | 100% | 7.16% | 16.7% | 8.701% | 17.64% | 0.9874% | 33.33% | 27.15% | 10.96% | 7.747% | 7.824% | 8.294% |  | 17.32% | 17.32% | 4.153% | 173.2% | 86.6% | 173.2% | 65.53% | 86.6% | 11.34% | 22.58% | 0.8405% | 1.932% | 3.406% | 5.569% | 1.434% | 2.02% | 2.232% | 114.6% | 46.63% | 7.807% | 6.589%
153 | MT 39 B | R | 6/3/2023 7:01:12 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 39 B.155 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 43.3% | 86.6% | 29.32% | 1.855% | 5.477% | 44.11% | 21.65% | 34.64% | 12.56% | 8.985% | 9.222% | 12.01% | 5.795% | 86.6% | 22.53% | 4.949% | 8.564% | 4.451% | 9.019% | 86.6% | 57.74% | 57.74% | 8.327% | 173.2% | 100% | 173.2% | 65.89% | 173.2% | 8.071% | 11.25% | 1.212% | 0.4638% | 2.127% | 15.68% | 0.6322% | 1.901% | 1.441% | 132.3% | 28.57% | 1.709% | 5.504%
154 | MT 40 B | R | 6/3/2023 7:07:15 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 40 B.156 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 43.3% | 86.6% | 13.92% | 2.284% | 1.615% | 17.32% | 50% | 100% | 18.32% | 32.2% | 5.888% | 32.13% | 1.745% | 44.1% | 21.31% | 2.112% | 4.575% | 5.012% | 6.94% | 173.2% | 41.93% | 41.93% | 4.734% | 173.2% | 100% | 86.6% | 74.98% | 173.2% | 44.39% | 7.692% | 1.492% | 0.04178% | 5.454% | 2.629% | 1.283% | 0.2662% | 2.496% | 86.6% | 8.66% | 1.859% | 2.73%
155 | MT 32 C | R | 6/3/2023 7:13:17 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 32 C.157 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 50% | 173.2% | 26.5% | 1.469% | 4.603% | 17.84% | 100% | 114.6% | 5.679% | 17.32% | 2.156% | 15.02% | 1.862% | 43.3% | 4.028% | 10.29% | 6.342% | 2.447% | 4.263% | 173.2% | 34.64% | 34.64% | 8.387% |  | 173.2% | 86.6% | 23.67% |  | 11.62% | 43.3% | 1.581% | 4.394% | 0.7933% | 0.516% | 2.115% | 2.579% | 2.367% | 34.64% | 12.5% | 5.214% | 4.847%
156 | LBLANK | R | 6/3/2023 7:19:20 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LBLANK.158 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0% |  | 66.14% | 2.063% | 2.943% | 32.97% |  | 114.6% | 15.8% | 18.83% | 2.192% | 62.22% | 2.627% | 0% | 9.091% | 14.78% | 23.08% | 4.217% | 1.43% | 173.2% | 45.83% | 45.83% | 5.39% |  | 100% | 0% | 79.83% | 86.6% | 27.74% | 50% | 2.229% | 3.111% | 3.885% | 1.139% | 1.64% | 1.283% | 6.081% | 173.2% | 37.75% | 60.4% | 22.22%
157 | CCV 1ppb TRACE | R | 6/3/2023 7:25:23 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.159 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.03% | 1.279% | 6.186% | 2.059% | 6.269% | 5.827% | 6.973% | 2.987% | 2.766% | 0.3495% | 4.013% | 0.9237% | 1.229% | 1.81% | 1.904% | 1.02% | 0.4796% | 1.382% | 3.693% | 11.24% | 1.517% | 1.517% | 3.82% | 23.7% | 14.58% | 4.293% | 2.637% | 12.95% | 1.338% | 2.168% | 2.265% | 1.394% | 4.541% | 0.2398% | 1.555% | 0.7395% | 0.3333% | 0.9112% | 1.967% | 1.608% | 2.879%
158 | CCV 1ppb TRACE | R | 6/3/2023 7:31:56 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.160 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.372% | 4.677% | 9.264% | 0.9793% | 4.869% | 2.561% | 3.929% | 2.056% | 3.55% | 2.926% | 3.145% | 2.613% | 3.092% | 1.059% | 1.29% | 5.035% | 2.792% | 2.913% | 6.981% | 13.77% | 5.906% | 5.906% | 3.201% | 20.02% | 2.857% | 13.57% | 9.117% | 11.12% | 1.665% | 3.61% | 0.2753% | 10.9% | 1.271% | 12.3% | 1.138% | 1.043% | 10.2% | 4.291% | 3.861% | 1.166% | 1.408%
159 | LFB 10ppb | R | 6/3/2023 7:37:58 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LFB 10ppb.161 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.731% | 1.807% | 4.517% | 1.88% | 3.896% | 1.836% | 2.908% | 1.038% | 0.6231% | 1.572% | 1.443% | 4.183% | 1.591% | 0.3476% | 0.4379% | 0.8568% | 0.9091% | 0.7355% | 2.222% | 4.188% | 4.905% | 4.905% | 1.803% | 8.964% | 4.111% | 2.591% | 3.459% | 4.01% | 1.741% | 5.083% | 0.3911% | 15.61% | 0.6502% | 0.5084% | 1.486% | 1.66% | 17.85% | 0.8252% | 0.2769% | 0.8166% | 1.557%
160 | LFB 10ppb | R | 6/3/2023 7:44:29 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LFB 10ppb.162 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.41% | 4.172% | 4.821% | 2.94% | 6.687% | 3.091% | 1.598% | 3.37% | 4.122% | 147.4% | 118.2% | 2.081% | 1.805% | 2.143% | 1.524% | 2.275% | 149.2% | 137.5% | 30.08% | 3.425% | 1.213% | 1.213% | 11.75% | 7.921% | 1.536% | 2.14% | 200.2% | 1.218% | 146% | 2.306% | 7.329% | 4.752% | 6.464% | 11.35% | 4.381% | 1.244% | 3.787% | 148.8% | 150.6% | 30.31% | 6.611%
161 | CCV 10 ppb MAJORS | R | 6/3/2023 7:50:31 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.163 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.44% | 1.61% | 2.076% | 0.8381% | 2.5% | 1.81% | 0.98% | 1.467% | 1.137% | 1.93% | 1.806% | 0.533% | 3.172% | 0.5913% | 2.944% | 0.9354% | 2.015% | 0.4605% | 3.644% | 2.821% | 2.03% | 2.03% | 4.836% | 3.583% | 2.192% | 1.533% | 3.604% | 2.349% | 1.487% | 1.55% | 11.22% | 0.9423% | 1.508% | 1.205% | 1.575% | 0.5462% | 5.632% | 1.243% | 1.263% | 3.816% | 1.392%
162 | CCV 10 ppb MAJORS | R | 6/3/2023 7:57:03 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.164 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.151% | 4.626% | 5.762% | 2.829% | 3.046% | 2.592% | 3.731% | 2.319% | 2.492% | 0.2867% | 1.591% | 1.538% | 1.503% | 1.444% | 4.189% | 1.896% | 1.261% | 0.6497% | 1.714% | 4.463% | 2.126% | 2.126% | 2.803% | 5.965% | 1.99% | 0.5406% | 5.029% | 0.8738% | 0.3589% | 1.62% | 5.403% | 7.973% | 3.675% | 0.9733% | 16.26% | 0.7824% | 6.295% | 0.4939% | 1.498% | 2.208% | 3.711%
163 | MT 32 C LD | R | 6/3/2023 8:03:06 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 32 C LD.165 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 22.91% | 173.2% | 10.75% | 0.8987% | 6.046% | 3.464% | 86.6% | 86.6% | 9.524% | 5.119% | 5.367% | 39.76% | 3.554% | 69.28% | 9.962% | 7.138% | 15.42% | 0.1607% | 3.668% | 173.2% | 62.5% | 62.5% | 3.583% |  | 43.3% | 86.6% | 51.19% | 173.2% | 17.32% | 15.75% | 8.649% | 7.805% | 3.406% | 0.9663% | 1.024% | 1.317% | 6.748% | 86.6% | 26.19% | 6.177% | 5.003%
164 | MT 33 C | R | 6/3/2023 8:09:08 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 33 C.166 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 108.3% |  | 24.07% | 0.6849% | 7.76% | 55.87% | 43.3% | 86.6% | 31.23% | 2.564% | 3.584% | 31.12% | 2.998% | 50% | 10% | 15.34% | 7.699% | 5.721% | 4.007% |  | 33.33% | 33.33% | 8.623% | 173.2% | 173.2% | 100% | 120% |  | 15.95% | 3.936% | 12.8% | 0.8302% | 4.486% | 0.7645% | 2.292% | 1.733% | 2.175% | 100% | 27.15% | 10.47% | 2.812%
165 | MT 33 C LSP | R | 6/3/2023 8:15:09 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 33 C LSP.167 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.051% | 3.398% | 1.796% | 0.9058% | 1.317% | 1.522% | 1.727% | 0.2145% | 0.7419% | 0.9381% | 3.109% | 2.949% | 0.3637% | 0.7274% | 0.4859% | 0.3536% | 1.165% | 1.026% | 1.824% | 1.193% | 1.558% | 1.558% | 0.7662% | 4.892% | 0.9689% | 1.581% | 1.49% | 2.341% | 6.327% | 1.623% | 11.4% | 3.145% | 3.735% | 2.278% | 1.312% | 0.9037% | 6.544% | 1.406% | 0.9766% | 2.233% | 0.4289%
166 | MT 33 C LSP | R | 6/3/2023 8:21:40 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 33 C LSP.168 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.429% | 1.881% | 1.725% | 3.592% | 1.48% | 1.553% | 2.631% | 1.398% | 1.601% | 0.8945% | 0.7509% | 1.016% | 2.219% | 1.744% | 2.566% | 1.552% | 0.7054% | 0.2782% | 1.137% | 2.785% | 0.2882% | 0.2882% | 3.536% | 2.327% | 0.6026% | 1.21% | 0.4404% | 2.065% | 0.4142% | 0.4779% | 2.904% | 0.81% | 3.376% | 1.277% | 1.147% | 0.9664% | 4.506% | 0.8223% | 0.8164% | 1.029% | 0.9906%
167 | MT 34 C | R | 6/3/2023 8:27:42 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 34 C.169 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 12.37% | 86.6% | 18.79% | 0.2178% | 1.268% | 21.57% | 86.6% | 173.2% | 25.75% | 13.5% | 3.616% | 35.39% | 3.612% | 53.91% | 8.636% | 17.96% | 2.84% | 1.897% | 1.61% | 100% | 48.04% | 48.04% | 4.593% | 173.2% | 0% | 173.2% | 138.4% |  | 5.175% | 4.683% | 8.495% | 3.462% | 0.4912% | 0.3255% | 1.149% | 1.153% | 2.364% | 86.6% | 34.64% | 7.456% | 7.032%
168 | MT 35 C | R | 6/3/2023 8:33:45 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 35 C.170 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 47.19% | 86.6% | 28.78% | 4.182% | 3.378% | 30.46% | 114.6% | 0% | 11.55% | 14.43% | 1.288% | 21.21% | 1.411% | 86.6% | 22.72% | 7.947% | 8.643% | 10.09% | 6.587% |  | 10.19% | 10.19% | 1.678% |  | 43.3% |  | 57.88% | 173.2% | 28.21% | 5.587% | 2.723% | 6.665% | 3.68% | 3.089% | 3.027% | 2.378% | 4.968% | 86.6% | 6.662% | 10.01% | 5.445%

## Net Intensities
| Sample Id | R | Acquisition Time | QC Status | Dataset File | Method File | Na KED 23/23 Helium KED_Low (cps) | Mg KED 24/24 Helium KED_Low (cps) | Al KED 27/27 Helium KED_Low (cps) | P Oshift 31/47 Oxygen DRC (cps) | S (34) -1 34/34 Helium KED_Low (cps) | K KED 39/39 Helium KED_Low (cps) | Ca KED 43/43 Helium KED_Low (cps) | V KED 51/51 Helium KED_Low (cps) | Cr (52) 52/52 Helium KED_Low (cps) | Cr (53) NH3 53/53 Ammonia DRC (cps) | Fe (54) NH3 54/54 Ammonia DRC (cps) | Mn Oshift 55/71 Oxygen DRC (cps) | Fe (56) KED 56/56 Helium KED_Low (cps) | Co KED 59/59 Helium KED_Low (cps) | Ni KED 60/60 Helium KED_Low (cps) | Cu (63) KED 63/63 Helium KED_Low (cps) | Cu (65) NH3 65/65 Ammonia DRC (cps) | Zn NH3 66/66 Ammonia DRC (cps) | As 75/75 (cps) | As KED 75/75 Helium KED_Low (cps) | As RH Oshift 75/91 Oxygen DRC (cps) | As Oshift 75/91 Oxygen DRC (cps) | Se (77) 77/77 (cps) | Se (77) Oshift 77/93 Oxygen DRC (cps) | Se (78) Oshift 78/94 Oxygen DRC (cps) | Se (80) Oshift 80/96 Oxygen DRC (cps) | Se (82) 82/82 (cps) | Se (82) Oshift 82/98 Oxygen DRC (cps) | Mo (95) NH3 95/95 Ammonia DRC (cps) | Mo (95) OShift 95/127 Oxygen DRC (cps) | Rh (IS1) 103/103 (IS) (cps) | Rh (IS2) 103/103 (IS) Oxygen DRC (cps) | Rh (IS3) 103/103 (IS) Helium KED_Low (cps) | Rh (IS4) 103/103 (IS) Ammonia DRC (cps) | Rh (IS5) 103/103 (IS) Ammonia DRC (cps) | Rh (IS6) 103/103 (IS) Oxygen DRC (cps) | Rh OShift (IS1) 103/119 (IS) Oxygen DRC (cps) | Cd (111) NH3 111/111 Ammonia DRC (cps) | Cd (114) NH3 114/114 Ammonia DRC (cps) | Pb 208/208 (cps) | Pb NH3 208/208 Ammonia DRC (cps)
1 | CAL BLANK | R | 06/02/2023 15:17:36 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CAL BLANK.004 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.8434e-07 | 0 | -8.9399e-07 | 5.9034e-06 | 1.0527e-05 | -8.6644e-07 | -1.1047e-07 | 1.1421e-07 | 6.9412e-07 | 1.0334e-08 | -1.9445e-07 | 2.6422e-08 | 1.1876e-05 | 1.1795e-07 | -1.685e-06 | 2.5232e-06 | 3.1103e-08 | -1.2499e-06 | -2.0732e-05 | -5.1039e-07 | 2.3447e-06 | -3.1561e-08 | 7.2139e-06 | 0 | 1.4687e-08 | 2.3149e-08 | -1.6992e-06 | 0 | 5.1542e-08 | 3.1611e-08 | 136153.5367 | 200107.0943 | 10262.6851 | 220973.2575 | 166481.7999 | 182757.8968 | 3107.6714 | 0 | 1.2114e-08 | -1.1501e-06 | -1.8501e-08
2 | LOW 0.01 | R | 06/02/2023 15:23:43 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LOW 0.01.005 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.0012106 | 0.00048834 | 0.00065877 | 0.004825 | -0.00060497 | 0.00086016 | 0.00013071 | 0.00068613 | 0.00088954 | 0.00034964 | 0.0020426 | 0.00069966 | 0.0064163 | 0.0019258 | 0.00076012 | 0.002763 | 0.0011216 | 0.0034672 | 0.00024194 | 2.0421e-06 | 0.022797 | 0.00035622 | -0.00038668 | 6.675e-06 | 3.443e-05 | 3.1096e-05 | 4.5528e-05 | 2.3622e-05 | 0.0015026 | 0.0014669 | 135983.4907 | 196439.5642 | 10202.9768 | 216608.7837 | 165740.6331 | 181647.6968 | 3052.3263 | 0.00040055 | 0.00087683 | 0.0023835 | 0.0045514
3 | LOW 0.05 | R | 06/02/2023 15:29:46 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LOW 0.05.006 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.0051649 | 0.0024662 | 0.0014124 | 0.012099 | 0.00018344 | 0.0033861 | 0.00095226 | 0.0033327 | 0.0060549 | 0.001966 | 0.0033788 | 0.0032822 | 0.016827 | 0.010838 | 0.0036324 | 0.0095589 | 0.0041343 | 0.0028949 | 0.0017998 | 0.00036958 | 0.12247 | 0.0019004 | -0.00036916 | 2.3777e-05 | 0.0001287 | 0.00029016 | 0.00022457 | 5.9625e-05 | 0.0035566 | 0.0035828 | 136570.7617 | 191527.9639 | 9835.0601 | 215509.8351 | 158594.455 | 176415.5452 | 2976.9775 | 0.0018146 | 0.0043644 | 0.011652 | 0.021995
4 | LOW 0.1 | R | 06/02/2023 15:35:48 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LOW 0.1.007 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.011634 | 0.0046541 | 0.00080229 | 0.0093731 | -0.0014633 | 0.0052619 | 0.0013913 | 0.0072289 | 0.010132 | 0.0034383 | 0.0073995 | 0.0062396 | 0.030445 | 0.020914 | 0.005328 | 0.015686 | 0.0077799 | 0.0066006 | 0.0036447 | 0.00067787 | 0.22393 | 0.0032415 | 0.00013418 | 6.0641e-05 | 0.00030015 | 0.00051948 | 0.00045436 | 9.1072e-05 | 0.0068009 | 0.0067405 | 136777.7305 | 197752.0331 | 10333.4036 | 216000.8763 | 171533.0741 | 178731.7931 | 2868.2881 | 0.0037066 | 0.0088892 | 0.022281 | 0.039589
5 | LOW 0.5 | R | 06/02/2023 15:41:51 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LOW 0.5.008 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.055996 | 0.021831 | 0.0030312 | 0.023222 | 0.0060597 | 0.029955 | 0.0089071 | 0.037736 | 0.052442 | 0.016785 | 0.023585 | 0.029293 | 0.092662 | 0.10106 | 0.027755 | 0.08462 | 0.036004 | 0.020829 | 0.017294 | 0.0045714 | 1.0758 | 0.016197 | 0.0014701 | 0.00033487 | 0.0010509 | 0.0023645 | 0.0016346 | 0.00037965 | 0.033379 | 0.031184 | 143511.248 | 210180.4162 | 10266.3611 | 220777.046 | 164559.0387 | 161616.1654 | 3165.6858 | 0.018162 | 0.043608 | 0.10827 | 0.21289
6 | STD 1 | R | 06/02/2023 15:47:53 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\STD 1.009 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.1139 | 0.044287 | 0.0042525 | 0.024893 | 0.016638 | 0.060755 | 0.01856 | 0.076041 | 0.10649 | 0.034859 | 0.047991 | 0.059699 | 0.19121 | 0.20434 | 0.0579 | 0.16219 | 0.072257 | 0.04146 | 0.036012 | 0.0079006 | 2.204 | 0.033104 | 0.0031966 | 0.0007581 | 0.0022346 | 0.0047427 | 0.0035746 | 0.00091102 | 0.068347 | 0.06631 | 149820.2053 | 204953.0789 | 10386.7758 | 225869.0355 | 141872.1783 | 188348.2289 | 3079.6654 | 0.035696 | 0.084585 | 0.21557 | 0.53685
7 | STD 5 | R | 06/02/2023 15:53:56 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\STD 5.010 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.53054 | 0.21514 | 0.013117 | 0.079664 | 0.071543 | 0.27785 | 0.091125 | 0.37794 | 0.5215 | 0.17092 | 0.1181 | 0.2962 | 0.47153 | 1.0145 | 0.27767 | 0.76205 | 0.33452 | 0.16238 | 0.1758 | 0.040801 | 10.9293 | 0.16241 | 0.01866 | 0.0033396 | 0.011232 | 0.024074 | 0.016803 | 0.0042258 | 0.33828 | 0.32647 | 155400.1918 | 199674.5214 | 10442.1492 | 229347.7605 | 170792.2626 | 189936.6217 | 2968.3085 | 0.17637 | 0.41713 | 1.0061 | 2.04
8 | STD 10 | R | 06/02/2023 15:59:58 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\STD 10.011 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.04 | 0.42109 | 0.023857 | 0.13764 | 0.14873 | 0.53719 | 0.18335 | 0.74842 | 1.0532 | 0.33602 | 0.22777 | 0.60107 | 0.90358 | 2.0213 | 0.56416 | 1.4847 | 0.64926 | 0.31527 | 0.35708 | 0.083161 | 22.1249 | 0.32461 | 0.036588 | 0.0069237 | 0.021884 | 0.046042 | 0.032695 | 0.0083206 | 0.68282 | 0.65972 | 151074.6348 | 188390.8451 | 10337.4056 | 218989.7884 | 165985.472 | 178791.366 | 2764.9344 | 0.34574 | 0.83659 | 1.922 | 4.1045
9 | STD 25 | R | 06/02/2023 16:06:00 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\STD 25.012 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.6621 | 1.0582 | 0.066153 | 0.32982 | 0.3936 | 1.389 | 0.48351 | 2.0907 | 2.8766 | 0.8314 | 0.56036 | 1.4792 | 2.4216 | 5.5445 | 1.5196 | 3.9855 | 1.5883 | 0.78279 | 0.8974 | 0.22938 | 55.9463 | 0.83833 | 0.091949 | 0.017923 | 0.057569 | 0.12201 | 0.082066 | 0.021731 | 1.7833 | 1.7049 | 139005.7051 | 159585.1454 | 8542.2228 | 191644.2411 | 162940.7854 | 153484.1099 | 2392.8672 | 0.88101 | 2.1105 | 4.3293 | 9.2284
10 | STD 50 | R | 06/02/2023 16:12:02 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\STD 50.013 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 4.6893 | 1.852 | 0.1091 | 0.62341 | 0.71747 | 2.4268 | 0.89674 | 3.8593 | 5.2542 | 1.642 | 1.0991 | 2.9686 | 4.3568 | 10.0095 | 2.7066 | 7.172 | 3.0853 | 1.5094 | 1.6719 | 0.40474 | 104.5128 | 1.6636 | 0.17233 | 0.036836 | 0.11526 | 0.24552 | 0.15018 | 0.044189 | 3.5697 | 3.5001 | 144558.1694 | 141863.5909 | 8202.3541 | 177018.0631 | 136971.5515 | 135937.0912 | 2258.8453 | 1.7407 | 4.1569 | 6.1368 | 20.4072
11 | STD 100 | R | 06/02/2023 16:18:04 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\STD 100.014 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 9.5983 | 3.6634 | 0.22821 | 1.2391 | 1.4876 | 4.7968 | 1.9285 | 8.2834 | 11.1021 | 3.3649 | 2.2313 | 6.0656 | 9.2524 | 21.683 | 5.6759 | 14.5686 | 6.0253 | 2.9202 | 3.8733 | 0.83333 | 202.5623 | 3.288 | 0.41222 | 0.07561 | 0.23331 | 0.49815 | 0.35965 | 0.089424 | 7.4644 | 7.2792 | 126049.1057 | 139734.0772 | 8070.2796 | 176024.5478 | 137735.472 | 134260.625 | 2268.18 | 3.4364 | 8.2245 | 10.1486 | 40.8034
12 | Calibration Curves | R | 06/02/2023 16:18:04
13 | LBLANK | R | 06/02/2023 16:24:07 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LBLANK.015 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.00029519 | 2.9986e-05 | -0.00015117 | -0.0039887 | 0.00409 | -0.00019233 | 5.7041e-05 | 0.00020406 | 0.00038309 | 4.9177e-05 | 0.00035552 | 7.8642e-05 | 0.00046777 | 0.00046252 | 0.00010495 | 0.00053702 | 0.0001843 | -0.00091093 | -0.00019518 | -3.8341e-05 | 0.0046404 | 8.7382e-05 | 0.00016259 | 0 | 9.8115e-06 | 2.2924e-05 | 0.00019638 | 3.2543e-06 | 0.00055264 | 0.00071741 | 152331.0678 | 202290.3113 | 11170.3657 | 259075.6478 | 197216.0209 | 180605.1451 | 3391.736 | 3.6096e-05 | 9.6806e-05 | 0.00023395 | 0.00015618
14 | ICV PE | R | 06/02/2023 16:30:09 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\ICV PE.016 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.99708 | 0.42213 | 0.027309 | 0.25006 | 0.1578 | 0.50319 | 0.1981 | 0.78052 | 1.0847 | 0.33507 | 0.2297 | 0.60469 | 0.9882 | 2.1566 | 0.58596 | 1.5408 | 0.65108 | 0.31258 | 0.34766 | 0.073905 | 19.743 | 0.32639 | 0.025859 | 0.0070813 | 0.023042 | 0.048942 | 0.032449 | 0.0087325 | 0.67116 | 0.65217 | 167954.8939 | 198192.0532 | 11340.5016 | 237648.5696 | 184397.8319 | 181738.1293 | 3276.3757 | 0.35085 | 0.83572 | 1.7095 | 3.8714
15 | CCV 1ppb TRACE | R | 06/02/2023 16:36:12 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.017 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.10656 | 0.045714 | 0.003509 | 0.010275 | 0.024337 | 0.051539 | 0.020224 | 0.082996 | 0.11377 | 0.035385 | 0.025625 | 0.063275 | 0.1182 | 0.22991 | 0.0629 | 0.16816 | 0.072794 | 0.033503 | 0.034245 | 0.0083723 | 1.8799 | 0.032402 | 0.0032821 | 0.00072896 | 0.002174 | 0.0047312 | 0.0034487 | 0.00083532 | 0.069995 | 0.06754 | 167751.1728 | 219473.7251 | 12405.7181 | 253755.3037 | 192724.737 | 178916.1584 | 3781.167 | 0.036012 | 0.086655 | 0.19055 | 0.39031
16 | CCV 1ppb TRACE | R | 06/02/2023 16:42:44 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.018 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.1137 | 0.044944 | 0.003181 | 0.0065678 | 0.028305 | 0.056225 | 0.019601 | 0.082532 | 0.1136 | 0.03506 | 0.02478 | 0.066944 | 0.11596 | 0.23102 | 0.065024 | 0.16488 | 0.073722 | 0.052554 | 0.034822 | 0.007833 | 1.9289 | 0.034129 | 0.0034784 | 0.00076298 | 0.0024128 | 0.0050271 | 0.0032109 | 0.00098497 | 0.070298 | 0.067849 | 159879.8628 | 210573.2893 | 12601.5564 | 242172.7819 | 184857.6825 | 191460.5238 | 3723.8187 | 0.035854 | 0.084485 | 0.19181 | 0.3853
17 | LFB 10ppb | R | 06/02/2023 16:48:46 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LFB 10ppb.019 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.1367 | 0.44915 | 0.026988 | 0.14467 | 0.17232 | 0.55486 | 0.20883 | 0.80487 | 1.1027 | 0.35322 | 0.25667 | 0.65764 | 1.0723 | 2.2777 | 0.6119 | 1.5427 | 0.67831 | 0.43847 | 0.37142 | 0.080507 | 19.6559 | 0.34053 | 0.038437 | 0.0078312 | 0.023705 | 0.050372 | 0.034108 | 0.0092628 | 0.70171 | 0.6688 | 154025.2739 | 193170.6946 | 11563.6791 | 222034.8778 | 170737.5701 | 178051.3927 | 3349.7264 | 0.34834 | 0.8271 | 1.6746 | 3.5428
18 | LFB 10ppb | R | 06/02/2023 16:55:16 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LFB 10ppb.020 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.1311 | 0.46026 | 0.029833 | 0.1502 | 0.17332 | 0.55377 | 0.21018 | 0.8142 | 1.117 | 0.34963 | 0.25599 | 0.66683 | 1.0691 | 2.2913 | 0.62123 | 1.5559 | 0.68247 | 0.44026 | 0.36545 | 0.08653 | 19.8382 | 0.33875 | 0.03949 | 0.0078689 | 0.023712 | 0.051176 | 0.033597 | 0.0090679 | 0.69711 | 0.67369 | 152595.3823 | 192459.8529 | 11437.2434 | 218308.8433 | 167749.982 | 175515.7858 | 3287.0449 | 0.34555 | 0.82122 | 1.6462 | 3.5321
19 | LLOQ 0.01 | R | 06/02/2023 17:01:19 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LLOQ 0.01.021 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.0010631 | 0.00041078 | 0.00062182 | 0.0053744 | 0.011228 | 0.00026875 | 0.00032255 | 0.00091741 | 0.0016303 | 0.00039137 | 0.0025396 | 0.00074151 | 0.026215 | 0.0024674 | 0.00081596 | 0.0026792 | 0.0013069 | 0.0024165 | -0.00068704 | -4.2978e-05 | 0.021551 | 0.00040169 | -0.00020307 | 1.9237e-05 | 1.8759e-05 | 5.3878e-05 | 7.4501e-05 | 6.4804e-06 | 0.001666 | 0.0017166 | 151811.7147 | 207220.2345 | 12185.5285 | 235343.3661 | 171146.8218 | 185712.7347 | 3680.4742 | 0.00035698 | 0.00091637 | 0.00203 | 0.004031
20 | CCV 10 ppb MAJORS | R | 06/02/2023 17:07:22 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.022 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.2365 | 0.49993 | 0.028771 | 0.14298 | 0.18358 | 0.59125 | 0.2202 | 0.83566 | 1.1599 | 0.36122 | 0.24573 | 0.69015 | 1.0832 | 2.3844 | 0.63998 | 1.5984 | 0.69208 | 0.33137 | 0.36209 | 0.08394 | 20.0831 | 0.34396 | 0.039382 | 0.0075261 | 0.023976 | 0.050904 | 0.033384 | 0.0090179 | 0.70073 | 0.68106 | 151695.1769 | 188879.1644 | 11114.9936 | 214186.4725 | 170593.2057 | 173564.123 | 3235.033 | 0.3473 | 0.82873 | 1.6727 | 3.4693
21 | CCV 10 ppb MAJORS | R | 06/02/2023 17:13:54 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.023 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.2302 | 0.48927 | 0.028394 | 0.13315 | 0.18156 | 0.59598 | 0.22093 | 0.83822 | 1.1536 | 0.36225 | 0.24758 | 0.73805 | 1.0571 | 2.3666 | 0.63674 | 1.5828 | 0.69522 | 0.33231 | 0.36606 | 0.082501 | 20.2886 | 0.35921 | 0.039437 | 0.0080328 | 0.025231 | 0.053787 | 0.034504 | 0.0094412 | 0.70138 | 0.71404 | 151280.8643 | 182241.17 | 11431.239 | 215573.8082 | 164422.9991 | 181938.5843 | 3227.6988 | 0.34611 | 0.82515 | 1.6859 | 3.6165
22 | LLOQ 0.1 | R | 06/02/2023 17:19:57 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LLOQ 0.1.024 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.012672 | 0.0042428 | 0.0010191 | 0.0035352 | 0.015348 | 0.0061493 | 0.0023722 | 0.0079488 | 0.011822 | 0.0033858 | 0.0075173 | 0.0073298 | 0.047574 | 0.02295 | 0.0070938 | 0.017459 | 0.0076976 | 0.0061023 | 0.0028271 | 0.00073184 | 0.19029 | 0.0034102 | -0.00015121 | 8.4083e-05 | 0.00023193 | 0.0005192 | 0.00040338 | 7.1143e-05 | 0.0067675 | 0.0070346 | 163654.5928 | 206499.229 | 12089.4481 | 244179.7047 | 181189.4878 | 184214.2113 | 3683.8085 | 0.0032729 | 0.0080241 | 0.017646 | 0.037116
23 | LBLANK | R | 06/02/2023 17:25:59 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LBLANK.025 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.00010343 | 0 | -7.9291e-06 | -0.0036671 | 0.015048 | 1.6267e-05 | -3.248e-05 | 4.3878e-05 | 0.00012352 | 4.2924e-05 | 0.00077148 | 0.00010006 | 0.021086 | -8.1399e-05 | -9.0016e-06 | 0.00023199 | 0.00024439 | -0.00028806 | -7.0698e-05 | -4.3739e-05 | -0.003836 | -4.8399e-05 | -0.00011868 | 6.521e-06 | -2.0295e-07 | 6.1317e-06 | -4.7203e-05 | 1.3018e-05 | 8.237e-05 | 5.1594e-05 | 150831.4389 | 204707.5949 | 12272.2697 | 230293.733 | 177548.5211 | 186473.8448 | 3747.1582 | 2.8823e-05 | 3.9692e-05 | 3.8746e-05 | 8.7339e-05
24 | LFB 6.25ppb & 2500ppb | R | 06/02/2023 17:32:01 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LFB 6.25ppb & 2500ppb.026 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 3.2427 | 1.2822 | 0.058234 | 0.12667 | 0.46155 | 1.589 | 0.578 | 0.55718 | 0.7637 | 0.24232 | 0.18072 | 0.44559 | 0.79715 | 1.5855 | 0.41878 | 1.0163 | 0.44934 | 0.22594 | 0.24148 | 0.053757 | 12.3004 | 0.22093 | 0.026331 | 0.005053 | 0.015522 | 0.032725 | 0.023143 | 0.0059489 | 0.46675 | 0.42816 | 138233.6768 | 181501.9314 | 10377.4352 | 196261.6271 | 152534.7005 | 159055.9435 | 3273.045 | 0.22793 | 0.53721 | 1.1131 | 2.2776
25 | MT 1 | R | 06/02/2023 17:38:02 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 1.027 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.4086 | 0.01174 | 0.96548 | 31.3695 | 0.046331 | 0.1813 | 0.012493 | 0.035016 | 0.060203 | 0.018246 | 2.0019 | 0.051905 | 8.7815 | 0.014994 | 0.021644 | 0.63586 | 0.2737 | 0.40574 | 0.055145 | 0.011417 | 2.9483 | 0.053345 | 0.0024506 | 0.00064961 | 0.0019294 | 0.0041682 | 0.0032718 | 0.00079822 | 0.0052546 | 0.0047921 | 144844.6375 | 197024.032 | 11442.9149 | 219361.3746 | 171321.164 | 160237.9986 | 3575.115 | 0.002474 | 0.0059379 | 0.031402 | 0.059662
26 | MT 1 LD | R | 06/02/2023 17:44:04 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 1 LD.028 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.43613 | 0.012153 | 1.0298 | 32.7299 | 0.049074 | 0.1959 | 0.013299 | 0.034939 | 0.060542 | 0.019825 | 2.0781 | 0.053938 | 8.5829 | 0.016159 | 0.021413 | 0.63353 | 0.29164 | 0.44546 | 0.060496 | 0.011582 | 3.1892 | 0.055881 | 0.0031117 | 0.00072748 | 0.0020423 | 0.0046507 | 0.0037322 | 0.00086148 | 0.0055735 | 0.0049322 | 134023.9811 | 194272.0411 | 11274.4472 | 206447.5151 | 167556.2411 | 161413.4497 | 3406.4065 | 0.0028358 | 0.0063985 | 0.033445 | 0.062299
27 | LBLANK | R | 06/02/2023 18:05:56 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LBLANK.029 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.00025056 | 3.1438e-05 | -2.8619e-05 | -0.0018928 | 0.0024266 | -7.796e-05 | -3.248e-05 | -5.861e-06 | 0.00050844 | 5.1527e-05 | 0.0011538 | 9.7539e-05 | 0.010171 | -7.5199e-05 | 0.00019788 | 0.00047105 | 0.00024304 | -0.00066722 | -0.00068984 | -9.744e-05 | -0.0013136 | -7.1438e-07 | -4.7305e-05 | 3.2356e-06 | 6.6077e-06 | 1.4197e-05 | -4.8249e-06 | 0 | 5.6994e-05 | 0.00012904 | 142124.1783 | 197414.6228 | 11533.0139 | 215951.5156 | 168258.4464 | 175336.8371 | 3644.4651 | 0 | 6.5355e-05 | 5.9722e-06 | 4.5925e-06
28 | STD 1 | R | 06/02/2023 18:11:58 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\STD 1.030 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.1405 | 0.054994 | 0.0050016 | 0.024117 | 0.03518 | 0.068256 | 0.023606 | 0.087914 | 0.1274 | 0.03818 | 0.054892 | 0.075401 | 0.24478 | 0.25192 | 0.066499 | 0.18026 | 0.077962 | 0.044216 | 0.038706 | 0.0077217 | 1.9741 | 0.0351 | 0.0039787 | 0.00090113 | 0.0025627 | 0.0054782 | 0.0034816 | 0.0010954 | 0.070827 | 0.068761 | 139194.8733 | 188660.9318 | 11125.9981 | 208790.0466 | 164122.7696 | 171734.3065 | 3353.0603 | 0.035752 | 0.084333 | 0.19112 | 0.37908
29 | STD 0.5 | R | 06/02/2023 18:18:06 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\STD 0.5.031 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.071931 | 0.025667 | 0.0042189 | 0.023708 | 0.028038 | 0.032428 | 0.010729 | 0.042887 | 0.061798 | 0.019336 | 0.027649 | 0.038211 | 0.13732 | 0.12062 | 0.034012 | 0.087788 | 0.04085 | 0.023354 | 0.017965 | 0.0043428 | 1.0776 | 0.019087 | 0.0011796 | 0.00039499 | 0.0013654 | 0.0026746 | 0.0014849 | 0.00050644 | 0.036545 | 0.035308 | 144190.7011 | 185612.7502 | 11343.5028 | 206054.1387 | 162058.7423 | 167877.3475 | 3285.7115 | 0.017858 | 0.043006 | 0.089252 | 0.18975
30 | ICV PE | R | 06/02/2023 18:24:08 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\ICV PE.032 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.4044 | 0.58284 | 0.0338 | 0.3503 | 0.207 | 0.68002 | 0.24612 | 0.89659 | 1.2283 | 0.3824 | 0.26796 | 0.77292 | 1.1494 | 2.4839 | 0.66755 | 1.6236 | 0.7067 | 0.33862 | 0.37661 | 0.073805 | 20.1814 | 0.36307 | 0.029058 | 0.0080315 | 0.025813 | 0.054971 | 0.035987 | 0.009645 | 0.67822 | 0.68479 | 138961.264 | 170011.6725 | 10382.7717 | 195913.7733 | 154115.5183 | 155238.2888 | 3058.9943 | 0.34545 | 0.82433 | 1.675 | 3.6315
31 | ICV PE | R | 06/02/2023 18:30:39 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\ICV PE.033 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.3807 | 0.56829 | 0.032905 | 0.32358 | 0.19787 | 0.65088 | 0.24576 | 0.8909 | 1.2112 | 0.38293 | 0.27171 | 0.76206 | 1.1291 | 2.4808 | 0.664 | 1.6212 | 0.71217 | 0.3416 | 0.36988 | 0.075038 | 19.7498 | 0.35478 | 0.028606 | 0.0082912 | 0.025235 | 0.053704 | 0.035081 | 0.0096734 | 0.68729 | 0.67894 | 140173.7185 | 176288.4629 | 10594.2605 | 195773.885 | 155613.7196 | 167457.2363 | 3167.6846 | 0.35193 | 0.83669 | 1.6504 | 3.5864
32 | CCV 1ppb TRACE | R | 06/02/2023 18:36:41 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.034 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.13457 | 0.054826 | 0.0028562 | 0.014959 | 0.030133 | 0.065782 | 0.023119 | 0.08706 | 0.12042 | 0.038173 | 0.027801 | 0.079285 | 0.13239 | 0.24682 | 0.068237 | 0.1668 | 0.075416 | 0.034087 | 0.037557 | 0.0084191 | 2.0382 | 0.037166 | 0.0041709 | 0.00089298 | 0.0025413 | 0.0057484 | 0.0034491 | 0.00094115 | 0.07138 | 0.073 | 140177.1374 | 182573.1799 | 11606.049 | 211623.3522 | 163291.273 | 172642.3142 | 3330.3889 | 0.035403 | 0.084464 | 0.18596 | 0.37688
33 | CCV 1ppb TRACE | R | 06/02/2023 18:43:13 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.035 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.14289 | 0.058274 | 0.0042944 | 0.010408 | 0.038725 | 0.065347 | 0.025571 | 0.090295 | 0.12841 | 0.039526 | 0.02772 | 0.077289 | 0.14714 | 0.2592 | 0.070799 | 0.17647 | 0.076834 | 0.034422 | 0.035866 | 0.007385 | 2.0066 | 0.0359 | 0.0038516 | 0.00077336 | 0.0024392 | 0.0054967 | 0.0035482 | 0.00097053 | 0.071162 | 0.068893 | 139511.253 | 188081.5374 | 11315.1461 | 206739.3439 | 161535.5529 | 168004.1199 | 3365.7301 | 0.036114 | 0.084621 | 0.18725 | 0.37734
34 | CCV 10 ppb MAJORS | R | 06/02/2023 18:49:16 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.036 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.4263 | 0.57795 | 0.032976 | 0.16395 | 0.21412 | 0.66027 | 0.24859 | 0.89941 | 1.2345 | 0.39163 | 0.26855 | 0.7684 | 1.137 | 2.5385 | 0.68564 | 1.6407 | 0.73156 | 0.34951 | 0.3781 | 0.082309 | 19.934 | 0.35618 | 0.041345 | 0.0080062 | 0.025638 | 0.052934 | 0.03537 | 0.0093258 | 0.7193 | 0.69285 | 139313.3907 | 176965.3537 | 10864.1297 | 194370.8024 | 154444.0407 | 160979.4707 | 3164.3507 | 0.35116 | 0.83347 | 1.6615 | 3.5535
35 | CCV 10 ppb MAJORS | R | 06/02/2023 18:55:47 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.037 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.4006 | 0.55126 | 0.031618 | 0.1624 | 0.2069 | 0.65297 | 0.23732 | 0.90155 | 1.2327 | 0.39119 | 0.26546 | 0.76002 | 1.1306 | 2.5139 | 0.67786 | 1.6544 | 0.72322 | 0.35039 | 0.38214 | 0.085367 | 20.0352 | 0.36153 | 0.041771 | 0.0083056 | 0.025333 | 0.052926 | 0.0354 | 0.00943 | 0.71956 | 0.70512 | 137711.9193 | 174914.3109 | 10709.6798 | 195365.2362 | 147201.2731 | 156556.6226 | 3156.3487 | 0.34914 | 0.82466 | 1.6635 | 3.701
36 | MT 1 | R | 06/02/2023 19:01:50 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 1.038 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.45386 | 0.012694 | 1.0972 | 31.9784 | 0.054081 | 0.19738 | 0.014956 | 0.036719 | 0.06291 | 0.019212 | 2.1182 | 0.063018 | 9.2517 | 0.017049 | 0.02165 | 0.63946 | 0.27672 | 0.41949 | 0.056241 | 0.011293 | 3.5285 | 0.062554 | 0.0026809 | 0.00075834 | 0.0023331 | 0.0050842 | 0.0032794 | 0.00083647 | 0.0054975 | 0.0058007 | 136255.3949 | 167334.0119 | 11002.9024 | 203483.6597 | 158164.14 | 165982.7923 | 2971.6451 | 0.002605 | 0.0062049 | 0.029757 | 0.058429
37 | MT 1 LD | R | 06/02/2023 19:07:53 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 1 LD.039 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.48855 | 0.01367 | 1.1947 | 33.4862 | 0.056284 | 0.2169 | 0.016293 | 0.036828 | 0.064325 | 0.019398 | 2.0965 | 0.058109 | 9.4936 | 0.017534 | 0.025185 | 0.68121 | 0.28978 | 0.44406 | 0.059574 | 0.011764 | 3.1544 | 0.057901 | 0.0025982 | 0.00072543 | 0.0022296 | 0.004825 | 0.0035941 | 0.00090809 | 0.0052154 | 0.0049809 | 133188.6655 | 180148.2406 | 10453.4981 | 199046.5042 | 157483.2956 | 160897.8818 | 3304.3822 | 0.0028547 | 0.0063771 | 0.030373 | 0.059844
38 | MT 2 | R | 06/02/2023 19:13:55 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 2.040 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.53064 | 0.012745 | 1.2795 | 36.0314 | 0.059004 | 0.22588 | 0.015773 | 0.03857 | 0.067442 | 0.021094 | 2.1802 | 0.056762 | 9.2846 | 0.017587 | 0.022997 | 0.70775 | 0.31517 | 0.46764 | 0.062467 | 0.013626 | 3.5023 | 0.062213 | 0.0028872 | 0.00067219 | 0.0024536 | 0.0052381 | 0.0036912 | 0.0008993 | 0.0073007 | 0.007065 | 132124.199 | 177085.044 | 10714.0171 | 194790.4249 | 155296.9091 | 159628.3091 | 3157.0175 | 0.0027525 | 0.0065032 | 0.030615 | 0.061103
39 | MT 2 LSP | R | 06/02/2023 19:19:58 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 2 LSP.041 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 4.3182 | 1.4826 | 1.3936 | 41.8948 | 0.56132 | 2.0279 | 0.64662 | 0.64594 | 0.89382 | 0.27407 | 2.5189 | 0.54682 | 10.8127 | 1.6521 | 0.46222 | 1.7729 | 0.75842 | 0.71414 | 0.30827 | 0.070874 | 16.7804 | 0.29239 | 0.029945 | 0.0059737 | 0.018737 | 0.039205 | 0.026742 | 0.0073506 | 0.46582 | 0.44301 | 124900.3051 | 160710.2982 | 9356.0679 | 178301.8517 | 138288.8266 | 143964.3133 | 2800.2744 | 0.22334 | 0.52911 | 1.092 | 2.2888
40 | MT 2 LSP | R | 06/02/2023 19:26:09 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 2 LSP.042 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 4.0308 | 1.3731 | 1.3115 | 40.2206 | 0.53379 | 1.8732 | 0.60861 | 0.6143 | 0.8334 | 0.27123 | 2.4785 | 0.54738 | 10.247 | 1.594 | 0.44588 | 1.711 | 0.75617 | 0.7065 | 0.3096 | 0.068438 | 16.6662 | 0.29734 | 0.02956 | 0.0061432 | 0.01862 | 0.040031 | 0.026552 | 0.0072185 | 0.46536 | 0.45929 | 128148.2593 | 160164.9291 | 9883.4199 | 183852.8584 | 143600.1409 | 148532.8516 | 2858.9529 | 0.22124 | 0.52496 | 1.0989 | 2.2753
41 | MT 3 | R | 06/02/2023 19:32:11 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 3 .043 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.41525 | 0.011055 | 0.99547 | 29.3032 | 0.052843 | 0.18153 | 0.011494 | 0.032329 | 0.07242 | 0.022677 | 1.8289 | 0.047654 | 7.5499 | 0.015117 | 0.024377 | 0.58334 | 0.26456 | 0.39126 | 0.055652 | 0.012465 | 2.9512 | 0.05309 | 0.0028717 | 0.00075076 | 0.0020575 | 0.0043284 | 0.003493 | 0.00080035 | 0.005558 | 0.0049409 | 135927.6343 | 188286.7732 | 11285.1229 | 206055.5583 | 157327.6248 | 166069.064 | 3387.0684 | 0.0023532 | 0.0058419 | 0.030077 | 0.058984
42 | MT 4 | R | 06/02/2023 19:38:14 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 4.044 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.54901 | 0.013037 | 1.3413 | 37.664 | 0.070501 | 0.24287 | 0.016055 | 0.04286 | 0.072316 | 0.021793 | 2.9308 | 0.06891 | 12.9222 | 0.020139 | 0.024585 | 0.75374 | 0.32564 | 0.49141 | 0.068765 | 0.01454 | 3.6737 | 0.066721 | 0.0030748 | 0.00080522 | 0.002466 | 0.0055018 | 0.0039668 | 0.00092042 | 0.0063557 | 0.0058867 | 133692.7936 | 179527.8694 | 10609.9386 | 199132.4504 | 156175.7518 | 161023.3211 | 3259.7052 | 0.0031111 | 0.0073076 | 0.032929 | 0.064724
43 | LBLANK | R | 06/02/2023 19:44:17 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LBLANK.045 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.8698e-05 | 3.1323e-05 | -0.00020352 | 0.0058835 | 0.019689 | 0.00051902 | -3.248e-05 | 2.8274e-05 | 0.00026038 | 3.194e-05 | 0.0013575 | 0.00010133 | 0.026725 | -3.7997e-05 | 0.00028719 | 0.00018643 | 0.00031588 | -0.00030014 | -0.0016129 | -6.6194e-05 | 0.00018019 | 2.478e-05 | -2.8555e-05 | 3.6904e-06 | 8.5336e-06 | 2.4101e-05 | 0.00010069 | 7.3695e-06 | 0.00013049 | 5.4384e-05 | 135236.4674 | 180037.4239 | 10735.6995 | 195111.13 | 155773.6461 | 174898.4378 | 3292.3794 | 3.6979e-06 | 0.00010174 | -3.3317e-05 | 1.6024e-05
44 | CCV 1ppb TRACE | R | 06/02/2023 19:50:20 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.046 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.1499 | 0.060241 | 0.0042183 | 0.012413 | 0.041435 | 0.071746 | 0.026489 | 0.097589 | 0.12891 | 0.041856 | 0.030213 | 0.080261 | 0.15292 | 0.26712 | 0.074414 | 0.17896 | 0.07829 | 0.036291 | 0.035423 | 0.0086034 | 1.986 | 0.036385 | 0.0048724 | 0.0008111 | 0.002543 | 0.005611 | 0.0033078 | 0.00098876 | 0.072125 | 0.070922 | 128252.7365 | 172602.4461 | 10644.9647 | 193088.909 | 151309.8731 | 159650.519 | 3160.3496 | 0.035905 | 0.085557 | 0.1859 | 0.37162
45 | CCV 1ppb TRACE | R | 06/02/2023 19:56:52 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.047 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.14897 | 0.061615 | 0.004537 | 0.010063 | 0.042651 | 0.067829 | 0.024983 | 0.09359 | 0.12937 | 0.041024 | 0.029049 | 0.078574 | 0.14808 | 0.26041 | 0.072177 | 0.18067 | 0.079025 | 0.035118 | 0.038063 | 0.0085226 | 1.965 | 0.035596 | 0.0043478 | 0.00078854 | 0.0025328 | 0.0054789 | 0.0035363 | 0.001096 | 0.071558 | 0.068791 | 128243.058 | 175163.3795 | 10753.0455 | 193599.0231 | 152742.968 | 159227.1506 | 3173.0192 | 0.035183 | 0.084354 | 0.18035 | 0.36722
46 | CCV 10 ppb MAJORS | R | 06/02/2023 20:02:55 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.048 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.5283 | 0.61943 | 0.035435 | 0.17922 | 0.23394 | 0.70955 | 0.25657 | 0.95747 | 1.299 | 0.40665 | 0.28197 | 0.79011 | 1.1963 | 2.6745 | 0.71221 | 1.6921 | 0.73865 | 0.35873 | 0.38558 | 0.086772 | 20.1942 | 0.36315 | 0.042545 | 0.0082517 | 0.025174 | 0.053572 | 0.035754 | 0.0095553 | 0.71781 | 0.69911 | 128424.0928 | 161333.3506 | 9884.0852 | 180318.8873 | 141062.5391 | 145793.7914 | 2900.2944 | 0.34954 | 0.82484 | 1.5835 | 3.4839
47 | CCV 10 ppb MAJORS | R | 06/02/2023 20:09:27 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.049 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.5039 | 0.60079 | 0.035007 | 0.17537 | 0.21752 | 0.69474 | 0.25244 | 0.94553 | 1.2716 | 0.41283 | 0.28283 | 0.80294 | 1.155 | 2.6395 | 0.70161 | 1.6742 | 0.75746 | 0.35954 | 0.38073 | 0.086488 | 20.2164 | 0.36643 | 0.043269 | 0.0082379 | 0.025323 | 0.055173 | 0.035674 | 0.009692 | 0.73237 | 0.71034 | 129573.0782 | 162414.0373 | 10226.9935 | 179926.8098 | 135480.1754 | 149038.7763 | 2944.3034 | 0.35191 | 0.83328 | 1.5958 | 3.7086
48 | MT 5 | R | 06/02/2023 20:15:30 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 5.050 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.55921 | 0.014894 | 1.3077 | 39.1437 | 0.066487 | 0.23382 | 0.017118 | 0.040847 | 0.069391 | 0.021194 | 2.4169 | 0.062785 | 10.6446 | 0.020294 | 0.024782 | 0.7175 | 0.30853 | 0.47594 | 0.063927 | 0.012673 | 3.2152 | 0.060078 | 0.0033837 | 0.0007471 | 0.0023125 | 0.004857 | 0.0037758 | 0.00079824 | 0.0058257 | 0.0052248 | 130798.729 | 178717.776 | 10548.2274 | 193578.1334 | 151337.4828 | 154341.3133 | 3337.7232 | 0.0029203 | 0.0067907 | 0.029526 | 0.060733
49 | MT 6 | R | 06/02/2023 20:21:32 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 6.051 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.34 | 0.032491 | 3.0416 | 89.3793 | 0.13545 | 0.56844 | 0.040172 | 0.094052 | 0.14321 | 0.043214 | 5.2959 | 0.14207 | 23.284 | 0.04222 | 0.054624 | 1.6864 | 0.7175 | 1.104 | 0.15036 | 0.03413 | 8.1568 | 0.14769 | 0.0071242 | 0.0016992 | 0.0057819 | 0.011588 | 0.0083905 | 0.0021869 | 0.013216 | 0.011563 | 127129.159 | 168286.7897 | 10156.2764 | 186880.4102 | 143389.9955 | 154815.7027 | 3046.9916 | 0.0068676 | 0.015638 | 0.055548 | 0.11611
50 | MT 7 | R | 06/02/2023 20:27:35 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 7.052 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.324 | 0.034112 | 3.1695 | 89.5858 | 0.135 | 0.57514 | 0.035797 | 0.091374 | 0.14358 | 0.045201 | 5.1794 | 0.13816 | 22.705 | 0.043505 | 0.059703 | 1.676 | 0.72053 | 1.1021 | 0.15133 | 0.032719 | 8.1107 | 0.14809 | 0.0075025 | 0.0016377 | 0.0057901 | 0.011731 | 0.0090365 | 0.0020197 | 0.014214 | 0.012086 | 129252.1457 | 167696.549 | 10167.2835 | 188551.2353 | 147612.9305 | 150956.3899 | 3067.6633 | 0.0063846 | 0.015587 | 0.065796 | 0.13163
51 | MT 8 | R | 06/02/2023 20:33:38 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 8.053 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.2942 | 0.032433 | 3.0633 | 84.5096 | 0.13419 | 0.55956 | 0.038063 | 0.089645 | 0.14809 | 0.043241 | 4.7276 | 0.13315 | 21.5059 | 0.042155 | 0.058079 | 1.6231 | 0.67009 | 1.0153 | 0.14024 | 0.031681 | 7.843 | 0.13799 | 0.0068983 | 0.0015829 | 0.0052912 | 0.011031 | 0.0079411 | 0.0019961 | 0.012525 | 0.011607 | 128172.4594 | 169297.3043 | 10030.871 | 189685.047 | 148808.5385 | 151524.0481 | 2979.6443 | 0.006232 | 0.014211 | 0.054533 | 0.11077
52 | MT 9 | R | 06/02/2023 20:39:40 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 9.054 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.2355 | 0.029699 | 2.9329 | 87.0153 | 0.12742 | 0.51874 | 0.035345 | 0.088137 | 0.13468 | 0.042347 | 4.8758 | 0.13136 | 20.7734 | 0.041589 | 0.052568 | 1.8193 | 0.79708 | 1.0588 | 0.15021 | 0.030859 | 7.6971 | 0.13875 | 0.0077553 | 0.0017793 | 0.005227 | 0.011152 | 0.0088527 | 0.002068 | 0.013033 | 0.011122 | 120959.5698 | 166886.0691 | 10207.646 | 183249.8391 | 146179.4366 | 147215.3498 | 3008.3167 | 0.0062579 | 0.014977 | 0.055883 | 0.10733
53 | MT 10 | R | 06/02/2023 20:45:42 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 10.055 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.2875 | 0.034194 | 3.1401 | 91.9285 | 0.13223 | 0.55915 | 0.036869 | 0.092018 | 0.16281 | 0.049224 | 5.3504 | 0.14637 | 23.1929 | 0.042679 | 0.057653 | 1.86 | 0.79487 | 1.1241 | 0.15274 | 0.031561 | 8.3637 | 0.14913 | 0.0072663 | 0.0016923 | 0.0055954 | 0.012151 | 0.0084888 | 0.0020594 | 0.013472 | 0.012757 | 123539.8879 | 161519.2994 | 9978.1514 | 183500.4057 | 144693.1356 | 147368.2118 | 2880.2904 | 0.0065197 | 0.014863 | 0.055186 | 0.10956
54 | MT 11 | R | 06/02/2023 20:51:44 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 11 .056 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.5391 | 0.063956 | 6.0339 | 182.8873 | 0.24076 | 1.094 | 0.080045 | 0.18557 | 0.27416 | 0.088152 | 10.3939 | 0.29231 | 44.9914 | 0.085357 | 0.11267 | 3.1251 | 1.3608 | 2.2057 | 0.29883 | 0.062084 | 16.4431 | 0.29106 | 0.014537 | 0.003586 | 0.010973 | 0.023905 | 0.017964 | 0.0043324 | 0.028053 | 0.023388 | 121253.7655 | 159139.6013 | 9787.6852 | 177527.5964 | 139817.6084 | 144517.4363 | 2816.9444 | 0.013066 | 0.030075 | 0.096588 | 0.19966
55 | MT 11 LD | R | 06/02/2023 20:57:47 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 11 LD.057 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.5769 | 0.061376 | 6.0207 | 174.5045 | 0.23714 | 1.0707 | 0.081598 | 0.18022 | 0.28114 | 0.088584 | 10.206 | 0.28312 | 44.1832 | 0.20918 | 0.11005 | 3.1248 | 1.3558 | 2.1822 | 0.30119 | 0.063969 | 16.2771 | 0.28258 | 0.014028 | 0.0034005 | 0.010836 | 0.023509 | 0.017396 | 0.0042191 | 0.026521 | 0.022602 | 120725.3317 | 161174.7755 | 9701.9611 | 178193.9609 | 140444.9968 | 149055.6547 | 2799.6078 | 0.012595 | 0.029435 | 0.09938 | 0.19704
56 | MT 11 LD | R | 06/02/2023 21:04:18 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 11 LD.058 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.7037 | 0.068009 | 6.4646 | 185.7532 | 0.25479 | 1.1662 | 0.086664 | 0.19382 | 0.30596 | 0.086693 | 10.2864 | 0.28997 | 47.8159 | 0.22209 | 0.11952 | 3.3659 | 1.34 | 2.1617 | 0.29745 | 0.068289 | 16.1864 | 0.28951 | 0.013331 | 0.0033264 | 0.011013 | 0.023167 | 0.017199 | 0.0040387 | 0.025316 | 0.023152 | 123525.8371 | 160919.9653 | 9337.7394 | 183494.2989 | 146405.128 | 144846.0727 | 2879.6237 | 0.012362 | 0.028736 | 0.095947 | 0.19056
57 | MT 12 | R | 06/02/2023 21:10:20 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 12.059 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.6486 | 0.067986 | 6.1164 | 190.5542 | 0.24332 | 1.1257 | 0.078598 | 0.18468 | 0.28091 | 0.092827 | 10.8538 | 0.28461 | 44.3453 | 0.090751 | 0.10715 | 3.2161 | 1.4746 | 2.3284 | 0.31147 | 0.064873 | 16.1393 | 0.29662 | 0.014822 | 0.0038283 | 0.011856 | 0.024633 | 0.017979 | 0.0044554 | 0.027512 | 0.023143 | 120311.7469 | 159872.3432 | 9710.637 | 174475.8863 | 140914.4996 | 143179.5855 | 2937.6353 | 0.013926 | 0.031464 | 0.10007 | 0.20152
58 | MT 12 LSP | R | 06/02/2023 21:16:23 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 12 LSP.060 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 6.8743 | 1.6271 | 6.2708 | 197.546 | 0.8009 | 3.1223 | 0.7574 | 0.83775 | 1.1336 | 0.36986 | 11.1152 | 0.81043 | 45.866 | 1.8166 | 0.57538 | 4.2867 | 1.8809 | 2.5402 | 0.59243 | 0.12484 | 29.7297 | 0.54719 | 0.046724 | 0.0094023 | 0.029502 | 0.061475 | 0.044305 | 0.011294 | 0.50587 | 0.48133 | 115192.398 | 148272.1867 | 8982.8234 | 164505.223 | 130282.0417 | 139379.3903 | 2731.5948 | 0.23457 | 0.5556 | 1.1484 | 2.378
59 | LBLANK | R | 06/02/2023 21:22:26 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LBLANK.061 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.00049658 | 0.00012517 | 0.00016931 | 0.0019718 | 0.022395 | 0.00017966 | -3.248e-05 | 0.00018451 | 0.00041593 | 3.3251e-05 | 0.001774 | 0.00012054 | 0.034837 | -3.8213e-05 | 0.00075 | 0.00090089 | 0.00043833 | 0.00076102 | -0.0016178 | -3.4858e-05 | -0.0028994 | -2.9684e-05 | 0.00029802 | 3.8491e-06 | 1.7147e-05 | 1.8797e-05 | 0.00018264 | 3.9464e-06 | 0.00088631 | 8.5954e-05 | 134720.4606 | 170805.145 | 10731.363 | 192305.7812 | 155442.5215 | 153069.7151 | 3167.0177 | 2.443e-05 | 8.3526e-05 | 1.8592e-05 | 7.5579e-06
60 | CCV 1ppb TRACE | R | 06/02/2023 21:28:29 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.062 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.16656 | 0.066306 | 0.0035183 | 0.015033 | 0.048933 | 0.076211 | 0.027269 | 0.10133 | 0.13419 | 0.043 | 0.030538 | 0.080748 | 0.16382 | 0.27699 | 0.076033 | 0.19206 | 0.082754 | 0.037859 | 0.037824 | 0.0084524 | 2.0056 | 0.036814 | 0.0045987 | 0.00070889 | 0.0025934 | 0.005465 | 0.0037757 | 0.00092964 | 0.073242 | 0.068748 | 127328.3794 | 168464.2953 | 10240.3378 | 184730.1891 | 146846.3703 | 151761.863 | 3091.668 | 0.036172 | 0.085559 | 0.1728 | 0.35532
61 | CCV 1ppb TRACE | R | 06/02/2023 21:35:00 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.063 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.1687 | 0.069525 | 0.0042116 | 0.014054 | 0.049718 | 0.076217 | 0.027765 | 0.099913 | 0.13474 | 0.044037 | 0.032193 | 0.082738 | 0.16274 | 0.27727 | 0.073678 | 0.18691 | 0.085832 | 0.038293 | 0.036389 | 0.0080449 | 2.0082 | 0.0375 | 0.004609 | 0.00079514 | 0.0027738 | 0.0055754 | 0.0034065 | 0.00098346 | 0.075099 | 0.070552 | 123219.8364 | 163585.5071 | 10398.4512 | 178047.8048 | 144301.8385 | 147804.7071 | 3052.9931 | 0.036636 | 0.086797 | 0.17291 | 0.35524
62 | LFB 10ppb | R | 06/02/2023 21:41:03 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LFB 10ppb.064 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.7109 | 0.6694 | 0.036605 | 0.20699 | 0.24526 | 0.75971 | 0.26724 | 0.97834 | 1.315 | 0.43521 | 0.30943 | 0.7263 | 1.2608 | 2.6587 | 0.72298 | 1.7198 | 0.77356 | 0.49662 | 0.3729 | 0.086887 | 18.0314 | 0.33397 | 0.04227 | 0.0075061 | 0.024078 | 0.049827 | 0.033398 | 0.0092496 | 0.72627 | 0.63878 | 124459.7536 | 165483.5535 | 9448.458 | 167732.4454 | 134285.6268 | 138930.3167 | 3067.6669 | 0.34663 | 0.82053 | 1.4659 | 3.2601
63 | LFB 10ppb | R | 06/02/2023 21:47:33 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LFB 10ppb.065 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.6762 | 0.66025 | 0.038436 | 0.22194 | 0.24752 | 0.74503 | 0.26487 | 0.97644 | 1.3179 | 0.42579 | 0.30424 | 0.81653 | 1.2729 | 2.7039 | 0.73255 | 1.7341 | 0.75479 | 0.48208 | 0.39692 | 0.086427 | 20.8691 | 0.37686 | 0.044526 | 0.0085467 | 0.025696 | 0.054211 | 0.037278 | 0.0098498 | 0.71325 | 0.71184 | 118736.0864 | 150460.2357 | 9509.4977 | 170571.7415 | 135062.8438 | 137357.2477 | 2718.2587 | 0.34256 | 0.81358 | 1.5257 | 3.2498
64 | CCV 10 ppb MAJORS | R | 06/02/2023 21:53:35 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.066 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.6964 | 0.67547 | 0.037704 | 0.20775 | 0.25355 | 0.7564 | 0.26621 | 0.97982 | 1.3328 | 0.42259 | 0.28857 | 0.82402 | 1.2167 | 2.715 | 0.73993 | 1.7369 | 0.75084 | 0.36763 | 0.46058 | 0.091876 | 20.4533 | 0.37405 | 0.051659 | 0.0080003 | 0.025672 | 0.052951 | 0.042369 | 0.009769 | 0.70997 | 0.70967 | 107144.1744 | 153110.1867 | 9617.5698 | 175211.1692 | 134292.2827 | 136320.006 | 2799.6076 | 0.33765 | 0.79785 | 1.7714 | 3.3401
65 | CCV 10 ppb MAJORS | R | 06/02/2023 22:00:06 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.067 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.7076 | 0.66402 | 0.036428 | 0.17312 | 0.24412 | 0.76171 | 0.27053 | 0.99114 | 1.3501 | 0.46243 | 0.31226 | 0.82286 | 1.2232 | 2.7359 | 0.72745 | 1.7427 | 0.82149 | 0.39253 | 0.39949 | 0.089492 | 20.1642 | 0.37531 | 0.04467 | 0.0083773 | 0.025769 | 0.055622 | 0.036903 | 0.009596 | 0.77327 | 0.71163 | 119054.0769 | 151671.5693 | 9761.3517 | 163865.8021 | 133803.7286 | 151602.3969 | 2822.9456 | 0.36724 | 0.86757 | 1.5471 | 3.3421
66 | MT 13 | R | 06/02/2023 22:06:09 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 13.068 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 3.088 | 0.078998 | 7.0741 | 223.8266 | 0.28591 | 1.2841 | 0.088716 | 0.21389 | 0.31751 | 0.09765 | 11.5378 | 0.31902 | 49.5177 | 0.087507 | 0.1183 | 3.5735 | 1.5174 | 2.3995 | 0.35463 | 0.06998 | 17.8023 | 0.3214 | 0.016989 | 0.003922 | 0.011857 | 0.025478 | 0.0194 | 0.0042798 | 0.030683 | 0.024159 | 112314.8915 | 154320.4279 | 9272.0083 | 171008.927 | 136888.0986 | 133641.6149 | 2789.606 | 0.014143 | 0.031526 | 0.10625 | 0.2019
67 | MT 14 | R | 06/02/2023 22:12:11 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 14.069 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 3.0847 | 0.076851 | 6.8657 | 213.5907 | 0.27718 | 1.2734 | 0.081895 | 0.20761 | 0.30918 | 0.10413 | 12.6687 | 0.34699 | 48.8297 | 0.093785 | 0.12249 | 3.5091 | 1.498 | 2.3431 | 0.33309 | 0.070029 | 17.337 | 0.3161 | 0.015155 | 0.00393 | 0.012003 | 0.025808 | 0.018419 | 0.004668 | 0.031497 | 0.024775 | 116802.4074 | 152026.7135 | 9298.3586 | 171074.9254 | 134229.0125 | 134756.6352 | 2774.2696 | 0.013696 | 0.031049 | 0.10236 | 0.20851
68 | MT 15 | R | 06/02/2023 22:18:14 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 15.070 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 3.0472 | 0.077096 | 6.8128 | 213.6721 | 0.28024 | 1.2799 | 0.082063 | 0.20658 | 0.31929 | 0.10128 | 11.2293 | 0.57582 | 48.4489 | 0.090113 | 0.15223 | 3.4951 | 1.4734 | 2.3588 | 0.33382 | 0.066473 | 17.4256 | 0.31796 | 0.015741 | 0.0038693 | 0.012313 | 0.02593 | 0.018712 | 0.0046517 | 0.032116 | 0.025827 | 114543.3948 | 151612.3087 | 9167.6077 | 169976.581 | 134488.0865 | 136857.7794 | 2768.9352 | 0.013983 | 0.032078 | 0.10177 | 0.20384
69 | MT 16 | R | 06/02/2023 22:24:17 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 16.071 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.90993 | 0.051603 | 0.058133 | 169.8521 | 0.19498 | 0.4986 | 0.0093444 | 0.019267 | 0.29959 | 0.094724 | 27.1131 | 0.48596 | 101.212 | 0.046405 | 0.12231 | 3.5097 | 1.5296 | 2.0608 | 0.2487 | 0.043676 | 11.4614 | 0.21183 | 0.016614 | 0.0034849 | 0.011194 | 0.024045 | 0.01855 | 0.0041112 | 0.14824 | 0.1409 | 106564.0197 | 157611.3892 | 9741.6538 | 174871.8643 | 140233.603 | 144563.1083 | 2915.631 | 0.41951 | 0.95475 | 0.043803 | 0.074086
70 | MT 17 | R | 06/02/2023 22:30:20 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 17.072 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.1375 | 0.11549 | 0.12206 | 379.4033 | 0.41939 | 1.1758 | 0.02096 | 0.044208 | 0.65964 | 0.20211 | 59.6502 | 1.171 | 243.1829 | 0.10655 | 0.27635 | 7.8513 | 3.279 | 4.4659 | 0.50008 | 0.10366 | 28.6597 | 0.51424 | 0.03199 | 0.0087356 | 0.027766 | 0.057992 | 0.036844 | 0.010186 | 0.31763 | 0.33447 | 111905.9737 | 144462.9091 | 9068.2106 | 169229.2388 | 133963.8739 | 141494.3302 | 2590.9025 | 0.91663 | 2.109 | 0.07676 | 0.15251
71 | MT 18 | R | 06/02/2023 22:36:32 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 18.073 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.8769 | 0.098436 | 0.10761 | 344.1271 | 0.36777 | 1.031 | 0.018602 | 0.039382 | 0.50928 | 0.1617 | 52.4835 | 0.92175 | 194.0661 | 0.092545 | 0.24763 | 6.6746 | 2.8779 | 3.8651 | 0.43971 | 0.087448 | 22.921 | 0.40825 | 0.027372 | 0.0068474 | 0.021835 | 0.045999 | 0.032086 | 0.0083635 | 0.28027 | 0.27104 | 110759.7225 | 151578.6203 | 8981.1566 | 167778.1924 | 132772.1802 | 134960.9842 | 2699.5884 | 0.79398 | 1.8065 | 0.06138 | 0.11933
72 | MT 19 | R | 06/02/2023 22:42:43 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 19.074 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.93615 | 0.05226 | 0.056815 | 173.09 | 0.19383 | 0.50849 | 0.0093988 | 0.019196 | 0.24745 | 0.076507 | 26.8005 | 0.464 | 96.3124 | 0.047769 | 0.19248 | 3.875 | 1.6964 | 1.9785 | 0.2089 | 0.043297 | 10.8856 | 0.20519 | 0.013489 | 0.0032271 | 0.010903 | 0.022886 | 0.01523 | 0.0039849 | 0.13643 | 0.13218 | 114056.8542 | 150063.4284 | 9047.1989 | 168498.4729 | 137121.638 | 135145.5984 | 2828.2799 | 0.399 | 0.91058 | 0.038951 | 0.075004
73 | MT 20 | R | 06/02/2023 22:48:45 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 20.075 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.0289 | 0.10263 | 0.11269 | 359.9405 | 0.38631 | 1.0792 | 0.018441 | 0.038236 | 0.51776 | 0.16482 | 54.2401 | 0.95424 | 201.203 | 0.091751 | 0.22378 | 7.0105 | 3.0172 | 4.0491 | 0.42927 | 0.091421 | 23.3243 | 0.41798 | 0.027361 | 0.0070146 | 0.022891 | 0.046938 | 0.031958 | 0.0083935 | 0.27731 | 0.26684 | 112145.1364 | 150600.4209 | 8730.6674 | 163576.5208 | 132230.5354 | 132943.8871 | 2700.2554 | 0.82805 | 1.8808 | 0.062446 | 0.1229
74 | MT 21 | R | 06/02/2023 22:54:57 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 21.076 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.0515 | 0.060796 | 0.061704 | 209.2486 | 0.20335 | 0.56428 | 0.01064 | 0.019677 | 0.30688 | 0.098775 | 29.4398 | 0.51016 | 107.4444 | 0.049644 | 0.13793 | 3.6693 | 1.599 | 2.2311 | 0.22146 | 0.046236 | 11.7145 | 0.21512 | 0.01413 | 0.0036349 | 0.011409 | 0.023946 | 0.016755 | 0.0040505 | 0.14687 | 0.14038 | 117200.7257 | 154500.8557 | 8900.105 | 166878.5796 | 140809.8097 | 125937.3574 | 2838.9491 | 0.43309 | 1.0019 | 0.043747 | 0.083393
75 | MT 21 LD | R | 06/02/2023 23:00:59 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 21 LD.077 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.0009575 | 0 | -0.00035719 | 0.0048578 | -0.031752 | 0.00013627 | -3.248e-05 | 4.2014e-06 | 0.00072376 | 4.9163e-05 | 0.0016434 | 4.4e-05 | 0.018608 | -3.0649e-05 | 7.8863e-05 | 0.0040733 | 0.00022362 | -0.0013512 | 0.021941 | -3.0284e-05 | -0.0039394 | -5.5076e-05 | -0.0060837 | 3.7373e-06 | 9.037e-06 | 9.9436e-06 | -0.00017467 | 0 | 0.00092053 | 2.5778e-05 | 133608.6325 | 175751.2744 | 10064.2137 | 204223.7019 | 159276.7328 | 158094.7596 | 3002.3155 | 1.9581e-05 | 5.0808e-05 | -0.00012463 | -0.00019767
76 | LBLANK | R | 06/02/2023 23:07:00 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LBLANK.078 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.00032367 | 0 | 0.00012243 | 0.0069238 | 0.02222 | 0.0005023 | 5.344e-06 | 6.9706e-05 | 0.00060665 | 0.00016391 | 0.0028726 | 9.4416e-05 | 0.040283 | -4.9167e-05 | 0.00049317 | 0.00075164 | 0.00039315 | -0.00047832 | -0.0034659 | -9.744e-05 | -0.0022176 | -2.0553e-05 | 0.00054238 | 8.4772e-06 | 3.0791e-06 | 2.1461e-05 | -6.1299e-05 | 0 | 0.00021396 | 4.5406e-05 | 113078.4538 | 155329.5126 | 9327.3939 | 171589.7764 | 137289.9279 | 136539.5429 | 2794.2733 | 1.9457e-05 | 6.1399e-05 | -2.522e-05 | 2.037e-05
77 | CCV 1ppb TRACE | R | 06/02/2023 23:13:03 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.079 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.20703 | 0.080884 | 0.0043059 | 0.027373 | 0.052745 | 0.088181 | 0.029448 | 0.10785 | 0.13788 | 0.042925 | 0.030652 | 0.086501 | 0.16224 | 0.27345 | 0.073524 | 0.19074 | 0.077036 | 0.035395 | 0.039648 | 0.0085292 | 2.1013 | 0.037474 | 0.0042582 | 0.00080686 | 0.0027651 | 0.005485 | 0.0039812 | 0.0010317 | 0.069082 | 0.070527 | 114126.3262 | 148660.8798 | 8806.7137 | 173714.0749 | 135349.0451 | 125986.4378 | 2653.5802 | 0.033527 | 0.079883 | 0.16889 | 0.3488
78 | CCV 1ppb TRACE | R | 06/02/2023 23:19:35 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.080 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.20855 | 0.08167 | 0.005352 | 0.024341 | 0.05903 | 0.091247 | 0.028038 | 0.10525 | 0.14331 | 0.046326 | 0.033081 | 0.086206 | 0.17415 | 0.2875 | 0.07552 | 0.19127 | 0.083414 | 0.037879 | 0.035241 | 0.0094716 | 2.0958 | 0.037908 | 0.0043428 | 0.00087505 | 0.0027618 | 0.0054802 | 0.0035285 | 0.0010148 | 0.073065 | 0.071293 | 116004.5935 | 147128.4808 | 8828.0607 | 166992.3963 | 137485.5039 | 130908.414 | 2659.5809 | 0.035754 | 0.083901 | 0.16949 | 0.34302
79 | CCV 10 ppb MAJORS | R | 06/02/2023 23:25:38 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.081 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.1212 | 0.79126 | 0.04194 | 0.22514 | 0.28029 | 0.90294 | 0.29816 | 1.0288 | 1.4057 | 0.44625 | 0.3085 | 0.8574 | 1.2608 | 2.7917 | 0.75854 | 1.7847 | 0.77383 | 0.37661 | 0.40144 | 0.093255 | 20.6711 | 0.37879 | 0.043487 | 0.0084037 | 0.027119 | 0.054617 | 0.037374 | 0.0096143 | 0.71983 | 0.71487 | 111569.9995 | 139171.9787 | 8299.0766 | 157317.5594 | 124034.8126 | 124546.6164 | 2550.8945 | 0.34556 | 0.82202 | 1.5001 | 3.2983
80 | CCV 10 ppb MAJORS | R | 06/02/2023 23:32:09 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.082 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.1219 | 0.79767 | 0.044493 | 0.21663 | 0.28084 | 0.88916 | 0.291 | 1.0392 | 1.4018 | 0.45478 | 0.30932 | 0.87572 | 1.262 | 2.8106 | 0.75379 | 1.8089 | 0.78212 | 0.37892 | 0.40383 | 0.092969 | 20.8963 | 0.38175 | 0.04365 | 0.0081967 | 0.026113 | 0.055905 | 0.039011 | 0.010013 | 0.73921 | 0.7171 | 112880.8874 | 139179.4175 | 8439.1594 | 156632.6845 | 127820.0602 | 129222.6311 | 2542.2262 | 0.34792 | 0.83644 | 1.5016 | 3.1994
81 | MT 22 | R | 06/02/2023 23:38:12 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 22.083 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.74976 | 0.044799 | 0.046927 | 136.5994 | 0.16066 | 0.40177 | 0.0069759 | 0.016855 | 0.1737 | 0.056533 | 20.4329 | 0.35683 | 73.0902 | 0.034355 | 0.082973 | 2.5253 | 1.0855 | 1.4477 | 0.15117 | 0.031057 | 8.401 | 0.14756 | 0.00985 | 0.0025028 | 0.008087 | 0.016642 | 0.011202 | 0.0028337 | 0.098766 | 0.09223 | 115211.3981 | 148398.208 | 8894.4354 | 167616.2884 | 135681.6873 | 131491.9249 | 2607.5715 | 0.28845 | 0.65354 | 0.03724 | 0.072804
82 | MT 22 LSP | R | 06/02/2023 23:44:15 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 22 LSP.084 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 6.2012 | 2.0545 | 0.075624 | 0.036095 | 0.77893 | 2.7145 | 0.77098 | 0.70998 | 1.1478 | 0.34872 | 21.5678 | 0.92859 | 75.4364 | 1.8216 | 0.56091 | 3.6547 | 1.5437 | 1.7183 | 0.42113 | 0.089275 | 21.8907 | 0.39818 | 0.038694 | 0.0081213 | 0.026103 | 0.053455 | 0.036865 | 0.0092999 | 0.56817 | 0.55543 | 105105.2427 | 132192.1374 | 7873.1691 | 152325.9554 | 120102.868 | 52527.0623 | 2404.8691 | 0.50565 | 1.1823 | 0.99813 | 2.1125
83 | MT 22 LSP | R | 06/02/2023 23:50:45 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 22 LSP.085 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 5.616 | 1.8128 | 0.062448 | 154.8855 | 0.68921 | 2.4688 | 0.68453 | 0.62174 | 0.9913 | 0.34983 | 21.5355 | 0.9348 | 67.3181 | 1.6302 | 0.49925 | 3.2862 | 1.537 | 1.712 | 0.42183 | 0.081514 | 22.0521 | 0.39674 | 0.038802 | 0.0083123 | 0.025935 | 0.053439 | 0.03651 | 0.0092324 | 0.57045 | 0.55564 | 104725.8797 | 134185.2674 | 8820.4072 | 153323.1957 | 123120.9683 | 121277.3737 | 2414.8708 | 0.51117 | 1.1804 | 1.0128 | 2.044
84 | MT 23 | R | 06/02/2023 23:56:47 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 23.086 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.43155 | 0.024542 | 0.029745 | 79.5026 | 0.10823 | 0.23754 | 0.0044495 | 0.008299 | 0.12197 | 0.037317 | 10.9623 | 0.20888 | 45.1483 | 0.021689 | 0.053345 | 1.4868 | 0.64398 | 0.8893 | 0.093037 | 0.018912 | 4.6818 | 0.086431 | 0.0070964 | 0.0015527 | 0.0045781 | 0.0093711 | 0.0071867 | 0.0017904 | 0.064823 | 0.05944 | 111197.7891 | 153920.4423 | 8925.1205 | 167676.4403 | 151266.3111 | 134601.3612 | 2847.6179 | 0.17307 | 0.39306 | 0.021381 | 0.035555
85 | MT 24 | R | 06/03/2023 00:02:49 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 24.087 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.41015 | 0.021584 | 0.024886 | 69.4838 | 0.098775 | 0.21548 | 0.0035782 | 0.0080888 | 0.11804 | 0.037935 | 9.832 | 0.19579 | 40.1594 | 0.019455 | 0.051533 | 1.4046 | 0.60494 | 0.81142 | 0.085983 | 0.016647 | 4.5215 | 0.082229 | 0.0055261 | 0.0015735 | 0.0045154 | 0.0088019 | 0.0065269 | 0.0017032 | 0.055419 | 0.054266 | 115136.194 | 150083.787 | 9041.8609 | 171106.6127 | 140053.8466 | 141044.2998 | 2728.9273 | 0.15948 | 0.36195 | 0.015685 | 0.030526
86 | MT 25 | R | 06/03/2023 00:08:52 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 25.088 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.47955 | 0.027142 | 0.032208 | 86.7795 | 0.10539 | 0.25875 | 0.0055399 | 0.0097984 | 0.08607 | 0.02947 | 11.3649 | 0.22595 | 45.7564 | 0.021964 | 0.038679 | 1.6291 | 0.72359 | 0.95397 | 0.10183 | 0.019548 | 5.311 | 0.095333 | 0.0082852 | 0.0017521 | 0.0048161 | 0.010824 | 0.0078118 | 0.0017222 | 0.058587 | 0.056977 | 106805.7954 | 148912.8855 | 8928.1234 | 168182.9478 | 137312.21 | 133182.643 | 2673.5836 | 0.18322 | 0.41783 | 0.022352 | 0.040806
87 | MT 26 | R | 06/03/2023 00:14:55 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 26.089 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.44328 | 0.024299 | 0.034499 | 80.5419 | 0.10392 | 0.22919 | 0.0038659 | 0.0090182 | 0.13052 | 0.041071 | 10.5637 | 0.2096 | 42.0958 | 0.020136 | 0.055602 | 1.5025 | 0.66429 | 0.87243 | 0.10813 | 0.01852 | 4.9652 | 0.087546 | 0.0094169 | 0.001509 | 0.0045052 | 0.0096703 | 0.0087014 | 0.0016617 | 0.059702 | 0.058402 | 96650.1443 | 148183.4215 | 8965.1459 | 166650.472 | 161035.9928 | 132369.2567 | 2612.2388 | 0.16841 | 0.388 | 0.030355 | 0.036012
88 | MT 27 | R | 06/03/2023 00:21:03 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 27.090 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.2437 | 0.074094 | 0.070167 | 216.9514 | 0.24041 | 0.64207 | 0.013606 | 0.023009 | 0.29508 | 0.089308 | 31.9477 | 0.57387 | 113.1681 | 0.053176 | 0.13201 | 3.852 | 1.6572 | 2.2396 | 0.23465 | 0.048004 | 13.075 | 0.23367 | 0.014698 | 0.0041133 | 0.011997 | 0.025926 | 0.017562 | 0.0044822 | 0.15329 | 0.14928 | 115912.501 | 148144.1742 | 8726.9986 | 166952.6005 | 136788.5712 | 132802.4223 | 2651.5798 | 0.44282 | 1.0127 | 0.03801 | 0.075195
89 | MT 28 | R | 06/03/2023 00:27:06 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 28.091 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.3895 | 0.073452 | 0.076139 | 235.4562 | 0.26789 | 0.70998 | 0.012782 | 0.024579 | 0.28031 | 0.090571 | 34.9077 | 0.59415 | 126.9657 | 0.058723 | 0.13295 | 4.3871 | 1.8135 | 2.4809 | 0.27358 | 0.057897 | 14.2062 | 0.24861 | 0.018421 | 0.0043304 | 0.013419 | 0.028196 | 0.020515 | 0.0050347 | 0.16428 | 0.154 | 105673.9642 | 150187.1507 | 8372.7894 | 163741.2064 | 133237.8398 | 131865.2749 | 2632.2459 | 0.48576 | 1.1279 | 0.040466 | 0.073254
90 | MT 29 | R | 06/03/2023 00:33:08 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 29.092 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.6075 | 0.14214 | 0.13708 | 436.2361 | 0.46334 | 1.37 | 0.02555 | 0.046488 | 0.59111 | 0.19508 | 68.2334 | 1.1764 | 233.2366 | 0.10861 | 0.27907 | 8.0439 | 3.4975 | 4.7834 | 0.52476 | 0.10696 | 28.2337 | 0.49079 | 0.033496 | 0.0087722 | 0.026147 | 0.055689 | 0.039528 | 0.0099048 | 0.33118 | 0.32019 | 107013.376 | 144635.2106 | 8323.7583 | 157815.3436 | 141360.3244 | 135902.1665 | 2519.5559 | 0.95426 | 2.1824 | 0.07006 | 0.12297
91 | LBLANK | R | 06/03/2023 00:39:20 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LBLANK.093 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.00031263 | 0 | 3.1125e-05 | 0.024234 | 0.031356 | 0.00077935 | 4.4222e-05 | 0.00016468 | 0.00041851 | 0.00012253 | 0.0034946 | 0.00014706 | 0.050132 | 6.8902e-05 | 0.00044624 | 0.00093417 | 0.00039158 | -8.1632e-05 | -0.0043519 | 5.633e-05 | -0.0029744 | -3.2714e-05 | 0.0013272 | 0 | 2.1758e-05 | 6.1347e-06 | 4.5658e-05 | 4.5598e-06 | 0.00065203 | 5.3702e-05 | 113464.8341 | 147087.3837 | 8695.9794 | 167910.1203 | 142143.6684 | 134062.1236 | 2688.2529 | 4.7861e-05 | 7.4676e-05 | 9.6675e-06 | 1.6148e-05
92 | CCV 1ppb TRACE | R | 06/03/2023 00:45:23 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.094 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.2196 | 0.084399 | 0.005112 | 0.027431 | 0.053756 | 0.092015 | 0.029574 | 0.10575 | 0.14 | 0.047268 | 0.035096 | 0.093599 | 0.16267 | 0.27737 | 0.074023 | 0.193 | 0.085202 | 0.038608 | 0.0365 | 0.0091633 | 2.2245 | 0.040757 | 0.0053143 | 0.00096805 | 0.0027988 | 0.0057364 | 0.003597 | 0.001018 | 0.074614 | 0.076612 | 113894.9341 | 141164.5371 | 8679.6363 | 161936.3549 | 133478.7491 | 129732.4821 | 2590.2364 | 0.035906 | 0.08573 | 0.16209 | 0.34489
93 | CCV 1ppb TRACE | R | 06/03/2023 00:51:55 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.095 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.22948 | 0.08695 | 0.0054523 | 0.017738 | 0.060512 | 0.09442 | 0.030222 | 0.10599 | 0.14493 | 0.044916 | 0.032658 | 0.089932 | 0.17434 | 0.28712 | 0.078383 | 0.19358 | 0.081677 | 0.036927 | 0.035729 | 0.0085243 | 2.1327 | 0.038097 | 0.005005 | 0.00083614 | 0.002688 | 0.0056977 | 0.0035375 | 0.00084099 | 0.070252 | 0.071888 | 113278.8358 | 144330.1355 | 8624.6029 | 167898.3355 | 132823.9645 | 130807.533 | 2578.8996 | 0.034235 | 0.081805 | 0.16345 | 0.33865
94 | LFB 10ppb | R | 06/03/2023 00:57:57 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LFB 10ppb.096 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.2887 | 0.86828 | 0.046697 | 0.25954 | 0.29813 | 0.94878 | 0.31877 | 1.1009 | 1.5169 | 0.45132 | 0.33308 | 0.80557 | 1.4186 | 2.9583 | 0.78165 | 1.9012 | 0.77418 | 0.49992 | 0.39636 | 0.097684 | 20.1846 | 0.36049 | 0.045361 | 0.0079493 | 0.025358 | 0.052678 | 0.038323 | 0.0097242 | 0.71843 | 0.66938 | 111714.6027 | 145491.3593 | 7693.7426 | 151509.1968 | 121769.4629 | 121544.9096 | 2584.9032 | 0.34479 | 0.81082 | 1.4093 | 3.1294
95 | LFB 10ppb | R | 06/03/2023 01:04:27 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LFB 10ppb.097 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.2266 | 0.82695 | 0.043014 | 0.25266 | 0.28822 | 0.94607 | 0.30798 | 1.0481 | 1.4305 | 0.43901 | 0.31682 | 0.84867 | 1.3334 | 2.8227 | 0.75613 | 1.8204 | 0.74979 | 0.49017 | 0.40378 | 0.089843 | 22.2071 | 0.37799 | 0.044844 | 0.0082042 | 0.02623 | 0.054819 | 0.038113 | 0.0097623 | 0.7007 | 0.69997 | 110317.0089 | 137368.6584 | 8059.6074 | 157916.7456 | 122369.9012 | 124286.3629 | 2338.1914 | 0.33024 | 0.78427 | 1.4342 | 3.1038
96 | CCV 10 ppb MAJORS | R | 06/03/2023 01:10:29 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.098 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.2734 | 0.86346 | 0.041504 | 0.24384 | 0.28937 | 0.94099 | 0.31098 | 1.0873 | 1.4816 | 0.47366 | 0.32252 | 0.89724 | 1.2666 | 2.8414 | 0.76979 | 1.8286 | 0.80498 | 0.3939 | 0.39706 | 0.091132 | 21.4354 | 0.39 | 0.044382 | 0.0090786 | 0.026975 | 0.055777 | 0.037125 | 0.010373 | 0.74748 | 0.73201 | 112283.5417 | 134529.8103 | 8043.5976 | 152503.2308 | 123634.0608 | 124774.0374 | 2454.2113 | 0.35331 | 0.83629 | 1.4475 | 3.1887
97 | CCV 10 ppb MAJORS | R | 06/03/2023 01:17:01 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.099 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.1913 | 0.81708 | 0.041154 | 0.23755 | 0.28112 | 0.93066 | 0.29571 | 1.0541 | 1.4174 | 0.46192 | 0.31434 | 0.8793 | 1.2425 | 2.7871 | 0.75781 | 1.8335 | 0.77982 | 0.37699 | 0.42589 | 0.08933 | 22.3305 | 0.38601 | 0.049384 | 0.0084557 | 0.026203 | 0.054706 | 0.038502 | 0.0097171 | 0.71572 | 0.71352 | 105776.0356 | 138432.1965 | 8312.7513 | 156152.9508 | 123959.4911 | 124070.475 | 2394.8676 | 0.34246 | 0.81585 | 1.5581 | 3.1815
98 | MT 30 | R | 06/03/2023 01:23:04 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 30.100 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.0037982 | 0.01143 | 1.663 | 35.9651 | 0.070562 | 0.0098545 | 0.037884 | 0.070754 | 0.066415 | 0.02218 | 16.6812 | 7.1913 | 52.0086 | 0.11907 | 0.056608 | 10.9965 | 5.3262 | 4.7874 | 0.049876 | 0.012038 | 2.8613 | 0.050723 | 0.0013721 | 0.00029243 | 0.0009064 | 0.0021789 | 0.001742 | 0.00038567 | 0.024284 | 0.020314 | 119644.9382 | 148429.9092 | 8605.9254 | 151186.7287 | 131847.6435 | 136521.7218 | 2630.9091 | 0.064586 | 0.14708 | 0.47195 | 1.0215
99 | MT 31 | R | 06/03/2023 01:29:15 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 31.101 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.021582 | 0.019936 | 1.2527 | 35.4388 | 0.054964 | 0.047559 | 0.021266 | 0.051277 | 0.05048 | 0.015538 | 9.3328 | 5.1952 | 38.6149 | 0.11769 | 0.036395 | 8.1244 | 3.3934 | 4.9613 | 0.044823 | 0.010268 | 2.5155 | 0.044605 | 0.0012897 | 0.00017156 | 0.00079551 | 0.0014632 | 0.0012895 | 0.00021909 | 0.012677 | 0.012879 | 116814.3859 | 151600.5842 | 8676.6345 | 169458.9726 | 132909.5109 | 129850.9095 | 2691.5874 | 0.063176 | 0.14442 | 0.34634 | 0.72117
100 | MT 31 LD | R | 06/03/2023 01:35:29 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 31 LD.102 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.022789 | 0.018888 | 1.297 | 36.403 | 0.059083 | 0.049638 | 0.021983 | 0.054516 | 0.050173 | 0.016345 | 9.6167 | 5.3098 | 39.3319 | 0.11725 | 0.036315 | 8.2886 | 3.4544 | 5.198 | 0.045973 | 0.0098629 | 2.5677 | 0.045842 | 0.0013973 | 0.00019348 | 0.00068089 | 0.0014992 | 0.0014375 | 0.00030398 | 0.01322 | 0.013248 | 115322.3341 | 144721.3635 | 8467.5111 | 161432.3987 | 131203.7645 | 129686.0012 | 2583.567 | 0.064802 | 0.14923 | 0.34733 | 0.71942
101 | MT 31 LD | R | 06/03/2023 01:42:00 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 31 LD.103 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.022727 | 0.017823 | 1.2576 | 36.31 | 0.056304 | 0.046028 | 0.020776 | 0.051115 | 0.052376 | 0.015975 | 9.5838 | 5.3237 | 37.9266 | 0.12029 | 0.034399 | 7.9173 | 3.4501 | 5.1233 | 0.046453 | 0.0085356 | 2.5516 | 0.045744 | 0.0012001 | 0.00018853 | 0.00074834 | 0.0014729 | 0.0011825 | 0.00022086 | 0.012956 | 0.012519 | 116583.199 | 145060.1186 | 8706.3217 | 164067.4556 | 122121.7896 | 127676.0124 | 2600.9037 | 0.064589 | 0.14759 | 0.34038 | 0.76705
102 | MT 32 | R | 06/03/2023 01:48:02 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 32.104 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.00039761 | 7.7815e-05 | 0.0049636 | 0.19326 | 0.032575 | 0.0011548 | 0.00016233 | -2.5687e-05 | 0.0019229 | 0.00038578 | 0.0045095 | 0.00048172 | 0.053626 | 7.3611e-05 | 0.0010853 | 0.0083031 | 0.0036715 | 0.00096384 | -0.0035719 | -9.744e-05 | -0.0011542 | -3.8405e-06 | 0.00058991 | 9.1242e-06 | -7.9323e-07 | 1.7619e-06 | 0.0001212 | 4.6007e-06 | 0.00093955 | 0.0010305 | 111292.5176 | 144762.6031 | 8541.5531 | 164784.467 | 124049.1002 | 123670.4919 | 2557.5623 | 1.2514e-05 | 9.7398e-05 | 0.0062026 | 0.011799
103 | MT 32 LSP | R | 06/03/2023 01:54:03 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 32 LSP.105 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 6.0025 | 2.1412 | 0.038163 | 0.40754 | 0.6878 | 2.4918 | 0.80609 | 0.72279 | 0.95091 | 0.31315 | 0.257 | 0.60885 | 1.0281 | 1.8401 | 0.71233 | 1.1708 | 0.51106 | 0.95294 | 0.26498 | 0.058993 | 13.858 | 0.24881 | 0.029873 | 0.0052965 | 0.017419 | 0.037033 | 0.025837 | 0.0064392 | 0.47076 | 0.46166 | 100484.1853 | 126873.6279 | 7525.6485 | 141266.3904 | 113790.8631 | 106976.9713 | 2278.1817 | 0.2189 | 0.51763 | 0.93933 | 1.9853
104 | MT 32 LSP | R | 06/03/2023 02:00:34 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 32 LSP.106 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 5.851 | 2.0924 | 0.039463 | 0.33918 | 0.67261 | 2.3998 | 0.77862 | 0.70809 | 0.9444 | 0.31034 | 0.25458 | 0.58495 | 1.0279 | 1.8425 | 0.71231 | 1.193 | 0.51515 | 0.96272 | 0.25594 | 0.061338 | 14.2958 | 0.24676 | 0.028366 | 0.0055398 | 0.017303 | 0.035584 | 0.02388 | 0.0062342 | 0.47533 | 0.45051 | 105605.9135 | 128938.7378 | 7598.6873 | 143298.1828 | 115624.0624 | 114485.6788 | 2227.5072 | 0.2193 | 0.52324 | 0.92092 | 1.9595
105 | MT 33 | R | 06/03/2023 02:06:37 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 33.107 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.00059732 | 7.4124e-05 | 0.0061034 | 0.16862 | 0.035199 | 0.00047019 | 0.00026703 | 4.8096e-05 | 0.0017733 | 0.00045128 | 0.0048437 | 0.00041341 | 0.054988 | 0.00050734 | 0.0013467 | 0.0083532 | 0.0045269 | 0.03002 | -0.0040985 | -9.744e-05 | -0.0018356 | -1.1277e-05 | 0.0011242 | 0 | 7.9322e-06 | 1.294e-06 | -1.3252e-05 | 8.926e-06 | 0.0010092 | 0.00018701 | 115843.1289 | 148915.4702 | 8934.7936 | 165698.852 | 115217.5325 | 131456.7634 | 2730.9277 | 2.0665e-05 | 0.0001022 | 0.00476 | 0.010507
106 | MT 34 | R | 06/03/2023 02:12:39 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 34.108 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.00065462 | 0 | 0.0068801 | 0.18837 | 0.038615 | 0.00045848 | 0.00015555 | 4.739e-05 | 0.0029162 | 0.00085735 | 0.092688 | 0.0014922 | 0.41154 | 0.00018041 | 0.0031667 | 0.0033533 | 0.001349 | 0.00055126 | -0.0057097 | -5.9067e-05 | -0.002026 | -1.9754e-05 | 0.0013702 | 4.5145e-06 | 1.705e-05 | 1.3672e-06 | 0.00012783 | 4.5145e-06 | 0.00034166 | 9.3838e-05 | 112212.5078 | 148035.0531 | 8771.0261 | 168553.2336 | 135223.2567 | 132593.2492 | 2617.5732 | 1.586e-05 | 0.00011036 | 0.0041983 | 0.00865
107 | MT 35 | R | 06/03/2023 02:18:42 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 35.109 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.00026886 | 3.7816e-05 | 0.0052849 | 0.18685 | 0.040295 | 0.00049282 | 0.00026886 | 1.0125e-05 | 0.0018811 | 0.00039446 | 0.016182 | 0.00039955 | 0.10043 | 0.00017835 | 0.0010758 | 0.0021577 | 0.001216 | 0.0012446 | -0.0071337 | -5.9125e-05 | -0.00072439 | 3.2553e-06 | 0.0012798 | 1.4249e-05 | 3.8635e-06 | 1.0877e-05 | 0.0001184 | 0 | 0.00025766 | 9.3236e-05 | 111357.9354 | 142431.9955 | 8819.3885 | 173481.1996 | 135276.9941 | 130614.6629 | 2499.5523 | 3.079e-05 | 8.5318e-05 | 0.0046021 | 0.0083522
108 | LBLANK | R | 06/03/2023 02:24:45 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LBLANK.110 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.00023847 | 0 | 0.00022496 | 0.018726 | 0.041538 | 0.0010135 | 4.4377e-05 | 5.0255e-05 | 0.00076888 | 5.4988e-05 | 0.0025574 | 0.00012217 | 0.048629 | -4.6997e-05 | 0.0010229 | 0.00081825 | 0.00026688 | -0.00029203 | -0.0035147 | -9.744e-05 | -0.0040453 | -5.4499e-05 | 0.00096466 | 0 | 8.2732e-06 | -1.2008e-05 | 5.2743e-05 | 0 | 0.00022845 | 2.718e-05 | 114932.4903 | 145070.9449 | 8665.6276 | 162453.8264 | 124689.2028 | 130952.9097 | 2567.5643 | 4.0819e-06 | 5.393e-05 | -4.1969e-06 | 1.7398e-05
109 | CCV 1ppb TRACE | R | 06/03/2023 02:30:48 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.111 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.22625 | 0.083862 | 0.0049515 | 0.023363 | 0.062796 | 0.097056 | 0.031312 | 0.11025 | 0.14573 | 0.048357 | 0.036158 | 0.09512 | 0.1796 | 0.28455 | 0.079364 | 0.20412 | 0.086748 | 0.04055 | 0.036045 | 0.009367 | 2.1904 | 0.039965 | 0.0042607 | 0.00094937 | 0.002548 | 0.0055899 | 0.0037317 | 0.00089604 | 0.073916 | 0.073894 | 113112.0466 | 141159.0472 | 8620.9347 | 156769.8035 | 135380.5013 | 127717.8207 | 2572.8984 | 0.035609 | 0.085573 | 0.16015 | 0.30767
110 | CCV 1ppb TRACE | R | 06/03/2023 02:37:20 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.112 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.23799 | 0.089182 | 0.0059695 | 0.020073 | 0.070024 | 0.099391 | 0.031981 | 0.11672 | 0.15495 | 0.048872 | 0.035543 | 0.092272 | 0.18401 | 0.29852 | 0.081699 | 0.20699 | 0.086836 | 0.039438 | 0.039314 | 0.010213 | 2.1219 | 0.037739 | 0.0049251 | 0.0009001 | 0.0027083 | 0.0055497 | 0.0041236 | 0.0010538 | 0.072814 | 0.070338 | 113799.0502 | 142998.1984 | 8411.1436 | 154916.2631 | 127859.4334 | 130729.0522 | 2542.2262 | 0.035663 | 0.085537 | 0.1588 | 0.33343
111 | CCV 10 ppb MAJORS | R | 06/03/2023 02:43:23 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.113 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.4144 | 0.90201 | 0.045842 | 0.24532 | 0.31356 | 1.0151 | 0.33999 | 1.1414 | 1.5429 | 0.47224 | 0.31624 | 0.89086 | 1.3415 | 3.0424 | 0.81423 | 1.9475 | 0.79171 | 0.3812 | 0.4 | 0.095844 | 20.8416 | 0.37922 | 0.046049 | 0.0084991 | 0.02598 | 0.054767 | 0.036371 | 0.0099057 | 0.72438 | 0.71004 | 112414.6413 | 137071.5047 | 7791.7923 | 149687.471 | 119628.1918 | 122170.862 | 2497.552 | 0.33896 | 0.80626 | 1.3877 | 3.1737
112 | CCV 10 ppb MAJORS | R | 06/03/2023 02:49:56 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.114 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.3005 | 0.87364 | 0.043691 | 0.24068 | 0.28788 | 0.9484 | 0.31305 | 1.0876 | 1.4674 | 0.49129 | 0.32972 | 0.89701 | 1.2836 | 2.9086 | 0.77009 | 1.8505 | 0.82503 | 0.39444 | 0.37935 | 0.089565 | 21.8513 | 0.38707 | 0.041327 | 0.0085834 | 0.026454 | 0.055577 | 0.034306 | 0.0097313 | 0.7574 | 0.71805 | 118179.1992 | 137302.0397 | 8287.7389 | 147867.6147 | 121829.2941 | 123835.1305 | 2432.2071 | 0.3568 | 0.84398 | 1.3344 | 3.1542
113 | MT 36 | R | 06/03/2023 02:55:59 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 36.115 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.00098975 | 0.00011263 | 0.005392 | 0.15502 | 0.034254 | 0.00059745 | 0.00044763 | 0.00015837 | 0.0037856 | 0.0009842 | 0.01186 | 0.00057576 | 0.079291 | 0.00024557 | 0.0010537 | 0.0033948 | 0.0014309 | 0.0017024 | -0.0032383 | -6.1369e-05 | -0.0029576 | -3.5526e-05 | 0.0008268 | 0 | -7.7433e-07 | -1.2069e-05 | 0.00017311 | 0 | 0.00049182 | 0.0002078 | 116851.3061 | 144317.3146 | 8875.0917 | 160959.7375 | 132710.3417 | 132676.6704 | 2543.5598 | 1.2437e-05 | 0.00011248 | 0.0042582 | 0.0087268
114 | MT 37 | R | 06/03/2023 03:02:01 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 37.116 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.00084823 | 0.00019943 | 0.0054278 | 0.21252 | 0.042375 | 0.0015729 | 0.00052743 | -2.7085e-05 | 0.0017896 | 0.0004022 | 0.0051561 | 0.00077405 | 0.068755 | -3.8723e-05 | 0.0017882 | 0.017211 | 0.0071046 | 0.0048525 | -0.0025117 | -5.5444e-05 | -0.0019542 | -1.3895e-05 | 0.00045082 | 4.438e-06 | -5.4003e-06 | 1.962e-05 | -0.00012994 | 0 | 0.00017209 | 0.00011255 | 118805.4331 | 146892.8744 | 8379.1276 | 164762.08 | 134050.6899 | 134252.0717 | 2651.5795 | 2.0187e-05 | 0.00012965 | 0.0044521 | 0.0093628
115 | MT 38 | R | 06/03/2023 03:08:04 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 38.117 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.00042282 | 7.6754e-05 | 0.0044287 | 0.17724 | 0.039901 | 0.00087165 | 0.0001558 | 8.7368e-05 | 0.0020473 | 0.00051477 | 0.0056065 | 0.00038863 | 0.058421 | 6.6681e-05 | 0.00066428 | 0.0047468 | 0.002197 | 0.0011815 | -0.0041275 | -5.8814e-05 | -0.0019723 | -1.7351e-05 | 0.0012897 | 0 | 2.23e-05 | -7.4748e-06 | 0.00021025 | 0 | 0.00018898 | 0.00010135 | 113606.6253 | 144764.5414 | 8798.0426 | 161297.5206 | 114891.818 | 137215.2503 | 2584.234 | 1.2477e-05 | 5.0316e-05 | 0.0043915 | 0.0103
116 | MT 39 | R | 06/03/2023 03:14:07 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 39.118 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.0013952 | 7.8358e-05 | 0.0056691 | 0.19482 | 0.036943 | 0.00032218 | 0.00034997 | 8.8325e-05 | 0.0015937 | 0.00059974 | 0.026689 | 0.00060112 | 0.14159 | 0.00010264 | 0.00053062 | 0.0047575 | 0.0021476 | 0.0036199 | -0.006137 | -6.0104e-05 | -0.0033413 | -3.9994e-05 | 0.0013831 | 4.6455e-06 | 1.2991e-05 | 2.0402e-05 | -4.1595e-05 | 1.41e-05 | 0.00014142 | 0.00011098 | 105375.9139 | 143915.977 | 8647.9515 | 159820.3609 | 128839.7196 | 128766.473 | 2602.2372 | 8.3605e-06 | 8.4359e-05 | 0.0053564 | 0.0089071
117 | MT 40 | R | 06/03/2023 03:20:10 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 40.119 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.001856 | 0.00067824 | 0.0056396 | 0.19709 | 0.034757 | 0.00097346 | 0.0015854 | 9.2968e-06 | 0.0011249 | 0.00031951 | 0.0071181 | 0.0024817 | 0.056918 | -1.2166e-05 | 0.00080525 | 0.0073062 | 0.0029936 | 0.00084592 | -0.00067838 | -2.305e-05 | 0.015907 | 0.00031431 | 6.1784e-05 | 0 | 9.0115e-06 | 6.8684e-06 | 2.4348e-05 | 4.7515e-06 | 0.0006488 | 0.00066753 | 123590.2816 | 142050.2008 | 8844.4046 | 155237.1126 | 127982.0705 | 125369.7924 | 2598.903 | 8.502e-06 | 6.1833e-05 | 0.0038641 | 0.0086352
118 | MT 31 | R | 06/03/2023 03:26:12 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 31 .120 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.023222 | 0.021814 | 1.3762 | 37.6833 | 0.068539 | 0.048186 | 0.023133 | 0.056221 | 0.051439 | 0.014868 | 9.4256 | 5.3543 | 40.3857 | 0.12209 | 0.036659 | 8.4423 | 3.3602 | 5.0085 | 0.045047 | 0.009637 | 2.6276 | 0.046297 | 0.0011281 | 0.00021122 | 0.00074178 | 0.0015038 | 0.001162 | 0.00022601 | 0.012215 | 0.012624 | 117943.6701 | 141716.9074 | 8287.0697 | 160814.8585 | 124708.6568 | 124744.9812 | 2495.5513 | 0.062602 | 0.14276 | 0.3308 | 0.71794
119 | MT 31 LD | R | 06/03/2023 03:32:23 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 31 LD.121 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.022183 | 0.01851 | 1.3022 | 38.998 | 0.058025 | 0.04616 | 0.022108 | 0.0502 | 0.048243 | 0.016479 | 9.6855 | 5.6258 | 38.7813 | 0.11651 | 0.037325 | 8.0065 | 3.4544 | 5.161 | 0.042228 | 0.0097941 | 2.6628 | 0.047958 | 0.0011683 | 0.00023474 | 0.00075232 | 0.0015264 | 0.00094755 | 0.00033501 | 0.013272 | 0.013374 | 114630.7847 | 137246.4134 | 8555.9069 | 155203.8268 | 142754.4613 | 122987.1447 | 2470.8804 | 0.063808 | 0.14886 | 0.33812 | 0.62075
120 | MT 31 LD | R | 06/03/2023 03:38:54 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 31 LD.122 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.021572 | 0.017948 | 1.2854 | 38.0674 | 0.060245 | 0.048807 | 0.021877 | 0.052536 | 0.051048 | 0.014223 | 8.4672 | 5.3912 | 39.9356 | 0.11935 | 0.034897 | 8.2464 | 3.0244 | 4.5416 | 0.043389 | 0.010059 | 2.5193 | 0.045273 | 0.00054314 | 0.00024898 | 0.00064104 | 0.0013907 | 0.00079036 | 0.00028708 | 0.011373 | 0.01265 | 108281.5072 | 139303.9149 | 8315.0863 | 166720.6075 | 123065.958 | 124537.8303 | 2503.5528 | 0.056412 | 0.13009 | 0.30246 | 0.67523
121 | MT 32 | R | 06/03/2023 03:44:56 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 32.123 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.00082194 | 0 | 0.0047001 | 0.18077 | 0.026316 | 0.00018788 | 0.00013081 | -2.3992e-05 | 0.0021452 | 0.00043236 | 0.0050431 | 0.00050689 | 0.053408 | 3.9771e-05 | 0.0011145 | 0.0089847 | 0.0040924 | 0.00060852 | -0.0030791 | -9.744e-05 | -0.0013987 | -7.0787e-06 | 0.00075016 | 4.8446e-06 | 4.5709e-06 | 2.6724e-06 | 0.00018041 | 0 | 0.00086742 | 0.0008312 | 108716.1944 | 138074.0731 | 8192.0153 | 152216.7577 | 125208.325 | 125646.8112 | 2438.875 | 1.3149e-05 | 9.8843e-05 | 0.0059328 | 0.011122
122 | MT 32 LSP | R | 06/03/2023 03:50:58 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 32 LSP.124 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 5.9336 | 2.1507 | 0.040454 | 0.26052 | 0.69914 | 2.4548 | 0.82672 | 0.72013 | 0.9704 | 0.31136 | 0.25644 | 0.61124 | 1.032 | 1.8368 | 0.72348 | 1.1874 | 0.51189 | 0.95467 | 0.26289 | 0.057446 | 13.6882 | 0.25165 | 0.029785 | 0.0056362 | 0.0176 | 0.036369 | 0.025741 | 0.006414 | 0.47124 | 0.46174 | 100580.1661 | 121543.6968 | 7423.2617 | 136585.8893 | 99201.9141 | 149819.6452 | 2236.1752 | 0.22026 | 0.51581 | 0.92508 | 2.2317
123 | MT 32 LSP | R | 06/03/2023 03:57:29 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 32 LSP.125 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 5.7136 | 2.1027 | 0.037864 | 0.3533 | 0.68976 | 2.4406 | 0.77984 | 0.70529 | 0.9611 | 0.28257 | 0.23678 | 0.59345 | 1.0231 | 1.8268 | 0.70895 | 1.181 | 0.47472 | 0.89099 | 0.23559 | 0.058823 | 13.6667 | 0.24469 | 0.0263 | 0.0054092 | 0.016638 | 0.03529 | 0.021815 | 0.0061684 | 0.43591 | 0.45395 | 104905.1045 | 125993.2018 | 7581.678 | 145551.9066 | 102666.3737 | 112465.0014 | 2258.8456 | 0.20193 | 0.48385 | 0.82694 | 2.0978
124 | LBLANK | R | 06/03/2023 04:03:32 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LBLANK.126 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.00041142 | 3.6394e-05 | 8.797e-05 | 0.0052442 | 0.033775 | 0.00011939 | 4.1167e-05 | 9.5459e-06 | 0.00097654 | 0.00017934 | 0.0022108 | 0.00014039 | 0.048233 | 9.5232e-05 | 0.00084862 | 0.00065069 | 0.00039626 | -0.0005034 | -0.0052933 | -2.4652e-05 | -8.8764e-05 | 1.9279e-05 | 0.0012917 | 0 | -8.8979e-07 | 2.0282e-05 | -8.5962e-06 | 0 | 0.0005675 | 8.711e-05 | 108541.5402 | 145195.574 | 9040.8602 | 164160.5262 | 133585.7036 | 129373.8503 | 2626.9086 | 8.0843e-06 | 7.3553e-05 | -2.0464e-05 | 4.5091e-06
125 | CCV 1ppb TRACE | R | 06/03/2023 04:09:35 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.127 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.23178 | 0.089822 | 0.0060913 | 0.020395 | 0.067301 | 0.094106 | 0.03109 | 0.10997 | 0.15225 | 0.048005 | 0.034287 | 0.093367 | 0.17826 | 0.30361 | 0.081509 | 0.20085 | 0.086297 | 0.038926 | 0.03714 | 0.009563 | 2.0876 | 0.038416 | 0.0048787 | 0.00082858 | 0.0025754 | 0.0056041 | 0.0036814 | 0.00074252 | 0.073327 | 0.072284 | 117299.6396 | 147169.5718 | 8555.2288 | 161020.3929 | 130402.5142 | 135110.6101 | 2706.9231 | 0.034243 | 0.082888 | 0.15425 | 0.32954
126 | CCV 1ppb TRACE | R | 06/03/2023 04:16:07 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.128 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.2285 | 0.085586 | 0.0059501 | 0.018723 | 0.064644 | 0.096873 | 0.031 | 0.11564 | 0.15512 | 0.050548 | 0.036791 | 0.09269 | 0.18017 | 0.29361 | 0.078052 | 0.19855 | 0.08866 | 0.040476 | 0.037359 | 0.0092044 | 2.1195 | 0.038641 | 0.0050043 | 0.00073433 | 0.0025768 | 0.0057114 | 0.0035356 | 0.0010757 | 0.07437 | 0.072074 | 114620.884 | 142520.7534 | 8779.7153 | 151688.9276 | 123995.8343 | 127335.1107 | 2596.2359 | 0.035255 | 0.087386 | 0.15439 | 0.33565
127 | LFB 10ppb | R | 06/03/2023 04:22:09 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LFB 10ppb.129 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.2237 | 0.82276 | 0.044255 | 0.26182 | 0.29116 | 0.9176 | 0.30803 | 1.0875 | 1.4641 | 0.45865 | 0.33222 | 0.86756 | 1.3511 | 2.8217 | 0.76703 | 1.8127 | 0.77955 | 0.50783 | 0.40637 | 0.088941 | 20.8728 | 0.37931 | 0.04445 | 0.0085332 | 0.025937 | 0.055131 | 0.037702 | 0.0096077 | 0.71738 | 0.69841 | 110426.5119 | 132264.148 | 7942.8742 | 145594.8476 | 114791.6429 | 117840.0348 | 2403.5355 | 0.33649 | 0.80394 | 1.3471 | 3.1288
128 | LFB 10ppb | R | 06/03/2023 04:28:39 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LFB 10ppb.130 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.1698 | 0.80231 | 0.046461 | 0.27029 | 0.28759 | 0.8955 | 0.29841 | 1.0706 | 1.4187 | 0.47744 | 0.34183 | 0.90053 | 1.3193 | 2.8176 | 0.76334 | 1.8055 | 0.79153 | 0.51106 | 0.40577 | 0.09339 | 21.9023 | 0.39168 | 0.045757 | 0.0080975 | 0.026922 | 0.057621 | 0.038079 | 0.0099235 | 0.73752 | 0.72699 | 109086.3127 | 130165.8418 | 8095.9602 | 144878.5144 | 116165.7947 | 120676.2441 | 2327.523 | 0.34491 | 0.81065 | 1.3818 | 3.0681
129 | CCV 10 ppb MAJORS | R | 06/03/2023 04:34:42 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.131 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.1238 | 0.80026 | 0.04402 | 0.25266 | 0.28403 | 0.89735 | 0.29259 | 1.0523 | 1.3822 | 0.49444 | 0.32767 | 1.2321 | 1.2139 | 2.7216 | 0.74336 | 1.7582 | 0.81709 | 0.39316 | 0.40119 | 0.087332 | 28.7638 | 0.53193 | 0.04555 | 0.011775 | 0.036165 | 0.074945 | 0.037417 | 0.013665 | 0.7519 | 0.98597 | 110081.9172 | 102892.7334 | 8550.9311 | 145162.4859 | 117233.8115 | 119906.4461 | 1911.463 | 0.34751 | 0.83245 | 1.3745 | 3.143
130 | CCV 10 ppb MAJORS | R | 06/03/2023 04:41:13 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.132 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.1906 | 0.82066 | 0.043398 | 0.24467 | 0.28807 | 0.93601 | 0.30661 | 1.0935 | 1.4458 | 0.48907 | 0.33637 | 0.8992 | 1.2801 | 2.8552 | 0.77313 | 1.8438 | 0.82112 | 0.38715 | 0.40351 | 0.091494 | 21.1295 | 0.38444 | 0.044982 | 0.0087835 | 0.025804 | 0.056255 | 0.036913 | 0.0098606 | 0.75283 | 0.72059 | 112791.543 | 132452.5288 | 8118.3064 | 146822.2091 | 113600.7444 | 118709.9045 | 2410.2033 | 0.34639 | 0.83553 | 1.34 | 3.3478
131 | MT 33 | R | 06/03/2023 04:47:16 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 33.133 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.00045398 | 3.753e-05 | 0.0057505 | 0.16018 | 0.037147 | 0.00036106 | 0.00022872 | -2.805e-05 | 0.0017923 | 0.00042114 | 0.0042964 | 0.00037033 | 0.056039 | 0.00047294 | 0.0018428 | 0.0093579 | 0.0046119 | -0.0001177 | -0.0048181 | -2.3619e-05 | -0.0030149 | -3.5794e-05 | 0.0010579 | 0 | 1.7557e-05 | 1.7661e-06 | 0.0001253 | 0 | 0.00096818 | 0.00026626 | 115843.1871 | 145028.4809 | 8904.1077 | 159531.1079 | 133381.1205 | 127813.2727 | 2576.8991 | 1.2555e-05 | 9.7562e-05 | 0.0045543 | 0.0087974
132 | MT 34 | R | 06/03/2023 04:53:19 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 34.134 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.00088879 | 7.2641e-05 | 0.004736 | 0.18916 | 0.033379 | 0.00052031 | 0.0002214 | 8.3523e-05 | 0.0027817 | 0.0007715 | 0.089653 | 0.0016972 | 0.39503 | 0.00021371 | 0.0022693 | 0.0029955 | 0.0011393 | 0.00050945 | -0.0035617 | -5.9107e-05 | -0.0039609 | -4.9869e-05 | 0.0010154 | 9.2217e-06 | -7.8634e-07 | 1.7025e-06 | 0.00010819 | 0 | 0.00036699 | 7.753e-05 | 115490.7872 | 145323.4411 | 9023.5196 | 163037.4842 | 129838.704 | 127702.9364 | 2676.2507 | 8.5799e-06 | 0.00011159 | 0.0041609 | 0.0082609
133 | MT 35 | R | 06/03/2023 04:59:21 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 35.135 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.00024294 | 0 | 0.0057043 | 0.19187 | 0.038029 | 0.00057406 | 0.00016358 | -2.5157e-05 | 0.0020793 | 0.00053575 | 0.017458 | 0.00039739 | 0.11592 | -4.3896e-05 | 0.0010887 | 0.003309 | 0.0013241 | 0.0015238 | -0.0023261 | 2.082e-05 | -0.0024486 | -2.3914e-05 | 0.00067571 | 0 | -5.5068e-07 | 2.595e-05 | 0.00014194 | 4.7766e-06 | 0.00029487 | 9.0171e-05 | 113440.0482 | 140648.3455 | 8522.5418 | 150450.5642 | 126431.7139 | 127141.3377 | 2561.563 | 1.3395e-05 | 9.4717e-05 | 0.0042583 | 0.0082446
134 | MT 36 | R | 06/03/2023 05:05:25 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 36.136 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.00055036 | 3.7988e-05 | 0.0050092 | 0.18651 | 0.037493 | 0.00061002 | 0.00020233 | -2.5548e-05 | 0.0030255 | 0.00099344 | 0.012346 | 0.00051946 | 0.14011 | 0.00010848 | 0.0008471 | 0.0029043 | 0.0015275 | 0.0014889 | -0.0028974 | -2.004e-05 | -0.0020575 | -1.4225e-05 | 0.0010216 | 4.6796e-06 | 9.1863e-06 | 2.2786e-06 | -1.7543e-05 | 9.4748e-06 | 0.0002711 | 0.00013339 | 105765.0366 | 140803.8994 | 8580.9104 | 153688.7898 | 127503.1423 | 127402.9563 | 2612.9056 | 8.7292e-06 | 7.752e-05 | 0.004334 | 0.0087485
135 | MT 37 | R | 06/03/2023 05:11:28 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 37.137 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.00064008 | 0 | 0.0058111 | 0.23294 | 0.03716 | 0.00084877 | 0.0005163 | 1.3172e-05 | 0.0015307 | 0.00035231 | 0.0048458 | 0.00081581 | 0.064406 | -0.00012394 | 0.0012517 | 0.017529 | 0.0064637 | 0.0044216 | -0.0039751 | -5.8374e-05 | -0.0034618 | -3.9672e-05 | 0.0010824 | 1.5007e-05 | -4.9849e-06 | -1.6621e-06 | 2.5077e-05 | 0 | 0.00017306 | 0.00011049 | 106779.4494 | 135159.9698 | 8501.5297 | 177351.56 | 127372.0955 | 125632.7287 | 2519.5557 | 4.2918e-06 | 0.00011804 | 0.0051837 | 0.0091527
136 | MT 38 | R | 06/03/2023 05:17:31 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 38.138 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.00027401 | 0.00011466 | 0.005545 | 0.20205 | 0.041928 | 0.00066078 | 8.2274e-05 | 1.1181e-05 | 0.0022968 | 0.00049765 | 0.00566 | 0.0004137 | 0.065337 | 0.0002976 | 0.0011718 | 0.0041513 | 0.00207 | 0.0011729 | -0.0061157 | -2.0757e-05 | -0.004758 | -6.7271e-05 | 0.00089176 | 0 | 3.9967e-06 | -2.6748e-06 | -9.3322e-05 | 4.6531e-06 | 0.00016975 | 9.7923e-05 | 110181.9876 | 142945.6135 | 8703.3171 | 151220.8616 | 124320.036 | 126385.9625 | 2529.5573 | 1.7641e-05 | 9.0401e-05 | 0.0041978 | 0.0084172
137 | MT 39 | R | 06/03/2023 05:23:33 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 39.139 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.0014157 | 8.0366e-05 | 0.0060869 | 0.19425 | 0.040639 | 0.00069883 | 0.00024936 | 1.5469e-05 | 0.0021933 | 0.00063623 | 0.029193 | 0.00064789 | 0.15101 | 0.00024027 | 0.0017008 | 0.0052418 | 0.0022283 | 0.0037184 | -0.0039898 | -5.7345e-05 | -0.0017161 | -9.5583e-06 | 0.00080084 | 5.9407e-06 | 5.3326e-06 | 2.0652e-05 | 0.00034402 | 0 | 0.00016483 | 8.8971e-05 | 110909.6427 | 125513.1679 | 8285.0685 | 148834.307 | 123948.0834 | 126617.5194 | 2293.5187 | 1.7827e-05 | 7.4318e-05 | 0.0043203 | 0.0085537
138 | MT 40 | R | 06/03/2023 05:29:36 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 40.140 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.0017021 | 0.00078512 | 0.0059864 | 0.19942 | 0.042812 | 0.00058068 | 0.0013237 | 1.6706e-05 | 0.002378 | 0.00048261 | 0.0076111 | 0.0023778 | 0.079537 | 8.0499e-05 | 0.00085265 | 0.0071268 | 0.0030658 | 0.00079278 | -0.0038317 | -1.2127e-05 | 0.015018 | 0.0002866 | 0.00052837 | 4.6402e-06 | -9.9946e-06 | 2.1253e-06 | 5.057e-05 | 9.5027e-06 | 0.00071265 | 0.00059915 | 109001.6313 | 141436.8137 | 8087.6293 | 151037.001 | 120807.4171 | 123997.8204 | 2527.5569 | 1.3463e-05 | 6.3984e-05 | 0.0042814 | 0.0086148
139 | MT 32 B | R | 06/03/2023 05:35:38 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 32 B.141 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.0003627 | 3.8223e-05 | 0.0039255 | 0.17135 | 0.040872 | 0.00067918 | 6.7424e-06 | 1.5135e-05 | 0.0021851 | 0.00052694 | 0.0057351 | 0.00033959 | 0.063967 | 3.4363e-05 | 0.0020739 | 0.0081773 | 0.0035652 | 0.0006063 | -0.0031693 | -5.9217e-05 | -0.0022533 | -1.8683e-05 | 0.00071246 | 0 | -9.9946e-06 | 1.2701e-05 | 0.00012112 | 0 | 0.0003166 | 0.00019688 | 109735.1973 | 134380.7531 | 8458.1717 | 147382.1668 | 124485.3752 | 125243.9642 | 2466.2131 | 1.3079e-05 | 0.00010863 | 0.0055027 | 0.011395
140 | MT 32 B LD | R | 06/03/2023 05:41:40 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 32 B LD.142 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.00050009 | 4.0236e-05 | 0.0038121 | 0.19262 | 0.040857 | 0.0012555 | 0.00021277 | -2.4724e-05 | 0.0017787 | 0.00047786 | 0.0038795 | 0.00021763 | 0.054829 | -4.0706e-05 | 0.00083428 | 0.011076 | 0.0042016 | 0.0028652 | -0.0021782 | -5.6582e-05 | -0.0030771 | -3.4539e-05 | 0.00043083 | 0 | -5.0916e-06 | -2.0035e-06 | 7.7335e-05 | 4.903e-06 | 0.00019652 | 0.00023221 | 110384.8308 | 135764.8934 | 8176.3394 | 148159.7448 | 120643.3072 | 112508.637 | 2491.5507 | 3.1466e-05 | 8.4069e-05 | 0.0064534 | 0.012482
141 | LBLANK | R | 06/03/2023 05:47:43 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LBLANK.143 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.00013171 | 0 | -6.2498e-05 | 0.017065 | 0.042346 | 0.0005268 | 8.7725e-06 | 5.8276e-05 | 0.0010951 | 0.00014414 | 0.0023313 | 0.000178 | 0.053911 | -8.0715e-05 | 0.00075466 | 0.00023075 | 0.00028361 | -0.00055786 | -0.0014801 | -9.744e-05 | -0.0046588 | -6.4259e-05 | 5.5871e-05 | 9.869e-06 | -1.2563e-07 | -1.1784e-05 | 0.00014724 | 4.977e-06 | 6.9902e-05 | 5.525e-05 | 118892.6303 | 135669.2228 | 8135.9833 | 156052.6038 | 124628.191 | 123071.2276 | 2460.2119 | 1.3315e-05 | 4.482e-05 | -7.8563e-05 | -1.3213e-05
142 | CCV 1ppb TRACE | R | 06/03/2023 05:53:46 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.144 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.23541 | 0.090243 | 0.0055371 | 0.026955 | 0.06846 | 0.09391 | 0.035624 | 0.11432 | 0.15835 | 0.050249 | 0.037147 | 0.095611 | 0.18519 | 0.30419 | 0.084107 | 0.20594 | 0.088219 | 0.039763 | 0.039267 | 0.0083228 | 2.1864 | 0.04014 | 0.0053642 | 0.00084827 | 0.0026337 | 0.005704 | 0.0035769 | 0.0010064 | 0.073233 | 0.071665 | 108962.6699 | 135163.7506 | 8107.9679 | 144719.4209 | 128114.0546 | 121753.3017 | 2481.5489 | 0.035368 | 0.084171 | 0.15867 | 0.31061
143 | CCV 1ppb TRACE | R | 06/03/2023 06:00:18 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.145 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.23132 | 0.088422 | 0.0059487 | 0.021201 | 0.069451 | 0.099435 | 0.032194 | 0.11419 | 0.15291 | 0.048609 | 0.036529 | 0.096385 | 0.18586 | 0.29428 | 0.083656 | 0.202 | 0.087163 | 0.039247 | 0.038052 | 0.0094283 | 2.2469 | 0.040784 | 0.0052591 | 0.0008777 | 0.0027548 | 0.005632 | 0.0036464 | 0.001002 | 0.073287 | 0.074644 | 109308.0459 | 132202.04 | 8255.7182 | 145354.418 | 118235.9765 | 120817.503 | 2398.2014 | 0.034745 | 0.082248 | 0.15313 | 0.33023
144 | CCV 10 ppb MAJORS | R | 06/03/2023 06:06:21 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.146 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.3566 | 0.90455 | 0.047061 | 0.255 | 0.30914 | 0.9936 | 0.32849 | 1.1477 | 1.5518 | 0.48319 | 0.32315 | 0.90439 | 1.3387 | 3.0489 | 0.81567 | 1.908 | 0.79769 | 0.38771 | 0.41757 | 0.096831 | 21.3467 | 0.39168 | 0.046148 | 0.0083833 | 0.026856 | 0.056659 | 0.038703 | 0.0096575 | 0.72749 | 0.72156 | 104410.9496 | 126236.6498 | 7482.2926 | 138433.5256 | 111639.2755 | 112936.7634 | 2317.5214 | 0.34073 | 0.81094 | 1.4159 | 3.1514
145 | CCV 10 ppb MAJORS | R | 06/03/2023 06:12:53 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.147 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.248 | 0.84715 | 0.047476 | 0.24439 | 0.31193 | 0.95032 | 0.31658 | 1.1037 | 1.5009 | 0.48153 | 0.31708 | 0.92067 | 1.3218 | 2.9083 | 0.78577 | 1.8834 | 0.80788 | 0.38825 | 0.39885 | 0.092549 | 21.9713 | 0.38706 | 0.044888 | 0.0088145 | 0.027236 | 0.055791 | 0.035607 | 0.0097016 | 0.73133 | 0.71725 | 108184.0893 | 128484.638 | 7761.1079 | 140634.1918 | 112430.1349 | 119538.1061 | 2263.5127 | 0.34518 | 0.80946 | 1.3739 | 3.1383
146 | MT 33 B | R | 06/03/2023 06:18:55 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 33 B.148 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.00028426 | 0.00011996 | 0.0043869 | 0.1623 | 0.036106 | 0.00016552 | 0.00021075 | 0.00012966 | 0.0019582 | 0.00044592 | 0.0027686 | 0.00032105 | 0.058723 | 0.00031309 | 0.0014933 | 0.0072874 | 0.0037459 | 0.00011142 | -0.0018705 | -1.8064e-05 | -0.0045473 | -6.1773e-05 | 0.00049301 | 4.8112e-06 | 4.4388e-06 | -1.6658e-05 | 6.1949e-05 | 0 | 0.00038041 | 0.00029873 | 113407.894 | 140808.6413 | 8418.1479 | 154598.0178 | 123486.8257 | 112713.6922 | 2576.2325 | 3.0176e-05 | 8.7973e-05 | 0.0043375 | 0.0091783
147 | MT 33 B LSP | R | 06/03/2023 06:24:56 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 33 B LSP.149 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 6.0175 | 2.2083 | 0.11757 | 0.80701 | 0.72615 | 2.5286 | 0.86192 | 2.9508 | 3.8811 | 1.1261 | 0.76125 | 2.3084 | 3.3315 | 7.5978 | 1.9994 | 4.716 | 1.7802 | 0.83925 | 0.98746 | 0.23004 | 52.0976 | 0.9403 | 0.10686 | 0.020521 | 0.062818 | 0.1324 | 0.087953 | 0.023285 | 1.7137 | 1.8717 | 110645.5313 | 125336.1184 | 7438.2704 | 151595.3223 | 109527.0168 | 111565.954 | 2262.846 | 0.75222 | 1.7845 | 3.0007 | 7.5795
148 | MT 34 B | R | 06/03/2023 06:30:59 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 34 B.150 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.00046337 | 0.00011491 | 0.0030951 | 0.15747 | 0.033284 | 0.001156 | 0.00019623 | 0.00020264 | 0.002178 | 0.00067083 | 0.0053395 | 0.00020279 | 0.061337 | 0.00021854 | 0.0033543 | 0.0043344 | 0.0023857 | 0.0008707 | -0.0031585 | 9.303e-05 | 0.0042106 | 0.00010014 | 0.0013062 | 9.4099e-06 | 8.9176e-06 | 6.9668e-06 | 9.5745e-05 | 0 | 0.0023824 | 0.0015179 | 111560.6606 | 140747.333 | 8704.6518 | 156238.4703 | 113549.1462 | 121911.3732 | 2562.2298 | 2.5604e-05 | 0.00015951 | 0.0045122 | 0.0098141
149 | MT 35 B | R | 06/03/2023 06:37:01 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 35 B.151 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.00064228 | 3.624e-05 | 0.0059437 | 0.16385 | 0.035097 | 0.00048838 | 0.00026849 | 8.3402e-05 | 0.0024534 | 0.00048582 | 0.0031431 | 0.00019867 | 0.054329 | 3.4435e-05 | 0.0012155 | 0.002739 | 0.001664 | -0.00021728 | -0.0042205 | -6.1369e-05 | -0.0018215 | -1.1437e-05 | 0.00074903 | 4.703e-06 | 8.7401e-06 | 1.1464e-05 | 0.00012984 | 0 | 0.0004567 | 0.00014934 | 113701.5025 | 143031.8332 | 8913.7866 | 157711.7943 | 127678.7591 | 128777.9023 | 2584.2341 | 2.1113e-05 | 6.8793e-05 | 0.0040968 | 0.0083031
150 | MT 36 B | R | 06/03/2023 06:43:03 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 36 B.152 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.0005618 | 0 | 0.0057132 | 0.15393 | 0.039379 | 0.00072663 | 4.9326e-05 | 9.4729e-05 | 0.0038173 | 0.0010144 | 0.010766 | 0.00038884 | 0.086247 | 0.00015121 | 0.0013027 | 0.0040205 | 0.0016266 | 0.13554 | -0.0042409 | -1.7134e-05 | -0.0031208 | -3.3487e-05 | 0.00068091 | 0 | -9.9946e-06 | 2.1989e-06 | 0.00022332 | 4.7635e-06 | 0.0040629 | 0.0038824 | 107517.2534 | 140482.7286 | 8423.8178 | 155251.8342 | 120076.0843 | 129946.5825 | 2604.2374 | 1.3008e-05 | 7.9257e-05 | 0.0042089 | 0.0087184
151 | MT 37 B | R | 06/03/2023 06:49:06 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 37 B.153 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.0018822 | 3.9628e-05 | 0.0040292 | 0.17316 | 0.036896 | 0.00030122 | 0.00020416 | 5.1672e-05 | 0.0017671 | 0.00058546 | 0.0030938 | 0.00029443 | 0.047636 | 7.418e-05 | 0.0012879 | 0.0045216 | 0.0020354 | 0.00030314 | -0.0040154 | -5.7812e-05 | -0.0052487 | -7.5297e-05 | 0.00083858 | 4.7767e-06 | 1.3523e-05 | -1.1881e-05 | -1.0057e-05 | 4.7036e-06 | 0.00023199 | 0.00014383 | 107769.8696 | 139907.8237 | 8521.8747 | 145882.8466 | 123289.7411 | 123509.6225 | 2501.5524 | 2.3015e-05 | 6.676e-05 | 0.0042375 | 0.0086087
152 | MT 38 B | R | 06/03/2023 06:55:09 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 38 B.154 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.00328 | 3.9756e-05 | 0.0040301 | 0.16718 | 0.034905 | 0.00073315 | 0.0001617 | 4.9623e-05 | 0.0020675 | 0.00053495 | 0.0033763 | 0.00017597 | 0.048456 | 0.00018833 | 0.0012427 | 0.0027056 | 0.0014335 | 0.00037291 | -0.0028468 | -9.744e-05 | -0.0053683 | -7.4543e-05 | 0.00060441 | 4.7894e-06 | -2.3627e-07 | -1.1689e-05 | 8.019e-05 | 9.7956e-06 | 0.00048182 | 0.00043007 | 111174.9387 | 137161.4801 | 8562.2338 | 159501.2907 | 127672.7171 | 123520.5176 | 2592.902 | 1.6053e-05 | 8.6613e-05 | 0.004483 | 0.0084733
153 | MT 39 B | R | 06/03/2023 07:01:12 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 39 B.155 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.00045372 | 8.1521e-05 | 0.0037725 | 0.16364 | 0.039325 | 0.0007422 | 0.00028925 | 0.00013738 | 0.0025164 | 0.00054162 | 0.0049424 | 0.00020343 | 0.061731 | -8.0879e-05 | 0.0013855 | 0.0035154 | 0.0014084 | 0.00054069 | -0.0031973 | -1.7426e-05 | -0.0056341 | -8.0583e-05 | 0.00084292 | 4.7207e-06 | 4.2086e-06 | -2.3725e-06 | 1.3397e-05 | 4.7617e-06 | 0.00098887 | 0.00093541 | 109188.3921 | 140477.3161 | 8279.3992 | 166545.4759 | 122897.0454 | 128990.5374 | 2613.5724 | 2.6572e-05 | 6.1375e-05 | 0.0044748 | 0.0083604
154 | MT 40 B | R | 06/03/2023 07:07:15 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 40 B.156 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.00028759 | 8.3739e-05 | 0.004508 | 0.16163 | 0.039374 | 0.00031324 | 0.00021135 | 6.0622e-05 | 0.0021028 | 0.00041696 | 0.0039882 | 0.00018804 | 0.05458 | 0.00056138 | 0.0012074 | 0.0081034 | 0.0030819 | 0.00048544 | -0.0030418 | -5.5597e-05 | -0.0017382 | -1.0059e-05 | 0.00047419 | 4.925e-06 | 4.7774e-06 | 1.2876e-05 | -3.6233e-05 | 4.922e-06 | 0.00015905 | 0.00016866 | 109975.7523 | 135427.5603 | 8220.0356 | 153551.8656 | 123180.851 | 122051.8341 | 2490.8839 | 1.7148e-05 | 6.2735e-05 | 0.0046307 | 0.0079897
155 | MT 32 C | R | 06/03/2023 07:13:17 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 32 C.157 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.00020805 | 3.9808e-05 | 0.0039436 | 0.16533 | 0.035838 | 0.00049529 | 8.8521e-05 | 9.6341e-05 | 0.0020289 | 0.00050885 | 0.0032334 | 0.00029912 | 0.04921 | -1.5911e-06 | 0.0011754 | 0.0079245 | 0.0036546 | 0.0014637 | -0.003676 | -1.6841e-05 | -0.004133 | -5.1656e-05 | 0.00033039 | 0 | -4.6717e-07 | 2.0157e-06 | 8.0559e-05 | 0 | 0.00025465 | 0.00014681 | 108777.3017 | 139746.3902 | 8298.4096 | 145988.16 | 121473.0822 | 120652.241 | 2636.2433 | 2.2808e-05 | 8.5418e-05 | 0.0055814 | 0.010992
156 | LBLANK | R | 06/03/2023 07:19:20 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LBLANK.158 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 8.6835e-05 | 0 | 8.8286e-05 | 0.015692 | 0.035841 | 0.00078888 | -3.248e-05 | 9.6042e-05 | 0.00073576 | 0.00012389 | 0.0019172 | 0.00011141 | 0.042189 | -4.3086e-05 | 0.00076101 | 0.00052106 | 0.00031354 | -0.00045627 | -0.0043 | -5.9374e-05 | -0.0052797 | -7.3809e-05 | 0.0004165 | 0 | 4.7811e-06 | -1.6853e-06 | -6.9747e-05 | 1.0122e-05 | 0.00013416 | 6.7365e-06 | 108185.5663 | 133665.4787 | 8389.4652 | 147146.6733 | 119876.9546 | 122861.4512 | 2468.8805 | 4.4744e-06 | 6.6533e-05 | -4.2353e-05 | -4.4267e-05
157 | CCV 1ppb TRACE | R | 06/03/2023 07:25:23 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.159 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.23764 | 0.093201 | 0.0053879 | 0.025644 | 0.064756 | 0.097302 | 0.032768 | 0.11525 | 0.15484 | 0.04965 | 0.036448 | 0.093392 | 0.18667 | 0.30379 | 0.080833 | 0.20363 | 0.087786 | 0.039731 | 0.037613 | 0.0092492 | 2.2016 | 0.039364 | 0.0053562 | 0.00083347 | 0.0027637 | 0.0055458 | 0.0036749 | 0.00097291 | 0.072677 | 0.071834 | 107282.1269 | 135732.5163 | 8100.9661 | 145758.489 | 118264.9241 | 118534.4169 | 2425.5392 | 0.034492 | 0.083893 | 0.1589 | 0.33511
158 | CCV 1ppb TRACE | R | 06/03/2023 07:31:56 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.160 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.24077 | 0.089141 | 0.0063737 | 0.02098 | 0.07234 | 0.098244 | 0.03325 | 0.11592 | 0.15247 | 0.055708 | 0.040626 | 0.093183 | 0.18414 | 0.30207 | 0.086977 | 0.21352 | 0.098017 | 0.044712 | 0.038688 | 0.0096581 | 2.0811 | 0.039166 | 0.0044014 | 0.00089502 | 0.0026007 | 0.0056667 | 0.0037993 | 0.0010466 | 0.081738 | 0.074014 | 106965.6327 | 133417.6497 | 8035.5929 | 130688.3426 | 114753.0851 | 119092.3798 | 2506.8881 | 0.039674 | 0.094809 | 0.15511 | 0.33203
159 | LFB 10ppb | R | 06/03/2023 07:37:58 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LFB 10ppb.161 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.2153 | 0.83585 | 0.048278 | 0.26928 | 0.29697 | 0.90121 | 0.31195 | 1.0963 | 1.4714 | 0.47149 | 0.33321 | 0.85247 | 1.3568 | 2.877 | 0.78428 | 1.8302 | 0.7789 | 0.50719 | 0.40259 | 0.091641 | 18.7284 | 0.35868 | 0.045362 | 0.0076784 | 0.02532 | 0.052535 | 0.03611 | 0.0097234 | 0.71502 | 0.66573 | 105510.8761 | 134505.7236 | 7576.3419 | 138475.9196 | 109901.8529 | 112617.5662 | 2588.2394 | 0.33715 | 0.78924 | 1.3336 | 3.0719
160 | LFB 10ppb | R | 06/03/2023 07:44:29 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LFB 10ppb.162 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.9673 | 0.75773 | 0.039726 | 0.25817 | 0.27074 | 0.83689 | 0.28035 | 0.99517 | 1.3461 | 0.023066 | 0.017763 | 0.88087 | 1.2396 | 2.5722 | 0.69847 | 1.6413 | 0.045602 | 0.023204 | 0.022208 | 0.084699 | 20.5375 | 0.38703 | -0.0058065 | 0.008252 | 0.027574 | 0.057437 | -0.00033844 | 0.0097677 | 0.040251 | 0.70544 | 107236.4077 | 121388.3141 | 8105.6386 | 157401.7136 | 135358.385 | 115099.1941 | 2286.1831 | 0.017099 | 0.039601 | 0.00029689 | 0.0003012
161 | CCV 10 ppb MAJORS | R | 06/03/2023 07:50:31 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.163 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.2503 | 0.84783 | 0.044761 | 0.2634 | 0.30231 | 0.94233 | 0.30879 | 1.1147 | 1.5042 | 0.49591 | 0.33017 | 0.91378 | 1.31 | 2.933 | 0.77391 | 1.8688 | 0.81478 | 0.39387 | 0.48007 | 0.090903 | 20.9098 | 0.39109 | 0.053794 | 0.0086477 | 0.026495 | 0.055982 | 0.043731 | 0.0095444 | 0.74633 | 0.72433 | 92513.7479 | 126923.3733 | 7568.6714 | 137014.649 | 102452.1228 | 112436.745 | 2377.5316 | 0.34558 | 0.82289 | 1.6104 | 3.443
162 | CCV 10 ppb MAJORS | R | 06/03/2023 07:57:03 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.164 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.2825 | 0.85997 | 0.043418 | 0.24994 | 0.30419 | 0.97239 | 0.32289 | 1.1486 | 1.5036 | 0.48834 | 0.32537 | 0.79015 | 1.3277 | 2.9651 | 0.79547 | 1.9115 | 0.80304 | 0.38535 | 0.38705 | 0.099219 | 18.6237 | 0.34133 | 0.042167 | 0.007617 | 0.02325 | 0.048742 | 0.035201 | 0.0087975 | 0.73597 | 0.63605 | 112585.4202 | 143947.2876 | 7557.3336 | 141370.0921 | 101985.7573 | 114759.1125 | 2632.2431 | 0.34063 | 0.80246 | 1.2724 | 3.4673
163 | MT 32 C LD | R | 06/03/2023 08:03:06 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 32 C LD.165 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.00079843 | 4.1896e-05 | 0.0036706 | 0.14602 | 0.038118 | 0.0011568 | 0.00013602 | 0.00010354 | 0.0021777 | 0.00045109 | 0.0035818 | 0.00023154 | 0.06254 | 4.1132e-05 | 0.0013529 | 0.0089734 | 0.0035945 | 0.00167 | -0.0014706 | -5.5544e-05 | -0.00075533 | 7.9539e-06 | 0.00051197 | 0 | 1.0784e-05 | 4.5769e-06 | 4.3054e-05 | 4.8941e-06 | 0.00057218 | 0.00026803 | 113745.265 | 126424.5215 | 8069.9471 | 151879.135 | 122677.8627 | 123253.4766 | 2297.5186 | 1.7575e-05 | 0.00012926 | 0.0055289 | 0.011158
164 | MT 33 C | R | 06/03/2023 08:09:08 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 33 C.166 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.00030163 | 0 | 0.0039691 | 0.15475 | 0.040496 | 0.00038911 | 0.00013373 | 1.9903e-05 | 0.0020826 | 0.00044629 | 0.0038212 | 0.0003144 | 0.06235 | 8.6991e-05 | 0.0019378 | 0.0083098 | 0.003608 | 0.00098354 | -0.0061636 | -9.744e-05 | -0.0055681 | -7.9208e-05 | 0.0011242 | 4.8834e-06 | -5.1113e-06 | -1.9539e-06 | -1.0173e-05 | 0 | 0.00030259 | 0.00019191 | 102830.0991 | 136268.1935 | 8059.6092 | 146983.639 | 123638.9099 | 121637.7582 | 2522.8895 | 2.7144e-05 | 8.0327e-05 | 0.0049804 | 0.0086118
165 | MT 33 C LSP | R | 06/03/2023 08:15:09 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 33 C LSP.167 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 5.9188 | 2.1599 | 0.12036 | 0.82796 | 0.69595 | 2.4295 | 0.84151 | 2.9485 | 3.8662 | 0.96172 | 0.60661 | 2.2629 | 3.2689 | 7.4253 | 1.9775 | 4.6104 | 1.497 | 0.80451 | 1.1194 | 0.23179 | 50.247 | 0.91152 | 0.12168 | 0.020436 | 0.062091 | 0.12818 | 0.097653 | 0.022725 | 0.0034528 | 1.8309 | 103961.4826 | 126128.0384 | 7388.2451 | 81125.7755 | 48003.1913 | 109728.4891 | 2293.518 | 0.92814 | 2.2406 | 2.5069 | 13.3088
166 | MT 33 C LSP | R | 06/03/2023 08:21:40 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 33 C LSP.168 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 5.8703 | 2.1768 | 0.11577 | 0.79715 | 0.70052 | 2.4799 | 0.83241 | 2.9625 | 3.8701 | 1.2601 | 0.85716 | 2.3081 | 3.2994 | 7.5777 | 1.996 | 4.6934 | 2.0137 | 0.93651 | 1.0161 | 0.23887 | 50.0561 | 0.9305 | 0.11113 | 0.0205 | 0.063007 | 0.13385 | 0.090997 | 0.022792 | 1.9163 | 1.8659 | 107237.2125 | 122716.8485 | 7409.5892 | 136288.4253 | 108838.3988 | 110484.9563 | 2284.1828 | 0.83073 | 1.9873 | 3.084 | 7.5928
167 | MT 34 C | R | 06/03/2023 08:27:42 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 34 C.169 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.00052605 | 7.9752e-05 | 0.0045583 | 0.15146 | 0.037417 | 0.00048582 | 4.7272e-05 | 9.5327e-05 | 0.002529 | 0.00069457 | 0.005599 | 0.00028382 | 0.068221 | 0.000516 | 0.0043138 | 0.0029562 | 0.0013802 | 7.5789e-06 | -0.0065806 | 2.2019e-05 | 0.0025443 | 6.6402e-05 | 0.0016949 | 4.979e-06 | 4.5234e-06 | -7.3444e-06 | 1.3753e-06 | 0 | 0.0023561 | 0.0016183 | 98199.9639 | 137868.6005 | 8358.4446 | 154848.0424 | 126369.1749 | 123458.5553 | 2480.2154 | 8.6166e-06 | 6.1907e-05 | 0.0045935 | 0.0076623
168 | MT 35 C | R | 06/03/2023 08:33:45 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 35 C.170 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.0006061 | 7.8735e-05 | 0.0038948 | 0.15744 | 0.036464 | 0.00073533 | 0.00013034 | 5.4528e-05 | 0.0013679 | 0.00046135 | 0.0032351 | 0.00019462 | 0.057774 | 7.4384e-05 | 0.0014286 | 0.0030689 | 0.0016504 | 0.00040336 | -0.0045963 | -9.744e-05 | -0.0035583 | -4.419e-05 | 0.0010526 | 0 | 8.826e-06 | -1.6658e-05 | 5.3118e-05 | 4.3138e-06 | 0.00060515 | 0.00012059 | 108502.6917 | 144322.8135 | 8376.4573 | 154816.2195 | 123058.7341 | 120223.9515 | 2600.237 | 1.7534e-05 | 8.7842e-05 | 0.0039732 | 0.0082232

## Net Intensities RSDs
| Sample Id | R | Acquisition Time | QC Status | Dataset File | Method File | Na KED 23/23 Helium KED_Low | Mg KED 24/24 Helium KED_Low | Al KED 27/27 Helium KED_Low | P Oshift 31/47 Oxygen DRC | S (34) -1 34/34 Helium KED_Low | K KED 39/39 Helium KED_Low | Ca KED 43/43 Helium KED_Low | V KED 51/51 Helium KED_Low | Cr (52) 52/52 Helium KED_Low | Cr (53) NH3 53/53 Ammonia DRC | Fe (54) NH3 54/54 Ammonia DRC | Mn Oshift 55/71 Oxygen DRC | Fe (56) KED 56/56 Helium KED_Low | Co KED 59/59 Helium KED_Low | Ni KED 60/60 Helium KED_Low | Cu (63) KED 63/63 Helium KED_Low | Cu (65) NH3 65/65 Ammonia DRC | Zn NH3 66/66 Ammonia DRC | As 75/75 | As KED 75/75 Helium KED_Low | As RH Oshift 75/91 Oxygen DRC | As Oshift 75/91 Oxygen DRC | Se (77) 77/77 | Se (77) Oshift 77/93 Oxygen DRC | Se (78) Oshift 78/94 Oxygen DRC | Se (80) Oshift 80/96 Oxygen DRC | Se (82) 82/82 | Se (82) Oshift 82/98 Oxygen DRC | Mo (95) NH3 95/95 Ammonia DRC | Mo (95) OShift 95/127 Oxygen DRC | Rh (IS1) 103/103 (IS) | Rh (IS2) 103/103 (IS) Oxygen DRC | Rh (IS3) 103/103 (IS) Helium KED_Low | Rh (IS4) 103/103 (IS) Ammonia DRC | Rh (IS5) 103/103 (IS) Ammonia DRC | Rh (IS6) 103/103 (IS) Oxygen DRC | Rh OShift (IS1) 103/119 (IS) Oxygen DRC | Cd (111) NH3 111/111 Ammonia DRC | Cd (114) NH3 114/114 Ammonia DRC | Pb 208/208 | Pb NH3 208/208 Ammonia DRC
1 | CAL BLANK | R | 6/2/2023 3:17:36 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CAL BLANK.004 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 19958% |  | 21532% | 9567% | 20105% | 27778% | 50751% | 49346% | 21942% | 134436% | 36290% | 58299% | 28034% | 48038% | 23171% | 20096% | 60020% | 40801% | 8103% | 32894% | 70215% | 79022% | 4647% |  | 68285% | 50156% | 8156% |  | 36940% | 48577% | 3.277% | 0.3095% | 0.7573% | 0.4141% | 0.5558% | 2.017% | 1.979% |  | 43731% | 4643% | 27324%
2 | LOW 0.01 | R | 6/2/2023 3:23:43 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LOW 0.01.005 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 37.93% | 38.37% | 34.34% | 14.63% | 720.6% | 45.48% | 157.2% | 32.78% | 47.54% | 16.06% | 3.985% | 12.08% | 64.71% | 37.97% | 67.08% | 40.79% | 2.628% | 8.752% | 361.5% | 4896% | 28.42% | 31.74% | 98.02% | 173.2% | 105.2% | 84.08% | 464.9% | 64.06% | 1.828% | 3.008% | 0.6899% | 1.464% | 2.25% | 1.661% | 1% | 1.345% | 3.114% | 14.25% | 8.617% | 6.584% | 5.353%
3 | LOW 0.05 | R | 6/2/2023 3:29:46 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LOW 0.05.006 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 22.65% | 8.42% | 21.23% | 12.05% | 1788% | 19.84% | 7.685% | 8.931% | 11.92% | 11.28% | 9.416% | 1.672% | 46.84% | 6.932% | 11.04% | 26.92% | 6.862% | 4.115% | 141.6% | 47.97% | 12% | 10.34% | 61.36% | 95.39% | 28.46% | 19.82% | 58.05% | 30.01% | 2.337% | 5.811% | 3.037% | 3.901% | 6.34% | 2.055% | 2.358% | 3.639% | 5.768% | 5.354% | 9.192% | 8.322% | 3.995%
4 | LOW 0.1 | R | 6/2/2023 3:35:48 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LOW 0.1.007 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 14.05% | 17.48% | 15.69% | 3.355% | 86.89% | 14.31% | 23.48% | 15.82% | 10.57% | 6.913% | 2.791% | 5.748% | 5.831% | 4.661% | 13.88% | 10.21% | 8.763% | 3.167% | 25.64% | 16.11% | 2.334% | 5.475% | 332.2% | 16.09% | 11.7% | 5.657% | 79.77% | 11.82% | 2.259% | 1.998% | 1.229% | 0.8802% | 2.017% | 1.21% | 0.9493% | 0.0713% | 2.965% | 0.9696% | 3.275% | 0.9616% | 1.512%
5 | LOW 0.5 | R | 6/2/2023 3:41:51 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LOW 0.5.008 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 6.21% | 8.498% | 11.12% | 4.773% | 24.44% | 5.172% | 14.51% | 15.17% | 9.428% | 0.3793% | 3.454% | 11.77% | 14.68% | 3.367% | 4.241% | 4.574% | 2.273% | 1.207% | 4.508% | 17.13% | 10.73% | 10.83% | 29.21% | 7.362% | 9.947% | 11.46% | 20.75% | 32.01% | 3.904% | 10.97% | 0.4323% | 8.264% | 5.283% | 1.121% | 0.4188% | 2.735% | 8.717% | 2.28% | 0.4935% | 0.6173% | 1.525%
6 | STD 1 | R | 6/2/2023 3:47:53 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\STD 1.009 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.199% | 3.568% | 28.76% | 3.714% | 3.95% | 2.192% | 4.261% | 1.265% | 1.605% | 1.7% | 3.296% | 0.6046% | 4.64% | 3.695% | 1.374% | 5.131% | 1.713% | 2.335% | 3.978% | 9.921% | 3.054% | 2.392% | 17.59% | 10.88% | 6.173% | 3.327% | 19.17% | 10.12% | 1.998% | 0.6261% | 0.6733% | 0.3953% | 2.17% | 1.01% | 9.13% | 0.9068% | 2.259% | 1.299% | 1.497% | 0.2841% | 10.91%
7 | STD 5 | R | 6/2/2023 3:53:56 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\STD 5.010 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.524% | 3.013% | 3.164% | 8.024% | 5.826% | 2.465% | 4.512% | 3.921% | 3.648% | 2.918% | 0.5387% | 0.6414% | 2.149% | 2.016% | 2.715% | 3.513% | 1.871% | 1.28% | 2.645% | 3.215% | 2.324% | 1.457% | 1.787% | 6.258% | 1.884% | 3.865% | 8.54% | 5.281% | 1.783% | 1.236% | 1.937% | 1.007% | 1.888% | 0.7846% | 0.8302% | 1.916% | 2.833% | 2.303% | 1.395% | 2.227% | 0.8798%
8 | STD 10 | R | 6/2/2023 3:59:58 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\STD 10.011 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.7704% | 1.719% | 1.791% | 1.031% | 6.078% | 0.5809% | 1.789% | 0.6704% | 1.042% | 1.087% | 2.541% | 1.834% | 1.001% | 1.088% | 0.6598% | 0.2344% | 1.273% | 1.497% | 2.184% | 2.554% | 2.103% | 1.925% | 4.638% | 3.616% | 3.432% | 1.756% | 1.729% | 5.913% | 1.019% | 1.426% | 0.4851% | 1.248% | 0.6937% | 0.8391% | 1.152% | 1.515% | 3.184% | 0.5491% | 1.025% | 0.5871% | 1.898%
9 | STD 25 | R | 6/2/2023 4:06:00 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\STD 25.012 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 6.527% | 5.257% | 5.674% | 1.049% | 5.631% | 6.483% | 7.079% | 6.003% | 5.528% | 1.229% | 1.138% | 1.613% | 6.77% | 6.813% | 6.776% | 5.293% | 1.078% | 0.22% | 2.767% | 4.545% | 3.126% | 1.529% | 4.806% | 0.2837% | 2.773% | 2.167% | 1.512% | 3.469% | 0.8785% | 0.9547% | 0.7074% | 1.095% | 4.204% | 0.3375% | 1.562% | 0.3715% | 3.595% | 0.1098% | 0.8436% | 1.016% | 0.9908%
10 | STD 50 | R | 6/2/2023 4:12:02 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\STD 50.013 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.94% | 0.7584% | 4.307% | 0.3753% | 1.687% | 1% | 1.001% | 1.347% | 1.022% | 1.506% | 1.068% | 1.439% | 0.8884% | 1.193% | 0.9179% | 0.9986% | 1.027% | 0.5775% | 3.197% | 1.582% | 2.128% | 0.7309% | 3.46% | 0.2135% | 0.8598% | 1.545% | 4.022% | 0.8252% | 0.7455% | 0.8752% | 3.075% | 0.651% | 0.2639% | 1.039% | 0.9888% | 1.18% | 2.279% | 1.101% | 1.062% | 6.202% | 1.389%
11 | STD 100 | R | 6/2/2023 4:18:04 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\STD 100.014 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.217% | 1.724% | 3.136% | 0.4734% | 2.52% | 2.265% | 1.373% | 1.824% | 2.448% | 0.6476% | 0.8827% | 0.5354% | 1.832% | 1.709% | 2.178% | 1.506% | 0.1297% | 0.3528% | 1.78% | 2.21% | 0.8052% | 0.2924% | 3.054% | 0.9717% | 0.9736% | 0.6654% | 1.798% | 0.4772% | 0.9044% | 0.2907% | 1.824% | 0.3891% | 2.22% | 0.9116% | 0.2953% | 0.2254% | 0.3844% | 0.5678% | 0.3087% | 2.301% | 0.3141%
13 | LBLANK | R | 6/2/2023 4:24:07 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LBLANK.015 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 45.58% | 173.2% | 33.62% | 11.13% | 41.47% | 204.9% | 157.7% | 44.73% | 107.3% | 47.51% | 38.62% | 47.79% | 394.8% | 87.78% | 103.3% | 129.7% | 12.3% | 5.24% | 543.6% | 267% | 58.37% | 50.3% | 645.9% |  | 102.4% | 44.25% | 87.11% | 173.2% | 12.56% | 15.93% | 17.26% | 1.191% | 0.8535% | 1.196% | 0.8029% | 0.3994% | 1.61% | 25.99% | 9.044% | 111% | 31.86%
14 | ICV PE | R | 6/2/2023 4:30:09 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\ICV PE.016 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 3.32% | 2.87% | 4.372% | 0.6954% | 1.464% | 1.72% | 4.526% | 2.516% | 4.133% | 1.668% | 1.343% | 0.3618% | 3.068% | 1.756% | 1.826% | 0.9874% | 1.693% | 1.275% | 2.892% | 1.307% | 1.206% | 1.215% | 1.759% | 2.008% | 2.472% | 1.51% | 1.604% | 1.215% | 0.4598% | 0.6169% | 0.5684% | 0.7782% | 2.672% | 1.203% | 0.7208% | 1.13% | 0.4998% | 1.822% | 1.62% | 0.7448% | 0.3811%
15 | CCV 1ppb TRACE | R | 6/2/2023 4:36:12 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.017 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 3.15% | 4.851% | 12.47% | 6.21% | 7.091% | 4.195% | 1.957% | 1.952% | 1.034% | 2.067% | 0.5884% | 0.9424% | 6.767% | 3.896% | 4.069% | 2.601% | 1.674% | 3.313% | 3.996% | 18.01% | 1.154% | 1.232% | 9.484% | 5.516% | 2.429% | 2.266% | 3.623% | 2.626% | 1.028% | 0.5735% | 0.3011% | 0.2138% | 1.258% | 0.8222% | 1.164% | 1.04% | 0.2663% | 1.518% | 2.082% | 1.767% | 1.069%
16 | CCV 1ppb TRACE | R | 6/2/2023 4:42:44 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.018 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.834% | 4.872% | 13.61% | 5.271% | 5.783% | 4.228% | 3.757% | 1.466% | 2.791% | 0.6202% | 0.4021% | 0.5035% | 6% | 3.253% | 5.018% | 1.431% | 1.468% | 4.347% | 4.087% | 9.221% | 2.404% | 3.099% | 3.573% | 8.644% | 8.127% | 1.598% | 15.16% | 9.317% | 1.06% | 0.4546% | 1.782% | 0.6125% | 1.518% | 0.4966% | 0.2556% | 0.8769% | 1.613% | 0.805% | 1.541% | 1.286% | 0.6484%
17 | LFB 10ppb | R | 6/2/2023 4:48:46 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LFB 10ppb.019 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.7295% | 1.678% | 3.565% | 0.7648% | 2.894% | 2.257% | 0.2621% | 2.241% | 2.804% | 0.185% | 1.872% | 0.6285% | 2.17% | 2.108% | 2.447% | 1.816% | 0.7211% | 0.9759% | 3.153% | 3.774% | 4.241% | 1.039% | 1.379% | 1.244% | 1.494% | 2.152% | 3.263% | 1.468% | 0.5984% | 2.54% | 1.052% | 1.135% | 1.665% | 0.5565% | 0.1487% | 0.4907% | 3.728% | 1.832% | 0.5359% | 0.8174% | 1.79%
18 | LFB 10ppb | R | 6/2/2023 4:55:16 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LFB 10ppb.020 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.223% | 1.548% | 1.709% | 2.581% | 3.633% | 1.835% | 1.567% | 0.578% | 1.619% | 0.5383% | 1.506% | 0.8246% | 1.584% | 0.9491% | 0.07163% | 1.904% | 0.993% | 0.1296% | 1.149% | 2.006% | 1.935% | 1.136% | 2.967% | 4.156% | 2.591% | 1.172% | 6.269% | 4.084% | 0.9583% | 1.105% | 1.041% | 1.084% | 0.8078% | 0.9704% | 1.585% | 1.534% | 2.448% | 0.2347% | 0.671% | 0.7906% | 1.882%
19 | LLOQ 0.01 | R | 6/2/2023 5:01:19 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LLOQ 0.01.021 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 20.43% | 40.57% | 46.07% | 12.67% | 13.17% | 111.5% | 37.86% | 39.83% | 9.784% | 8.361% | 12.2% | 16.62% | 10.06% | 27.12% | 7.324% | 20.62% | 10.33% | 9.673% | 207.5% | 109.8% | 18.03% | 20.65% | 205.1% | 48.95% | 133.4% | 55.62% | 125.2% | 86.6% | 3.669% | 4.832% | 0.3595% | 1.23% | 0.9884% | 0.27% | 2.44% | 0.9701% | 2.43% | 8.49% | 12.39% | 13.97% | 7.34%
20 | CCV 10 ppb MAJORS | R | 6/2/2023 5:07:22 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.022 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 4.232% | 4.446% | 3.712% | 2.383% | 5.344% | 4.901% | 7.113% | 3.61% | 5.447% | 1.3% | 1.008% | 1.795% | 5.183% | 4.844% | 4.781% | 5.094% | 1.102% | 1.452% | 1.851% | 6.908% | 1.474% | 1.008% | 3.066% | 3.833% | 1.518% | 1.447% | 1.841% | 5.272% | 0.9151% | 0.8814% | 0.7594% | 0.8657% | 4.005% | 0.9703% | 0.4442% | 1.381% | 1.44% | 0.8732% | 0.3795% | 0.6224% | 2.545%
21 | CCV 10 ppb MAJORS | R | 6/2/2023 5:13:54 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.023 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.13% | 2.288% | 1.205% | 0.9557% | 3.272% | 1.391% | 1.325% | 1.687% | 1.97% | 1.009% | 2.127% | 7.042% | 1.484% | 2.895% | 0.8611% | 2.111% | 0.9568% | 1.184% | 2.755% | 1.866% | 6.622% | 6.274% | 3.371% | 11.33% | 5.725% | 7.028% | 6.34% | 5.856% | 0.3589% | 6.572% | 1.208% | 5.284% | 1.377% | 1.12% | 1.755% | 0.5646% | 5.952% | 0.4711% | 1.547% | 0.788% | 2.224%
22 | LLOQ 0.1 | R | 6/2/2023 5:19:57 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LLOQ 0.1.024 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 14.18% | 7.66% | 19.14% | 17.95% | 17.24% | 17.12% | 20.51% | 16.9% | 8.578% | 14.15% | 10.51% | 4.059% | 17.16% | 2.27% | 16.21% | 2.826% | 13.32% | 16.26% | 62.84% | 22.03% | 7.096% | 5.627% | 387.3% | 74.3% | 11.29% | 1.621% | 40.98% | 21.96% | 12.42% | 9.497% | 9.201% | 1.188% | 2.087% | 9.1% | 2.183% | 1.131% | 3.106% | 13.42% | 9.921% | 9.686% | 1.985%
23 | LBLANK | R | 6/2/2023 5:25:59 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LBLANK.025 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 45.84% |  | 2216% | 4.171% | 22.14% | 1692% | 0% | 214.9% | 237.5% | 49.59% | 32.96% | 28.78% | 28.65% | 99.79% | 1368% | 206.5% | 15.73% | 15.98% | 1805% | 212.7% | 22.73% | 30.33% | 354.5% | 86.6% | 4835% | 91.15% | 405.4% | 114.3% | 40.76% | 66.43% | 0.7912% | 0.3726% | 1.346% | 6.44% | 0.4287% | 0.6385% | 1.826% | 10.79% | 47.75% | 246.5% | 56.94%
24 | LFB 6.25ppb & 2500ppb | R | 6/2/2023 5:32:01 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LFB 6.25ppb & 2500ppb.026 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.542% | 0.3859% | 5.616% | 3.43% | 1.013% | 1.797% | 2.386% | 1.968% | 2.035% | 0.6873% | 1.595% | 11.53% | 2.344% | 1.249% | 1.262% | 3.165% | 0.5916% | 1.694% | 1.988% | 2.988% | 12.86% | 10.38% | 5.974% | 14.82% | 9.068% | 8.694% | 2.888% | 10.34% | 1.704% | 11.28% | 1.492% | 8.57% | 1.654% | 0.2609% | 0.5971% | 0.9686% | 11.66% | 2.06% | 1.463% | 1.194% | 1.399%
25 | MT 1 | R | 6/2/2023 5:38:02 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 1.027 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.9523% | 10.06% | 0.7099% | 2.795% | 2.587% | 1.688% | 1.571% | 7.041% | 0.7907% | 1.406% | 0.6614% | 2.38% | 0.8748% | 10.58% | 0.6556% | 1.331% | 1.126% | 0.5997% | 5.388% | 9.099% | 8.295% | 2.104% | 8.876% | 21.63% | 7.741% | 7.101% | 9.04% | 12.21% | 4.435% | 1.999% | 0.9372% | 0.5176% | 1.204% | 0.4976% | 0.2489% | 2.413% | 5.928% | 1.426% | 5.123% | 1.358% | 1.416%
26 | MT 1 LD | R | 6/2/2023 5:44:04 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 1 LD.028 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.5946% | 12.08% | 0.7948% | 15.59% | 5.052% | 2.602% | 15.22% | 3.566% | 1.658% | 5.364% | 4.593% | 0.9719% | 0.9228% | 9.191% | 6.913% | 1.529% | 5.018% | 3.958% | 8.84% | 9.276% | 4.563% | 2.171% | 13.03% | 7.18% | 3.954% | 4.829% | 19.62% | 10.39% | 10.42% | 2.894% | 4.661% | 0.2501% | 0.2011% | 4.336% | 0.3626% | 12.94% | 3.78% | 12.86% | 4.926% | 6.176% | 2.374%
27 | LBLANK | R | 6/2/2023 6:05:56 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LBLANK.029 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 99.63% | 173.2% | 924.6% | 19.53% | 749.2% | 487.1% | 0% | 1747% | 68.13% | 40.65% | 10.72% | 53.54% | 274.8% | 10.55% | 19.61% | 43.32% | 41.06% | 10.9% | 142.6% | 0% | 156.3% | 5443% | 1371% | 173.2% | 226.8% | 129% | 1864% |  | 134.8% | 37.93% | 0.8819% | 3.788% | 9.354% | 1.37% | 0.708% | 0.9342% | 2.821% |  | 32.41% | 389.3% | 1476%
28 | STD 1 | R | 6/2/2023 6:11:58 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\STD 1.030 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.991% | 3.767% | 14.21% | 0.8094% | 5.794% | 2.459% | 6.318% | 1.727% | 3.107% | 1.901% | 2.689% | 2.499% | 4.27% | 2.132% | 4.93% | 1.034% | 1.399% | 1.081% | 2.971% | 8.526% | 1.204% | 1.964% | 15.82% | 1.509% | 3.133% | 1.716% | 6.416% | 12.86% | 1.243% | 0.9041% | 0.8004% | 0.4601% | 1.337% | 0.4157% | 0.9085% | 1.846% | 2.49% | 2.047% | 0.6988% | 0.9751% | 2.111%
29 | STD 0.5 | R | 6/2/2023 6:18:06 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\STD 0.5.031 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.583% | 4.824% | 18.14% | 0.5496% | 14.78% | 4.286% | 8.34% | 4.388% | 7.668% | 2.021% | 3.828% | 2.181% | 1.182% | 2.032% | 4.971% | 4.484% | 1.668% | 1.13% | 7.147% | 6.683% | 1.855% | 3.286% | 15.31% | 28.65% | 5.765% | 7.7% | 5.894% | 4.256% | 0.9006% | 1.759% | 5.292% | 0.7411% | 1.748% | 0.8597% | 0.9448% | 0.8309% | 3.994% | 5.048% | 0.3471% | 4.922% | 0.7829%
30 | ICV PE | R | 6/2/2023 6:24:08 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\ICV PE.032 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.8744% | 2.565% | 4.119% | 1.345% | 0.826% | 4.14% | 1.83% | 1.824% | 0.7845% | 1.226% | 2.773% | 0.3675% | 1.493% | 1.4% | 1.43% | 0.1444% | 1.477% | 2.458% | 2.325% | 2.718% | 1.109% | 1.556% | 2.286% | 1.76% | 1.84% | 0.6234% | 1.674% | 7.121% | 2.012% | 0.6537% | 0.4835% | 0.5159% | 0.3361% | 1.553% | 0.3284% | 0.9607% | 2.663% | 2.4% | 2.538% | 0.6701% | 1.604%
31 | ICV PE | R | 6/2/2023 6:30:39 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\ICV PE.033 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.505% | 1.362% | 2.616% | 1.658% | 2.309% | 1.649% | 3.18% | 1.385% | 0.3292% | 2.573% | 1.653% | 1.353% | 0.7721% | 0.9123% | 1.391% | 1.557% | 0.8237% | 1.311% | 2.133% | 6.16% | 2.526% | 1.221% | 2.112% | 2.531% | 1.477% | 0.8706% | 0.712% | 4.539% | 0.9972% | 0.2323% | 0.9306% | 1.079% | 0.9531% | 1.59% | 1.321% | 1.55% | 2.529% | 1.426% | 0.7498% | 0.4837% | 1.799%
32 | CCV 1ppb TRACE | R | 6/2/2023 6:36:41 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.034 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 5.24% | 0.3862% | 15.18% | 6.495% | 8.454% | 3.162% | 6.565% | 3.847% | 7.067% | 1.462% | 2.749% | 5.853% | 5.261% | 3.666% | 2.131% | 1.779% | 0.7186% | 2.08% | 4.946% | 11.79% | 7.059% | 5.74% | 16.37% | 9.737% | 7.902% | 5.335% | 3.868% | 18.83% | 0.6244% | 4.93% | 1.422% | 4.075% | 3.134% | 0.9271% | 0.1638% | 0.9257% | 5.339% | 0.6597% | 0.8619% | 0.5581% | 0.9821%
33 | CCV 1ppb TRACE | R | 6/2/2023 6:43:13 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.035 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.693% | 4.741% | 7.186% | 3.659% | 1.973% | 1.354% | 6.682% | 3.776% | 1.996% | 1.233% | 0.8493% | 1.603% | 3.045% | 1.963% | 1.226% | 1.819% | 1.524% | 1.887% | 3.57% | 6.523% | 3.828% | 1.579% | 7.383% | 12.62% | 5.082% | 2.875% | 6.31% | 7.986% | 0.9657% | 0.7475% | 0.4997% | 1.523% | 0.2567% | 1.24% | 1.106% | 0.315% | 3.524% | 2.592% | 1.272% | 0.5398% | 1.048%
34 | CCV 10 ppb MAJORS | R | 6/2/2023 6:49:16 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.036 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.719% | 3.159% | 7.193% | 0.5549% | 0.7488% | 2.675% | 2.845% | 1.841% | 1.286% | 1.276% | 1.126% | 0.6143% | 0.843% | 0.9014% | 1.574% | 0.2794% | 0.4116% | 1.43% | 1.949% | 4.532% | 3.447% | 0.4628% | 2.612% | 2.337% | 0.3932% | 1.4% | 0.5407% | 5.186% | 1.466% | 1.706% | 1.213% | 0.4581% | 0.9478% | 1.095% | 2.5% | 1.903% | 3.479% | 0.9398% | 1.432% | 0.2887% | 1.552%
35 | CCV 10 ppb MAJORS | R | 6/2/2023 6:55:47 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.037 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.2967% | 1.362% | 6.933% | 1.373% | 4.063% | 1.507% | 2.864% | 1.312% | 1.801% | 1.107% | 2.787% | 0.9143% | 1.299% | 0.6006% | 0.7309% | 1.242% | 1.539% | 2.758% | 1.074% | 2.836% | 1.371% | 0.9293% | 1.458% | 2.875% | 3.755% | 1.603% | 0.6947% | 1.416% | 0.9571% | 0.3417% | 0.4744% | 0.1205% | 0.8444% | 2.254% | 1.46% | 1.43% | 1.266% | 1.093% | 2.006% | 0.9762% | 0.7456%
36 | MT 1 | R | 6/2/2023 7:01:50 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 1.038 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.618% | 2.146% | 0.8829% | 0.8532% | 6.4% | 0.9956% | 13.69% | 3.899% | 8.83% | 3.571% | 1.405% | 12.64% | 0.2735% | 6.226% | 5.744% | 0.6275% | 1.923% | 0.6785% | 6.155% | 8.522% | 13.77% | 12.27% | 13.12% | 4.858% | 13.09% | 17.48% | 8.395% | 16.57% | 2.537% | 7.712% | 0.6626% | 10.46% | 0.5676% | 0.6905% | 1.108% | 1.269% | 11.49% | 12.66% | 2.494% | 2.261% | 3.264%
37 | MT 1 LD | R | 6/2/2023 7:07:53 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 1 LD.039 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 6.614% | 6.627% | 6.589% | 0.9083% | 14.73% | 6.425% | 18.23% | 4.551% | 12.44% | 0.3385% | 0.3624% | 2.289% | 7.38% | 6.627% | 7.199% | 7.134% | 1.446% | 1.316% | 3.23% | 9.314% | 1.097% | 3.339% | 16.43% | 7.353% | 15.16% | 4.734% | 2.208% | 17.87% | 4.92% | 3.065% | 0.4137% | 1.447% | 5.661% | 0.9494% | 0.5818% | 1.856% | 1.371% | 5.555% | 2.308% | 1.273% | 0.8294%
38 | MT 2 | R | 6/2/2023 7:13:55 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 2.040 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.673% | 6.939% | 3.273% | 0.9013% | 2.438% | 1.695% | 13.19% | 1.393% | 3.435% | 1.494% | 1.935% | 7.32% | 3.103% | 7.214% | 10.61% | 2.961% | 2.013% | 1.378% | 5.04% | 2.11% | 12.44% | 9.284% | 11.68% | 18.92% | 11.31% | 11.85% | 5.786% | 8.799% | 5.765% | 11.03% | 0.7604% | 6.356% | 1.884% | 0.8882% | 0.8442% | 1.233% | 9.465% | 4.7% | 3.209% | 1.106% | 2.073%
39 | MT 2 LSP | R | 6/2/2023 7:19:58 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 2 LSP.041 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 3.86% | 5.129% | 6.004% | 5.762% | 5.774% | 6.554% | 4.947% | 4.334% | 5.808% | 1.094% | 0.3682% | 0.5327% | 5.606% | 6.298% | 6.59% | 6.789% | 0.6402% | 0.6013% | 2.794% | 12.02% | 1.875% | 1.452% | 2.487% | 6.369% | 3.318% | 2.365% | 2.042% | 1.022% | 1.198% | 1.469% | 0.5755% | 0.9017% | 5.028% | 0.8139% | 1.077% | 5.053% | 0.6228% | 1.1% | 0.5801% | 0.3512% | 1.546%
40 | MT 2 LSP | R | 6/2/2023 7:26:09 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 2 LSP.042 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 4.585% | 4.819% | 2.95% | 1.36% | 3.243% | 2.205% | 2.883% | 1.681% | 2.185% | 0.7603% | 0.7882% | 2.081% | 3.037% | 4.249% | 3.64% | 2.768% | 1.087% | 0.7477% | 1.627% | 2.508% | 3.32% | 1.47% | 4.741% | 2.453% | 1.906% | 1.686% | 3.517% | 3.133% | 1.65% | 0.6608% | 1.223% | 1.209% | 3.109% | 0.5526% | 0.5652% | 0.4011% | 3.085% | 1.917% | 0.506% | 1.017% | 1.437%
41 | MT 3 | R | 6/2/2023 7:32:11 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 3 .043 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.129% | 7.466% | 3.622% | 1.546% | 7.44% | 3.837% | 6.059% | 4.96% | 4.512% | 4.708% | 0.6642% | 1.517% | 2.214% | 8.263% | 1.708% | 2.222% | 0.4454% | 1.059% | 0.8146% | 7.91% | 1.997% | 1.878% | 19.24% | 4.132% | 3.657% | 2.807% | 10.92% | 3.491% | 3.564% | 1.06% | 0.7661% | 0.9923% | 1.467% | 1.192% | 3.551% | 0.8537% | 3.301% | 7.487% | 7.375% | 0.4957% | 3.643%
42 | MT 4 | R | 6/2/2023 7:38:14 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 4.044 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.605% | 3.951% | 0.3797% | 1.231% | 4.638% | 2.288% | 10.18% | 5.864% | 2.775% | 2.206% | 1.035% | 1.501% | 1.352% | 7.346% | 10.42% | 1.421% | 1.709% | 0.799% | 4.387% | 6.278% | 1.615% | 0.8044% | 4.613% | 8.739% | 8.395% | 1.116% | 4.224% | 6.232% | 0.2755% | 3.632% | 1.747% | 1.595% | 0.5574% | 1.393% | 0.9051% | 2.019% | 1.509% | 3.402% | 4.245% | 4.618% | 1.114%
43 | LBLANK | R | 6/2/2023 7:44:17 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LBLANK.045 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 369.2% | 173.2% | 46.02% | 35.33% | 25.93% | 45.69% | 0% | 332.4% | 39.63% | 79.34% | 27.79% | 39.39% | 10.85% | 143.8% | 204.5% | 34.89% | 25.28% | 65.63% | 53.97% | 81.76% | 1708% | 219.3% | 1022% | 173.2% | 75.71% | 27.51% | 112% | 86.6% | 32.25% | 19.29% | 0.9507% | 0.8704% | 1.308% | 6.585% | 2.3% | 11.33% | 1.138% | 173.2% | 19.49% | 72.33% | 461.2%
44 | CCV 1ppb TRACE | R | 6/2/2023 7:50:20 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.046 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.3229% | 2.811% | 24.23% | 3.469% | 12.15% | 2.776% | 1.739% | 5.481% | 2.119% | 3.575% | 1.61% | 0.5419% | 5.076% | 2.31% | 0.2077% | 1.438% | 1.58% | 0.6894% | 5.986% | 12.45% | 1.362% | 2.067% | 14.1% | 5.078% | 11.6% | 3.18% | 7.284% | 10.85% | 3.455% | 1.616% | 8.804% | 0.1298% | 0.7926% | 1.001% | 1.473% | 0.9488% | 1.764% | 1.335% | 0.5268% | 9.501% | 1.803%
45 | CCV 1ppb TRACE | R | 6/2/2023 7:56:52 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.047 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 3.22% | 5.252% | 11.8% | 6.354% | 5.102% | 4.768% | 2.726% | 1.293% | 2.343% | 2.696% | 2.474% | 1.147% | 0.8245% | 2.178% | 3.164% | 1.589% | 0.828% | 3.414% | 2.624% | 7.211% | 3.392% | 0.2297% | 3.893% | 16.52% | 3.074% | 2.466% | 5.436% | 25.63% | 0.8997% | 1.196% | 0.5208% | 0.8058% | 0.5806% | 0.6088% | 1.214% | 0.795% | 2.535% | 2.136% | 1.308% | 0.6368% | 1.157%
46 | CCV 10 ppb MAJORS | R | 6/2/2023 8:02:55 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.048 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.289% | 1.347% | 3.768% | 1.226% | 4.75% | 2.135% | 3.51% | 2.011% | 1.835% | 0.8877% | 1.956% | 1.928% | 2.668% | 1.481% | 2.075% | 2.361% | 0.9361% | 0.6499% | 2.661% | 6.695% | 1.357% | 2.887% | 0.9557% | 6.167% | 2.973% | 1.233% | 2.47% | 5.537% | 1.559% | 2.024% | 1.593% | 1.513% | 1.337% | 0.1982% | 1.956% | 2.399% | 0.9019% | 1.083% | 0.9206% | 0.8158% | 0.4519%
47 | CCV 10 ppb MAJORS | R | 6/2/2023 8:09:27 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.049 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 3.267% | 2.924% | 1.789% | 0.9745% | 2.896% | 2.005% | 4.507% | 2.445% | 2.559% | 2.789% | 2.488% | 1.129% | 3.787% | 2.111% | 2.143% | 3.88% | 1.946% | 2.05% | 3.611% | 5.005% | 1.993% | 0.4077% | 3.544% | 3.656% | 1.384% | 0.9998% | 4.401% | 4.981% | 1.665% | 0.707% | 1.685% | 0.492% | 1.812% | 1.45% | 4.316% | 0.8608% | 1.801% | 0.3462% | 0.8331% | 0.279% | 5.2%
48 | MT 5 | R | 6/2/2023 8:15:30 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 5.050 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.29% | 8.265% | 2.911% | 1.672% | 2.473% | 1.101% | 3.944% | 3.493% | 2.708% | 0.6871% | 2.844% | 1.394% | 2.541% | 12.66% | 6.535% | 1.404% | 1.86% | 2.26% | 4.009% | 3.222% | 1.681% | 2.204% | 3.516% | 15.19% | 6.825% | 8.633% | 9.786% | 11.8% | 3.133% | 2.911% | 0.09245% | 1.465% | 2.185% | 1.019% | 1.38% | 0.9517% | 1.248% | 1.986% | 2.152% | 1.839% | 0.7303%
49 | MT 6 | R | 6/2/2023 8:21:32 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 6.051 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.772% | 1.374% | 3.424% | 1.208% | 3.779% | 1.644% | 2.286% | 4.576% | 1.639% | 1.1% | 1.148% | 1.776% | 1.788% | 2.741% | 1.052% | 2.077% | 1.645% | 1.105% | 1.981% | 3.572% | 1.547% | 0.1607% | 1.708% | 2.919% | 4.218% | 0.6309% | 3.357% | 1.212% | 5.087% | 3.625% | 0.1585% | 0.8318% | 1.771% | 0.2038% | 3.655% | 0.5005% | 0.8938% | 6.287% | 1.275% | 0.166% | 6.217%
50 | MT 7 | R | 6/2/2023 8:27:35 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 7.052 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.263% | 13.11% | 1.852% | 2.419% | 1.792% | 0.4407% | 4.391% | 0.9809% | 0.9718% | 16.57% | 15.95% | 0.9843% | 1.305% | 2.216% | 2.589% | 1.088% | 16.9% | 15.38% | 2.366% | 1.599% | 5.984% | 1.43% | 14.77% | 3.733% | 5.016% | 1.004% | 4.62% | 3.226% | 17.78% | 1.559% | 1.298% | 0.5251% | 0.5266% | 13.62% | 2.073% | 2.553% | 5.283% | 16.31% | 14.09% | 1.655% | 0.834%
51 | MT 8 | R | 6/2/2023 8:33:38 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 8.053 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 7.704% | 3.68% | 10.87% | 1.416% | 10.69% | 9.75% | 13.76% | 9.425% | 7.974% | 3.452% | 0.3203% | 0.5276% | 8.437% | 13.33% | 11.01% | 10.33% | 0.8455% | 0.348% | 1.701% | 10.45% | 2.566% | 1.782% | 10.59% | 5.559% | 2.443% | 1.845% | 3.73% | 4.11% | 2.673% | 2.913% | 1.041% | 0.941% | 8.569% | 0.6776% | 1.889% | 0.8354% | 3.476% | 8.588% | 2.741% | 2.371% | 2.747%
52 | MT 9 | R | 6/2/2023 8:39:40 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 9.054 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.01031% | 6.595% | 0.6398% | 1.861% | 2.752% | 1.745% | 10.01% | 5.337% | 2.849% | 0.9476% | 0.9148% | 2.498% | 0.8676% | 7.247% | 6.14% | 1.859% | 0.6768% | 0.7404% | 3.663% | 8.377% | 2.855% | 0.8472% | 4.518% | 7.52% | 5.713% | 7.062% | 7.397% | 8.634% | 1.803% | 3.582% | 2.996% | 2.092% | 1.329% | 0.3894% | 0.5578% | 0.8843% | 1.274% | 3.763% | 3.761% | 4.922% | 1.817%
53 | MT 10 | R | 6/2/2023 8:45:42 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 10.055 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.323% | 9.876% | 4.124% | 1.021% | 2.231% | 2.348% | 6.217% | 3.281% | 1.199% | 2.399% | 1.384% | 0.8286% | 1.62% | 2.526% | 2.28% | 2.334% | 1.901% | 1.341% | 2.221% | 8.512% | 1.665% | 0.3359% | 3.833% | 16.62% | 1.011% | 4.927% | 3.844% | 13.04% | 2.343% | 2.998% | 0.6311% | 0.1809% | 2.343% | 0.8768% | 1.155% | 0.2671% | 1.945% | 4.796% | 2.232% | 1.972% | 1.588%
54 | MT 11 | R | 6/2/2023 8:51:44 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 11 .056 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.658% | 2.042% | 1.275% | 0.6638% | 0.1183% | 3.305% | 5.785% | 3.47% | 0.5175% | 0.5407% | 0.05031% | 0.8457% | 1.358% | 0.9129% | 4.981% | 0.3242% | 0.4376% | 0.812% | 1.858% | 4.618% | 2.151% | 1.501% | 4.778% | 6.538% | 3.262% | 3.007% | 0.8276% | 1.663% | 1.053% | 3.38% | 0.771% | 1.447% | 0.4026% | 0.4229% | 1.155% | 1.645% | 1.41% | 0.4153% | 1.898% | 3.219% | 1.377%
55 | MT 11 LD | R | 6/2/2023 8:57:47 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 11 LD.057 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.235% | 3.977% | 2.833% | 1.815% | 1.963% | 1.898% | 3.68% | 1.749% | 2.283% | 1.676% | 1.401% | 2.174% | 2.384% | 0.8659% | 3.085% | 1.922% | 2.277% | 1.718% | 2.335% | 2.908% | 3.427% | 0.9455% | 4.482% | 8.463% | 3.215% | 1.836% | 5.59% | 2.314% | 1.356% | 2.321% | 0.9997% | 0.233% | 2.233% | 1.473% | 0.3604% | 1.211% | 2.723% | 4.002% | 2.309% | 0.9331% | 0.5449%
56 | MT 11 LD | R | 6/2/2023 9:04:18 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 11 LD.058 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 10.95% | 13.19% | 14% | 1.038% | 14.23% | 12.72% | 13.43% | 15.17% | 13.3% | 0.7542% | 1.632% | 1.062% | 13.74% | 14.67% | 10.24% | 11.69% | 0.405% | 1.264% | 1.612% | 14.09% | 3.07% | 0.5866% | 2.559% | 2.191% | 3.409% | 0.8897% | 6.877% | 3.8% | 1.762% | 2.508% | 1.546% | 0.8353% | 10.39% | 0.6253% | 1.273% | 1.641% | 2.827% | 1.867% | 1.75% | 3.153% | 2.174%
57 | MT 12 | R | 6/2/2023 9:10:20 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 12.059 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 4.835% | 7.835% | 4.964% | 1.033% | 5.627% | 5.071% | 3.78% | 4.713% | 2.903% | 23.13% | 22.97% | 1.943% | 5.905% | 6.119% | 1.72% | 4.316% | 22.62% | 22.47% | 5.693% | 6.557% | 1.133% | 1.974% | 7.861% | 3.611% | 9.517% | 1.027% | 4.655% | 7.271% | 23.24% | 1.531% | 4.179% | 1.193% | 4.493% | 18.38% | 2.352% | 2.59% | 0.208% | 20.97% | 20.77% | 3.738% | 2.087%
58 | MT 12 LSP | R | 6/2/2023 9:16:23 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 12 LSP.060 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.005% | 2.551% | 1.04% | 0.8617% | 2.5% | 1.005% | 0.3533% | 1.005% | 1.549% | 1.743% | 1.32% | 1.265% | 0.5242% | 1.192% | 2.064% | 0.5378% | 1.379% | 1.922% | 7.501% | 4.494% | 3.674% | 1.117% | 7.863% | 4.588% | 2.182% | 0.3149% | 8.43% | 3.45% | 0.7928% | 1.338% | 6.188% | 0.7726% | 0.7554% | 1.518% | 0.2408% | 1.442% | 4.271% | 1.465% | 1.311% | 7.244% | 1.223%
59 | LBLANK | R | 6/2/2023 9:22:26 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LBLANK.061 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 29.51% | 86.61% | 95.21% | 58.81% | 9.168% | 126.6% | 0% | 77.89% | 58.49% | 68.54% | 1.698% | 8.016% | 16.48% | 281.5% | 34.84% | 6.632% | 14.31% | 87.26% | 56.58% | 155.5% | 55.74% | 102.8% | 236.7% | 173.2% | 140.8% | 227.2% | 180.7% | 173.2% | 21.97% | 15.52% | 9.207% | 1.278% | 1.399% | 0.9841% | 1.208% | 0.6588% | 1.331% | 108.1% | 36.96% | 346.5% | 957.1%
60 | CCV 1ppb TRACE | R | 6/2/2023 9:28:29 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.062 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 4.813% | 5.025% | 7.704% | 3.35% | 6.037% | 2.724% | 10.05% | 6.206% | 7.747% | 1.108% | 0.9418% | 2.279% | 2.926% | 4.751% | 4.045% | 8.519% | 1.876% | 1.295% | 2.415% | 9.112% | 3.908% | 2.524% | 13.04% | 10.71% | 5.047% | 1.164% | 6.48% | 3.696% | 1.137% | 2.719% | 1.834% | 1.485% | 2.968% | 1.363% | 1.804% | 0.514% | 2.523% | 1.684% | 0.9932% | 1.253% | 0.8743%
61 | CCV 1ppb TRACE | R | 6/2/2023 9:35:00 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.063 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.639% | 3.916% | 29.43% | 8.822% | 3.464% | 5.188% | 4.002% | 1.844% | 1.858% | 10.44% | 10.08% | 2.137% | 4.14% | 2.38% | 4.329% | 3.225% | 10.56% | 10.76% | 5.676% | 10.5% | 0.9758% | 1.938% | 7.203% | 16.51% | 1.279% | 0.9182% | 12.16% | 10.41% | 9.593% | 2.346% | 1.099% | 1.942% | 2.372% | 8.414% | 0.7711% | 0.6928% | 3.014% | 8.052% | 10.68% | 1.077% | 0.9474%
62 | LFB 10ppb | R | 6/2/2023 9:41:03 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LFB 10ppb.064 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.9% | 2.928% | 5.207% | 0.9751% | 6.144% | 4.188% | 1.3% | 2.582% | 0.9722% | 0.8373% | 4.245% | 16.94% | 2.491% | 2.418% | 3.179% | 1.698% | 1.809% | 2.21% | 6.724% | 2.29% | 16.23% | 15.76% | 7.141% | 13.43% | 14.74% | 15.42% | 6.692% | 13.67% | 1.263% | 17.29% | 4.147% | 13.4% | 2.311% | 1.59% | 1.582% | 0.8024% | 13.85% | 2.455% | 1.38% | 5.79% | 1.497%
63 | LFB 10ppb | R | 6/2/2023 9:47:33 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LFB 10ppb.065 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.156% | 1.928% | 7.385% | 1.051% | 5.617% | 1.355% | 2.2% | 2.323% | 1.208% | 1.293% | 1.659% | 2.424% | 2.186% | 0.73% | 0.1394% | 1.025% | 1.84% | 0.9525% | 0.9149% | 3.019% | 2.977% | 1.153% | 5.266% | 2.013% | 3.157% | 1.548% | 2.179% | 0.7443% | 0.4031% | 1.096% | 0.8385% | 0.9454% | 1.375% | 1.43% | 1.04% | 1.192% | 2.932% | 0.8785% | 1.11% | 0.6741% | 2.21%
64 | CCV 10 ppb MAJORS | R | 6/2/2023 9:53:35 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.066 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.6134% | 1.299% | 10.93% | 1.589% | 3.949% | 2.307% | 4.198% | 0.6948% | 2.04% | 15.54% | 15.61% | 1.516% | 3.002% | 1.758% | 0.8807% | 0.2247% | 14.61% | 14.59% | 13.22% | 5.703% | 0.8608% | 1.762% | 15.13% | 6.944% | 3.464% | 2.729% | 11.55% | 4.359% | 16.43% | 1.002% | 9.805% | 1.096% | 0.7243% | 13.02% | 0.8641% | 0.6459% | 0.7573% | 16.08% | 15.38% | 10.94% | 0.7985%
65 | CCV 10 ppb MAJORS | R | 6/2/2023 10:00:06 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.067 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 11.2% | 11.25% | 13.46% | 13.1% | 13.23% | 11.64% | 12.42% | 10.23% | 11.5% | 17.38% | 18.86% | 0.1295% | 12.26% | 12.46% | 13.24% | 11.99% | 16.29% | 16.36% | 3.625% | 13.18% | 1.068% | 0.3854% | 2.454% | 5.472% | 1.962% | 2.293% | 1.674% | 1.452% | 16.69% | 1.641% | 1.235% | 0.3058% | 8.985% | 13.17% | 0.8413% | 9.628% | 0.7577% | 15.62% | 16.48% | 2.197% | 1.146%
66 | MT 13 | R | 6/2/2023 10:06:09 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 13.068 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.946% | 6.794% | 2.892% | 1.587% | 5.363% | 3.025% | 4.263% | 0.4162% | 1.619% | 1.913% | 2.02% | 2.175% | 1.236% | 2.622% | 3.609% | 0.8119% | 1.248% | 1.977% | 13.54% | 1.428% | 4.577% | 0.3386% | 20.16% | 2.661% | 1.74% | 1.272% | 15.07% | 6.429% | 2.641% | 0.4981% | 12.36% | 0.9991% | 1.253% | 1.454% | 1.079% | 1.472% | 4.465% | 1.828% | 1.342% | 15.48% | 1.432%
67 | MT 14 | R | 6/2/2023 10:12:11 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 14.069 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.224% | 3.367% | 1.881% | 0.07655% | 2.449% | 0.883% | 7.067% | 1.416% | 3.08% | 11% | 7.299% | 0.6105% | 1.057% | 4.193% | 2.133% | 1.816% | 0.911% | 0.9486% | 2.027% | 2.772% | 4.157% | 1.788% | 6.77% | 3.319% | 1.864% | 0.4896% | 1.602% | 9.842% | 3.349% | 5.609% | 2.437% | 1.154% | 0.7913% | 0.5921% | 0.6574% | 0.9968% | 3.947% | 2.11% | 2.918% | 4.002% | 0.4273%
68 | MT 15 | R | 6/2/2023 10:18:14 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 15.070 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 3.595% | 7.022% | 3.292% | 3.522% | 2.503% | 1.178% | 1.559% | 3.707% | 2.76% | 0.4388% | 0.8351% | 2.258% | 1.034% | 3.369% | 1.848% | 2.299% | 0.6572% | 1.641% | 0.5682% | 7.542% | 4.19% | 0.7784% | 0.8318% | 10.23% | 2.603% | 0.5271% | 2.458% | 11.17% | 1.393% | 3.152% | 0.8478% | 0.8153% | 1.37% | 0.3174% | 0.3711% | 0.6713% | 3.628% | 3.79% | 2.035% | 3.689% | 1.051%
69 | MT 16 | R | 6/2/2023 10:24:17 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 16.071 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.305% | 1.591% | 6.219% | 0.3291% | 5.029% | 1.636% | 5.336% | 4.596% | 1.418% | 1.816% | 1.119% | 0.9035% | 0.3392% | 1.834% | 4.681% | 0.4744% | 0.9308% | 1.226% | 16.45% | 5.99% | 4.699% | 1.312% | 26.77% | 5.607% | 1.679% | 3.808% | 22.36% | 4.495% | 2.505% | 1.379% | 13.4% | 0.8907% | 0.5334% | 0.9951% | 0.3905% | 0.4574% | 3.255% | 2.056% | 2.648% | 15.76% | 0.9135%
70 | MT 17 | R | 6/2/2023 10:30:20 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 17.072 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.15% | 2.981% | 1.542% | 3.091% | 1.876% | 1.309% | 4.036% | 8.476% | 1.345% | 1.816% | 0.984% | 10.45% | 0.7676% | 0.9796% | 3.876% | 1.048% | 1.616% | 1.017% | 1.474% | 4.44% | 9.779% | 10.26% | 4.69% | 5.037% | 8.194% | 10.52% | 1.293% | 6.488% | 1.402% | 9.864% | 1.278% | 7.663% | 0.6128% | 1.382% | 0.6488% | 2.926% | 7.486% | 0.6166% | 0.8778% | 1.583% | 0.651%
71 | MT 18 | R | 6/2/2023 10:36:32 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 18.073 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.904% | 6.103% | 3.868% | 1.868% | 4.377% | 3.94% | 14.92% | 1.456% | 1.523% | 1.305% | 1.414% | 2.313% | 1.495% | 4.218% | 3.117% | 2.324% | 1.031% | 1.319% | 1.544% | 5.133% | 0.3677% | 0.955% | 7.432% | 4.521% | 2.536% | 2.36% | 0.8564% | 4.992% | 0.3195% | 1.785% | 0.6094% | 0.8108% | 2.343% | 0.8281% | 0.9725% | 3.367% | 0.6683% | 0.8716% | 0.7466% | 2.119% | 2.246%
72 | MT 19 | R | 6/2/2023 10:42:43 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 19.074 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 3.167% | 5.141% | 2.388% | 0.03002% | 4.025% | 3.56% | 16.24% | 8.464% | 17.4% | 21.82% | 20.05% | 2.728% | 3.939% | 8.535% | 1.88% | 3.94% | 18.29% | 18.74% | 1.054% | 7.58% | 1.892% | 1.739% | 8.703% | 9.806% | 2.496% | 5.499% | 1.777% | 0.9763% | 19.84% | 1.975% | 0.7626% | 0.952% | 2.992% | 16.28% | 0.8677% | 0.7387% | 0.3675% | 18.13% | 19.46% | 0.8586% | 1.542%
73 | MT 20 | R | 6/2/2023 10:48:45 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 20.075 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.7333% | 5.179% | 6.041% | 2.169% | 2.188% | 2.742% | 5.267% | 4.111% | 5.744% | 2.825% | 0.9872% | 1.259% | 0.9007% | 8.321% | 1.712% | 0.835% | 0.9955% | 0.8756% | 0.7999% | 4.604% | 3.249% | 1.371% | 5.357% | 4.494% | 1.502% | 4.228% | 3.569% | 0.964% | 0.6723% | 0.7504% | 0.6466% | 1.382% | 1.375% | 0.476% | 1.106% | 1.98% | 3.274% | 0.2579% | 0.4491% | 2.25% | 1.336%
74 | MT 21 | R | 6/2/2023 10:54:57 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 21.076 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.471% | 4.876% | 4.389% | 0.5244% | 2.313% | 2.163% | 1.863% | 4.389% | 1.866% | 7.807% | 1.306% | 3.11% | 0.8421% | 2.199% | 2.135% | 0.7356% | 1.314% | 1.405% | 3.467% | 3.011% | 5.291% | 3.949% | 7.235% | 2.235% | 1.094% | 6.647% | 1.43% | 4.49% | 1.907% | 2.755% | 4.339% | 3.162% | 0.9335% | 1.814% | 9.275% | 0.7241% | 4.421% | 2.339% | 0.5921% | 6.624% | 8.423%
75 | MT 21 LD | R | 6/2/2023 11:00:59 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 21 LD.077 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 126.3% |  | 15.8% | 12.31% | 3.211% | 604.8% | 0% | 2851% | 112% | 36.66% | 12.59% | 17.37% | 590.5% | 175.6% | 159.3% | 151.9% | 53.38% | 6.8% | 1.097% | 192.3% | 29.4% | 35.08% | 1.187% | 173.2% | 147.6% | 290.8% | 96.2% |  | 5.481% | 108.3% | 4.263% | 1.299% | 3.683% | 1.791% | 2.635% | 5.031% | 1.434% | 86.52% | 123.7% | 40.28% | 8.664%
76 | LBLANK | R | 6/2/2023 11:07:00 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LBLANK.078 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 70.75% |  | 210% | 9.38% | 37.27% | 61.83% | 1226% | 210.7% | 78.07% | 39.95% | 8.074% | 33.75% | 26.22% | 230.8% | 30.14% | 30.74% | 27.55% | 27.46% | 37.46% | 0% | 54.12% | 98.91% | 65.64% | 86.61% | 429.1% | 156.3% | 56.93% |  | 18.6% | 43.36% | 0.7277% | 2.293% | 9.042% | 1.143% | 1.085% | 1.805% | 1.055% | 35.12% | 12.63% | 384% | 415.9%
77 | CCV 1ppb TRACE | R | 6/2/2023 11:13:03 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.079 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.096% | 1.334% | 17.63% | 4.023% | 1.316% | 4.134% | 6.79% | 0.6522% | 2.399% | 6.598% | 5.938% | 2.418% | 1.997% | 3.199% | 1.266% | 0.1207% | 5.992% | 9.088% | 6.891% | 11.07% | 4.437% | 2.357% | 15.35% | 8.505% | 4.781% | 2.735% | 3.878% | 4.825% | 5.006% | 1.496% | 3.006% | 0.9754% | 0.5034% | 6.526% | 2.304% | 0.2884% | 5.331% | 7.001% | 8.411% | 3.179% | 0.6862%
78 | CCV 1ppb TRACE | R | 6/2/2023 11:19:35 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.080 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.28% | 1.981% | 9.191% | 11.57% | 5.889% | 2.548% | 5.119% | 3.398% | 2.035% | 1.033% | 3.449% | 1.267% | 8.244% | 0.7813% | 1.487% | 5.462% | 1.347% | 0.6232% | 5.62% | 16.19% | 0.5722% | 1.174% | 6.397% | 23.51% | 7.636% | 3.654% | 8.418% | 5.536% | 0.306% | 2.771% | 2.323% | 1.079% | 1.707% | 0.8071% | 3.131% | 0.7261% | 0.7534% | 2.157% | 2.727% | 2.952% | 2.935%
79 | CCV 10 ppb MAJORS | R | 6/2/2023 11:25:38 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.081 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.5004% | 2.148% | 3.077% | 1.104% | 3.354% | 1.757% | 1.661% | 1.853% | 1.541% | 1.461% | 0.6921% | 2.159% | 0.7714% | 0.8112% | 0.2965% | 1.295% | 0.9045% | 3.404% | 0.9113% | 1.514% | 1.944% | 1.944% | 4.23% | 4.665% | 1.693% | 0.4509% | 2.574% | 1.663% | 1.803% | 1.907% | 0.8125% | 0.7045% | 0.6476% | 1.288% | 0.5617% | 1.271% | 3.113% | 0.755% | 0.7672% | 1.002% | 1.455%
80 | CCV 10 ppb MAJORS | R | 6/2/2023 11:32:09 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.082 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.5222% | 1.003% | 10.35% | 3.008% | 2.012% | 0.7594% | 2.643% | 1.566% | 2.106% | 1.345% | 0.4048% | 1.643% | 2.339% | 1.557% | 1.504% | 1.058% | 0.5872% | 0.3129% | 1.517% | 3.254% | 1.054% | 1.852% | 2.107% | 7.294% | 1.292% | 1.908% | 4.262% | 4.815% | 0.352% | 0.2996% | 0.2871% | 1.117% | 2.235% | 0.5438% | 1.721% | 1.967% | 0.5508% | 1.544% | 1.244% | 0.8924% | 0.3412%
81 | MT 22 | R | 6/2/2023 11:38:12 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 22.083 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.442% | 6.086% | 3.062% | 2.379% | 1.244% | 1.514% | 6.421% | 8.599% | 4.866% | 4.977% | 1.845% | 2.386% | 1.252% | 8.705% | 2.386% | 2.335% | 1.227% | 0.7747% | 1.361% | 3.31% | 3.272% | 1.824% | 13.42% | 6.779% | 1.847% | 0.5008% | 3.314% | 6.85% | 2.236% | 1.431% | 0.7876% | 1.307% | 1.936% | 0.6126% | 3.071% | 2.294% | 3.163% | 1.531% | 0.895% | 0.8883% | 3.224%
82 | MT 22 LSP | R | 6/2/2023 11:44:15 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 22 LSP.084 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.311% | 1.161% | 5.849% | 4.14% | 1.339% | 2.665% | 2.237% | 1.04% | 6.186% | 1.623% | 1.064% | 0.5053% | 0.8324% | 1.522% | 1.839% | 0.8891% | 1.745% | 2.449% | 0.5421% | 2.033% | 2.794% | 2.211% | 5.046% | 8.701% | 2.844% | 1.944% | 1.551% | 2.122% | 1.906% | 0.7511% | 0.6553% | 0.3218% | 0.9114% | 1.679% | 1.875% | 0.8518% | 1.923% | 2.304% | 2.332% | 0.3175% | 2.223%
83 | MT 22 LSP | R | 6/2/2023 11:50:45 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 22 LSP.085 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 10.96% | 11.99% | 11.66% | 1.35% | 13.36% | 9.591% | 12.16% | 11.82% | 12.86% | 1.351% | 2.66% | 0.8742% | 12.72% | 11.79% | 12.31% | 11.76% | 1.767% | 1.76% | 3.187% | 9.157% | 2.295% | 0.7862% | 3.893% | 2.583% | 2.529% | 1.054% | 4.788% | 2.799% | 0.3584% | 0.6887% | 1.837% | 0.5864% | 10.06% | 1.614% | 0.776% | 1.682% | 2.474% | 0.4868% | 0.9798% | 2.285% | 3.131%
84 | MT 23 | R | 6/2/2023 11:56:47 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 23.086 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.065% | 7.48% | 7.17% | 2.205% | 2.653% | 1.641% | 2.737% | 16.08% | 3.958% | 3.135% | 0.761% | 4.763% | 0.7502% | 5.787% | 2.899% | 1.691% | 0.8419% | 1.587% | 3.262% | 5.779% | 7.215% | 3.696% | 1.635% | 8.971% | 7.251% | 4.517% | 4.313% | 9.743% | 2.243% | 2.759% | 0.974% | 2.535% | 0.398% | 1.489% | 18.35% | 2.303% | 6.36% | 2.116% | 1.347% | 1.445% | 21.2%
85 | MT 24 | R | 6/3/2023 12:02:49 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 24.087 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 4.049% | 10.71% | 10.66% | 1.326% | 9.625% | 2.953% | 8.295% | 12.06% | 4.968% | 5.009% | 0.9555% | 1.856% | 2.518% | 2.789% | 9.513% | 3.015% | 0.5653% | 0.8025% | 1.735% | 19.42% | 2.248% | 1.657% | 15.51% | 11.42% | 1.864% | 5.117% | 4.763% | 11.54% | 0.7254% | 0.7754% | 1.66% | 1.933% | 1.351% | 0.139% | 0.1458% | 1.582% | 1.749% | 1.484% | 0.5748% | 2.467% | 0.3134%
86 | MT 25 | R | 6/3/2023 12:08:52 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 25.088 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.146% | 3.845% | 2.588% | 2.359% | 2.03% | 3.748% | 10.66% | 10.86% | 4.295% | 3.604% | 0.8139% | 2.516% | 2.698% | 4.606% | 6.7% | 1.122% | 1.889% | 1.084% | 4.405% | 7.863% | 3.129% | 1.676% | 11.95% | 10.89% | 7.647% | 3.319% | 3.618% | 9.711% | 1.533% | 2.893% | 5.11% | 1.511% | 2.406% | 1.124% | 1.667% | 0.9929% | 2.834% | 0.5898% | 0.745% | 6.023% | 1.551%
87 | MT 26 | R | 6/3/2023 12:14:55 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 26.089 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.6795% | 9.857% | 4.548% | 0.8579% | 4.355% | 1.484% | 17.57% | 13.92% | 10.71% | 0.8777% | 1.014% | 2.13% | 0.6885% | 2.717% | 6.106% | 1.081% | 1.307% | 1.084% | 24.19% | 8.89% | 2.876% | 2.389% | 48.11% | 13.01% | 7.598% | 3.946% | 32.86% | 14.02% | 1.685% | 1.929% | 21.85% | 1.883% | 1.35% | 1.088% | 4.918% | 2.001% | 1.104% | 2.538% | 1.031% | 28.24% | 8.194%
88 | MT 27 | R | 6/3/2023 12:21:03 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 27.090 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.019% | 1.076% | 4.455% | 2.607% | 0.5514% | 1.269% | 5.481% | 3.194% | 1.678% | 1.913% | 0.7254% | 1.515% | 1.345% | 3.768% | 4.139% | 1.617% | 0.9864% | 1.011% | 3.992% | 4.552% | 4.857% | 0.4476% | 9.502% | 4.41% | 5.241% | 2.694% | 8.351% | 6.537% | 0.6554% | 0.2399% | 2.21% | 0.2049% | 1.669% | 0.7935% | 1.777% | 0.1771% | 4.994% | 1.774% | 0.8392% | 3.795% | 0.2598%
89 | MT 28 | R | 6/3/2023 12:27:06 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 28.091 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 3.824% | 6.133% | 2.623% | 0.8543% | 6.381% | 4.594% | 10.91% | 9.661% | 5.525% | 1.111% | 1.082% | 15.38% | 5.881% | 5.326% | 6.392% | 5.056% | 1.694% | 1.143% | 5.545% | 11.31% | 15.79% | 15.08% | 2.72% | 23.71% | 13.47% | 16.33% | 6.942% | 13.7% | 2.107% | 15.17% | 4.661% | 13.76% | 4.393% | 1.624% | 1.204% | 1.882% | 14.63% | 0.5277% | 0.8633% | 7.606% | 1.303%
90 | MT 29 | R | 6/3/2023 12:33:08 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 29.092 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.094% | 2.983% | 2.959% | 0.6309% | 1.447% | 1.388% | 13.91% | 8.191% | 3.782% | 4.423% | 3.25% | 2.08% | 3.353% | 1.73% | 0.9326% | 2.717% | 2.868% | 3.202% | 2.569% | 3.823% | 6.422% | 1.976% | 4.255% | 3.582% | 2.345% | 0.5316% | 1.314% | 1.766% | 2.735% | 1.793% | 1.462% | 1.066% | 2.094% | 2.753% | 16.85% | 1.757% | 5.311% | 3.484% | 3.874% | 1.851% | 20.19%
91 | LBLANK | R | 6/3/2023 12:39:20 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LBLANK.093 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 63.79% |  | 412.3% | 5.276% | 14.02% | 39.41% | 150.2% | 68.79% | 102.9% | 74.65% | 13.16% | 18.59% | 6.54% | 294% | 56.22% | 19.31% | 39.46% | 97.81% | 54.12% | 121.2% | 29.83% | 45.04% | 25.59% |  | 36.85% | 340.6% | 239.8% | 173.2% | 8.07% | 63.88% | 0.7584% | 0.9775% | 1.084% | 2.596% | 3.789% | 0.669% | 0.5209% | 27.24% | 74.78% | 412.5% | 272.7%
92 | CCV 1ppb TRACE | R | 6/3/2023 12:45:23 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.094 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 5.545% | 3.414% | 33.54% | 3.939% | 6.98% | 3.627% | 1.134% | 4.318% | 1.004% | 3.803% | 4.835% | 8.978% | 3.102% | 3.026% | 4.518% | 1.882% | 2.658% | 4.012% | 4.821% | 9.493% | 10.8% | 9.389% | 10.03% | 2.937% | 12.35% | 9.735% | 14.68% | 13.96% | 5.807% | 8.365% | 1.268% | 8.323% | 1.44% | 4.163% | 4.235% | 1.407% | 10.07% | 4.547% | 2.588% | 2.194% | 5.229%
93 | CCV 1ppb TRACE | R | 6/3/2023 12:51:55 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.095 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.572% | 2.801% | 15% | 5.958% | 4% | 2.921% | 3.668% | 2.917% | 4.852% | 2.203% | 2.377% | 1.65% | 4.102% | 1.162% | 2.011% | 5.065% | 2.947% | 4.579% | 7.173% | 8.906% | 4.337% | 1.977% | 20.75% | 5.554% | 6.536% | 3.598% | 0.9943% | 8.293% | 3.13% | 1.806% | 2.273% | 0.9341% | 1.173% | 1.371% | 0.5272% | 1.209% | 3.074% | 1.876% | 2.028% | 0.6808% | 0.5736%
94 | LFB 10ppb | R | 6/3/2023 12:57:57 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LFB 10ppb.096 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 6.436% | 6.756% | 10% | 0.6184% | 8.406% | 5.182% | 5.562% | 6.565% | 6.596% | 2.646% | 0.706% | 16.25% | 9.487% | 6.173% | 5.825% | 6.386% | 1.22% | 0.7511% | 1.977% | 10.2% | 15.18% | 17.43% | 0.3837% | 14.61% | 15.35% | 16.7% | 6.605% | 14.31% | 1.043% | 17.34% | 2.397% | 15.87% | 5.813% | 0.4013% | 0.8652% | 1.072% | 13.21% | 1.089% | 0.4799% | 0.9494% | 0.6715%
95 | LFB 10ppb | R | 6/3/2023 1:04:27 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LFB 10ppb.097 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 3.204% | 4.897% | 5.568% | 1.612% | 5.039% | 2.58% | 4.782% | 4.941% | 4.362% | 1.094% | 2.123% | 0.7556% | 2.408% | 3.242% | 1.442% | 4.106% | 1.083% | 1.041% | 1.5% | 1.934% | 1.977% | 1.979% | 3.642% | 6.418% | 3.77% | 0.6598% | 1.483% | 6.419% | 1.073% | 1.512% | 1.132% | 0.8321% | 2.777% | 0.9989% | 0.9717% | 1.255% | 1.482% | 1.331% | 1.483% | 1.53% | 1.077%
96 | CCV 10 ppb MAJORS | R | 6/3/2023 1:10:29 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.098 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.607% | 3.226% | 3.696% | 3.1% | 3.57% | 5.076% | 5.18% | 3.606% | 2.152% | 2.706% | 5.459% | 0.2545% | 2.751% | 1.923% | 2.329% | 2.024% | 3.588% | 3.968% | 2.885% | 2.295% | 6.944% | 1.066% | 2.297% | 3.466% | 3.01% | 1.252% | 2.611% | 0.9662% | 1.438% | 0.8496% | 0.3343% | 0.3827% | 1.653% | 2.011% | 1.12% | 2.201% | 5.925% | 2.72% | 2.504% | 1.391% | 0.5987%
97 | CCV 10 ppb MAJORS | R | 6/3/2023 1:17:01 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.099 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.09% | 3.478% | 5.267% | 0.5864% | 1.816% | 2.273% | 1.805% | 3.306% | 2.421% | 2.26% | 1.322% | 1.179% | 2.421% | 0.4378% | 0.39% | 1.041% | 2.022% | 1.789% | 8.917% | 0.9251% | 3.255% | 1.483% | 9.409% | 3.712% | 4.161% | 2.944% | 7.791% | 5.038% | 2.599% | 0.4088% | 6.855% | 0.9804% | 0.7228% | 1.684% | 0.4096% | 0.9582% | 4.193% | 1.87% | 3.021% | 7.995% | 1.179%
98 | MT 30 | R | 6/3/2023 1:23:04 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 30.100 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 14.88% | 7.613% | 3.095% | 4.249% | 1.94% | 5.551% | 7.269% | 1.828% | 2.215% | 2.045% | 7.808% | 3.373% | 2.747% | 5.819% | 3.98% | 2.346% | 5.244% | 4.763% | 0.9252% | 8.062% | 3.934% | 4.183% | 19.75% | 16.24% | 2.953% | 12.4% | 7.579% | 22.94% | 6.386% | 5.402% | 0.4139% | 1.587% | 1.933% | 4.209% | 0.3621% | 5.189% | 3.401% | 6.175% | 4.42% | 0.9674% | 0.5171%
99 | MT 31 | R | 6/3/2023 1:29:15 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 31.101 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 6.957% | 3.888% | 3.387% | 2.847% | 10.13% | 5.087% | 7.86% | 3.256% | 9.753% | 5.672% | 5.803% | 3.127% | 1.491% | 1.274% | 3.456% | 2.506% | 6.197% | 5.599% | 2.901% | 7.224% | 5.772% | 2.276% | 42.08% | 7.933% | 7.322% | 10.79% | 22.36% | 19.72% | 3.461% | 6.456% | 1.469% | 2.868% | 1.476% | 5.268% | 1.575% | 1.621% | 5.433% | 5.357% | 4.595% | 1.477% | 1.366%
100 | MT 31 LD | R | 6/3/2023 1:35:29 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 31 LD.102 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 11.82% | 6.76% | 4.379% | 1.819% | 16.03% | 1.189% | 4.952% | 1.645% | 11.7% | 0.5161% | 1.177% | 0.2318% | 5.047% | 5.606% | 6.566% | 4.972% | 2.199% | 1.945% | 5.859% | 13.26% | 2.638% | 0.513% | 8.626% | 0.4173% | 6.984% | 8.371% | 16.44% | 20.66% | 1.401% | 3.843% | 1.459% | 0.4163% | 3.757% | 1.122% | 0.5867% | 0.6212% | 1.875% | 2.098% | 1.48% | 1.577% | 0.3001%
101 | MT 31 LD | R | 6/3/2023 1:42:00 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 31 LD.103 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 9.566% | 4.264% | 5.921% | 0.08368% | 7.445% | 3.393% | 12.79% | 7.769% | 5.789% | 1.205% | 1.509% | 1.088% | 4.331% | 7.073% | 12.24% | 5.135% | 1.651% | 0.9086% | 4.836% | 15.23% | 4.574% | 4.361% | 24.16% | 15.74% | 2.041% | 7.682% | 31.57% | 27.76% | 1.982% | 0.1365% | 1.708% | 0.6476% | 4.019% | 0.7374% | 0.9438% | 0.5114% | 3.925% | 1.826% | 1.262% | 1.641% | 3.396%
102 | MT 32 | R | 6/3/2023 1:48:02 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 32.104 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 62.23% | 86.62% | 20.73% | 2.673% | 5.197% | 60.27% | 40.95% | 264.8% | 23.28% | 10.65% | 3.317% | 2.45% | 8.47% | 320.9% | 16.6% | 4.128% | 2.134% | 21.67% | 62.07% | 0% | 143.8% | 690.4% | 70.5% | 86.61% | 2009% | 451.7% | 146.5% | 173.2% | 12.95% | 7.24% | 0.8806% | 1.86% | 1.304% | 3.723% | 7.835% | 2.144% | 1.085% | 173.2% | 33.96% | 2.338% | 12.33%
103 | MT 32 LSP | R | 6/3/2023 1:54:03 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 32 LSP.105 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.91% | 1.294% | 9.194% | 3.626% | 1.863% | 1.622% | 1.25% | 2.221% | 3.07% | 1.455% | 3.675% | 1.854% | 0.8854% | 1.585% | 2.215% | 0.4507% | 1.748% | 1.033% | 3.079% | 1.015% | 1.842% | 0.6074% | 0.5537% | 7.562% | 3.083% | 3.098% | 5.443% | 4.329% | 1.027% | 1.724% | 1.062% | 1.271% | 0.8434% | 0.7635% | 2.126% | 4.725% | 2.145% | 0.5836% | 0.7866% | 1.876% | 3.978%
104 | MT 32 LSP | R | 6/3/2023 2:00:34 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 32 LSP.106 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 3.167% | 0.8122% | 5.591% | 2.136% | 2.752% | 1.714% | 2.152% | 0.3693% | 2.416% | 7.758% | 7.988% | 1.055% | 2.104% | 3.158% | 1.164% | 2.661% | 8.953% | 7.458% | 14% | 3.555% | 4.404% | 2.743% | 14.88% | 5.577% | 3.216% | 4.293% | 17.23% | 2.205% | 8.426% | 1.996% | 10.81% | 1.89% | 1.409% | 6.059% | 1.08% | 1.811% | 4.381% | 6.391% | 8.134% | 13.03% | 1.272%
105 | MT 33 | R | 6/3/2023 2:06:37 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 33.107 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 86.43% | 86.61% | 6.814% | 1.712% | 14.62% | 47.84% | 50.04% | 235.9% | 40.2% | 24.1% | 18.51% | 9.004% | 14.62% | 57.62% | 34.66% | 5.178% | 20.73% | 6.075% | 26.96% | 0% | 114.9% | 346% | 11.76% |  | 98.81% | 1593% | 1482% | 173.2% | 17.04% | 17.66% | 0.4069% | 0.386% | 1.525% | 18.74% | 1.845% | 0.9518% | 0.1119% | 101.9% | 16.53% | 6.376% | 3.66%
106 | MT 34 | R | 6/3/2023 2:12:39 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 34.108 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 36.46% |  | 10.7% | 1.122% | 4.603% | 65.77% | 80.65% | 234.2% | 12.91% | 17.56% | 0.8279% | 5.276% | 4.482% | 129% | 20.3% | 29.77% | 8.68% | 65.16% | 49.05% | 112.5% | 160.8% | 275.1% | 89.02% | 173.2% | 79.64% | 574.1% | 47.7% | 173.2% | 13.93% | 44.32% | 12.01% | 0.2628% | 2.413% | 0.8015% | 2.602% | 1.5% | 2.47% | 115.1% | 6.569% | 17.31% | 5.315%
107 | MT 35 | R | 6/3/2023 2:18:42 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 35.109 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 47% | 173.2% | 14.59% | 2.244% | 11.72% | 84.81% | 47% | 642.3% | 32.72% | 28.99% | 8.369% | 25.5% | 1.778% | 64.56% | 49.04% | 10.58% | 8.315% | 21.21% | 43.34% | 112.2% | 48.72% | 162% | 29.02% | 98.63% | 363.7% | 210% | 98.21% |  | 26.77% | 28.55% | 7.479% | 4.168% | 1.385% | 7.165% | 2.089% | 2.639% | 4.96% | 22.01% | 52.31% | 10.52% | 6.94%
108 | LBLANK | R | 6/3/2023 2:24:45 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LBLANK.110 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 156.7% |  | 76.24% | 5.998% | 2.689% | 16.8% | 150% | 230.1% | 33.3% | 55.64% | 11.86% | 78.28% | 6.931% | 2% | 41.38% | 131.3% | 17.91% | 21.03% | 18.26% | 0% | 32.57% | 42.2% | 29.73% |  | 251% | 67.08% | 184.5% |  | 27.97% | 27.59% | 0.5983% | 1.21% | 0.8168% | 0.6004% | 7.613% | 1.681% | 4.429% | 173.2% | 48.24% | 1523% | 416.1%
109 | CCV 1ppb TRACE | R | 6/3/2023 2:30:48 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.111 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 3.216% | 6.017% | 23.02% | 3.446% | 4.832% | 3.403% | 5.509% | 3.359% | 4.775% | 13.35% | 11.35% | 4.473% | 6.325% | 2.524% | 3.271% | 3.266% | 12.54% | 13.04% | 4.153% | 21.76% | 0.6057% | 2.771% | 10.28% | 0.5598% | 6.772% | 5.107% | 3.651% | 7.311% | 10.81% | 1.429% | 0.601% | 2.968% | 2.31% | 8.74% | 5.73% | 1.739% | 0.9405% | 11.34% | 11.9% | 0.3262% | 6.63%
110 | CCV 1ppb TRACE | R | 6/3/2023 2:37:20 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.112 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.9594% | 5.198% | 17.21% | 38.95% | 13.34% | 8.133% | 5.01% | 9.575% | 3.684% | 3.553% | 4.676% | 2.877% | 1.354% | 2.061% | 2.484% | 5.278% | 1.129% | 1.514% | 10.21% | 7.645% | 1.201% | 1.301% | 8.33% | 6.599% | 1.489% | 3.629% | 7.244% | 2.487% | 3.345% | 2.04% | 2.934% | 1.181% | 3.025% | 1.702% | 3.526% | 23.81% | 1.665% | 3.318% | 1.461% | 4.862% | 2.717%
111 | CCV 10 ppb MAJORS | R | 6/3/2023 2:43:23 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.113 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 3.5% | 1.554% | 12.18% | 0.31% | 2.576% | 1.767% | 4.087% | 4.164% | 3.817% | 0.8657% | 0.9997% | 3.067% | 1.351% | 4.097% | 2.309% | 3.633% | 0.7128% | 0.9865% | 2.384% | 5.945% | 6.482% | 4.226% | 4.972% | 4.072% | 1.917% | 0.7355% | 2.101% | 1.742% | 0.2948% | 3.176% | 1.635% | 2.953% | 3.071% | 0.2421% | 1.407% | 0.7519% | 5.327% | 0.8993% | 0.4301% | 0.8666% | 0.7729%
112 | CCV 10 ppb MAJORS | R | 6/3/2023 2:49:56 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.114 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 4.16% | 5.721% | 3.611% | 1.775% | 3.52% | 2.921% | 5.582% | 5.237% | 4.433% | 2.651% | 3.465% | 2.717% | 5.186% | 3.912% | 3.247% | 3.921% | 3.287% | 3.438% | 10.32% | 1.574% | 1.811% | 1.554% | 10.49% | 2.435% | 2.492% | 3.135% | 5.776% | 7.996% | 2.781% | 1.281% | 8.621% | 1.044% | 3.708% | 2.385% | 0.6965% | 1.192% | 1.563% | 4.618% | 2.197% | 10.49% | 0.8465%
113 | MT 36 | R | 6/3/2023 2:55:59 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 36.115 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 52.32% | 102% | 14.78% | 2.289% | 4.858% | 49.32% | 83.92% | 66.35% | 15.15% | 4.034% | 2.927% | 9.308% | 3.401% | 110.2% | 38.65% | 28.54% | 4.835% | 17.23% | 19.91% | 101.8% | 14.57% | 21.03% | 69.4% |  | 1031% | 65.86% | 157.3% |  | 22.18% | 25.74% | 0.8226% | 0.5796% | 3.575% | 0.7782% | 1.6% | 1.374% | 2.5% | 100.5% | 18.6% | 7.857% | 5.897%
114 | MT 37 | R | 6/3/2023 3:02:01 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 37.116 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 20.57% | 35.45% | 16.76% | 2.652% | 16.95% | 36.68% | 17.62% | 242.2% | 30.44% | 47.02% | 4.827% | 17.36% | 14.96% | 325.5% | 17.02% | 9.572% | 4.35% | 2.723% | 69.67% | 131.2% | 192% | 519% | 93.3% | 173.2% | 147.4% | 106.7% | 121% |  | 20.94% | 22.03% | 4.949% | 1.962% | 5.157% | 1.024% | 0.7732% | 2.17% | 2.238% | 33.93% | 30.13% | 10.56% | 8.876%
115 | MT 38 | R | 6/3/2023 3:08:04 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 38.117 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 28.04% | 86.61% | 18.61% | 8.5% | 2.789% | 7.973% | 111.7% | 78.76% | 22.11% | 20.46% | 4.076% | 28.6% | 7.258% | 177.6% | 52.97% | 12.55% | 7.85% | 19.91% | 9.406% | 113.8% | 70.49% | 122.1% | 33.96% |  | 95.92% | 106.4% | 139.4% |  | 23.76% | 37.44% | 0.8482% | 1.332% | 2.295% | 0.8309% | 5.724% | 4.168% | 4.579% | 173.2% | 25.5% | 7.01% | 8.048%
116 | MT 39 | R | 6/3/2023 3:14:07 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 39.118 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 6.9% | 86.6% | 7.035% | 4.033% | 10.3% | 30.26% | 46.95% | 150.4% | 42.94% | 8.794% | 3.034% | 15.19% | 5.82% | 248.9% | 54.36% | 10.75% | 6.143% | 5.972% | 67.57% | 107.6% | 17.63% | 30.71% | 134.1% | 173.2% | 159.2% | 157.7% | 815.8% | 98.86% | 27.63% | 4.316% | 19.11% | 3.826% | 2.82% | 0.4397% | 0.3909% | 1.465% | 3.923% | 86.6% | 17.67% | 35.02% | 6.925%
117 | MT 40 | R | 6/3/2023 3:20:10 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 40.119 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 17.29% | 49.24% | 9.24% | 2.797% | 8.295% | 31.6% | 16.29% | 1383% | 51.6% | 36.36% | 7.544% | 3.495% | 6.206% | 516.7% | 14.69% | 9.298% | 4.703% | 41.85% | 217% | 279.5% | 35.96% | 35.82% | 888.5% |  | 365.3% | 239.2% | 329.7% | 173.2% | 5.075% | 13.32% | 10.63% | 1.623% | 2.306% | 1.998% | 0.6156% | 1.489% | 0.5008% | 173.2% | 13.73% | 7.851% | 4.314%
118 | MT 31 | R | 6/3/2023 3:26:12 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 31 .120 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 6.413% | 11.91% | 2.136% | 1.99% | 4.238% | 7.528% | 3.933% | 4.569% | 4.9% | 17.72% | 15.88% | 1.563% | 1.021% | 4.965% | 3.727% | 1.854% | 15.15% | 14.13% | 1.318% | 1.587% | 1.046% | 2.527% | 46.57% | 28.18% | 11.8% | 16.85% | 15.49% | 22.96% | 18.96% | 5.203% | 2.015% | 1.575% | 0.7747% | 12.16% | 0.8676% | 1.007% | 0.9116% | 15.22% | 14.59% | 2.64% | 0.918%
119 | MT 31 LD | R | 6/3/2023 3:32:23 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 31 LD.121 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 6.968% | 27.2% | 11.93% | 1.656% | 15.15% | 14.2% | 15.4% | 9.171% | 12.3% | 5.46% | 8.417% | 3.006% | 11.52% | 9.305% | 7.538% | 10.53% | 8.141% | 7.735% | 2.339% | 14.83% | 4.439% | 4.579% | 29.5% | 39.73% | 17.21% | 14.2% | 37.2% | 12.41% | 6.759% | 3.885% | 1.263% | 2.565% | 8.462% | 5.954% | 6.666% | 0.5211% | 2.805% | 7.88% | 8.098% | 1.253% | 10.1%
120 | MT 31 LD | R | 6/3/2023 3:38:54 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 31 LD.122 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 8.957% | 7.125% | 3.886% | 1.411% | 9.71% | 4.583% | 10.44% | 9.784% | 0.6969% | 14.17% | 9.568% | 2.72% | 1.883% | 0.9481% | 8.259% | 2.519% | 10.09% | 10.42% | 3.481% | 16.32% | 3.912% | 2.494% | 81.32% | 32.1% | 9.468% | 10.68% | 6.692% | 21.64% | 8.469% | 2.478% | 0.8365% | 0.5805% | 1.543% | 7.989% | 6.517% | 0.6336% | 2.476% | 11.86% | 10.63% | 3.035% | 8.727%
121 | MT 32 | R | 6/3/2023 3:44:56 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 32.123 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 14.79% |  | 9.156% | 3.894% | 9.252% | 107.3% | 55.57% | 295.8% | 31.38% | 1.621% | 4.375% | 16.99% | 0.6909% | 462.2% | 43.64% | 14.99% | 2.843% | 14.16% | 41.16% | 0% | 269.1% | 945.4% | 13.8% | 173.2% | 551.9% | 315.6% | 48.25% |  | 16.2% | 2.221% | 1.818% | 0.7765% | 1.688% | 1.42% | 0.739% | 0.6329% | 3.134% | 99.22% | 73.68% | 5.685% | 4.339%
122 | MT 32 LSP | R | 6/3/2023 3:50:58 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 32 LSP.124 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 3.158% | 1.13% | 6.28% | 14.91% | 2.028% | 2.261% | 1.652% | 2.567% | 2.021% | 1.788% | 0.2695% | 1.337% | 1.584% | 2.782% | 1.579% | 2.012% | 0.435% | 1.11% | 1.375% | 7.536% | 3.723% | 2.195% | 1.331% | 9.749% | 2.841% | 2.95% | 5.752% | 9.779% | 0.252% | 1.708% | 0.2819% | 2.234% | 1.325% | 1.038% | 7.669% | 12.56% | 4.573% | 0.6667% | 1.669% | 0.4209% | 7.914%
123 | MT 32 LSP | R | 6/3/2023 3:57:29 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 32 LSP.125 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.15% | 1.385% | 5.386% | 3.198% | 1.049% | 2.236% | 1.515% | 1.399% | 1.914% | 2.12% | 2.438% | 1.427% | 0.6846% | 1.278% | 1.564% | 1.988% | 1.552% | 1.962% | 2.745% | 6.991% | 4.151% | 1.782% | 6.221% | 8.178% | 5.01% | 3.162% | 2.139% | 4.269% | 0.6904% | 1.926% | 0.3078% | 1.06% | 0.3866% | 1.909% | 7.714% | 1.227% | 5.496% | 2.246% | 1.994% | 0.8453% | 10.81%
124 | LBLANK | R | 6/3/2023 4:03:32 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LBLANK.126 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 72.39% | 173.2% | 256.7% | 26.5% | 2.96% | 430.2% | 155% | 1352% | 30.88% | 15.9% | 9.773% | 29.9% | 5.448% | 63.59% | 29.84% | 44.2% | 15.12% | 20.02% | 24.95% | 511.4% | 952.8% | 120% | 57.69% |  | 886.2% | 157.8% | 1262% |  | 23.41% | 33.33% | 5.275% | 1.506% | 1.192% | 0.8449% | 0.8595% | 0.4695% | 4.999% | 86.6% | 60.65% | 501.9% | 1940%
125 | CCV 1ppb TRACE | R | 6/3/2023 4:09:35 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.127 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.7674% | 3.401% | 13.27% | 3.575% | 2.767% | 1.165% | 6.761% | 6.319% | 3.296% | 2.499% | 0.6508% | 2.445% | 2.04% | 5.896% | 0.8745% | 0.7132% | 2.575% | 2.494% | 5.17% | 2.074% | 1.755% | 0.4967% | 13.07% | 10.82% | 6.506% | 4.378% | 1.517% | 9.155% | 5.253% | 0.6136% | 1.071% | 1.119% | 2.577% | 0.4788% | 1.127% | 0.8431% | 1.445% | 2.704% | 1.173% | 1.688% | 0.8471%
126 | CCV 1ppb TRACE | R | 6/3/2023 4:16:07 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.128 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 10.71% | 10.46% | 9.407% | 1.845% | 20.75% | 10.45% | 17.45% | 13.99% | 12.71% | 3.524% | 8.221% | 1.42% | 15.49% | 13.5% | 8.928% | 11.92% | 4.827% | 3.731% | 1.64% | 26% | 3.419% | 4.731% | 24.93% | 6.999% | 0.4658% | 11.14% | 4.821% | 4.854% | 3.306% | 2.093% | 5.431% | 0.7013% | 10.07% | 3.118% | 1.005% | 1.87% | 1.716% | 3.601% | 4.25% | 5.669% | 1.472%
127 | LFB 10ppb | R | 6/3/2023 4:22:09 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LFB 10ppb.129 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.521% | 2.697% | 5.372% | 1.622% | 0.72% | 3.013% | 1.268% | 1.101% | 2.181% | 0.8263% | 0.6636% | 1.434% | 1.43% | 2.019% | 2.896% | 0.5609% | 0.9435% | 1.923% | 1.518% | 5.713% | 1.902% | 1.629% | 2.457% | 5.348% | 3.214% | 0.7628% | 2% | 1.243% | 1.876% | 2.14% | 1.054% | 0.6958% | 0.4979% | 0.5974% | 7.309% | 0.6011% | 1.086% | 1.546% | 1.357% | 0.4231% | 5.574%
128 | LFB 10ppb | R | 6/3/2023 4:28:39 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LFB 10ppb.130 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.393% | 1.11% | 5.533% | 0.9092% | 1.301% | 0.7382% | 5.048% | 1.053% | 1.281% | 3.475% | 5.281% | 2.944% | 0.4687% | 1.09% | 2.487% | 1.393% | 3.899% | 3.645% | 4.12% | 4.299% | 2.966% | 3.155% | 6.071% | 6.114% | 2.809% | 4.774% | 3.101% | 5.201% | 3.309% | 4.519% | 1.593% | 2.194% | 0.86% | 4.347% | 1.464% | 0.8585% | 2.151% | 3.61% | 3.13% | 1.837% | 1.404%
129 | CCV 10 ppb MAJORS | R | 6/3/2023 4:34:42 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.131 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 16.94% | 17.73% | 11.47% | 0.6695% | 19.56% | 15.31% | 17.94% | 16.06% | 16.74% | 3.114% | 1.311% | 14.14% | 18% | 17.84% | 17.47% | 15.58% | 2.197% | 2.127% | 3.628% | 14.91% | 16.23% | 13.96% | 2.741% | 10.54% | 12.11% | 11.73% | 4.708% | 11.87% | 1.188% | 13.27% | 1.699% | 12% | 15.21% | 0.6787% | 1.577% | 0.1978% | 14.5% | 1.788% | 1.813% | 0.8235% | 0.2823%
130 | CCV 10 ppb MAJORS | R | 6/3/2023 4:41:13 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.132 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.874% | 3.719% | 6.286% | 1.519% | 3.443% | 2.19% | 4.437% | 3.547% | 2.756% | 2.549% | 4.719% | 1.89% | 2.318% | 1.387% | 2.795% | 1.823% | 4.019% | 3.873% | 1.24% | 5.409% | 2.656% | 1.698% | 2.49% | 1.587% | 3.464% | 1.606% | 4.376% | 0.08693% | 2.914% | 1.427% | 0.9461% | 0.08694% | 1.411% | 2.23% | 13.62% | 1.327% | 1.117% | 0.8086% | 2.53% | 0.7921% | 17.07%
131 | MT 33 | R | 6/3/2023 4:47:16 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 33.133 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 37.78% | 173.2% | 11.6% | 1.402% | 4.432% | 73.45% | 55.89% | 227.9% | 52.52% | 33.94% | 10.9% | 6.893% | 10.47% | 59.42% | 69.48% | 7.969% | 14.36% | 69.94% | 37.16% | 541.3% | 54.62% | 81.47% | 45.63% |  | 78.39% | 459.6% | 139.5% |  | 3.841% | 4.917% | 1.426% | 0.847% | 1.316% | 4.714% | 0.5451% | 1.164% | 0.8964% | 4.698% | 77.37% | 5.688% | 0.9552%
132 | MT 34 | R | 6/3/2023 4:53:19 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 34.134 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 24.63% | 86.74% | 26.2% | 2.02% | 10.47% | 125% | 70.15% | 79.6% | 35.18% | 7.187% | 5.132% | 32.2% | 7.154% | 175.3% | 20.25% | 18.25% | 15.97% | 23.37% | 56.27% | 112.3% | 27.1% | 42.26% | 35.33% | 173.2% | 1014% | 469.3% | 168.1% |  | 13.59% | 26.32% | 4.029% | 0.6439% | 4.583% | 4.222% | 0.4155% | 0.8662% | 2.321% | 173.2% | 77.41% | 3.971% | 2.985%
133 | MT 35 | R | 6/3/2023 4:59:21 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 35.135 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 75.89% |  | 20.18% | 0.9326% | 8.976% | 47.76% | 42.44% | 274% | 13.08% | 32.09% | 3.509% | 8.942% | 7.371% | 272.1% | 16.55% | 21.87% | 14.51% | 20.96% | 73.37% | 573.6% | 88.16% | 153.4% | 59.05% |  | 1485% | 53.66% | 126.1% | 173.2% | 4.215% | 55.94% | 1.775% | 0.8806% | 1.538% | 4.688% | 1.014% | 1.391% | 2.195% | 97.69% | 43.97% | 1.377% | 8.154%
134 | MT 36 | R | 6/3/2023 5:05:25 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 36.136 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.076% | 173.2% | 1.951% | 2.589% | 8.546% | 122.7% | 101.9% | 267.2% | 26.01% | 24.66% | 3.917% | 21.14% | 54.79% | 122% | 50.41% | 16.16% | 9.609% | 29.05% | 130.9% | 334.7% | 96.24% | 257% | 122.1% | 173.2% | 361.7% | 359.2% | 333.3% | 86.62% | 10.56% | 47.48% | 10.71% | 1.222% | 1.98% | 5.821% | 0.4248% | 1.06% | 1.019% | 87.04% | 72.48% | 24.73% | 6.631%
135 | MT 37 | R | 6/3/2023 5:11:28 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 37.137 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 66.76% |  | 8.174% | 1.529% | 10.91% | 51.1% | 25.97% | 1027% | 35.65% | 23.66% | 17.26% | 19.38% | 4.226% | 53.74% | 19.44% | 7.171% | 18.97% | 26.73% | 4.048% | 115.9% | 31.16% | 53% | 49.4% | 99.93% | 174.1% | 1563% | 243.3% |  | 49.66% | 71.45% | 4.519% | 2.484% | 2.159% | 18.06% | 1.801% | 1.231% | 3.304% | 173.2% | 66.69% | 6.512% | 2.791%
136 | MT 38 | R | 6/3/2023 5:17:31 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 38.138 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 24.41% | 173.2% | 15.43% | 1.18% | 8.638% | 81.9% | 138.8% | 1180% | 2.077% | 7.838% | 7.493% | 18.52% | 8.942% | 39.53% | 29.6% | 16.12% | 8.012% | 26.15% | 20.19% | 320% | 34.13% | 41.78% | 74.27% |  | 0.8981% | 521.9% | 69.45% | 173.2% | 55.53% | 35.84% | 1.347% | 0.2564% | 0.7302% | 0.8453% | 1.715% | 0.7716% | 1.678% | 43.43% | 36.48% | 6.219% | 6.062%
137 | MT 39 | R | 6/3/2023 5:23:33 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 39.139 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 8.296% | 86.6% | 15.24% | 1.557% | 9.466% | 143.7% | 74.08% | 450.3% | 20.98% | 7.658% | 3.896% | 16.31% | 5.033% | 77.21% | 41.59% | 17.3% | 15.19% | 6.936% | 33.26% | 121.1% | 115.2% | 378.5% | 98.29% | 173.2% | 294.2% | 43.63% | 90.23% |  | 36.65% | 12.82% | 1.948% | 9.875% | 0.3066% | 2.624% | 0.9186% | 2.377% | 10.02% | 41.18% | 40.85% | 3.96% | 5.131%
138 | MT 40 | R | 6/3/2023 5:29:36 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 40.140 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 8.11% | 10.35% | 12.08% | 1.096% | 16.19% | 44% | 34.24% | 846.7% | 17.73% | 46.12% | 1.8% | 7.234% | 14.95% | 404.1% | 31.16% | 9.551% | 11.88% | 21.62% | 22.87% | 610.9% | 21.74% | 18.95% | 44.6% | 173.2% | 0% | 369.1% | 498.8% | 86.61% | 1.524% | 15.23% | 1.063% | 1.543% | 6.92% | 1.427% | 1.69% | 0.3667% | 0.7927% | 173.2% | 58.61% | 3.743% | 8.499%
139 | MT 32 B | R | 6/3/2023 5:35:38 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 32 B.141 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 49.92% | 173.2% | 22.31% | 1.286% | 7.746% | 81.25% | 1008% | 458.6% | 13.38% | 9.056% | 3.156% | 15.67% | 11.43% | 393.9% | 39.35% | 13.87% | 4.339% | 70.16% | 41.92% | 111.8% | 11.21% | 43.11% | 12.4% |  | 0% | 109.6% | 146.9% |  | 16.37% | 19.76% | 0.6845% | 7.508% | 3.367% | 6.031% | 0.6136% | 0.6488% | 4.381% | 99.36% | 38.61% | 2.064% | 1.999%
140 | MT 32 B LD | R | 6/3/2023 5:41:40 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 32 B LD.142 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 103.3% | 173.2% | 10.79% | 0.9923% | 0.8091% | 51.31% | 58.05% | 281.9% | 13.36% | 10.1% | 4.341% | 29.03% | 4.81% | 296.6% | 30.91% | 13.46% | 6.335% | 5.587% | 67.35% | 125.1% | 83.41% | 132% | 75.11% |  | 166.8% | 722.3% | 592.2% | 173.2% | 28.52% | 14.09% | 0.3621% | 1.906% | 1.226% | 1.253% | 0.9606% | 0.4011% | 3.324% | 24.01% | 33.68% | 8.277% | 1.264%
141 | LBLANK | R | 6/3/2023 5:47:43 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LBLANK.143 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 54.81% |  | 108.3% | 4.157% | 9.461% | 39.65% | 814.5% | 213.9% | 26.9% | 31.69% | 23.83% | 17.73% | 5.064% | 87.68% | 46.3% | 166.8% | 34.99% | 52.33% | 122.2% | 0% | 36.23% | 45.77% | 710% | 86.61% | 6804% | 71.64% | 51.62% | 173.2% | 44.29% | 39.98% | 6.251% | 1.114% | 1.838% | 7.988% | 2.69% | 1.065% | 2.821% | 98.49% | 69.94% | 59.57% | 220.2%
142 | CCV 1ppb TRACE | R | 6/3/2023 5:53:46 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.144 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 3.325% | 5.941% | 23.03% | 2.878% | 17.85% | 1.481% | 5.373% | 4.903% | 2.569% | 3.758% | 6.534% | 0.8238% | 0.8095% | 2.842% | 3.501% | 4.11% | 6.408% | 2.888% | 3.645% | 23.62% | 4.941% | 3.199% | 12.48% | 11.72% | 7.804% | 7.76% | 8.439% | 11.83% | 6.141% | 1.283% | 1.078% | 0.3289% | 2.55% | 3.953% | 4.668% | 0.1957% | 2.008% | 2.57% | 5.654% | 1.993% | 4.93%
143 | CCV 1ppb TRACE | R | 6/3/2023 6:00:18 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.145 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 5.425% | 3.9% | 4.261% | 4.009% | 7.528% | 3.746% | 8.462% | 5.316% | 2.463% | 1.034% | 3.743% | 5.763% | 3.93% | 1.84% | 4.915% | 2.521% | 0.7733% | 2.373% | 7.903% | 7.746% | 5.395% | 4.97% | 25.87% | 7.519% | 2.563% | 2.869% | 4.588% | 6.776% | 0.845% | 3.102% | 8.766% | 3.463% | 0.7633% | 0.6012% | 1.045% | 1.7% | 2.406% | 0.8383% | 0.8331% | 8.699% | 0.5934%
144 | CCV 10 ppb MAJORS | R | 6/3/2023 6:06:21 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.146 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.742% | 3.461% | 3.746% | 0.4774% | 2.892% | 1.851% | 2.546% | 0.6528% | 1.112% | 1.267% | 1.02% | 1.464% | 1.029% | 2.983% | 2.912% | 1.602% | 0.5304% | 1.42% | 2.364% | 3.026% | 3.873% | 1.959% | 5.166% | 6.391% | 1.945% | 1.728% | 2.456% | 4.276% | 0.8201% | 2.33% | 2.134% | 0.7508% | 1.599% | 0.7586% | 1.771% | 1.899% | 2.694% | 1.577% | 0.6331% | 4.736% | 0.3781%
145 | CCV 10 ppb MAJORS | R | 6/3/2023 6:12:53 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.147 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 3.178% | 2.116% | 5.982% | 4.136% | 4.908% | 0.5362% | 0.9492% | 2.338% | 1.839% | 2.444% | 1.849% | 2.812% | 1.25% | 2.309% | 2.159% | 1.688% | 0.8841% | 0.2825% | 2.187% | 1.008% | 1.571% | 1.954% | 2.721% | 7.186% | 1.097% | 1.798% | 0.3126% | 3.386% | 0.8524% | 1.452% | 0.2414% | 1.628% | 1.252% | 0.6619% | 2.003% | 1.375% | 2.372% | 1.003% | 1.227% | 1.18% | 0.5548%
146 | MT 33 B | R | 6/3/2023 6:18:55 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 33 B.148 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 47.97% | 98.8% | 22.19% | 2.181% | 10.59% | 164.7% | 155% | 184.1% | 14.95% | 27.67% | 17.38% | 28.16% | 17.32% | 4.906% | 44.09% | 12.08% | 4.804% | 217.9% | 96.07% | 381.1% | 41.35% | 52.74% | 64.9% | 173.2% | 563.2% | 0% | 132.9% |  | 31.67% | 14.72% | 2.552% | 1.458% | 3.217% | 0.1093% | 1.454% | 2.021% | 3.766% | 89.18% | 16.83% | 12.67% | 6.253%
147 | MT 33 B LSP | R | 6/3/2023 6:24:56 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 33 B LSP.149 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.299% | 3.165% | 5.094% | 0.6773% | 3.606% | 2.629% | 2.155% | 3.987% | 2.861% | 17.68% | 17.63% | 1.785% | 2.012% | 0.7066% | 2.496% | 2.187% | 16.57% | 16.71% | 3.439% | 2.6% | 2.199% | 1.16% | 3.672% | 3.211% | 0.4541% | 1.901% | 0.7583% | 2.751% | 16.52% | 1.852% | 1.071% | 0.9567% | 2.891% | 13.69% | 1.78% | 0.411% | 2.661% | 16.45% | 16.34% | 3.08% | 0.9971%
148 | MT 34 B | R | 6/3/2023 6:30:59 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 34 B.150 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 55.62% | 1.865% | 15.21% | 1.791% | 7.899% | 23.73% | 57.27% | 31% | 15.04% | 15.21% | 12.45% | 33.84% | 13.07% | 118.5% | 12.25% | 24.3% | 2.781% | 21.79% | 35.11% | 139.2% | 141.6% | 116.5% | 24.32% | 173.2% | 241.4% | 114.8% | 117.7% |  | 6.426% | 16.38% | 1.729% | 1.099% | 1.847% | 1.225% | 20.37% | 1.395% | 2.369% | 1.221% | 13.6% | 5.579% | 19.7%
149 | MT 35 B | R | 6/3/2023 6:37:01 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 35 B.151 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 6.466% | 173.2% | 21.94% | 0.4262% | 15.31% | 76% | 63.43% | 196.8% | 5.567% | 20.75% | 5.515% | 46.3% | 10.1% | 733.8% | 33.63% | 17.18% | 2.202% | 36.27% | 29.69% | 101.8% | 127.2% | 422.1% | 62.87% | 173.2% | 246.9% | 246.2% | 162.3% |  | 19.89% | 24.57% | 2.23% | 0.8355% | 5.945% | 0.5605% | 1.321% | 1.406% | 4.894% | 34.3% | 37.97% | 3.394% | 4.636%
150 | MT 36 B | R | 6/3/2023 6:43:03 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 36 B.152 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 55.68% |  | 8.026% | 1.612% | 4.847% | 101.1% | 287.3% | 78.39% | 16.57% | 4.144% | 2.493% | 16.05% | 7.967% | 156.7% | 18.68% | 17.81% | 15.61% | 0.9003% | 24.87% | 406.1% | 123.2% | 215% | 45.91% |  | 0% | 978.1% | 63.01% | 173.2% | 3.579% | 1.976% | 2.537% | 0.8917% | 3.069% | 1.455% | 2.594% | 2.508% | 2.122% | 100.4% | 44.6% | 5.775% | 6.017%
151 | MT 37 B | R | 6/3/2023 6:49:06 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 37 B.153 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 24.74% | 173.2% | 9.743% | 1.09% | 19.39% | 93.17% | 154.4% | 223.6% | 33.27% | 22.52% | 14.56% | 25.39% | 7.671% | 276.2% | 20.43% | 15.84% | 6.597% | 127.2% | 47.9% | 118.7% | 49.68% | 61.91% | 75.63% | 173.2% | 301.2% | 69.64% | 323.9% | 173.2% | 8.602% | 45.52% | 7.634% | 1.203% | 1.467% | 4.744% | 0.639% | 0.4963% | 2.766% | 36.75% | 25.87% | 10.62% | 7.33%
152 | MT 38 B | R | 6/3/2023 6:55:09 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 38 B.154 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 7.084% | 173.2% | 22.59% | 1.439% | 19.14% | 101.2% | 111.8% | 226.6% | 4.989% | 25.29% | 6.557% | 24.56% | 11.61% | 63.78% | 41.18% | 9.453% | 15.14% | 89.84% | 56.54% | 0% | 7.822% | 12.67% | 56.14% | 173.2% | 3579% | 73.63% | 239.2% | 86.63% | 14.66% | 25.78% | 0.8405% | 1.932% | 3.406% | 5.569% | 1.434% | 2.02% | 2.232% | 112.4% | 66% | 8.745% | 7.579%
153 | MT 39 B | R | 6/3/2023 7:01:12 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 39 B.155 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 47.8% | 86.6% | 34.08% | 0.1078% | 12.16% | 99.59% | 23.3% | 53.24% | 14.46% | 25.17% | 10.28% | 14.38% | 19.37% | 87.29% | 33.69% | 7.466% | 9.543% | 72.64% | 64.51% | 397.8% | 23.76% | 30.54% | 89.7% | 173.2% | 336.5% | 1043% | 1130% | 173.2% | 19.83% | 11.66% | 1.212% | 0.4638% | 2.127% | 15.68% | 0.6322% | 1.901% | 1.441% | 132.9% | 48.82% | 0.6249% | 5.094%
154 | MT 40 B | R | 6/3/2023 7:07:15 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 40 B.156 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 41.47% | 86.6% | 20.1% | 2.17% | 13.11% | 80.54% | 59.82% | 207.1% | 25.34% | 41.36% | 15.29% | 39.48% | 18.83% | 53.59% | 37.41% | 6.623% | 3.164% | 34.98% | 58.92% | 130.4% | 154.9% | 471.6% | 52.32% | 173.2% | 309.3% | 198.6% | 369.5% | 173.2% | 59.38% | 8.742% | 1.492% | 0.04178% | 5.454% | 2.629% | 1.283% | 0.2662% | 2.496% | 86.63% | 12.54% | 3.133% | 1.92%
155 | MT 32 C | R | 6/3/2023 7:13:17 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 32 C.157 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 57.05% | 173.2% | 28.34% | 1.271% | 10.45% | 48.38% | 136.6% | 191.7% | 7.771% | 19.63% | 5.16% | 12.59% | 6.116% | 4409% | 6.73% | 11.69% | 6.237% | 8.824% | 27.86% | 828.9% | 33.05% | 47.91% | 165.5% |  | 3532% | 802.8% | 83.38% |  | 13.77% | 45.36% | 1.581% | 4.394% | 0.7933% | 0.516% | 2.115% | 2.579% | 2.367% | 34.29% | 15.4% | 6.956% | 3.73%
156 | LBLANK | R | 6/3/2023 7:19:20 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LBLANK.158 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 5.248% |  | 364.8% | 4.553% | 13.5% | 65.44% | 0% | 196.6% | 30.39% | 31.37% | 5.21% | 85.89% | 17.33% | 10.58% | 18.71% | 40.5% | 40.15% | 29.24% | 7.081% | 111% | 20.13% | 28.87% | 115.9% |  | 311.4% | 27.85% | 165% | 86.65% | 35.06% | 223.8% | 2.229% | 3.111% | 3.885% | 1.139% | 1.64% | 1.283% | 6.081% | 173.2% | 52.02% | 214.5% | 146.1%
157 | CCV 1ppb TRACE | R | 6/3/2023 7:25:23 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.159 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 6.244% | 4.518% | 11.16% | 2.424% | 16.46% | 3.108% | 8.523% | 7.491% | 1.864% | 0.5694% | 4.531% | 1.405% | 5.457% | 3.992% | 6.461% | 5.394% | 0.6277% | 1.745% | 3.037% | 11.36% | 1.85% | 2.095% | 12.04% | 22.34% | 15.7% | 5.421% | 1.736% | 13.44% | 1.521% | 3.339% | 2.265% | 1.394% | 4.541% | 0.2398% | 1.555% | 0.7395% | 0.3333% | 0.8441% | 2.202% | 0.6885% | 2.267%
158 | CCV 1ppb TRACE | R | 6/3/2023 7:31:56 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.160 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.213% | 4.249% | 10.13% | 3.113% | 5.905% | 2.749% | 5.152% | 0.855% | 2.704% | 16.17% | 17.63% | 14.04% | 3.718% | 2.099% | 0.08896% | 4.332% | 15.98% | 17.17% | 3.751% | 14.63% | 17.09% | 17.46% | 8.626% | 31.65% | 12.27% | 22.36% | 9.87% | 22.72% | 14.72% | 15.04% | 0.2753% | 10.9% | 1.271% | 12.3% | 1.138% | 1.043% | 10.2% | 17.58% | 16.72% | 1.296% | 0.9829%
159 | LFB 10ppb | R | 6/3/2023 7:37:58 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LFB 10ppb.161 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 3.302% | 2.416% | 4.926% | 1.87% | 4.759% | 1.897% | 3.103% | 1.559% | 0.9392% | 1.066% | 1.741% | 17.07% | 2.13% | 0.7619% | 1.081% | 0.889% | 0.726% | 0.4657% | 2.462% | 4.26% | 20.51% | 18.93% | 1.769% | 23.04% | 14.3% | 16.8% | 3.806% | 18.47% | 2.254% | 18.78% | 0.3911% | 15.61% | 0.6502% | 0.5084% | 1.486% | 1.66% | 17.85% | 1.076% | 0.3317% | 1.201% | 1.687%
160 | LFB 10ppb | R | 6/3/2023 7:44:29 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LFB 10ppb.162 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 8.998% | 10.94% | 9.559% | 2.493% | 14.58% | 9.442% | 8.342% | 9.889% | 10.9% | 150.2% | 146.5% | 6.757% | 9.118% | 8.816% | 8.161% | 8.99% | 152% | 159.5% | 5.433% | 10.14% | 5.018% | 5.777% | 3.402% | 11.6% | 6.302% | 5.551% | 69.34% | 4.486% | 148.5% | 6.97% | 7.329% | 4.752% | 6.464% | 11.35% | 4.381% | 1.244% | 3.787% | 150.9% | 152.5% | 41.02% | 13.59%
161 | CCV 10 ppb MAJORS | R | 6/3/2023 7:50:31 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.163 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.71% | 2.429% | 2.746% | 0.5512% | 1.215% | 0.8791% | 2.328% | 2.324% | 1.701% | 3.105% | 1.972% | 1.467% | 3.308% | 1.156% | 3.431% | 0.6153% | 3.207% | 1.677% | 14.42% | 2.934% | 4.245% | 1.094% | 18.12% | 4.246% | 1.323% | 0.6731% | 14.49% | 2.253% | 2.587% | 2.008% | 11.22% | 0.9423% | 1.508% | 1.205% | 1.575% | 0.5462% | 5.632% | 2.311% | 0.1801% | 15.23% | 1.535%
162 | CCV 10 ppb MAJORS | R | 6/3/2023 7:57:03 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.164 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 5.5% | 7.774% | 9.32% | 2.522% | 7.376% | 5.856% | 5.715% | 5.823% | 5.847% | 0.9588% | 0.808% | 9.358% | 4.055% | 3.824% | 4.359% | 3.209% | 0.2879% | 0.4014% | 6.591% | 1.381% | 8.007% | 10.3% | 7.819% | 14.42% | 6.274% | 8.505% | 10.18% | 8.427% | 1.172% | 9.634% | 5.403% | 7.973% | 3.675% | 0.9733% | 16.26% | 0.7824% | 6.295% | 1.159% | 1.335% | 7.415% | 18.6%
163 | MT 32 C LD | R | 6/3/2023 8:03:06 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 32 C LD.165 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 26.84% | 173.2% | 15.33% | 2.045% | 10.47% | 7.057% | 107.3% | 140.9% | 8.601% | 6.249% | 11.2% | 38.55% | 11.7% | 324.9% | 18.16% | 11.16% | 16.78% | 2.706% | 160.3% | 130.6% | 642% | 1127% | 138.9% |  | 69.24% | 405.5% | 334.5% | 173.2% | 18.33% | 19.86% | 8.649% | 7.805% | 3.406% | 0.9663% | 1.024% | 1.317% | 6.748% | 86.61% | 30.01% | 4.532% | 4.441%
164 | MT 33 C | R | 6/3/2023 8:09:08 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 33 C.166 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 121% |  | 21.21% | 1.196% | 10.17% | 197.9% | 55.6% | 369.3% | 41.02% | 3.958% | 5.221% | 36.37% | 9.532% | 150.1% | 17.03% | 17.26% | 8.933% | 23.13% | 58.6% | 0% | 13.38% | 18.6% | 107.3% | 173.2% | 165.5% | 749.8% | 2237% |  | 18.56% | 3.565% | 12.8% | 0.8302% | 4.486% | 0.7645% | 2.292% | 1.733% | 2.175% | 99.49% | 36.3% | 22.7% | 1.587%
165 | MT 33 C LSP | R | 6/3/2023 8:15:09 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 33 C LSP.167 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 4.156% | 6.754% | 5.237% | 1.02% | 5.385% | 2.248% | 5.508% | 3.836% | 4.29% | 2.925% | 5.343% | 5.557% | 4.244% | 4.39% | 4.269% | 4.137% | 3.398% | 3.223% | 13.17% | 2.789% | 7.603% | 4.359% | 12.7% | 7.738% | 3.338% | 4.712% | 13.13% | 5.475% | 7.552% | 4.66% | 11.4% | 3.145% | 3.735% | 2.278% | 1.312% | 0.9037% | 6.544% | 3.536% | 3.125% | 14.01% | 1.622%
166 | MT 33 C LSP | R | 6/3/2023 8:21:40 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 33 C LSP.168 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 5.876% | 4.423% | 1.832% | 2.763% | 2.039% | 4.584% | 5.495% | 4.723% | 4.965% | 1.809% | 1.348% | 1.817% | 4.351% | 5.213% | 5.667% | 4.897% | 1.937% | 1.236% | 3.41% | 5.812% | 4.529% | 0.5198% | 6.897% | 2.633% | 1.175% | 1.509% | 3.195% | 1.816% | 0.8772% | 0.7919% | 2.904% | 0.81% | 3.376% | 1.277% | 1.147% | 0.9664% | 4.506% | 0.5424% | 0.5% | 1.98% | 0.2707%
167 | MT 34 C | R | 6/3/2023 8:27:42 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 34 C.169 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 13.56% | 86.61% | 20.92% | 1.269% | 1.739% | 61.36% | 146.1% | 291.2% | 30.07% | 14.82% | 5.634% | 37.59% | 9.408% | 70.95% | 9.993% | 22.84% | 3.63% | 890.2% | 27.59% | 543.5% | 195.6% | 140.4% | 21.31% | 173.2% | 10.97% | 219.6% | 22184% |  | 5.598% | 7.411% | 8.495% | 3.462% | 0.4912% | 0.3255% | 1.149% | 1.153% | 2.364% | 86.6% | 47.82% | 2.311% | 8.215%
168 | MT 35 C | R | 6/3/2023 8:33:45 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 35 C.170 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 51.44% | 86.72% | 34.59% | 1.926% | 13.56% | 74.34% | 143.3% | 7.899% | 13.94% | 13.82% | 4.022% | 31.22% | 8.005% | 276.1% | 26.12% | 10.77% | 7.475% | 80.55% | 40.92% | 0% | 17.88% | 28.65% | 27.72% |  | 105.3% | 0% | 295.1% | 173.2% | 31.74% | 13.99% | 2.723% | 6.665% | 3.68% | 3.089% | 3.027% | 2.378% | 4.968% | 86.61% | 8.119% | 13.18% | 3.262%

## Concentrations
| Sample Id | R | Acquisition Time | QC Status | Dataset File | Method File | Na KED 23/23 Helium KED_Low (ppb) | Mg KED 24/24 Helium KED_Low (ppb) | Al KED 27/27 Helium KED_Low (ppb) | P Oshift 31/47 Oxygen DRC (ppb) | S (34) -1 34/34 Helium KED_Low (ppb) | K KED 39/39 Helium KED_Low (ppb) | Ca KED 43/43 Helium KED_Low (ppb) | V KED 51/51 Helium KED_Low (ppb) | Cr (52) 52/52 Helium KED_Low (ppb) | Cr (53) NH3 53/53 Ammonia DRC (ppb) | Fe (54) NH3 54/54 Ammonia DRC (ppb) | Mn Oshift 55/71 Oxygen DRC (ppb) | Fe (56) KED 56/56 Helium KED_Low (ppb) | Co KED 59/59 Helium KED_Low (ppb) | Ni KED 60/60 Helium KED_Low (ppb) | Cu (63) KED 63/63 Helium KED_Low (ppb) | Cu (65) NH3 65/65 Ammonia DRC (ppb) | Zn NH3 66/66 Ammonia DRC (ppb) | As 75/75 (ppb) | As KED 75/75 Helium KED_Low (ppb) | As RH Oshift 75/91 Oxygen DRC (ppb) | As Oshift 75/91 Oxygen DRC (ppb) | Se (77) 77/77 (ppb) | Se (77) Oshift 77/93 Oxygen DRC (ppb) | Se (78) Oshift 78/94 Oxygen DRC (ppb) | Se (80) Oshift 80/96 Oxygen DRC (ppb) | Se (82) 82/82 (ppb) | Se (82) Oshift 82/98 Oxygen DRC (ppb) | Mo (95) NH3 95/95 Ammonia DRC (ppb) | Mo (95) OShift 95/127 Oxygen DRC (ppb) | Rh (IS1) 103/103 (IS) | Rh (IS2) 103/103 (IS) Oxygen DRC | Rh (IS3) 103/103 (IS) Helium KED_Low | Rh (IS4) 103/103 (IS) Ammonia DRC | Rh (IS5) 103/103 (IS) Ammonia DRC | Rh (IS6) 103/103 (IS) Oxygen DRC | Rh OShift (IS1) 103/119 (IS) Oxygen DRC | Cd (111) NH3 111/111 Ammonia DRC (ppb) | Cd (114) NH3 114/114 Ammonia DRC (ppb) | Pb 208/208 (ppb) | Pb NH3 208/208 Ammonia DRC (ppb)
1 | CAL BLANK | R | 06/02/2023 15:17:36 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CAL BLANK.004 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | [30 empty] | 100% | 100% | 100% | 100% | 100% | 100% | 100%
2 | LOW 0.01 | R | 06/02/2023 15:23:43 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LOW 0.01.005 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | -422.3993 | -327.5683 | -0.93583 | 0.38721 | 154.9971 | -303.9575 | -295.0181 | 0.017673 | 0.0098479 | 0.0097913 | -0.69414 | 0.0087781 | -1.0252 | 0.010013 | 0.0047174 | 0.015504 | 0.0047898 | -0.18306 | 0.0060177 | 0.0083017 | 0.0039041 | 0.0057828 | -0.041074 | 0.024948 | -0.016212 | -0.00031501 | -0.026953 | 0.022713 | 0.010451 | 0.010344 | 99.88% | 98.17% | 99.42% | 98.02% | 99.55% | 99.39% | 98.22% | 0.0099663 | 0.0099165 | 0.0096244 | 0.009868
3 | LOW 0.05 | R | 06/02/2023 15:29:46 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LOW 0.05.006 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | -382.0795 | -279.2759 | -0.59744 | 0.97096 | 207.3922 | -255.2057 | -249.3962 | 0.050063 | 0.055303 | 0.055384 | -0.63168 | 0.048923 | -0.90336 | 0.050103 | 0.05348 | 0.058267 | 0.048548 | -0.2017 | 0.049533 | 0.052476 | 0.05064 | 0.051525 | -0.036241 | 0.04765 | 0.024099 | 0.051803 | 0.029596 | 0.06202 | 0.040989 | 0.042998 | 100.3% | 95.71% | 95.83% | 97.53% | 95.26% | 96.53% | 95.79% | 0.049938 | 0.05099 | 0.057051 | 0.053985
4 | LOW 0.1 | R | 06/02/2023 15:35:48 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LOW 0.1.007 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | -316.1199 | -225.8575 | -0.87139 | 0.75218 | 97.956 | -219.0009 | -225.018 | 0.097748 | 0.091177 | 0.096914 | -0.44375 | 0.094897 | -0.74407 | 0.095425 | 0.082266 | 0.096822 | 0.1015 | -0.081046 | 0.10107 | 0.089531 | 0.098214 | 0.09125 | 0.10263 | 0.096584 | 0.097405 | 0.097935 | 0.10217 | 0.096353 | 0.089224 | 0.091731 | 100.5% | 98.82% | 100.7% | 97.75% | 103% | 97.8% | 92.3% | 0.10342 | 0.10428 | 0.11144 | 0.09848
5 | LOW 0.5 | R | 06/02/2023 15:41:51 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LOW 0.5.008 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 136.2129 | 193.5249 | 0.1294 | 1.8636 | 597.907 | 257.6033 | 192.3505 | 0.48108 | 0.47365 | 0.48243 | 0.50147 | 0.47199 | 0.50256 | 0.46838 | 0.47691 | 0.53363 | 0.52159 | 0.44981 | 0.48482 | 0.55327 | 0.49777 | 0.48646 | 0.46062 | 0.47137 | 0.43518 | 0.47763 | 0.46859 | 0.42841 | 0.4904 | 0.47815 | 105.4% | 105% | 100% | 99.91% | 98.85% | 88.43% | 101.9% | 0.5109 | 0.51221 | 0.53569 | 0.5287
6 | STD 1 | R | 06/02/2023 15:47:53 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\STD 1.009 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 990.5164 | 999.0291 | 1.0046 | 1.9976 | 1300.8923 | 993.6008 | 1002.1128 | 0.9878 | 0.98469 | 1.0152 | 1.6265 | 1.0029 | 1.6127 | 0.97921 | 1.0292 | 1.0428 | 1.0734 | 1.1343 | 1.0225 | 0.94708 | 1.0103 | 1.0098 | 0.95732 | 1.0784 | 0.98316 | 0.98738 | 1.0685 | 1.0665 | 1.022 | 1.0373 | 110% | 102.4% | 101.2% | 102.2% | 85.22% | 103.1% | 99.1% | 1.0056 | 0.99349 | 1.049 | 1.2687
7 | STD 5 | R | 06/02/2023 15:53:56 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\STD 5.010 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 5157.6412 | 5044.1152 | 4.8911 | 6.3929 | 4949.6683 | 5129.6499 | 4932.2957 | 4.8816 | 4.8161 | 4.9794 | 4.8582 | 4.9955 | 4.7704 | 4.8589 | 4.9788 | 4.9116 | 5.0183 | 5.0507 | 4.9976 | 4.8557 | 5.0325 | 4.9796 | 5.2796 | 4.6996 | 5.08 | 5.0466 | 5.1723 | 4.9666 | 5.1069 | 5.1578 | 114.1% | 99.78% | 101.7% | 103.8% | 102.6% | 103.9% | 95.52% | 4.9739 | 4.9009 | 4.9029 | 4.8232
8 | STD 10 | R | 06/02/2023 15:59:58 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\STD 10.011 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 10252.7747 | 9920.0692 | 9.5999 | 11.0451 | 10079.4311 | 10070.508 | 9927.3956 | 9.6599 | 9.7247 | 9.7897 | 9.9137 | 10.1422 | 9.6373 | 9.6803 | 10.1272 | 9.5721 | 9.7527 | 10.0025 | 10.153 | 9.8883 | 10.1937 | 9.9592 | 10.2911 | 9.7273 | 9.9298 | 9.6592 | 10.1021 | 9.7844 | 10.321 | 10.436 | 111% | 94.15% | 100.7% | 99.1% | 99.7% | 97.83% | 88.97% | 9.7515 | 9.8295 | 9.368 | 9.7051
9 | STD 25 | R | 06/02/2023 16:06:00 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\STD 25.012 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 26477.3189 | 25003.5244 | 28.1438 | 26.4679 | 26352.1901 | 26299.7528 | 26184.4979 | 26.9715 | 26.5589 | 24.2227 | 25.2453 | 24.9675 | 26.7371 | 26.5526 | 27.2977 | 25.7008 | 23.8781 | 25.1448 | 25.5185 | 27.26 | 25.7851 | 25.7309 | 25.7657 | 25.1569 | 26.178 | 25.6105 | 25.4178 | 25.563 | 26.9742 | 26.9903 | 102.1% | 79.75% | 83.24% | 86.73% | 97.87% | 83.98% | 77% | 24.851 | 24.7974 | 21.1033 | 21.8212
10 | STD 50 | R | 06/02/2023 16:12:02 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\STD 50.013 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 46753.0659 | 43797.3242 | 46.9741 | 50.028 | 47875.4863 | 46071.4173 | 48565.4265 | 49.7824 | 48.5088 | 47.8411 | 50.0775 | 50.1112 | 48.5368 | 47.9347 | 48.6293 | 46.2521 | 46.3961 | 48.6774 | 47.544 | 48.0934 | 48.1739 | 51.0684 | 48.2345 | 51.687 | 52.4448 | 51.5451 | 46.5481 | 51.9872 | 54.0085 | 55.4237 | 106.2% | 70.89% | 79.92% | 80.11% | 82.27% | 74.38% | 72.69% | 49.1032 | 48.8427 | 29.915 | 48.2553
11 | STD 100 | R | 06/02/2023 16:18:04 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\STD 100.014 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 95852.3792 | 86684.2364 | 99.1938 | 99.4349 | 99052.5906 | 91223.4826 | 104446.7551 | 106.8424 | 102.4972 | 98.0373 | 102.2693 | 102.3935 | 103.685 | 103.8368 | 101.9906 | 93.9563 | 90.6201 | 94.3714 | 110.1467 | 99.0107 | 93.374 | 100.9361 | 115.2894 | 106.0795 | 106.1982 | 104.5918 | 111.5283 | 105.2095 | 112.948 | 115.2779 | 92.58% | 69.83% | 78.64% | 79.66% | 82.73% | 73.46% | 72.99% | 96.9352 | 96.6369 | 49.4726 | 96.4856
12 | Calibration Curves | R | 06/02/2023 16:18:04
13 | LBLANK | R | 06/02/2023 16:24:07 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LBLANK.015 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | -145.773 | -48.7752 | -0.92607 | -0.32009 | 467.0084 | -167.5555 | -0.01244 | 0.0096825 | 0.0050838 | 0.00097923 | -0.56933 | -0.0036322 | -0.53599 | 0.00289 | -0.0094525 | 0.00025038 | -0.010745 | -0.23801 | -0.0071944 | 0.0039143 | -0.0036233 | -0.0038431 | 0.10924 | 0.014903 | -0.029849 | -0.0036589 | 0.020548 | -0.0016216 | -0.0039439 | -0.0015604 | 111.9% | 101.1% | 108.8% | 117.2% | 118.5% | 98.82% | 109.1% | -0.00031202 | 0.00075458 | -0.00075595 | -0.00039769
14 | ICV PE | R | 06/02/2023 16:30:09 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\ICV PE.016 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 9823.9524 | 9944.6776 | 11.1131 | 20.0669 | 10681.7877 | 9422.7641 | 10726.211 | 10.0739 | 10.0157 | 9.7619 | 10.0029 | 10.2034 | 10.5905 | 10.3283 | 10.5191 | 9.9341 | 9.7802 | 9.9155 | 9.885 | 8.7886 | 9.0956 | 10.0139 | 7.2921 | 9.9484 | 10.457 | 10.2681 | 10.0259 | 10.2691 | 10.1446 | 10.3165 | 123.4% | 99.04% | 110.5% | 107.5% | 110.8% | 99.44% | 105.4% | 9.8959 | 9.8193 | 8.3318 | 9.1537
15 | CCV 1ppb TRACE | R | 06/02/2023 16:36:12 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.017 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 917.1242 | 1032.8229 | 0.67866 | 0.82457 | 1812.5312 | 818.0184 | 1092.2421 | 1.0775 | 1.0519 | 1.0305 | 0.5955 | 1.0632 | 0.79028 | 1.1017 | 1.119 | 1.0813 | 1.0815 | 0.87659 | 0.97221 | 1.0031 | 0.86087 | 0.98823 | 0.98123 | 1.0375 | 0.95556 | 0.98496 | 1.0295 | 0.97738 | 1.0469 | 1.0568 | 123.2% | 109.7% | 120.9% | 114.8% | 115.8% | 97.9% | 121.7% | 1.0145 | 1.0178 | 0.92704 | 0.92218
16 | CCV 1ppb TRACE | R | 06/02/2023 16:42:44 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.018 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 988.4925 | 1014.6011 | 0.53485 | 0.52706 | 2076.2749 | 907.2816 | 1058.5224 | 1.0715 | 1.0503 | 1.021 | 0.55656 | 1.1252 | 0.76502 | 1.107 | 1.1572 | 1.0602 | 1.0954 | 1.4936 | 0.98863 | 0.93905 | 0.88344 | 1.0413 | 1.0361 | 1.0852 | 1.0643 | 1.0471 | 0.95568 | 1.1535 | 1.0515 | 1.0617 | 117.4% | 105.2% | 122.8% | 109.6% | 111% | 104.8% | 119.8% | 1.0101 | 0.99231 | 0.93319 | 0.91035
17 | LFB 10ppb | R | 06/02/2023 16:48:46 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LFB 10ppb.019 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 11220.7527 | 10584.4469 | 10.9725 | 11.6099 | 11646.8755 | 10407.1031 | 11307.4482 | 10.388 | 10.1814 | 10.2906 | 11.2461 | 11.0973 | 11.538 | 10.9081 | 10.9852 | 9.9463 | 10.1897 | 13.9928 | 10.5608 | 9.573 | 9.0555 | 10.448 | 10.8077 | 11.0004 | 10.7592 | 10.5684 | 10.5406 | 10.893 | 10.6068 | 10.5799 | 113.1% | 96.53% | 112.7% | 100.5% | 102.6% | 97.42% | 107.8% | 9.825 | 9.718 | 8.1616 | 8.3769
18 | LFB 10ppb | R | 06/02/2023 16:55:16 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LFB 10ppb.020 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 11164.1162 | 10847.4002 | 12.22 | 12.0531 | 11713.2675 | 10386.4177 | 11380.2089 | 10.5083 | 10.3136 | 10.1861 | 11.2148 | 11.2524 | 11.5015 | 10.9732 | 11.1528 | 10.0312 | 10.2523 | 14.0507 | 10.3909 | 10.2885 | 9.1395 | 10.3933 | 11.1021 | 11.0533 | 10.7623 | 10.7371 | 10.3819 | 10.6638 | 10.5372 | 10.6573 | 112.1% | 96.18% | 111.4% | 98.79% | 100.8% | 96.04% | 105.8% | 9.7462 | 9.649 | 8.0232 | 8.3515
19 | LLOQ 0.01 | R | 06/02/2023 17:01:19 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LLOQ 0.01.021 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | -138.0924 | -39.7596 | -0.58716 | 0.43129 | 941.3472 | -158.7711 | 14.3678 | 0.018883 | 0.016598 | 0.010949 | -0.46865 | 0.0075582 | -0.24595 | 0.012491 | 0.003325 | 0.014066 | 0.006141 | -0.13024 | -0.021182 | 0.0033634 | 0.0041725 | 0.0058062 | 0.0070255 | 0.041888 | -0.025776 | 0.0028406 | -0.017262 | 0.0021742 | 0.012904 | 0.014265 | 111.5% | 103.6% | 118.7% | 106.5% | 102.8% | 101.6% | 118.4% | 0.0087397 | 0.010384 | 0.0079997 | 0.0087649
20 | CCV 10 ppb MAJORS | R | 06/02/2023 17:07:22 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.022 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 12218.3718 | 11786.596 | 11.7543 | 11.4738 | 12395.218 | 11100.4799 | 11922.9999 | 10.785 | 10.7097 | 10.5238 | 10.7419 | 11.6461 | 11.6607 | 11.419 | 11.4898 | 10.3056 | 10.3968 | 10.5241 | 10.2955 | 9.9808 | 9.2524 | 10.5532 | 11.0721 | 10.5724 | 10.8825 | 10.68 | 10.3157 | 10.6049 | 10.592 | 10.774 | 111.4% | 94.39% | 108.3% | 96.93% | 102.5% | 94.97% | 104.1% | 9.7957 | 9.7371 | 8.1524 | 8.203
21 | CCV 10 ppb MAJORS | R | 06/02/2023 17:13:54 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.023 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 12155.8073 | 11534.2219 | 11.5892 | 10.6853 | 12260.8312 | 11190.5038 | 11962.7689 | 10.818 | 10.652 | 10.5538 | 10.8271 | 12.4547 | 11.3666 | 11.3338 | 11.4317 | 10.2051 | 10.4441 | 10.5546 | 10.4084 | 9.8099 | 9.3471 | 11.0214 | 11.0873 | 11.2832 | 11.4539 | 11.2855 | 10.6633 | 11.103 | 10.6018 | 11.2965 | 111.1% | 91.07% | 111.4% | 97.56% | 98.76% | 99.55% | 103.9% | 9.762 | 9.6951 | 8.2167 | 8.5509
22 | LLOQ 0.1 | R | 06/02/2023 17:19:57 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LLOQ 0.1.024 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | -21.9855 | 50.9662 | -0.41297 | 0.28369 | 1215.184 | -46.7358 | 125.3788 | 0.10957 | 0.11069 | 0.098192 | -0.2392 | 0.11878 | -0.0053462 | 0.11058 | 0.11615 | 0.10939 | 0.10227 | -0.010861 | 0.078753 | 0.095414 | 0.081962 | 0.09817 | 0.02152 | 0.13285 | 0.071288 | 0.10055 | 0.084762 | 0.078256 | 0.090107 | 0.098494 | 120.2% | 103.2% | 117.8% | 110.5% | 108.8% | 100.8% | 118.5% | 0.090996 | 0.0939 | 0.084129 | 0.087
23 | LBLANK | R | 06/02/2023 17:25:59 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LBLANK.025 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | -147.691 | -49.4851 | -0.86327 | -0.29428 | 1195.2686 | -163.5813 | -4.861 | 0.0076166 | 0.0026875 | 0.00079703 | -0.55016 | -0.0032707 | -0.30372 | 0.00028526 | -0.0115 | -0.0017169 | -0.0098416 | -0.21784 | -0.0036543 | 0.003273 | -0.0075308 | -0.0080116 | 0.030613 | 0.02405 | -0.034409 | -0.0071849 | -0.055017 | 0.0098666 | -0.011061 | -0.012106 | 110.8% | 102.3% | 119.6% | 104.2% | 106.6% | 102% | 120.6% | -0.0005172 | 8.349e-05 | -0.0017076 | -0.00056047
24 | LFB 6.25ppb & 2500ppb | R | 06/02/2023 17:32:01 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LFB 6.25ppb & 2500ppb.026 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 32284.453 | 30307.3697 | 24.6718 | 10.165 | 30868.3027 | 30109.4028 | 31301.7661 | 7.1933 | 7.052 | 7.0597 | 7.7449 | 7.5174 | 8.4385 | 7.5935 | 7.5147 | 6.5514 | 6.7455 | 7.1095 | 6.8655 | 6.395 | 5.6646 | 6.7761 | 7.4238 | 7.1031 | 7.0332 | 6.863 | 7.1388 | 6.994 | 7.0511 | 6.7685 | 101.5% | 90.7% | 101.1% | 88.82% | 91.62% | 87.03% | 105.3% | 6.4283 | 6.3118 | 5.4246 | 5.385
25 | MT 1 | R | 06/02/2023 17:38:02 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 1.027 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 3938.0267 | 228.4614 | 422.4351 | 2517.3629 | 3274.1776 | 3290.2344 | 673.5447 | 0.45867 | 0.55734 | 0.53116 | 91.6986 | 0.8713 | 98.3796 | 0.072481 | 0.37763 | 4.0977 | 4.1035 | 12.9328 | 1.5666 | 1.3648 | 1.3534 | 1.6312 | 0.7488 | 0.92616 | 0.84418 | 0.86674 | 0.97459 | 0.93372 | 0.067212 | 0.062977 | 106.4% | 98.46% | 111.5% | 99.27% | 102.9% | 87.68% | 115% | 0.068459 | 0.069387 | 0.15119 | 0.14031
26 | MT 1 LD | R | 06/02/2023 17:44:04 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 1 LD.028 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 4213.4507 | 238.2359 | 450.616 | 2626.5344 | 3456.5058 | 3568.3228 | 717.1952 | 0.45768 | 0.56047 | 0.57715 | 95.2097 | 0.90562 | 96.1427 | 0.078058 | 0.37349 | 4.0827 | 4.3734 | 14.2192 | 1.7187 | 1.3844 | 1.4645 | 1.709 | 0.93357 | 1.0354 | 0.89558 | 0.96807 | 1.1174 | 1.0082 | 0.072039 | 0.065195 | 98.44% | 97.08% | 109.9% | 93.43% | 100.6% | 88.32% | 109.6% | 0.078664 | 0.074799 | 0.16115 | 0.14655
27 | LBLANK | R | 06/02/2023 18:05:56 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LBLANK.029 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | -146.2195 | -48.7408 | -0.87234 | -0.15189 | 356.4664 | -165.3765 | -4.861 | 0.0069751 | 0.006241 | 0.0010477 | -0.53254 | -0.0033132 | -0.42668 | 0.00031495 | -0.0077825 | -0.00017509 | -0.009862 | -0.23012 | -0.021261 | -0.0031069 | -0.006368 | -0.0065477 | 0.050564 | 0.019441 | -0.031308 | -0.0054913 | -0.04187 | -0.0054505 | -0.011445 | -0.010879 | 104.4% | 98.65% | 112.4% | 97.73% | 101.1% | 95.94% | 117.3% | -0.0013303 | 0.00038503 | -0.0018674 | -0.00075614
28 | STD 1 | R | 06/02/2023 18:11:58 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\STD 1.030 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1256.5884 | 1252.5356 | 1.3331 | 1.9354 | 2533.1423 | 1136.5104 | 1275.4261 | 1.1409 | 1.1777 | 1.1119 | 1.9447 | 1.2679 | 2.2161 | 1.2071 | 1.1837 | 1.1593 | 1.1592 | 1.2236 | 1.0991 | 0.92583 | 0.90431 | 1.0711 | 1.1759 | 1.279 | 1.1325 | 1.1418 | 1.0397 | 1.2834 | 1.0595 | 1.0761 | 102.2% | 94.28% | 108.4% | 94.49% | 98.58% | 93.97% | 107.9% | 1.0072 | 0.99053 | 0.9298 | 0.89562
29 | STD 0.5 | R | 06/02/2023 18:18:06 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\STD 0.5.031 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 570.7201 | 558.1965 | 0.98988 | 1.9026 | 2058.4941 | 453.9212 | 578.0099 | 0.56019 | 0.57207 | 0.5629 | 0.68882 | 0.64012 | 1.0056 | 0.57829 | 0.59989 | 0.56297 | 0.60095 | 0.54788 | 0.50923 | 0.5244 | 0.49099 | 0.57946 | 0.39351 | 0.56899 | 0.58736 | 0.55312 | 0.42028 | 0.59043 | 0.54074 | 0.54631 | 105.9% | 92.76% | 110.5% | 93.25% | 97.34% | 91.86% | 105.7% | 0.50244 | 0.50494 | 0.43321 | 0.44793
30 | ICV PE | R | 06/02/2023 18:24:08 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\ICV PE.032 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 13897.8579 | 13749.6052 | 13.9589 | 28.1108 | 13951.4949 | 12791.5569 | 13327.2364 | 11.5708 | 11.3416 | 11.1409 | 11.7665 | 13.0434 | 12.4062 | 11.8958 | 11.9854 | 10.4681 | 10.6167 | 10.7588 | 10.7083 | 8.7767 | 9.2977 | 11.1401 | 8.1861 | 11.2813 | 11.719 | 11.5341 | 11.1233 | 11.3428 | 10.2514 | 10.8331 | 102.1% | 84.96% | 101.2% | 88.66% | 92.57% | 84.94% | 98.43% | 9.7436 | 9.6854 | 8.1639 | 8.5864
31 | ICV PE | R | 06/02/2023 18:30:39 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\ICV PE.033 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 13661.1271 | 13405.2844 | 13.5666 | 25.9672 | 13345.0393 | 12236.5207 | 13307.5124 | 11.4975 | 11.1834 | 11.1563 | 11.9394 | 12.86 | 12.1774 | 11.8808 | 11.9216 | 10.4523 | 10.6991 | 10.8552 | 10.5169 | 8.9232 | 9.0987 | 10.8854 | 8.0598 | 11.6457 | 11.4559 | 11.268 | 10.8422 | 11.3762 | 10.3886 | 10.7405 | 103% | 88.1% | 103.2% | 88.6% | 93.47% | 91.63% | 101.9% | 9.9262 | 9.8307 | 8.0438 | 8.4798
32 | CCV 1ppb TRACE | R | 06/02/2023 18:36:41 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.034 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1197.1876 | 1248.5498 | 0.39246 | 1.2004 | 2197.7058 | 1089.3757 | 1249.0218 | 1.1299 | 1.1132 | 1.1117 | 0.69584 | 1.3335 | 0.95012 | 1.1826 | 1.215 | 1.0725 | 1.1209 | 0.89552 | 1.0664 | 1.0087 | 0.93385 | 1.1345 | 1.2297 | 1.2676 | 1.1228 | 1.1986 | 1.0296 | 1.1019 | 1.0679 | 1.1433 | 103% | 91.24% | 113.1% | 95.77% | 98.08% | 94.47% | 107.2% | 0.99734 | 0.99206 | 0.90466 | 0.89042
33 | CCV 1ppb TRACE | R | 06/02/2023 18:43:13 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.035 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1280.3979 | 1330.1937 | 1.023 | 0.83521 | 2768.7563 | 1081.081 | 1381.8236 | 1.1716 | 1.187 | 1.1511 | 0.69211 | 1.2998 | 1.1162 | 1.2419 | 1.261 | 1.1349 | 1.1422 | 0.90637 | 1.0183 | 0.88583 | 0.91925 | 1.0956 | 1.1404 | 1.0998 | 1.0763 | 1.1457 | 1.0603 | 1.1365 | 1.0646 | 1.0782 | 102.5% | 93.99% | 110.3% | 93.56% | 97.03% | 91.93% | 108.3% | 1.0174 | 0.99391 | 0.91094 | 0.89151
34 | CCV 10 ppb MAJORS | R | 06/02/2023 18:49:16 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.036 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 14116.8842 | 13633.8579 | 13.5979 | 13.1568 | 14424.6605 | 12415.4571 | 13460.7718 | 11.6073 | 11.3983 | 11.4099 | 11.7938 | 12.9671 | 12.2666 | 12.157 | 12.3103 | 10.5781 | 10.9908 | 11.1114 | 10.7506 | 9.7871 | 9.1837 | 10.9285 | 11.6208 | 11.2458 | 11.6391 | 11.1064 | 10.9319 | 10.9671 | 10.873 | 10.9608 | 102.3% | 88.44% | 105.9% | 87.96% | 92.77% | 88.08% | 101.8% | 9.9046 | 9.7928 | 8.0981 | 8.4022
35 | CCV 10 ppb MAJORS | R | 06/02/2023 18:55:47 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.037 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 13859.7015 | 13002.0386 | 13.0026 | 13.0325 | 13944.7964 | 12276.3185 | 12850.3773 | 11.6348 | 11.3822 | 11.3969 | 11.6512 | 12.8255 | 12.1949 | 12.0396 | 12.1705 | 10.6666 | 10.8653 | 11.14 | 10.8657 | 10.1504 | 9.2303 | 11.0926 | 11.7398 | 11.6659 | 11.5002 | 11.1047 | 10.9412 | 11.0898 | 10.877 | 11.1551 | 101.1% | 87.41% | 104.4% | 88.41% | 88.42% | 85.66% | 101.6% | 9.8475 | 9.6894 | 8.1076 | 8.7508
36 | MT 1 | R | 06/02/2023 19:01:50 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 1.038 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 4390.7358 | 251.0427 | 480.1772 | 2566.2268 | 3789.2436 | 3596.5431 | 806.9309 | 0.48063 | 0.58234 | 0.55929 | 97.0559 | 1.0589 | 103.677 | 0.082321 | 0.37773 | 4.1209 | 4.1489 | 13.3781 | 1.5977 | 1.3501 | 1.6209 | 1.9139 | 0.81317 | 1.0787 | 1.028 | 1.0591 | 0.97694 | 0.97873 | 0.070888 | 0.078951 | 100.1% | 83.62% | 107.2% | 92.09% | 95% | 90.82% | 95.62% | 0.072153 | 0.072524 | 0.14317 | 0.1374
37 | MT 1 LD | R | 06/02/2023 19:07:53 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 1 LD.039 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 4737.7152 | 274.1657 | 522.9345 | 2687.2307 | 3935.6455 | 3968.5109 | 879.3574 | 0.48204 | 0.5954 | 0.56471 | 96.0594 | 0.97603 | 106.4012 | 0.084645 | 0.44127 | 4.3902 | 4.3454 | 14.174 | 1.6925 | 1.4061 | 1.4484 | 1.7711 | 0.79004 | 1.0325 | 0.98088 | 1.0047 | 1.0746 | 1.063 | 0.066619 | 0.065967 | 97.82% | 90.03% | 101.9% | 90.08% | 94.59% | 88.04% | 106.3% | 0.079197 | 0.074548 | 0.14617 | 0.14074
38 | MT 2 | R | 06/02/2023 19:13:55 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 2.040 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 5158.7022 | 252.262 | 560.1251 | 2891.4774 | 4116.3693 | 4139.4865 | 851.201 | 0.50451 | 0.62417 | 0.61414 | 99.9173 | 0.95328 | 104.0469 | 0.084898 | 0.40195 | 4.5614 | 4.7273 | 14.9376 | 1.7748 | 1.6273 | 1.6088 | 1.9035 | 0.87084 | 0.95784 | 1.0829 | 1.0914 | 1.1047 | 1.0527 | 0.098176 | 0.098976 | 97.04% | 88.5% | 104.4% | 88.15% | 93.28% | 87.34% | 101.6% | 0.076314 | 0.076029 | 0.14735 | 0.14372
39 | MT 2 LSP | R | 06/02/2023 19:19:58 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 2 LSP.041 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 43041.2575 | 35051.0008 | 610.1414 | 3362.0097 | 37498.7812 | 38471.4225 | 35018.4305 | 8.3381 | 8.2533 | 7.9847 | 115.5293 | 9.2264 | 121.2611 | 7.9125 | 8.2952 | 11.4312 | 11.3947 | 22.9213 | 8.7648 | 8.4286 | 7.7299 | 8.9702 | 8.4342 | 8.3947 | 8.4973 | 8.2237 | 8.2555 | 8.6432 | 7.037 | 7.0037 | 91.73% | 80.31% | 91.17% | 80.69% | 83.07% | 78.77% | 90.11% | 6.2989 | 6.2167 | 5.3218 | 5.4114
40 | MT 2 LSP | R | 06/02/2023 19:26:09 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 2 LSP.042 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 40167.212 | 32459.88 | 574.1295 | 3227.6575 | 35669.1004 | 35523.6786 | 32959.6604 | 7.93 | 7.6956 | 7.9018 | 113.666 | 9.2359 | 114.8888 | 7.634 | 8.0017 | 11.0315 | 11.3609 | 22.6737 | 8.8028 | 8.1391 | 7.6772 | 9.1221 | 8.3266 | 8.6324 | 8.444 | 8.397 | 8.1963 | 8.4878 | 7.03 | 7.2615 | 94.12% | 80.04% | 96.3% | 83.2% | 86.26% | 81.27% | 92% | 6.2396 | 6.1679 | 5.3555 | 5.3794
41 | MT 3 | R | 06/02/2023 19:32:11 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 3 .043 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 4004.6025 | 212.2501 | 435.5836 | 2351.5468 | 3706.9383 | 3294.5661 | 619.4182 | 0.42402 | 0.67014 | 0.66024 | 83.7203 | 0.79953 | 84.5065 | 0.073066 | 0.42675 | 3.759 | 3.9661 | 12.4638 | 1.581 | 1.4894 | 1.3547 | 1.6234 | 0.8665 | 1.0681 | 0.90249 | 0.90039 | 1.0432 | 0.93624 | 0.071804 | 0.065333 | 99.83% | 94.09% | 110% | 93.25% | 94.5% | 90.87% | 109% | 0.065051 | 0.068259 | 0.14473 | 0.13871
42 | MT 4 | R | 06/02/2023 19:38:14 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 4.044 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 5342.3849 | 259.1695 | 587.2254 | 3022.489 | 4880.4377 | 4463.1677 | 866.4666 | 0.55984 | 0.66918 | 0.63449 | 134.5167 | 1.1584 | 145.0237 | 0.097116 | 0.43049 | 4.8579 | 4.8848 | 15.7073 | 1.9539 | 1.7358 | 1.6878 | 2.0419 | 0.92327 | 1.1444 | 1.0885 | 1.1468 | 1.1902 | 1.0775 | 0.083875 | 0.080313 | 98.19% | 89.72% | 103.4% | 90.12% | 93.81% | 88.11% | 104.9% | 0.086431 | 0.085481 | 0.15863 | 0.15228
43 | LBLANK | R | 06/02/2023 19:44:17 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LBLANK.045 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | -148.4385 | -48.7435 | -0.94902 | 0.47215 | 1503.6932 | -154.0031 | -4.861 | 0.0074153 | 0.0039509 | 0.000477 | -0.52315 | -0.0032491 | -0.2402 | 0.0004931 | -0.0061775 | -0.0020108 | -0.0087663 | -0.21823 | -0.047512 | 0.00060526 | -0.0056794 | -0.005765 | 0.055805 | 0.02008 | -0.030431 | -0.0034118 | -0.0091382 | 0.0032203 | -0.010332 | -0.012062 | 99.33% | 89.97% | 104.6% | 88.3% | 93.57% | 95.7% | 105.9% | -0.0012259 | 0.00081252 | -0.0020589 | -0.00072911
44 | CCV 1ppb TRACE | R | 06/02/2023 19:50:20 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.046 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1350.6062 | 1376.7676 | 0.98963 | 0.99616 | 2948.8417 | 1202.9902 | 1431.5703 | 1.2657 | 1.1917 | 1.219 | 0.80702 | 1.35 | 1.1813 | 1.2799 | 1.326 | 1.151 | 1.1641 | 0.96691 | 1.0057 | 1.0306 | 0.90977 | 1.1105 | 1.4257 | 1.1527 | 1.1236 | 1.1697 | 0.98577 | 1.1579 | 1.0792 | 1.1104 | 94.2% | 86.26% | 103.7% | 87.38% | 90.89% | 87.36% | 101.7% | 1.0115 | 1.0049 | 0.90438 | 0.87798
45 | CCV 1ppb TRACE | R | 06/02/2023 19:56:52 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.047 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1341.2632 | 1409.3016 | 1.1294 | 0.80754 | 3029.6237 | 1128.3677 | 1350.0194 | 1.2141 | 1.1959 | 1.1948 | 0.75335 | 1.3215 | 1.1268 | 1.2477 | 1.2858 | 1.162 | 1.1752 | 0.9289 | 1.0808 | 1.021 | 0.90009 | 1.0863 | 1.2791 | 1.121 | 1.1189 | 1.142 | 1.0566 | 1.2841 | 1.0706 | 1.0766 | 94.19% | 87.53% | 104.8% | 87.61% | 91.75% | 87.12% | 102.1% | 0.99114 | 0.99077 | 0.87733 | 0.86759
46 | CCV 10 ppb MAJORS | R | 06/02/2023 20:02:55 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.048 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 15136.6738 | 14616.0052 | 14.6761 | 14.3822 | 15742.228 | 13354.1633 | 13892.798 | 12.3561 | 11.9936 | 11.8475 | 12.4123 | 13.3335 | 12.9346 | 12.8082 | 12.7879 | 10.9096 | 11.0973 | 11.41 | 10.9634 | 10.3172 | 9.3036 | 11.1423 | 11.956 | 11.5902 | 11.4282 | 11.2403 | 11.051 | 11.2373 | 10.8506 | 11.0599 | 94.32% | 80.62% | 96.31% | 81.6% | 84.73% | 79.77% | 93.33% | 9.8588 | 9.6914 | 7.7179 | 8.2375
47 | CCV 10 ppb MAJORS | R | 06/02/2023 20:09:27 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.049 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 14892.7634 | 14174.5086 | 14.4885 | 14.0731 | 14651.0299 | 13072.0905 | 13669.1522 | 12.2021 | 11.7414 | 12.0275 | 12.452 | 13.5502 | 12.469 | 12.641 | 12.5974 | 10.7941 | 11.3803 | 11.4364 | 10.8255 | 10.2835 | 9.3139 | 11.2431 | 12.1584 | 11.5709 | 11.4959 | 11.5765 | 11.0262 | 11.3981 | 11.0708 | 11.2378 | 95.17% | 81.16% | 99.65% | 81.42% | 81.38% | 81.55% | 94.74% | 9.9257 | 9.7906 | 7.7776 | 8.7688
48 | MT 5 | R | 06/02/2023 20:15:30 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 5.050 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 5444.3989 | 303.142 | 572.47 | 3141.2391 | 4613.6915 | 4290.7112 | 924.0429 | 0.53388 | 0.64217 | 0.61705 | 110.827 | 1.055 | 119.3673 | 0.097861 | 0.43402 | 4.6243 | 4.6274 | 15.2062 | 1.8163 | 1.514 | 1.4764 | 1.8379 | 1.0096 | 1.0629 | 1.0186 | 1.0114 | 1.1309 | 0.93376 | 0.075855 | 0.06983 | 96.07% | 89.31% | 102.8% | 87.6% | 90.9% | 84.45% | 107.4% | 0.08105 | 0.079407 | 0.14204 | 0.14285
49 | MT 6 | R | 06/02/2023 20:21:32 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 6.051 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 13254.0143 | 719.7689 | 1332.6762 | 7172.5836 | 9196.9916 | 10665.9048 | 2172.6291 | 1.2201 | 1.3237 | 1.2586 | 243.5425 | 2.3934 | 261.7472 | 0.20286 | 0.97031 | 10.8734 | 10.7793 | 35.549 | 4.2744 | 4.0632 | 3.7545 | 4.5278 | 2.0552 | 2.3986 | 2.5983 | 2.4247 | 2.5625 | 2.5676 | 0.18769 | 0.17022 | 93.37% | 84.1% | 98.96% | 84.57% | 86.13% | 84.71% | 98.05% | 0.1924 | 0.18337 | 0.2689 | 0.27379
50 | MT 7 | R | 06/02/2023 20:27:35 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 7.052 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 13093.7643 | 758.1468 | 1388.7267 | 7189.1611 | 9166.6726 | 10793.4359 | 1935.6723 | 1.1856 | 1.3271 | 1.3165 | 238.1719 | 2.3275 | 255.2242 | 0.20901 | 1.0616 | 10.8057 | 10.8248 | 35.487 | 4.3018 | 3.8956 | 3.7332 | 4.5399 | 2.1609 | 2.3122 | 2.6021 | 2.4548 | 2.7629 | 2.3709 | 0.2028 | 0.17851 | 94.93% | 83.8% | 99.07% | 85.33% | 88.67% | 82.6% | 98.71% | 0.17877 | 0.18277 | 0.31886 | 0.31049
51 | MT 8 | R | 06/02/2023 20:33:38 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 8.053 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 12795.8848 | 718.3922 | 1342.167 | 6781.8018 | 9112.9461 | 10496.7611 | 2058.432 | 1.1633 | 1.3687 | 1.2594 | 217.3419 | 2.2428 | 241.7167 | 0.20255 | 1.0324 | 10.4651 | 10.0661 | 32.6747 | 3.9864 | 3.7723 | 3.6098 | 4.2297 | 1.992 | 2.2354 | 2.3749 | 2.3077 | 2.4231 | 2.3432 | 0.17724 | 0.17091 | 94.14% | 84.6% | 97.74% | 85.84% | 89.38% | 82.91% | 95.88% | 0.17447 | 0.1666 | 0.26395 | 0.26116
52 | MT 9 | R | 06/02/2023 20:39:40 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 9.054 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 12208.9519 | 653.6601 | 1285.0152 | 6982.8798 | 8663.2891 | 9719.0445 | 1911.2318 | 1.1438 | 1.2449 | 1.2333 | 224.1746 | 2.2126 | 233.4657 | 0.19984 | 0.93338 | 11.7302 | 11.9763 | 34.085 | 4.27 | 3.6746 | 3.5425 | 4.2531 | 2.2316 | 2.5109 | 2.3457 | 2.3332 | 2.7059 | 2.4277 | 0.18493 | 0.16324 | 88.84% | 83.4% | 99.46% | 82.93% | 87.81% | 80.55% | 96.8% | 0.1752 | 0.1756 | 0.27053 | 0.25304
53 | MT 10 | R | 06/02/2023 20:45:42 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 10.055 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 12729.0397 | 760.0704 | 1375.8636 | 7377.1544 | 8982.8171 | 10488.9016 | 1993.7665 | 1.1939 | 1.5046 | 1.4337 | 246.053 | 2.4661 | 260.7206 | 0.20506 | 1.0248 | 11.9928 | 11.9431 | 36.2005 | 4.3419 | 3.758 | 3.8499 | 4.572 | 2.0949 | 2.3888 | 2.5134 | 2.543 | 2.593 | 2.4176 | 0.19157 | 0.18912 | 90.74% | 80.72% | 97.23% | 83.04% | 86.91% | 80.64% | 92.68% | 0.18258 | 0.17426 | 0.26714 | 0.25831
54 | MT 11 | R | 06/02/2023 20:51:44 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 11 .056 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 25247.4894 | 1464.7179 | 2644.5768 | 14676.4934 | 16195.242 | 20678.5465 | 4332.207 | 2.4004 | 2.5326 | 2.5679 | 478.5436 | 4.9298 | 506.2743 | 0.40944 | 2.0135 | 20.152 | 20.4554 | 71.2293 | 8.4965 | 7.3843 | 7.5744 | 8.9291 | 4.1273 | 5.0453 | 4.9622 | 5.0111 | 5.5323 | 5.092 | 0.41223 | 0.35751 | 89.06% | 79.53% | 95.37% | 80.34% | 83.98% | 79.08% | 90.64% | 0.36724 | 0.353 | 0.46897 | 0.47136
55 | MT 11 LD | R | 06/02/2023 20:57:47 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 11 LD.057 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 25625.0538 | 1403.6442 | 2638.7866 | 14003.7903 | 15954.7379 | 20235.4264 | 4416.3157 | 2.3315 | 2.5971 | 2.5805 | 469.8813 | 4.7747 | 497.1705 | 1.0024 | 1.9664 | 20.1502 | 20.3807 | 70.4693 | 8.5636 | 7.6082 | 7.4979 | 8.6688 | 3.985 | 4.785 | 4.8997 | 4.9278 | 5.3562 | 4.9587 | 0.38904 | 0.34505 | 88.67% | 80.54% | 94.54% | 80.64% | 84.36% | 81.56% | 90.09% | 0.35397 | 0.34548 | 0.48258 | 0.46516
56 | MT 11 LD | R | 06/02/2023 21:04:18 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 11 LD.058 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 26893.5705 | 1560.6703 | 2833.4009 | 14906.4788 | 17127.6831 | 22054.5885 | 4690.7287 | 2.5068 | 2.8262 | 2.5254 | 473.5868 | 4.8903 | 538.092 | 1.0642 | 2.1366 | 21.7048 | 20.1428 | 69.8057 | 8.4572 | 8.1214 | 7.456 | 8.8817 | 3.7901 | 4.6811 | 4.98 | 4.856 | 5.295 | 4.7465 | 0.37081 | 0.35376 | 90.73% | 80.42% | 90.99% | 83.04% | 87.94% | 79.26% | 92.66% | 0.34739 | 0.33726 | 0.46585 | 0.44984
57 | MT 12 | R | 06/02/2023 21:10:20 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 12.059 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 26342.5624 | 1560.1194 | 2680.7375 | 15291.7572 | 16365.7034 | 21281.8487 | 4253.8227 | 2.389 | 2.5949 | 2.7041 | 499.7453 | 4.7999 | 498.9965 | 0.43527 | 1.9143 | 20.7387 | 22.1678 | 75.2042 | 8.8558 | 7.7155 | 7.4343 | 9.0999 | 4.2069 | 5.3851 | 5.364 | 5.1638 | 5.537 | 5.2368 | 0.40404 | 0.35362 | 88.36% | 79.89% | 94.62% | 78.96% | 84.64% | 78.34% | 94.53% | 0.39152 | 0.36931 | 0.48594 | 0.47577
58 | MT 12 LSP | R | 06/02/2023 21:16:23 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 12 LSP.060 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 68607.0345 | 38472.7417 | 2748.4494 | 15852.8372 | 53419.8926 | 59320.7928 | 41018.542 | 10.8119 | 10.4666 | 10.7756 | 511.793 | 13.6767 | 516.1271 | 8.6999 | 10.3288 | 27.6434 | 28.2797 | 82.0636 | 16.8457 | 14.8397 | 13.6994 | 16.7925 | 13.1244 | 13.2042 | 13.3986 | 12.8997 | 13.7037 | 13.2825 | 7.6431 | 7.6106 | 84.6% | 74.1% | 87.53% | 74.45% | 78.26% | 76.26% | 87.9% | 6.6157 | 6.5279 | 5.5967 | 5.6225
59 | LBLANK | R | 06/02/2023 21:22:26 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LBLANK.061 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | -143.7588 | -46.5218 | -0.78556 | 0.15823 | 1683.4894 | -160.4684 | -4.861 | 0.0094303 | 0.005387 | 0.00051521 | -0.50395 | -0.0029249 | -0.14883 | 0.00049207 | 0.0021397 | 0.0025971 | -0.0069244 | -0.18386 | -0.047651 | 0.0043281 | -0.0070991 | -0.0074371 | 0.14709 | 0.020302 | -0.026509 | -0.0045254 | 0.016283 | -0.00080727 | 0.0011055 | -0.011562 | 98.95% | 85.36% | 104.6% | 87.03% | 93.37% | 83.76% | 101.9% | -0.0006411 | 0.00059854 | -0.0018058 | -0.00074913
60 | CCV 1ppb TRACE | R | 06/02/2023 21:28:29 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.062 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1517.2125 | 1520.3576 | 0.68273 | 1.2064 | 3447.1305 | 1288.0561 | 1473.8406 | 1.3139 | 1.2404 | 1.2524 | 0.82199 | 1.3582 | 1.3041 | 1.3271 | 1.3551 | 1.2355 | 1.2313 | 1.0177 | 1.074 | 1.0126 | 0.91882 | 1.1237 | 1.3492 | 1.0093 | 1.1465 | 1.1391 | 1.1309 | 1.0883 | 1.0961 | 1.0759 | 93.52% | 84.19% | 99.78% | 83.6% | 88.21% | 83.04% | 99.49% | 1.0191 | 1.0049 | 0.8405 | 0.83944
61 | CCV 1ppb TRACE | R | 06/02/2023 21:35:00 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.063 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1538.606 | 1596.5673 | 0.98671 | 1.1278 | 3499.2738 | 1288.1703 | 1500.661 | 1.2957 | 1.2455 | 1.2826 | 0.89831 | 1.3918 | 1.292 | 1.3285 | 1.3127 | 1.2022 | 1.2776 | 1.0317 | 1.0332 | 0.96423 | 0.91999 | 1.1448 | 1.3521 | 1.1303 | 1.2286 | 1.1622 | 1.0164 | 1.1517 | 1.1242 | 1.1045 | 90.5% | 81.75% | 101.3% | 80.57% | 86.68% | 80.87% | 98.24% | 1.0321 | 1.0195 | 0.84105 | 0.83925
62 | LFB 10ppb | R | 06/02/2023 21:41:03 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LFB 10ppb.064 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 16963.3012 | 15798.9625 | 15.1889 | 16.611 | 16494.2433 | 14309.8878 | 14470.7536 | 12.6252 | 12.1419 | 12.6794 | 13.6782 | 12.2563 | 13.6614 | 12.7325 | 12.9815 | 11.0887 | 11.6226 | 15.8761 | 10.6027 | 10.3309 | 8.3066 | 10.2467 | 11.8793 | 10.5443 | 10.9292 | 10.454 | 10.3203 | 10.8775 | 10.9785 | 10.1044 | 91.41% | 82.7% | 92.07% | 75.91% | 80.66% | 76.02% | 98.71% | 9.7768 | 9.6408 | 7.1442 | 7.7084
63 | LFB 10ppb | R | 06/02/2023 21:47:33 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LFB 10ppb.065 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 16616.0665 | 15582.4334 | 15.9918 | 17.8106 | 16644.221 | 14030.1672 | 14342.6585 | 12.6007 | 12.1688 | 12.405 | 13.4389 | 13.7796 | 13.7973 | 12.9494 | 13.1534 | 11.1804 | 11.3402 | 15.4052 | 11.2858 | 10.2763 | 9.6147 | 11.5632 | 12.5099 | 12.0041 | 11.6658 | 11.3745 | 11.5239 | 11.5838 | 10.7815 | 11.2615 | 87.21% | 75.19% | 92.66% | 77.19% | 81.13% | 75.16% | 87.47% | 9.6619 | 9.5592 | 7.4358 | 7.6839
64 | CCV 10 ppb MAJORS | R | 06/02/2023 21:53:35 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.066 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 16818.1446 | 15942.7348 | 15.6709 | 16.672 | 17044.9012 | 14246.7741 | 14415.2684 | 12.6443 | 12.306 | 12.3118 | 12.7164 | 13.9061 | 13.1644 | 13.0024 | 13.286 | 11.1987 | 11.2807 | 11.6984 | 13.0962 | 10.9236 | 9.4231 | 11.477 | 14.5036 | 11.2376 | 11.6545 | 11.1099 | 13.1031 | 11.4886 | 10.7319 | 11.2272 | 78.69% | 76.51% | 93.71% | 79.29% | 80.66% | 74.59% | 90.09% | 9.5234 | 9.3744 | 8.6335 | 7.8975
65 | CCV 10 ppb MAJORS | R | 06/02/2023 22:00:06 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.067 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 16930.2435 | 15671.7282 | 15.1116 | 13.8925 | 16418.3915 | 14348.0151 | 14649.3146 | 12.7903 | 12.4661 | 13.4725 | 13.8084 | 13.8865 | 13.2379 | 13.1027 | 13.0618 | 11.236 | 12.3435 | 12.5048 | 11.359 | 10.6404 | 9.2898 | 11.5157 | 12.5503 | 11.7664 | 11.6988 | 11.6708 | 11.4076 | 11.2852 | 11.6897 | 11.2583 | 87.44% | 75.8% | 95.11% | 74.16% | 80.37% | 82.95% | 90.84% | 10.358 | 10.1935 | 7.5404 | 7.9021
66 | MT 13 | R | 06/02/2023 22:06:09 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 13.068 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 30737.5572 | 1820.8405 | 3100.6238 | 17961.8263 | 19195.8508 | 24300.9309 | 4801.868 | 2.7657 | 2.9328 | 2.8446 | 531.2763 | 5.3806 | 557.2625 | 0.41973 | 2.1147 | 23.0437 | 22.8116 | 77.5063 | 10.0831 | 8.3222 | 8.201 | 9.8608 | 4.8127 | 5.5166 | 5.3644 | 5.3412 | 5.9779 | 5.0301 | 0.45202 | 0.36972 | 82.49% | 77.12% | 90.35% | 77.39% | 82.22% | 73.12% | 89.77% | 0.39763 | 0.37004 | 0.51607 | 0.47666
67 | MT 14 | R | 06/02/2023 22:12:11 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 14.069 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 30703.9049 | 1770.0111 | 3009.2567 | 17140.4128 | 18615.777 | 24097.4289 | 4432.4352 | 2.6847 | 2.8559 | 3.0333 | 583.4065 | 5.853 | 549.5117 | 0.44979 | 2.19 | 22.6285 | 22.5197 | 75.6795 | 9.4708 | 8.3281 | 7.9865 | 9.6979 | 4.3 | 5.5278 | 5.431 | 5.4105 | 5.6735 | 5.4869 | 0.46434 | 0.37948 | 85.79% | 75.97% | 90.6% | 77.42% | 80.63% | 73.74% | 89.27% | 0.38501 | 0.36444 | 0.49709 | 0.49228
68 | MT 15 | R | 06/02/2023 22:18:14 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 15.070 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 30329.0572 | 1775.8103 | 2986.0743 | 17146.9435 | 18819.1582 | 24220.4616 | 4441.4852 | 2.6714 | 2.9493 | 2.9503 | 517.054 | 9.7159 | 545.2228 | 0.43221 | 2.7244 | 22.5383 | 22.1493 | 76.1901 | 9.4913 | 7.9056 | 8.0273 | 9.7551 | 4.4638 | 5.4427 | 5.5721 | 5.4362 | 5.7642 | 5.4677 | 0.47372 | 0.39614 | 84.13% | 75.77% | 89.33% | 76.92% | 80.78% | 74.88% | 89.1% | 0.39311 | 0.37653 | 0.49425 | 0.48125
69 | MT 16 | R | 06/02/2023 22:24:17 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 16.071 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 8952.299 | 1172.257 | 24.6276 | 13630.4333 | 13152.7928 | 9335.3707 | 502.9978 | 0.25554 | 2.7674 | 2.7594 | 1249.2524 | 8.199 | 1139.5852 | 0.2229 | 2.1867 | 22.6325 | 22.9954 | 66.5366 | 7.0708 | 5.1974 | 5.2779 | 6.4969 | 4.7078 | 4.9034 | 5.0627 | 5.0403 | 5.7142 | 4.8317 | 2.231 | 2.2188 | 78.27% | 78.76% | 94.92% | 79.14% | 84.23% | 79.1% | 93.82% | 11.8326 | 11.2179 | 0.21164 | 0.17442
70 | MT 17 | R | 06/02/2023 22:30:20 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 17.072 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 21229.9516 | 2684.8274 | 52.6549 | 30446.6843 | 28066.5028 | 22237.4235 | 1132.122 | 0.57723 | 6.0914 | 5.8882 | 2749.1227 | 19.764 | 2738.849 | 0.51094 | 4.9551 | 50.6334 | 49.3095 | 144.4342 | 14.2194 | 12.3232 | 13.2062 | 15.7811 | 9.0058 | 12.269 | 12.6082 | 12.1684 | 11.3891 | 11.9798 | 4.7945 | 5.2846 | 82.19% | 72.19% | 88.36% | 76.58% | 80.47% | 77.42% | 83.37% | 25.856 | 24.7805 | 0.37231 | 0.35988
71 | MT 18 | R | 06/02/2023 22:36:32 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 18.073 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 18623.5557 | 2281.0516 | 46.3191 | 27615.8101 | 24636.082 | 19477.6248 | 1004.3962 | 0.51498 | 4.7032 | 4.7108 | 2418.7587 | 15.5559 | 2185.5614 | 0.44386 | 4.4388 | 43.0445 | 43.2757 | 124.9763 | 12.5026 | 10.3976 | 10.5607 | 12.5271 | 7.7149 | 9.6203 | 9.9076 | 9.6502 | 9.9133 | 9.835 | 4.2291 | 4.2799 | 81.35% | 75.75% | 87.51% | 75.93% | 79.75% | 73.85% | 86.87% | 22.396 | 21.2256 | 0.29733 | 0.28141
72 | MT 19 | R | 06/02/2023 22:42:43 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 19.074 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 9214.5213 | 1187.8096 | 24.0498 | 13890.2763 | 13076.5655 | 9523.619 | 505.9479 | 0.25464 | 2.286 | 2.2286 | 1234.8446 | 7.8282 | 1084.3919 | 0.22943 | 3.4478 | 24.9884 | 25.5039 | 63.8711 | 5.9391 | 5.1523 | 5.0124 | 6.293 | 3.8344 | 4.5419 | 4.9302 | 4.797 | 4.6841 | 4.6831 | 2.0523 | 2.0806 | 83.77% | 74.99% | 88.16% | 76.25% | 82.36% | 73.95% | 91.01% | 11.254 | 10.6989 | 0.18799 | 0.17659
73 | MT 20 | R | 06/02/2023 22:48:45 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 20.075 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 20144.4177 | 2380.2906 | 48.5462 | 28884.8142 | 25867.958 | 20396.1114 | 995.6703 | 0.50021 | 4.7815 | 4.8015 | 2499.7317 | 16.1044 | 2265.957 | 0.44006 | 4.0102 | 45.2108 | 45.3714 | 130.9351 | 12.2058 | 10.8696 | 10.7466 | 12.8257 | 7.7118 | 9.8549 | 10.3885 | 9.8474 | 9.8735 | 9.8703 | 4.1843 | 4.2133 | 82.37% | 75.26% | 85.07% | 74.03% | 79.43% | 72.74% | 86.89% | 23.357 | 22.0995 | 0.30253 | 0.28984
74 | MT 21 | R | 06/02/2023 22:54:57 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 21.076 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 10368.6275 | 1389.894 | 26.1933 | 16791.9613 | 13708.9997 | 10586.5301 | 573.1915 | 0.26084 | 2.8347 | 2.8774 | 1356.5094 | 8.6074 | 1209.7909 | 0.23841 | 2.4675 | 23.6617 | 24.0391 | 72.0528 | 6.2961 | 5.5014 | 5.3945 | 6.5977 | 4.0134 | 5.1139 | 5.1603 | 5.0197 | 5.1573 | 4.7603 | 2.2103 | 2.2105 | 86.08% | 77.21% | 86.72% | 75.52% | 84.58% | 68.91% | 91.35% | 12.2157 | 11.7715 | 0.21137 | 0.19643
75 | MT 21 LD | R | 06/02/2023 23:00:59 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 21 LD.077 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | -139.1487 | -49.4851 | -1.0164 | 0.38984 | -1914.8967 | -161.295 | -4.861 | 0.0071048 | 0.008229 | 0.00097882 | -0.50997 | -0.004217 | -0.33164 | 0.00052829 | -0.0099214 | 0.023057 | -0.010154 | -0.25227 | 0.62232 | 0.0048715 | -0.0075785 | -0.0082167 | -1.6368 | 0.020145 | -0.030202 | -0.0063845 | -0.09456 | -0.0054505 | 0.0016234 | -0.012515 | 98.13% | 87.83% | 98.07% | 92.42% | 95.67% | 86.51% | 96.61% | -0.00077791 | 0.00021411 | -0.0025041 | -0.0012344
76 | LBLANK | R | 06/02/2023 23:07:00 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LBLANK.078 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | -145.4882 | -49.4851 | -0.80611 | 0.55563 | 1671.8764 | -154.3215 | -2.8124 | 0.0079497 | 0.0071477 | 0.0043221 | -0.4533 | -0.0033659 | -0.087474 | 0.00043961 | -0.0024758 | 0.0016345 | -0.007604 | -0.224 | -0.10021 | -0.0031069 | -0.0067848 | -0.0071568 | 0.2154 | 0.026794 | -0.032915 | -0.0039661 | -0.059389 | -0.0054505 | -0.0090693 | -0.012204 | 83.05% | 77.62% | 90.89% | 77.65% | 82.47% | 74.71% | 89.92% | -0.00078141 | 0.00033854 | -0.0020194 | -0.00071883
77 | CCV 1ppb TRACE | R | 06/02/2023 23:13:03 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.079 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1921.9981 | 1865.4923 | 1.028 | 2.1966 | 3700.4663 | 1516.1001 | 1591.8276 | 1.398 | 1.2745 | 1.2502 | 0.82723 | 1.4553 | 1.2864 | 1.3102 | 1.31 | 1.227 | 1.1453 | 0.93786 | 1.1259 | 1.0218 | 0.96293 | 1.144 | 1.2541 | 1.1468 | 1.2247 | 1.1432 | 1.1947 | 1.2084 | 1.0331 | 1.1041 | 83.82% | 74.29% | 85.81% | 78.61% | 81.3% | 68.94% | 85.39% | 0.94444 | 0.93824 | 0.82146 | 0.82402
78 | CCV 1ppb TRACE | R | 06/02/2023 23:19:35 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.080 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1937.1723 | 1884.1061 | 1.4867 | 1.9533 | 4118.1488 | 1574.5214 | 1515.4807 | 1.3645 | 1.3246 | 1.3493 | 0.93922 | 1.4504 | 1.4205 | 1.3774 | 1.3459 | 1.2303 | 1.2412 | 1.0183 | 1.0005 | 1.1337 | 0.9604 | 1.1573 | 1.2777 | 1.2424 | 1.2232 | 1.1422 | 1.0542 | 1.1886 | 1.0934 | 1.1163 | 85.2% | 73.52% | 86.02% | 75.57% | 82.58% | 71.63% | 85.58% | 1.0072 | 0.98545 | 0.82437 | 0.81035
79 | CCV 10 ppb MAJORS | R | 06/02/2023 23:25:38 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.081 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 21067.7647 | 18684.2381 | 17.5281 | 18.0676 | 18822.0334 | 17038.6265 | 16145.4684 | 13.2759 | 12.979 | 13.0011 | 13.6353 | 14.4695 | 13.6609 | 13.3696 | 13.6205 | 11.5073 | 11.6265 | 11.9894 | 11.4145 | 11.0874 | 9.5235 | 11.6226 | 12.2196 | 11.8034 | 12.3136 | 11.4598 | 11.5536 | 11.3066 | 10.8811 | 11.3095 | 81.94% | 69.55% | 80.87% | 71.19% | 74.5% | 68.15% | 82.08% | 9.7467 | 9.6583 | 7.311 | 7.7987
80 | CCV 10 ppb MAJORS | R | 06/02/2023 23:32:09 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.082 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 21074.6872 | 18835.8147 | 18.6471 | 17.3845 | 18858.5816 | 16776.0468 | 15757.751 | 13.4103 | 12.9431 | 13.2497 | 13.6733 | 14.7788 | 13.6744 | 13.4601 | 13.5352 | 11.6628 | 11.7513 | 12.0642 | 11.4824 | 11.0534 | 9.6273 | 11.7135 | 12.2649 | 11.513 | 11.8556 | 11.7302 | 12.0614 | 11.7759 | 11.1743 | 11.3448 | 82.91% | 69.55% | 82.23% | 70.88% | 76.78% | 70.71% | 81.8% | 9.8131 | 9.8278 | 7.3186 | 7.5647
81 | MT 22 | R | 06/02/2023 23:38:12 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 22.083 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 7350.3081 | 1011.1728 | 19.7143 | 10961.948 | 10871.874 | 7490.4282 | 374.7182 | 0.22444 | 1.6051 | 1.6466 | 941.3146 | 6.019 | 822.8003 | 0.16519 | 1.4798 | 16.2837 | 16.3145 | 46.6807 | 4.2972 | 3.6982 | 3.867 | 4.5236 | 2.8171 | 3.5258 | 3.6479 | 3.486 | 3.4347 | 3.3286 | 1.4823 | 1.4479 | 84.62% | 74.16% | 86.67% | 75.85% | 81.5% | 71.95% | 83.91% | 8.1357 | 7.6787 | 0.17965 | 0.17139
82 | MT 22 LSP | R | 06/02/2023 23:44:15 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 22 LSP.084 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 61874.9849 | 48592.044 | 32.2961 | 2.8966 | 51960.4154 | 51552.8568 | 41753.7778 | 9.1641 | 10.5983 | 10.1595 | 993.632 | 15.6714 | 849.2299 | 8.7239 | 10.0689 | 23.5674 | 23.2071 | 55.4439 | 11.9743 | 10.6146 | 10.0857 | 12.2178 | 10.8796 | 11.4073 | 11.8509 | 11.2157 | 11.3957 | 10.9367 | 8.586 | 8.7843 | 77.2% | 66.06% | 76.72% | 68.93% | 72.14% | 28.74% | 77.38% | 14.2626 | 13.8916 | 4.864 | 4.9946
83 | MT 22 LSP | R | 06/02/2023 23:50:45 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 22 LSP.085 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 56021.5458 | 42869.2132 | 26.5192 | 12429.388 | 45997.6668 | 46870.1974 | 37071.8348 | 8.026 | 9.1533 | 10.1918 | 992.1416 | 15.7762 | 757.7792 | 7.8074 | 8.9608 | 21.1906 | 23.1063 | 55.2389 | 11.9942 | 9.6926 | 10.1601 | 12.1738 | 10.9098 | 11.6752 | 11.7744 | 11.2124 | 11.2857 | 10.8573 | 8.6205 | 8.7875 | 76.92% | 67.06% | 85.95% | 69.39% | 73.95% | 66.36% | 77.71% | 14.4183 | 13.8689 | 4.9356 | 4.8327
84 | MT 23 | R | 06/02/2023 23:56:47 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 23.086 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 4167.6084 | 531.5582 | 12.1815 | 6379.9953 | 7387.8538 | 4361.5988 | 237.885 | 0.11409 | 1.1276 | 1.0868 | 504.7472 | 3.5213 | 508.0426 | 0.10454 | 0.94733 | 9.5858 | 9.6734 | 28.5944 | 2.6441 | 2.2552 | 2.1525 | 2.647 | 2.0474 | 2.193 | 2.0502 | 1.9592 | 2.189 | 2.1011 | 0.96868 | 0.92851 | 81.67% | 76.92% | 86.97% | 75.88% | 90.86% | 73.65% | 91.63% | 4.8808 | 4.6181 | 0.10234 | 0.083308
85 | MT 24 | R | 06/03/2023 00:02:49 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 24.087 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 3953.5362 | 461.5362 | 10.051 | 5575.9935 | 6759.4653 | 3941.4553 | 190.6957 | 0.11138 | 1.0913 | 1.1048 | 452.6408 | 3.3004 | 451.843 | 0.093842 | 0.91476 | 9.0554 | 9.0861 | 26.0721 | 2.4435 | 1.9862 | 2.0786 | 2.518 | 1.6085 | 2.2221 | 2.0217 | 1.8397 | 1.9844 | 1.9986 | 0.82636 | 0.84657 | 84.56% | 75% | 88.1% | 77.43% | 84.13% | 77.18% | 87.81% | 4.4974 | 4.2525 | 0.074568 | 0.071418
86 | MT 25 | R | 06/03/2023 00:08:52 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 25.088 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 4647.6853 | 593.1143 | 13.2612 | 6963.9529 | 7198.9212 | 4765.6722 | 296.9423 | 0.13343 | 0.79615 | 0.85817 | 523.3037 | 3.8095 | 514.8923 | 0.10586 | 0.68377 | 10.5038 | 10.8708 | 30.6891 | 2.894 | 2.3309 | 2.4426 | 2.9203 | 2.3797 | 2.4727 | 2.1586 | 2.2642 | 2.383 | 2.0209 | 0.87431 | 0.8895 | 78.45% | 74.42% | 87% | 76.11% | 82.48% | 72.87% | 86.03% | 5.1671 | 4.909 | 0.10707 | 0.095725
87 | MT 26 | R | 06/03/2023 00:14:55 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 26.089 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 4284.898 | 525.8116 | 14.2656 | 6463.3928 | 7101.0936 | 4202.6383 | 206.2797 | 0.12336 | 1.2065 | 1.1962 | 486.3739 | 3.5335 | 473.6568 | 0.097103 | 0.9879 | 9.6872 | 9.9788 | 28.048 | 3.0733 | 2.2087 | 2.2832 | 2.6812 | 2.696 | 2.1316 | 2.017 | 2.0221 | 2.6589 | 1.9497 | 0.89117 | 0.91208 | 70.99% | 74.05% | 87.36% | 75.42% | 96.73% | 72.43% | 84.06% | 4.7494 | 4.5586 | 0.14609 | 0.08439
88 | MT 27 | R | 06/03/2023 00:21:03 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 27.090 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 12290.216 | 1704.7488 | 29.9034 | 17410.0987 | 16171.658 | 12068.5986 | 733.8392 | 0.30382 | 2.7257 | 2.6016 | 1472.115 | 9.683 | 1274.2666 | 0.25533 | 2.361 | 24.8398 | 24.9139 | 72.3287 | 6.6712 | 5.7115 | 6.0217 | 7.1672 | 4.1722 | 5.7849 | 5.4283 | 5.4354 | 5.4076 | 5.2683 | 2.3074 | 2.3514 | 85.13% | 74.03% | 85.04% | 75.55% | 82.16% | 72.67% | 85.32% | 12.4902 | 11.8986 | 0.1834 | 0.17704
89 | MT 28 | R | 06/03/2023 00:27:06 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 28.091 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 13748.68 | 1689.5292 | 32.522 | 18895.0879 | 17998.3442 | 13362.4792 | 689.1893 | 0.32406 | 2.5894 | 2.6384 | 1608.5639 | 10.0255 | 1429.693 | 0.28189 | 2.3779 | 28.2912 | 27.2659 | 80.1424 | 7.7783 | 6.8868 | 6.5432 | 7.6258 | 5.2128 | 6.0895 | 6.0755 | 5.9121 | 6.3238 | 5.9183 | 2.4737 | 2.4263 | 77.61% | 75.05% | 81.58% | 74.1% | 80.03% | 72.15% | 84.7% | 13.7015 | 13.2527 | 0.19538 | 0.17245
90 | MT 29 | R | 06/03/2023 00:33:08 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 29.092 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 25930.7988 | 3315.7679 | 59.2385 | 35007.4459 | 30986.87 | 25936.1238 | 1380.7219 | 0.60663 | 5.4587 | 5.6832 | 3144.7884 | 19.8541 | 2626.8061 | 0.52079 | 5.0038 | 51.8753 | 52.5969 | 154.718 | 14.9214 | 12.7158 | 13.0098 | 15.0609 | 9.4267 | 12.3204 | 11.871 | 11.6849 | 12.2219 | 11.6484 | 4.9995 | 5.0585 | 78.6% | 72.28% | 81.11% | 71.42% | 84.91% | 74.36% | 81.08% | 26.9175 | 25.6433 | 0.33965 | 0.29001
91 | LBLANK | R | 06/03/2023 00:39:20 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LBLANK.093 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | -145.5986 | -49.4851 | -0.84614 | 1.9448 | 2279.0338 | -149.0433 | -0.70675 | 0.0091747 | 0.0054108 | 0.0031163 | -0.42463 | -0.0024772 | 0.023472 | 0.001005 | -0.0033192 | 0.0028118 | -0.0076276 | -0.21115 | -0.1254 | 0.015162 | -0.0071337 | -0.0075301 | 0.43478 | 0.014903 | -0.02441 | -0.0071842 | -0.02621 | -8.5498e-05 | -0.0024399 | -0.012072 | 83.34% | 73.5% | 84.73% | 75.99% | 85.38% | 73.36% | 86.5% | 1.9841e-05 | 0.00049455 | -0.0018493 | -0.00072881
92 | CCV 1ppb TRACE | R | 06/03/2023 00:45:23 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.094 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2047.6878 | 1948.7255 | 1.3815 | 2.2013 | 3767.647 | 1589.153 | 1598.6672 | 1.3709 | 1.2941 | 1.3767 | 1.0321 | 1.5752 | 1.2912 | 1.329 | 1.3189 | 1.2415 | 1.2681 | 1.042 | 1.0363 | 1.0971 | 1.0197 | 1.2447 | 1.5493 | 1.3729 | 1.24 | 1.196 | 1.0755 | 1.1923 | 1.1168 | 1.2005 | 83.65% | 70.54% | 84.57% | 73.28% | 80.18% | 70.99% | 83.35% | 1.0115 | 1.0069 | 0.78832 | 0.81479
93 | CCV 1ppb TRACE | R | 06/03/2023 00:51:55 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.095 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2146.4917 | 2009.1125 | 1.5307 | 1.4234 | 4216.6121 | 1634.9721 | 1633.7414 | 1.3741 | 1.3396 | 1.3082 | 0.91974 | 1.5133 | 1.4226 | 1.3757 | 1.3973 | 1.2453 | 1.2151 | 0.98749 | 1.0144 | 1.0212 | 0.97742 | 1.1631 | 1.4628 | 1.1878 | 1.1896 | 1.1879 | 1.057 | 0.98405 | 1.0508 | 1.1257 | 83.2% | 72.13% | 84.04% | 75.98% | 79.78% | 71.57% | 82.98% | 0.96441 | 0.96082 | 0.79491 | 0.80002
94 | LFB 10ppb | R | 06/03/2023 00:57:57 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LFB 10ppb.096 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 22743.1014 | 20507.6626 | 19.6136 | 20.8277 | 20007.7152 | 17912.0399 | 17261.8613 | 14.2062 | 14.006 | 13.1488 | 14.7685 | 13.5946 | 15.4391 | 14.1674 | 14.0358 | 12.2584 | 11.6318 | 15.9831 | 11.2699 | 11.6136 | 9.2992 | 11.0607 | 12.7432 | 11.166 | 11.5117 | 11.0525 | 11.8479 | 11.436 | 10.8598 | 10.589 | 82.05% | 72.71% | 74.97% | 68.56% | 73.14% | 66.51% | 83.18% | 9.7247 | 9.5267 | 6.8686 | 7.3993
95 | LFB 10ppb | R | 06/03/2023 01:04:27 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LFB 10ppb.097 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 22121.7484 | 19529.2105 | 17.9986 | 20.2756 | 19349.5194 | 17860.4334 | 16677.1287 | 13.5251 | 13.2079 | 12.7902 | 14.0186 | 14.3222 | 14.4788 | 13.5183 | 13.5772 | 11.7374 | 11.265 | 15.6674 | 11.4809 | 10.6821 | 10.2316 | 11.5979 | 12.5989 | 11.5237 | 11.9089 | 11.5022 | 11.783 | 11.4808 | 10.5915 | 11.0736 | 81.02% | 68.65% | 78.53% | 71.46% | 73.5% | 68.01% | 75.24% | 9.3144 | 9.2147 | 6.9899 | 7.3386
96 | CCV 10 ppb MAJORS | R | 06/03/2023 01:10:29 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.098 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 22589.6947 | 20393.5272 | 17.3366 | 19.5679 | 19425.7924 | 17763.5482 | 16839.7548 | 14.0308 | 13.6793 | 13.7998 | 14.2817 | 15.1421 | 13.7263 | 13.6078 | 13.8226 | 11.7902 | 12.0951 | 12.5491 | 11.2899 | 10.8352 | 9.8758 | 11.9667 | 12.4697 | 12.7502 | 12.2478 | 11.7033 | 11.4765 | 12.1988 | 11.2995 | 11.581 | 82.47% | 67.23% | 78.38% | 69.01% | 74.26% | 68.27% | 78.97% | 9.9651 | 9.826 | 7.0548 | 7.5393
97 | CCV 10 ppb MAJORS | R | 06/03/2023 01:17:01 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.099 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 21768.7225 | 19295.4051 | 17.1835 | 19.0631 | 18877.5652 | 17566.8589 | 16012.6827 | 13.6025 | 13.0867 | 13.4579 | 13.9046 | 14.8392 | 13.4556 | 13.3476 | 13.6075 | 11.8217 | 11.7167 | 12.0015 | 12.1097 | 10.6211 | 10.2884 | 11.8442 | 13.8678 | 11.8764 | 11.8966 | 11.4785 | 11.9035 | 11.4277 | 10.8188 | 11.2882 | 77.69% | 69.18% | 81% | 70.67% | 74.46% | 67.89% | 77.06% | 9.6592 | 9.5858 | 7.5939 | 7.5223
98 | MT 30 | R | 06/03/2023 01:23:04 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 30.100 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | -110.7365 | 221.1253 | 728.2531 | 2886.1545 | 4884.4676 | 23.8532 | 2048.7213 | 0.91961 | 0.6147 | 0.64576 | 768.3711 | 121.3975 | 585.3216 | 0.57087 | 1.006 | 70.9182 | 80.1046 | 154.8468 | 1.4167 | 1.4386 | 1.3133 | 1.5507 | 0.44731 | 0.42512 | 0.37839 | 0.44905 | 0.50002 | 0.44832 | 0.35519 | 0.30881 | 87.88% | 74.18% | 83.86% | 68.42% | 79.2% | 74.7% | 84.66% | 1.8206 | 1.7278 | 2.2988 | 2.4146
99 | MT 31 | R | 06/03/2023 01:29:15 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 31.101 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 67.1371 | 422.5248 | 548.3727 | 2843.9193 | 3847.9022 | 742.1862 | 1148.704 | 0.66841 | 0.46758 | 0.45224 | 429.6306 | 87.6994 | 434.4449 | 0.56428 | 0.64272 | 52.3942 | 51.031 | 160.4809 | 1.273 | 1.2283 | 1.1539 | 1.3629 | 0.4243 | 0.25557 | 0.3279 | 0.29876 | 0.35966 | 0.25233 | 0.17954 | 0.19105 | 85.8% | 75.76% | 84.55% | 76.69% | 79.83% | 71.05% | 86.61% | 1.7808 | 1.6966 | 1.6865 | 1.7046
100 | MT 31 LD | R | 06/03/2023 01:35:29 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 31 LD.102 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 79.2036 | 397.7084 | 567.784 | 2921.2983 | 4121.6521 | 781.8006 | 1187.5135 | 0.71018 | 0.46475 | 0.47577 | 442.7182 | 89.6347 | 442.5217 | 0.56217 | 0.64128 | 53.4535 | 51.9474 | 168.1468 | 1.3057 | 1.1802 | 1.1779 | 1.4009 | 0.45437 | 0.28631 | 0.27571 | 0.30633 | 0.40555 | 0.35221 | 0.18775 | 0.19691 | 84.7% | 72.32% | 82.51% | 73.06% | 78.81% | 70.96% | 83.14% | 1.8267 | 1.7531 | 1.6913 | 1.7004
101 | MT 31 LD | R | 06/03/2023 01:42:00 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 31 LD.103 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 78.5828 | 372.4812 | 550.5283 | 2913.8358 | 3936.957 | 713.0268 | 1122.1404 | 0.66631 | 0.48509 | 0.46498 | 441.1994 | 89.8688 | 426.6917 | 0.57672 | 0.60685 | 51.0588 | 51.8841 | 165.7256 | 1.3194 | 1.0225 | 1.1705 | 1.3978 | 0.39925 | 0.27937 | 0.30642 | 0.30079 | 0.32646 | 0.25441 | 0.18376 | 0.18537 | 85.63% | 72.49% | 84.83% | 74.25% | 73.35% | 69.86% | 83.69% | 1.8206 | 1.7338 | 1.6575 | 1.813
102 | MT 32 | R | 06/03/2023 01:48:02 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 32.104 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | -144.7487 | -47.6428 | 1.3164 | 15.509 | 2360.0329 | -141.8897 | 5.6901 | 0.0067194 | 0.0193 | 0.010786 | -0.37785 | 0.0031724 | 0.062826 | 0.0010276 | 0.0081659 | 0.050337 | 0.04171 | -0.17729 | -0.10322 | -0.0031069 | -0.0062945 | -0.0066437 | 0.22868 | 0.027702 | -0.034678 | -0.0081024 | -0.0027766 | -3.7398e-05 | 0.0019113 | 0.003399 | 81.74% | 72.34% | 83.23% | 74.57% | 74.51% | 67.67% | 82.3% | -0.00097726 | 0.00076153 | 0.028341 | 0.027134
103 | MT 32 LSP | R | 06/03/2023 01:54:03 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 32 LSP.105 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 59887.2643 | 50644.7811 | 15.872 | 32.7045 | 45903.8404 | 47309.0756 | 43655.7692 | 9.3293 | 8.7804 | 9.1232 | 11.2615 | 10.2735 | 11.0396 | 8.8125 | 12.7901 | 7.548 | 7.6739 | 30.6556 | 7.5337 | 7.017 | 6.3827 | 7.6321 | 8.4141 | 7.4448 | 7.897 | 7.7676 | 7.9748 | 7.5708 | 7.1118 | 7.299 | 73.8% | 63.4% | 73.33% | 63.93% | 68.35% | 58.53% | 73.31% | 6.1736 | 6.0817 | 4.5774 | 4.6939
104 | MT 32 LSP | R | 06/03/2023 02:00:34 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 32 LSP.106 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 58372.3945 | 49490.2435 | 16.4419 | 27.2185 | 44894.4671 | 45555.581 | 42167.8164 | 9.1397 | 8.7203 | 9.0414 | 11.1499 | 9.8701 | 11.038 | 8.8242 | 12.7897 | 7.6911 | 7.7355 | 30.9724 | 7.2768 | 7.2956 | 6.5845 | 7.5691 | 7.9927 | 7.786 | 7.8443 | 7.4632 | 7.3677 | 7.3296 | 7.181 | 7.1225 | 77.56% | 64.43% | 74.04% | 64.85% | 69.45% | 62.64% | 71.68% | 6.1848 | 6.1477 | 4.4876 | 4.6328
105 | MT 33 | R | 06/03/2023 02:06:37 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 33.107 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | -142.7512 | -47.7302 | 1.8161 | 13.5315 | 2534.3942 | -154.9333 | 11.3607 | 0.007671 | 0.017918 | 0.012695 | -0.36244 | 0.0020194 | 0.078175 | 0.0031046 | 0.012863 | 0.05066 | 0.054577 | 0.7638 | -0.1182 | -0.0031069 | -0.0066087 | -0.006872 | 0.37803 | 0.014903 | -0.030705 | -0.0082007 | -0.044485 | 0.0050517 | 0.0029657 | -0.0099611 | 85.08% | 74.42% | 87.06% | 74.99% | 69.21% | 71.93% | 87.88% | -0.00074733 | 0.00081791 | 0.021308 | 0.024079
106 | MT 34 | R | 06/03/2023 02:12:39 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 34.108 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | -142.178 | -49.4851 | 2.1566 | 15.1166 | 2761.4014 | -155.1564 | 5.323 | 0.0076619 | 0.028469 | 0.024526 | 3.6869 | 0.020232 | 4.0946 | 0.001539 | 0.045571 | 0.018414 | 0.006774 | -0.19065 | -0.16402 | 0.001452 | -0.0066964 | -0.0071322 | 0.4468 | 0.021236 | -0.026554 | -0.0081853 | -0.00071836 | -0.00013879 | -0.0071368 | -0.011437 | 82.42% | 73.98% | 85.47% | 76.28% | 81.22% | 72.55% | 84.23% | -0.00088286 | 0.00091382 | 0.01857 | 0.019687
107 | MT 35 | R | 06/03/2023 02:18:42 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 35.109 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | -146.0364 | -48.5898 | 1.4573 | 14.9942 | 2873.0389 | -154.5021 | 11.4599 | 0.0071812 | 0.018914 | 0.011039 | 0.16023 | 0.0017853 | 0.59007 | 0.0015291 | 0.0079942 | 0.010703 | 0.0047741 | -0.1682 | -0.20451 | 0.0014451 | -0.0060964 | -0.0064258 | 0.42152 | 0.034891 | -0.032558 | -0.0061885 | -0.0036452 | -0.0054505 | -0.008408 | -0.011446 | 81.79% | 71.18% | 85.94% | 78.51% | 81.26% | 71.47% | 80.43% | -0.00046169 | 0.00061959 | 0.020539 | 0.018983
108 | LBLANK | R | 06/03/2023 02:24:45 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LBLANK.110 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | -146.3403 | -49.4851 | -0.76116 | 1.5027 | 2955.6642 | -144.5824 | -0.69835 | 0.0076988 | 0.0086455 | 0.0011485 | -0.46784 | -0.0028974 | 0.0065413 | 0.00045 | 0.0070436 | 0.0020642 | -0.0095033 | -0.21796 | -0.10159 | -0.0031069 | -0.0076273 | -0.0081989 | 0.33343 | 0.014903 | -0.03055 | -0.010994 | -0.024012 | -0.0054505 | -0.00885 | -0.012493 | 84.41% | 72.5% | 84.44% | 73.52% | 74.9% | 71.65% | 82.62% | -0.0012151 | 0.00025078 | -0.0019169 | -0.00072586
109 | CCV 1ppb TRACE | R | 06/03/2023 02:30:48 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.111 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2114.1886 | 1935.993 | 1.3111 | 1.8749 | 4368.4181 | 1685.1854 | 1692.7601 | 1.429 | 1.3469 | 1.4085 | 1.0811 | 1.6008 | 1.4819 | 1.3633 | 1.4149 | 1.3133 | 1.2914 | 1.1048 | 1.0234 | 1.1213 | 1.004 | 1.2204 | 1.2548 | 1.3467 | 1.1259 | 1.1653 | 1.1173 | 1.0488 | 1.1063 | 1.1574 | 83.08% | 70.54% | 84% | 70.95% | 81.32% | 69.88% | 82.79% | 1.0032 | 1.0051 | 0.77883 | 0.72678
110 | CCV 1ppb TRACE | R | 06/03/2023 02:37:20 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.112 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2231.5756 | 2061.9625 | 1.7574 | 1.6109 | 4848.718 | 1729.6854 | 1729.0406 | 1.5125 | 1.432 | 1.4234 | 1.0527 | 1.5528 | 1.5316 | 1.4302 | 1.4569 | 1.3318 | 1.2927 | 1.0688 | 1.1164 | 1.2217 | 0.9724 | 1.1521 | 1.4405 | 1.2776 | 1.1988 | 1.1568 | 1.2388 | 1.2345 | 1.0896 | 1.1011 | 83.58% | 71.46% | 81.96% | 70.11% | 76.8% | 71.53% | 81.8% | 1.0047 | 1.0047 | 0.77226 | 0.78767
111 | CCV 10 ppb MAJORS | R | 06/03/2023 02:43:23 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.113 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 24000.1796 | 21306.2792 | 19.2387 | 19.6868 | 21033.2194 | 19175.1195 | 18411.3094 | 14.7288 | 14.2458 | 13.7585 | 13.992 | 15.0344 | 14.5708 | 14.5704 | 14.6214 | 12.557 | 11.8956 | 12.138 | 11.3734 | 11.395 | 9.6021 | 11.6359 | 12.9356 | 11.9372 | 11.7948 | 11.4913 | 11.2424 | 11.6494 | 10.9499 | 11.233 | 82.56% | 68.5% | 75.92% | 67.74% | 71.86% | 66.85% | 80.37% | 9.5604 | 9.4731 | 6.7632 | 7.5041
112 | CCV 10 ppb MAJORS | R | 06/03/2023 02:49:56 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.114 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 22860.7575 | 20634.5933 | 18.2955 | 19.3139 | 19326.5307 | 17904.7891 | 16952.1775 | 14.0346 | 13.5491 | 14.3133 | 14.6135 | 15.1383 | 13.9184 | 13.9297 | 13.828 | 11.9315 | 12.3968 | 12.5668 | 10.7862 | 10.649 | 10.0675 | 11.8767 | 11.6157 | 12.0555 | 12.0109 | 11.6614 | 10.6018 | 11.4443 | 11.4496 | 11.3598 | 86.8% | 68.61% | 80.76% | 66.92% | 73.18% | 67.76% | 78.26% | 10.0635 | 9.9163 | 6.5032 | 7.4578
113 | MT 36 | R | 06/03/2023 02:55:59 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 36.115 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | -138.8261 | -46.8185 | 1.5042 | 12.4404 | 2471.587 | -152.5087 | 21.1422 | 0.0090932 | 0.036496 | 0.028221 | -0.038993 | 0.0047601 | 0.35193 | 0.001851 | 0.0075976 | 0.018681 | 0.0080063 | -0.15337 | -0.093735 | 0.0011785 | -0.0071259 | -0.0076165 | 0.2949 | 0.014903 | -0.034669 | -0.011007 | 0.013328 | -0.0054505 | -0.0048643 | -0.0096318 | 85.82% | 72.12% | 86.48% | 72.84% | 79.71% | 72.6% | 81.85% | -0.00097943 | 0.00093873 | 0.018862 | 0.019869
114 | MT 37 | R | 06/03/2023 03:02:01 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 37.116 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | -140.2416 | -44.7635 | 1.5199 | 17.0547 | 3011.2968 | -133.9243 | 25.4643 | 0.0067013 | 0.018068 | 0.011265 | -0.34804 | 0.0081076 | 0.23325 | 0.00048963 | 0.020798 | 0.10779 | 0.093351 | -0.051341 | -0.073072 | 0.0018824 | -0.0066633 | -0.0069524 | 0.1898 | 0.021128 | -0.036776 | -0.0043526 | -0.080683 | -0.0054505 | -0.0097029 | -0.01114 | 87.26% | 73.41% | 81.65% | 74.56% | 80.52% | 73.46% | 85.32% | -0.00076081 | 0.0011405 | 0.019808 | 0.021373
115 | MT 38 | R | 06/03/2023 03:08:04 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 38.117 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | -144.4965 | -47.6679 | 1.0819 | 14.2231 | 2846.886 | -147.2849 | 5.3362 | 0.0081775 | 0.020448 | 0.014544 | -0.32728 | 0.001601 | 0.11684 | 0.00099439 | 0.0005993 | 0.027401 | 0.01953 | -0.17024 | -0.11902 | 0.0014821 | -0.0066717 | -0.0070585 | 0.4243 | 0.014903 | -0.024163 | -0.010042 | 0.024851 | -0.0054505 | -0.0094473 | -0.011318 | 83.44% | 72.34% | 85.73% | 72.99% | 69.01% | 75.08% | 83.16% | -0.00097828 | 0.00020832 | 0.019512 | 0.02359
116 | MT 39 | R | 06/03/2023 03:14:07 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 39.118 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | -134.7712 | -47.6299 | 1.6257 | 15.6344 | 2650.2677 | -157.7532 | 15.8529 | 0.0081898 | 0.01626 | 0.01702 | 0.64455 | 0.0051882 | 1.0537 | 0.0011666 | -0.0018028 | 0.02747 | 0.018787 | -0.091264 | -0.17617 | 0.0013288 | -0.0073028 | -0.0077536 | 0.45038 | 0.021419 | -0.028402 | -0.0041884 | -0.053277 | 0.01114 | -0.010167 | -0.011165 | 77.39% | 71.92% | 84.27% | 72.33% | 77.39% | 70.46% | 83.74% | -0.0010944 | 0.00060832 | 0.024216 | 0.020295
117 | MT 40 | R | 06/03/2023 03:20:10 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 40.119 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | -130.1622 | -33.4273 | 1.6128 | 15.8166 | 2505.0298 | -145.3452 | 82.7629 | 0.0071706 | 0.011932 | 0.0088556 | -0.2576 | 0.036936 | 0.099912 | 0.00061681 | 0.0031326 | 0.043908 | 0.031512 | -0.18111 | -0.020935 | 0.0057309 | 0.0015706 | 0.0031237 | 0.081057 | 0.014903 | -0.030214 | -0.0070302 | -0.03282 | 0.00014006 | -0.0024888 | -0.0023504 | 90.77% | 70.99% | 86.18% | 70.25% | 76.87% | 68.6% | 83.63% | -0.0010904 | 0.00034364 | 0.016941 | 0.019652
118 | MT 31 | R | 06/03/2023 03:26:12 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 31 .120 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 83.5412 | 466.9858 | 602.5057 | 3024.0423 | 4750.0796 | 754.125 | 1249.7838 | 0.73216 | 0.47644 | 0.43272 | 433.9076 | 90.3855 | 454.3924 | 0.58535 | 0.64747 | 54.4451 | 50.5312 | 162.0076 | 1.2794 | 1.1534 | 1.2055 | 1.4148 | 0.37912 | 0.3112 | 0.30343 | 0.3073 | 0.32009 | 0.26047 | 0.17255 | 0.18702 | 86.63% | 70.82% | 80.75% | 72.78% | 74.91% | 68.26% | 80.3% | 1.7646 | 1.677 | 1.6107 | 1.6969
119 | MT 31 LD | R | 06/03/2023 03:32:23 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 31 LD.121 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 73.1438 | 388.7412 | 570.0477 | 3129.5429 | 4051.3574 | 715.5443 | 1194.2733 | 0.65451 | 0.44693 | 0.47966 | 445.8888 | 94.9684 | 436.3192 | 0.55864 | 0.65944 | 51.6342 | 51.9486 | 166.9469 | 1.1992 | 1.172 | 1.2218 | 1.4658 | 0.39034 | 0.34419 | 0.30823 | 0.31203 | 0.25357 | 0.38872 | 0.18854 | 0.19889 | 84.19% | 68.59% | 83.37% | 70.24% | 85.75% | 67.3% | 79.51% | 1.7986 | 1.7487 | 1.6464 | 1.4671
120 | MT 31 LD | R | 06/03/2023 03:38:54 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 31 LD.122 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 67.038 | 375.452 | 562.7079 | 3054.865 | 4198.8439 | 765.9647 | 1181.7597 | 0.68464 | 0.47282 | 0.41395 | 389.7312 | 91.0084 | 449.3221 | 0.57222 | 0.61579 | 53.1815 | 45.4802 | 146.8858 | 1.2322 | 1.2036 | 1.1556 | 1.3834 | 0.21561 | 0.36417 | 0.25757 | 0.28355 | 0.20481 | 0.33233 | 0.15981 | 0.18743 | 79.53% | 69.61% | 81.02% | 75.45% | 73.92% | 68.14% | 80.56% | 1.59 | 1.5282 | 1.4726 | 1.5959
121 | MT 32 | R | 06/03/2023 03:44:56 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 32.123 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | -140.5046 | -49.4851 | 1.2009 | 14.5063 | 1944.0591 | -160.3118 | 3.9832 | 0.0067412 | 0.021352 | 0.012143 | -0.35325 | 0.0035975 | 0.06037 | 0.00086552 | 0.0086911 | 0.054733 | 0.048041 | -0.1888 | -0.089208 | -0.0031069 | -0.0064072 | -0.0067431 | 0.27347 | 0.021699 | -0.032235 | -0.0079112 | 0.015592 | -0.0054505 | 0.00081964 | 0.00024197 | 79.85% | 69% | 79.82% | 68.88% | 75.21% | 68.75% | 78.48% | -0.00095934 | 0.00077851 | 0.027026 | 0.025532
122 | MT 32 LSP | R | 06/03/2023 03:50:58 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 32 LSP.124 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 59198.2805 | 50868.7813 | 16.8765 | 20.9067 | 46657.696 | 46604.017 | 44772.8246 | 9.2949 | 8.9604 | 9.0711 | 11.2355 | 10.314 | 11.0842 | 8.797 | 12.9905 | 7.6546 | 7.6865 | 30.7115 | 7.4744 | 6.8332 | 6.3044 | 7.7194 | 8.3895 | 7.9213 | 7.9794 | 7.628 | 7.9449 | 7.5412 | 7.1191 | 7.3003 | 73.87% | 60.74% | 72.33% | 61.81% | 59.59% | 81.98% | 71.96% | 6.2121 | 6.0604 | 4.5079 | 5.2764
123 | MT 32 LSP | R | 06/03/2023 03:57:29 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 32 LSP.125 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 56997.9224 | 49732.8893 | 15.7411 | 28.3522 | 46033.9925 | 46333.5436 | 42233.8606 | 9.1036 | 8.8745 | 8.2322 | 10.3291 | 10.0136 | 10.9836 | 8.749 | 12.7293 | 7.6139 | 7.1273 | 28.649 | 6.6981 | 6.9968 | 6.2945 | 7.5056 | 7.4153 | 7.6028 | 7.5414 | 7.4016 | 6.727 | 7.2523 | 6.5845 | 7.1769 | 77.05% | 62.96% | 73.88% | 65.87% | 61.67% | 61.54% | 72.69% | 5.695 | 5.6849 | 4.0294 | 4.9599
124 | LBLANK | R | 06/03/2023 04:03:32 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LBLANK.126 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | -144.6105 | -48.6235 | -0.82122 | 0.42084 | 2439.7664 | -161.6167 | -0.8722 | 0.0071738 | 0.010563 | 0.0047714 | -0.48381 | -0.0025897 | 0.002076 | 0.0011311 | 0.0039121 | 0.00098345 | -0.0075572 | -0.22481 | -0.15217 | 0.0055406 | -0.0058034 | -0.0059339 | 0.42485 | 0.014903 | -0.034722 | -0.0042137 | -0.04304 | -0.0054505 | -0.0037191 | -0.011543 | 79.72% | 72.56% | 88.09% | 74.29% | 80.24% | 70.79% | 84.53% | -0.0011022 | 0.00048136 | -0.0019962 | -0.00075634
125 | CCV 1ppb TRACE | R | 06/03/2023 04:09:35 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.127 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2169.5302 | 2077.1071 | 1.8108 | 1.6366 | 4667.7849 | 1628.9883 | 1680.7396 | 1.4254 | 1.4072 | 1.3982 | 0.99482 | 1.5712 | 1.4668 | 1.4546 | 1.4535 | 1.2922 | 1.2846 | 1.0523 | 1.0545 | 1.1446 | 0.95661 | 1.1729 | 1.4275 | 1.1772 | 1.1383 | 1.1683 | 1.1016 | 0.86819 | 1.0974 | 1.132 | 86.15% | 73.55% | 83.36% | 72.87% | 78.33% | 73.93% | 87.1% | 0.96462 | 0.97355 | 0.75007 | 0.77849
126 | CCV 1ppb TRACE | R | 06/03/2023 04:16:07 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.128 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2136.6886 | 1976.8154 | 1.7489 | 1.5025 | 4491.1948 | 1681.6968 | 1675.864 | 1.4986 | 1.4337 | 1.4723 | 1.1102 | 1.5598 | 1.4883 | 1.4067 | 1.3914 | 1.2773 | 1.3201 | 1.1024 | 1.0608 | 1.102 | 0.97129 | 1.1798 | 1.4626 | 1.045 | 1.139 | 1.1908 | 1.0564 | 1.2602 | 1.1131 | 1.1286 | 84.19% | 71.22% | 85.55% | 68.65% | 74.48% | 69.67% | 83.54% | 0.99318 | 1.0264 | 0.75077 | 0.79293
127 | LFB 10ppb | R | 06/03/2023 04:22:09 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LFB 10ppb.129 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 22092.3553 | 19429.9035 | 18.5431 | 21.0107 | 19544.9564 | 17317.922 | 16680.2598 | 14.0329 | 13.5186 | 13.3624 | 14.7285 | 14.641 | 14.6787 | 13.5133 | 13.7731 | 11.6876 | 11.7126 | 16.2393 | 11.5546 | 10.5749 | 9.6165 | 11.6384 | 12.4886 | 11.9852 | 11.7756 | 11.5678 | 11.6554 | 11.2988 | 10.8439 | 11.0488 | 81.1% | 66.1% | 77.4% | 65.89% | 68.95% | 64.48% | 77.34% | 9.4908 | 9.4459 | 6.5651 | 7.3977
128 | LFB 10ppb | R | 06/03/2023 04:28:39 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LFB 10ppb.130 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 21552.9202 | 18945.8218 | 19.5101 | 21.6908 | 19307.5939 | 16896.9813 | 16158.8223 | 13.8152 | 13.0991 | 13.9098 | 15.1718 | 15.1977 | 14.3208 | 13.4938 | 13.7069 | 11.6415 | 11.8928 | 16.344 | 11.5375 | 11.1034 | 10.0911 | 12.0182 | 12.8541 | 11.374 | 12.2237 | 12.0905 | 11.7723 | 11.6705 | 11.1487 | 11.5015 | 80.12% | 65.05% | 78.89% | 65.56% | 69.78% | 66.03% | 74.9% | 9.7282 | 9.5247 | 6.7343 | 7.2543
129 | CCV 10 ppb MAJORS | R | 06/03/2023 04:34:42 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.131 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 21093.3047 | 18897.1161 | 18.4399 | 20.276 | 19070.8214 | 16932.2362 | 15843.9077 | 13.5792 | 12.7624 | 14.4052 | 14.519 | 20.7953 | 13.133 | 13.0339 | 13.3477 | 11.3365 | 12.2772 | 12.5252 | 11.4074 | 10.3838 | 13.2541 | 16.3241 | 12.7961 | 16.5329 | 16.4322 | 15.7281 | 11.567 | 16.0732 | 11.3664 | 15.6033 | 80.85% | 51.42% | 83.32% | 65.69% | 70.42% | 65.61% | 61.51% | 9.8015 | 9.7809 | 6.6986 | 7.4315
130 | CCV 10 ppb MAJORS | R | 06/03/2023 04:41:13 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.132 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 21761.6891 | 19380.299 | 18.1671 | 19.6343 | 19339.3999 | 17668.7693 | 16603.32 | 14.1108 | 13.3492 | 14.2487 | 14.9199 | 15.1752 | 13.8782 | 13.6736 | 13.8828 | 11.8883 | 12.3379 | 12.3305 | 11.4732 | 10.8783 | 9.7348 | 11.7961 | 12.6374 | 12.3362 | 11.7149 | 11.8037 | 11.4107 | 11.5964 | 11.3804 | 11.4001 | 82.84% | 66.19% | 79.11% | 66.44% | 68.24% | 64.95% | 77.56% | 9.77 | 9.817 | 6.5307 | 7.9157
131 | MT 33 | R | 06/03/2023 04:47:16 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 33.133 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | -144.1848 | -48.5966 | 1.6614 | 12.8542 | 2663.8684 | -157.0125 | 9.2859 | 0.0066889 | 0.018093 | 0.011816 | -0.38767 | 0.001292 | 0.090013 | 0.0029399 | 0.021779 | 0.05714 | 0.055856 | -0.21232 | -0.13866 | 0.0056633 | -0.0071523 | -0.0076247 | 0.35949 | 0.014903 | -0.026322 | -0.0081015 | -0.0015029 | -0.0054505 | 0.0023444 | -0.0087058 | 85.08% | 72.48% | 86.76% | 72.19% | 80.12% | 69.94% | 82.92% | -0.00097609 | 0.00076346 | 0.020306 | 0.020036
132 | MT 34 | R | 06/03/2023 04:53:19 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 34.134 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | -139.8359 | -47.7653 | 1.2166 | 15.18 | 2413.4705 | -153.9784 | 8.8892 | 0.0081279 | 0.027228 | 0.022024 | 3.547 | 0.023692 | 3.9087 | 0.0016985 | 0.029443 | 0.016106 | 0.0036195 | -0.19201 | -0.10293 | 0.0014472 | -0.0075884 | -0.0080568 | 0.34761 | 0.027839 | -0.034675 | -0.0081149 | -0.0068125 | -0.0054505 | -0.0067535 | -0.011695 | 84.82% | 72.62% | 87.93% | 73.78% | 77.99% | 69.88% | 86.12% | -0.0010882 | 0.00092831 | 0.018388 | 0.018767
133 | MT 35 | R | 06/03/2023 04:59:21 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 35.135 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | -146.2957 | -49.4851 | 1.6412 | 15.3975 | 2722.4985 | -152.9544 | 5.7575 | 0.0067262 | 0.020743 | 0.015156 | 0.21904 | 0.0017489 | 0.7646 | 0.00046485 | 0.0082274 | 0.018128 | 0.0064002 | -0.15915 | -0.067794 | 0.010943 | -0.0068913 | -0.0072599 | 0.25266 | 0.014903 | -0.034567 | -0.0030236 | 0.0036586 | 0.00016959 | -0.0078448 | -0.011495 | 83.32% | 70.29% | 83.04% | 68.09% | 75.94% | 69.57% | 82.43% | -0.00095241 | 0.00073003 | 0.018863 | 0.018729
134 | MT 36 | R | 06/03/2023 05:05:25 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 36.136 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | -143.2209 | -48.5857 | 1.3364 | 14.9673 | 2686.8262 | -152.2694 | 7.8563 | 0.0067211 | 0.029478 | 0.028491 | -0.016615 | 0.0038097 | 1.037 | 0.0011945 | 0.0038847 | 0.015518 | 0.0094586 | -0.16028 | -0.084038 | 0.0060886 | -0.006711 | -0.0069625 | 0.34936 | 0.021467 | -0.030134 | -0.0079939 | -0.045816 | 0.0056974 | -0.0082046 | -0.01081 | 77.68% | 70.36% | 83.61% | 69.55% | 76.59% | 69.71% | 84.08% | -0.001084 | 0.00052797 | 0.019232 | 0.01992
135 | MT 37 | R | 06/03/2023 05:11:28 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 37.137 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | -142.3234 | -49.4851 | 1.688 | 18.6929 | 2664.7185 | -147.7208 | 24.8616 | 0.0072205 | 0.015679 | 0.0098111 | -0.36234 | 0.0088126 | 0.18426 | 8.152e-05 | 0.011156 | 0.10984 | 0.083711 | -0.065297 | -0.11469 | 0.0015343 | -0.0073583 | -0.0077437 | 0.36634 | 0.035954 | -0.036586 | -0.0088214 | -0.032594 | -0.0054505 | -0.0096882 | -0.011173 | 78.43% | 67.54% | 82.84% | 80.26% | 76.51% | 68.74% | 81.08% | -0.0012092 | 0.001004 | 0.023374 | 0.020876
136 | MT 38 | R | 06/03/2023 05:17:31 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 38.138 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | -145.9849 | -46.7705 | 1.5713 | 16.214 | 2981.6137 | -151.3022 | 1.3542 | 0.0071949 | 0.022751 | 0.014046 | -0.32481 | 0.0020243 | 0.19475 | 0.0021002 | 0.0097205 | 0.02356 | 0.017619 | -0.17052 | -0.17556 | 0.0060034 | -0.0079559 | -0.008591 | 0.31306 | 0.014903 | -0.032497 | -0.009034 | -0.069323 | 2.4275e-05 | -0.0097384 | -0.011372 | 80.92% | 71.43% | 84.81% | 68.43% | 74.67% | 69.15% | 81.4% | -0.00083264 | 0.00067931 | 0.018568 | 0.019137
137 | MT 39 | R | 06/03/2023 05:23:33 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 39.139 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | -134.5654 | -47.5824 | 1.8089 | 15.5881 | 2895.9019 | -150.5773 | 10.404 | 0.0072502 | 0.021796 | 0.018083 | 0.75999 | 0.0059778 | 1.1598 | 0.0018257 | 0.019226 | 0.030593 | 0.020001 | -0.088074 | -0.1151 | 0.0016566 | -0.0065536 | -0.0068192 | 0.28764 | 0.023236 | -0.031889 | -0.004136 | 0.066347 | -0.0054505 | -0.0098128 | -0.011514 | 81.46% | 62.72% | 80.73% | 67.35% | 74.45% | 69.28% | 73.8% | -0.00082737 | 0.00049034 | 0.019165 | 0.01946
138 | MT 40 | R | 06/03/2023 05:29:36 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 40.140 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | -131.7014 | -30.8969 | 1.7648 | 16.0031 | 3040.3577 | -152.8283 | 68.5884 | 0.0072661 | 0.023501 | 0.013608 | -0.23487 | 0.035183 | 0.35471 | 0.0010606 | 0.0039844 | 0.042751 | 0.032599 | -0.18283 | -0.11061 | 0.0070286 | 0.0011609 | 0.0022731 | 0.21148 | 0.021412 | -0.038868 | -0.0080261 | -0.024686 | 0.0057302 | -0.0015225 | -0.0034334 | 80.06% | 70.68% | 78.81% | 68.35% | 72.56% | 67.85% | 81.33% | -0.00095047 | 0.00036892 | 0.018976 | 0.019604
139 | MT 32 B | R | 06/03/2023 05:35:38 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 32 B.141 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | -145.0978 | -48.5802 | 0.86127 | 13.7504 | 2911.3754 | -150.9518 | -2.7367 | 0.0072459 | 0.02172 | 0.014899 | -0.32135 | 0.00077306 | 0.17932 | 0.00083963 | 0.025931 | 0.049526 | 0.040111 | -0.18887 | -0.091773 | 0.0014342 | -0.0068012 | -0.0070993 | 0.26294 | 0.014903 | -0.038868 | -0.0058054 | -0.0027996 | -0.0054505 | -0.007516 | -0.0098047 | 80.6% | 67.15% | 82.42% | 66.7% | 74.77% | 68.53% | 79.36% | -0.00096132 | 0.00089346 | 0.024929 | 0.026179
140 | MT 32 B LD | R | 06/03/2023 05:41:40 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 32 B LD.142 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | -143.7237 | -48.5325 | 0.81155 | 15.4575 | 2910.3949 | -139.9721 | 8.4222 | 0.0067318 | 0.017968 | 0.013469 | -0.40689 | -0.0012858 | 0.076383 | 0.00048013 | 0.0036544 | 0.068218 | 0.049683 | -0.11571 | -0.063588 | 0.0017472 | -0.007181 | -0.0075861 | 0.18421 | 0.014903 | -0.036635 | -0.008893 | -0.016383 | 0.00031832 | -0.0093333 | -0.0092452 | 81.07% | 67.85% | 79.67% | 67.05% | 72.47% | 61.56% | 80.17% | -0.00044264 | 0.00060491 | 0.029564 | 0.02875
141 | LBLANK | R | 06/03/2023 05:47:43 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LBLANK.143 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | -147.4082 | -49.4851 | -0.88719 | 1.3695 | 3009.3505 | -153.8548 | -2.6267 | 0.0078023 | 0.011657 | 0.003746 | -0.47825 | -0.0019548 | 0.066039 | 0.00028853 | 0.0022234 | -0.0017249 | -0.0092517 | -0.22657 | -0.043734 | -0.0031069 | -0.0079101 | -0.0084986 | 0.079404 | 0.028747 | -0.034374 | -0.010947 | 0.005303 | 0.00040536 | -0.011249 | -0.012048 | 87.32% | 67.8% | 79.28% | 70.62% | 74.86% | 67.34% | 79.17% | -0.00095466 | 0.00014374 | -0.0022795 | -0.00079824
142 | CCV 1ppb TRACE | R | 06/03/2023 05:53:46 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.144 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2205.8349 | 2087.0744 | 1.5678 | 2.1631 | 4744.8351 | 1625.2449 | 1926.3046 | 1.4815 | 1.4635 | 1.4636 | 1.1267 | 1.6091 | 1.5448 | 1.4574 | 1.5002 | 1.325 | 1.3135 | 1.0794 | 1.115 | 0.99725 | 1.0021 | 1.2258 | 1.5632 | 1.2048 | 1.1649 | 1.1892 | 1.0693 | 1.1787 | 1.0959 | 1.1221 | 80.03% | 67.55% | 79% | 65.49% | 76.95% | 66.62% | 79.85% | 0.99635 | 0.98862 | 0.77162 | 0.73371
143 | CCV 1ppb TRACE | R | 06/03/2023 06:00:18 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.145 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2164.876 | 2043.9683 | 1.7483 | 1.7014 | 4810.6832 | 1730.5175 | 1740.5768 | 1.4799 | 1.4132 | 1.4158 | 1.0982 | 1.6222 | 1.5524 | 1.4099 | 1.4921 | 1.2996 | 1.2976 | 1.0626 | 1.0805 | 1.1286 | 1.0301 | 1.2456 | 1.5338 | 1.2461 | 1.22 | 1.1741 | 1.0908 | 1.1735 | 1.0968 | 1.1693 | 80.28% | 66.07% | 80.44% | 65.78% | 71.02% | 66.11% | 77.17% | 0.9788 | 0.96602 | 0.74461 | 0.78011
144 | CCV 10 ppb MAJORS | R | 06/03/2023 06:06:21 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.146 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 23421.8549 | 21366.2416 | 19.7732 | 20.4635 | 20739.4614 | 18765.838 | 17788.2699 | 14.8099 | 14.3279 | 14.0774 | 14.3105 | 15.2628 | 14.539 | 14.6014 | 14.6472 | 12.3025 | 11.9855 | 12.3486 | 11.8731 | 11.5123 | 9.8349 | 12.0184 | 12.9632 | 11.7748 | 12.1936 | 11.8886 | 11.966 | 11.3574 | 10.997 | 11.4155 | 76.69% | 63.08% | 72.91% | 62.65% | 67.06% | 61.8% | 74.57% | 9.6105 | 9.5282 | 6.9007 | 7.4512
145 | CCV 10 ppb MAJORS | R | 06/03/2023 06:12:53 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.147 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 22335.7108 | 20007.3012 | 19.9551 | 19.6122 | 20925.1471 | 17941.292 | 17143.2219 | 14.2424 | 13.8579 | 14.0292 | 14.0306 | 15.5376 | 14.3482 | 13.928 | 14.1099 | 12.1434 | 12.1387 | 12.3662 | 11.3409 | 11.0035 | 10.1228 | 11.8766 | 12.6111 | 12.3797 | 12.3669 | 11.7063 | 11.0055 | 11.4094 | 11.055 | 11.3472 | 79.46% | 64.21% | 75.62% | 63.64% | 67.53% | 65.41% | 72.84% | 9.736 | 9.5107 | 6.6961 | 7.4202
146 | MT 33 B | R | 06/03/2023 06:18:55 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 33 B.148 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | -145.8824 | -46.6451 | 1.0636 | 13.0248 | 2594.679 | -160.7379 | 8.3126 | 0.0087229 | 0.019625 | 0.012539 | -0.4581 | 0.00046016 | 0.12024 | 0.0021744 | 0.015497 | 0.043786 | 0.042829 | -0.2049 | -0.054838 | 0.0063233 | -0.0078587 | -0.0084222 | 0.20159 | 0.021652 | -0.032296 | -0.01197 | -0.021156 | -0.0054505 | -0.0065505 | -0.0081916 | 83.29% | 70.37% | 82.03% | 69.96% | 74.17% | 61.67% | 82.9% | -0.00047903 | 0.00065078 | 0.019249 | 0.020937
147 | MT 33 B LSP | R | 06/03/2023 06:24:56 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 33 B LSP.149 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 60037.5792 | 52232.7355 | 50.6883 | 64.7614 | 48452.8172 | 48010.6355 | 46679.3948 | 38.0646 | 35.8324 | 32.8088 | 34.5058 | 38.9653 | 36.9867 | 36.3855 | 35.9194 | 30.4122 | 26.7644 | 26.9735 | 28.0795 | 27.3384 | 24.0109 | 28.8613 | 29.9328 | 28.8019 | 28.5682 | 27.7929 | 27.2441 | 27.3911 | 25.922 | 29.6321 | 81.27% | 62.63% | 72.48% | 68.6% | 65.79% | 61.05% | 72.81% | 21.218 | 20.9679 | 14.6263 | 17.9222
148 | MT 34 B | R | 06/03/2023 06:30:59 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 34 B.150 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | -144.091 | -46.7646 | 0.49718 | 12.6368 | 2407.141 | -141.8676 | 7.526 | 0.0096642 | 0.021654 | 0.019091 | -0.33959 | -0.0015364 | 0.14969 | 0.0017216 | 0.048942 | 0.024742 | 0.022368 | -0.18031 | -0.091465 | 0.019522 | -0.0038214 | -0.0034514 | 0.42889 | 0.028103 | -0.030256 | -0.0070095 | -0.010672 | -0.0054505 | 0.023746 | 0.011118 | 81.94% | 70.34% | 84.82% | 70.7% | 68.21% | 66.71% | 82.45% | -0.00060798 | 0.0014914 | 0.0201 | 0.02244
149 | MT 35 B | R | 06/03/2023 06:37:01 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 35 B.151 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | -142.3015 | -48.6271 | 1.7461 | 13.149 | 2527.6116 | -154.5867 | 11.4399 | 0.0081263 | 0.024197 | 0.013701 | -0.44083 | -0.0016059 | 0.070743 | 0.00083997 | 0.010506 | 0.014452 | 0.011512 | -0.21554 | -0.12167 | 0.0011785 | -0.0066022 | -0.0068769 | 0.27316 | 0.0215 | -0.030337 | -0.0060652 | -9.4583e-05 | -0.0054505 | -0.0053959 | -0.010558 | 83.51% | 71.48% | 86.86% | 71.37% | 76.69% | 70.46% | 83.16% | -0.00073469 | 0.00042543 | 0.018076 | 0.018867
150 | MT 36 B | R | 06/03/2023 06:43:03 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 36 B.152 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | -143.1065 | -49.4851 | 1.645 | 12.3528 | 2812.1803 | -150.0478 | -0.4303 | 0.0082724 | 0.036789 | 0.029102 | -0.089439 | 0.0016045 | 0.43029 | 0.0013992 | 0.012072 | 0.022717 | 0.01095 | 4.1816 | -0.12225 | 0.0064338 | -0.0072011 | -0.0075538 | 0.25412 | 0.014903 | -0.038868 | -0.0080107 | 0.028903 | 0.00015418 | 0.049177 | 0.048569 | 78.97% | 70.2% | 82.08% | 70.26% | 72.13% | 71.1% | 83.8% | -0.0009633 | 0.00054837 | 0.018622 | 0.019849
151 | MT 37 B | R | 06/03/2023 06:49:06 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 37 B.153 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | -129.8997 | -48.5469 | 0.90672 | 13.896 | 2647.189 | -158.1526 | 7.9558 | 0.0077171 | 0.017862 | 0.016604 | -0.44311 | 1.0653e-05 | -0.004652 | 0.0010303 | 0.011807 | 0.025949 | 0.017099 | -0.19869 | -0.11583 | 0.0016011 | -0.0081821 | -0.0088374 | 0.29819 | 0.021603 | -0.028159 | -0.010967 | -0.043493 | 8.3653e-05 | -0.0087965 | -0.010645 | 79.15% | 69.92% | 83.04% | 66.02% | 74.06% | 67.58% | 80.5% | -0.00068102 | 0.00040154 | 0.018762 | 0.01959
152 | MT 38 B | R | 06/03/2023 06:55:09 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 38 B.154 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | -115.919 | -48.5439 | 0.90711 | 13.416 | 2514.8662 | -149.9235 | 5.6562 | 0.0076907 | 0.020634 | 0.015132 | -0.43008 | -0.0019892 | 0.0045858 | 0.001577 | 0.010993 | 0.014237 | 0.0080447 | -0.19643 | -0.082601 | -0.0031069 | -0.0082372 | -0.0088143 | 0.23274 | 0.021621 | -0.034424 | -0.010927 | -0.015497 | 0.0060749 | -0.0050157 | -0.0061113 | 81.65% | 68.54% | 83.43% | 72.18% | 76.69% | 67.59% | 83.44% | -0.00087741 | 0.0006348 | 0.019958 | 0.019269
153 | MT 39 B | R | 06/03/2023 07:01:12 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 39 B.155 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | -144.1875 | -47.5551 | 0.79419 | 13.132 | 2808.6213 | -149.7511 | 12.5645 | 0.0088225 | 0.024779 | 0.015327 | -0.35789 | -0.0015255 | 0.15413 | 0.00028775 | 0.013561 | 0.019459 | 0.007667 | -0.19099 | -0.092568 | 0.0063991 | -0.0083597 | -0.0089997 | 0.2994 | 0.021525 | -0.0324 | -0.0089705 | -0.036218 | 0.00015208 | 0.0026576 | 0.0018925 | 80.2% | 70.2% | 80.67% | 75.37% | 73.82% | 70.58% | 84.1% | -0.00058069 | 0.00033827 | 0.019918 | 0.019003
154 | MT 40 B | R | 06/03/2023 07:07:15 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 40 B.156 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | -145.849 | -47.5025 | 1.1167 | 12.9709 | 2811.8769 | -157.9235 | 8.345 | 0.0078325 | 0.02096 | 0.011695 | -0.40188 | -0.0017854 | 0.07357 | 0.0033634 | 0.010359 | 0.049049 | 0.03284 | -0.19278 | -0.088147 | 0.0018642 | -0.0065638 | -0.0068346 | 0.19634 | 0.021811 | -0.032141 | -0.0057686 | -0.051613 | 0.00034068 | -0.0099003 | -0.010252 | 80.77% | 67.68% | 80.1% | 69.49% | 73.99% | 66.78% | 80.15% | -0.00084652 | 0.00035425 | 0.020678 | 0.018126
155 | MT 32 C | R | 06/03/2023 07:13:17 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 32 C.157 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | -146.6446 | -48.5426 | 0.8692 | 13.2677 | 2576.878 | -154.4551 | 1.6925 | 0.0082932 | 0.020278 | 0.014372 | -0.43667 | 8.9921e-05 | 0.01308 | 0.00066745 | 0.0097843 | 0.047896 | 0.041455 | -0.1611 | -0.10618 | 0.0064685 | -0.0076678 | -0.0081116 | 0.15614 | 0.014903 | -0.034529 | -0.0080491 | -0.015383 | -0.0054505 | -0.0084536 | -0.010598 | 79.89% | 69.84% | 80.86% | 66.07% | 72.96% | 66.02% | 84.83% | -0.00068688 | 0.00062077 | 0.025313 | 0.025226
156 | LBLANK | R | 06/03/2023 07:19:20 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LBLANK.158 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | -147.857 | -49.4851 | -0.82108 | 1.2593 | 2577.0718 | -148.8617 | -4.861 | 0.0082894 | 0.0083397 | 0.0031559 | -0.49735 | -0.0030791 | -0.066008 | 0.00046874 | 0.0023377 | 0.00014746 | -0.0088015 | -0.22328 | -0.12393 | 0.0014155 | -0.0081964 | -0.0087918 | 0.18021 | 0.014903 | -0.03214 | -0.0088262 | -0.06201 | 0.0064591 | -0.010277 | -0.012816 | 79.46% | 66.8% | 81.75% | 66.59% | 72.01% | 67.23% | 79.44% | -0.001204 | 0.00039887 | -0.0021029 | -0.00087167
157 | CCV 1ppb TRACE | R | 06/03/2023 07:25:23 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.159 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2228.141 | 2157.1104 | 1.5024 | 2.0579 | 4498.6333 | 1689.8757 | 1771.6199 | 1.4935 | 1.431 | 1.4461 | 1.0944 | 1.5717 | 1.5615 | 1.4555 | 1.4413 | 1.3101 | 1.307 | 1.0783 | 1.068 | 1.1073 | 1.0091 | 1.202 | 1.561 | 1.1841 | 1.2241 | 1.156 | 1.0996 | 1.1393 | 1.0875 | 1.1248 | 78.79% | 67.83% | 78.94% | 65.96% | 71.04% | 64.86% | 78.05% | 0.97166 | 0.98536 | 0.77275 | 0.79166
158 | CCV 1ppb TRACE | R | 06/03/2023 07:31:56 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.160 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2259.4002 | 2060.9943 | 1.9346 | 1.6836 | 5002.6423 | 1707.8261 | 1797.7708 | 1.5021 | 1.4091 | 1.6226 | 1.287 | 1.5681 | 1.533 | 1.4472 | 1.5517 | 1.3739 | 1.4609 | 1.2397 | 1.0986 | 1.1559 | 0.9536 | 1.1959 | 1.2941 | 1.2704 | 1.1499 | 1.1814 | 1.1382 | 1.226 | 1.2247 | 1.1593 | 78.56% | 66.67% | 78.3% | 59.14% | 68.93% | 65.16% | 80.67% | 1.1178 | 1.1136 | 0.75428 | 0.78438
159 | LFB 10ppb | R | 06/03/2023 07:37:58 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LFB 10ppb.161 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 22008.8046 | 19739.8604 | 20.3066 | 21.6092 | 19930.9088 | 17005.7439 | 16892.2286 | 14.1463 | 13.5856 | 13.7366 | 14.7745 | 14.3862 | 14.7431 | 13.7783 | 14.0832 | 11.8005 | 11.7028 | 16.2186 | 11.447 | 10.8957 | 8.6279 | 11.0052 | 12.7434 | 10.786 | 11.4947 | 11.0227 | 11.1616 | 11.435 | 10.8083 | 10.5312 | 77.49% | 67.22% | 73.82% | 62.67% | 66.01% | 61.62% | 83.29% | 9.5094 | 9.2732 | 6.4994 | 7.2632
160 | LFB 10ppb | R | 06/03/2023 07:44:29 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LFB 10ppb.162 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 19528.3549 | 17890.3419 | 16.5575 | 20.7182 | 18187.6028 | 15780.3209 | 15180.9207 | 12.8424 | 12.4285 | 0.67159 | 0.23309 | 14.8657 | 13.4224 | 12.3185 | 12.5409 | 10.5825 | 0.67244 | 0.54305 | 0.62992 | 10.071 | 9.4619 | 11.8757 | -1.5593 | 11.5906 | 12.5208 | 12.0519 | -0.14536 | 11.4872 | 0.59682 | 11.1602 | 78.76% | 60.66% | 78.98% | 71.23% | 81.31% | 62.98% | 73.57% | 0.48101 | 0.46492 | -0.00044911 | -5.4767e-05
161 | CCV 10 ppb MAJORS | R | 06/03/2023 07:50:31 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.163 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 22358.4488 | 20023.5587 | 18.7649 | 21.1372 | 20285.393 | 17789.0904 | 16721.3055 | 14.3841 | 13.8884 | 14.4481 | 14.6342 | 15.4213 | 14.215 | 14.0465 | 13.8967 | 12.0496 | 12.2425 | 12.5482 | 13.6506 | 10.808 | 9.6335 | 12.0002 | 15.1006 | 12.1458 | 12.0296 | 11.7464 | 13.5256 | 11.2244 | 11.2821 | 11.4594 | 67.95% | 63.43% | 73.75% | 62.01% | 61.54% | 61.52% | 76.51% | 9.7471 | 9.6685 | 7.8487 | 8.1409
162 | CCV 10 ppb MAJORS | R | 06/03/2023 07:57:03 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.164 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 22680.5735 | 20310.9018 | 18.1759 | 20.0575 | 20410.5381 | 18361.8529 | 17484.8905 | 14.8212 | 13.8826 | 14.2275 | 14.4131 | 13.3342 | 14.4153 | 14.1999 | 14.2842 | 12.3249 | 12.066 | 12.2722 | 11.0053 | 11.796 | 8.5796 | 10.4726 | 11.8505 | 10.6998 | 10.5519 | 10.2262 | 10.8794 | 10.3456 | 11.1253 | 10.0612 | 82.69% | 71.94% | 73.64% | 63.98% | 61.26% | 62.79% | 84.7% | 9.6075 | 9.4285 | 6.2012 | 8.1981
163 | MT 32 C LD | R | 06/03/2023 08:03:06 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 32 C LD.165 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | -140.7397 | -48.4932 | 0.74952 | 11.7177 | 2728.4124 | -141.8514 | 4.2651 | 0.0083861 | 0.021652 | 0.012689 | -0.42061 | -0.001051 | 0.16325 | 0.00087204 | 0.012975 | 0.05466 | 0.040552 | -0.15442 | -0.043464 | 0.0018705 | -0.0061107 | -0.0062816 | 0.2069 | 0.014903 | -0.029407 | -0.0075113 | -0.027017 | 0.00030781 | -0.0036483 | -0.0086778 | 83.54% | 63.18% | 78.63% | 68.73% | 73.69% | 67.44% | 73.93% | -0.00083449 | 0.0011359 | 0.025057 | 0.025618
164 | MT 33 C | R | 06/03/2023 08:09:08 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 33 C.166 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | -145.7086 | -49.4851 | 0.88039 | 12.4185 | 2886.3951 | -156.478 | 4.1413 | 0.0073074 | 0.020773 | 0.012549 | -0.40958 | 0.0003478 | 0.1611 | 0.0010916 | 0.023486 | 0.050381 | 0.040755 | -0.17665 | -0.17692 | -0.0031069 | -0.0083293 | -0.0089575 | 0.37803 | 0.021753 | -0.036644 | -0.0088826 | -0.043529 | -0.0054505 | -0.0077281 | -0.0098835 | 75.53% | 68.1% | 78.53% | 66.52% | 74.27% | 66.56% | 81.18% | -0.00056455 | 0.00056094 | 0.022383 | 0.019597
165 | MT 33 C LSP | R | 06/03/2023 08:15:09 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 33 C LSP.167 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 59050.3916 | 51087.8021 | 51.9107 | 66.4428 | 46445.6396 | 46122.6891 | 45574.0222 | 38.0361 | 35.6948 | 28.0196 | 27.3774 | 38.1977 | 36.282 | 35.5592 | 35.5266 | 29.7314 | 22.5045 | 25.8482 | 31.8323 | 27.5463 | 23.1578 | 27.9778 | 34.0757 | 28.6826 | 28.2372 | 26.906 | 30.253 | 26.7324 | 0.039944 | 28.9852 | 76.36% | 63.03% | 71.99% | 36.71% | 28.83% | 60.04% | 73.8% | 26.1806 | 26.3261 | 12.2191 | 31.4701
166 | MT 33 C LSP | R | 06/03/2023 08:21:40 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 33 C LSP.168 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 58565.1544 | 51486.5813 | 49.8975 | 63.9704 | 46749.5255 | 47082.3799 | 45081.2455 | 38.2164 | 35.7307 | 36.7121 | 38.9269 | 38.9595 | 36.6258 | 36.289 | 35.8593 | 30.2663 | 30.2764 | 30.1235 | 28.8949 | 28.3867 | 23.0697 | 28.5605 | 31.1265 | 28.7723 | 28.654 | 28.0977 | 28.1884 | 26.8109 | 28.9871 | 29.5395 | 78.76% | 61.33% | 72.2% | 61.68% | 65.38% | 60.45% | 73.5% | 23.4326 | 23.3502 | 15.0327 | 17.9537
167 | MT 34 C | R | 06/03/2023 08:27:42 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 34 C.169 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | -143.464 | -47.5969 | 1.1387 | 12.1541 | 2681.8217 | -154.6356 | -0.54156 | 0.0082801 | 0.024895 | 0.019783 | -0.32762 | -0.00016841 | 0.22723 | 0.0031461 | 0.066185 | 0.015852 | 0.0072428 | -0.20826 | -0.18878 | 0.011085 | -0.0045896 | -0.0044872 | 0.53756 | 0.021887 | -0.032257 | -0.010015 | -0.039947 | -0.0054505 | 0.023349 | 0.012708 | 72.12% | 68.9% | 81.45% | 70.08% | 75.91% | 67.55% | 79.81% | -0.0010872 | 0.00034451 | 0.020497 | 0.017352
168 | MT 35 C | R | 06/03/2023 08:33:45 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 35 C.170 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | -142.6634 | -47.621 | 0.84783 | 12.634 | 2618.4619 | -149.882 | 3.9574 | 0.0077539 | 0.014176 | 0.012988 | -0.43659 | -0.0016743 | 0.10955 | 0.0010313 | 0.014336 | 0.016579 | 0.011307 | -0.19544 | -0.13235 | -0.0031069 | -0.0074028 | -0.0078824 | 0.358 | 0.014903 | -0.030298 | -0.01197 | -0.023896 | -0.00037493 | -0.0031493 | -0.011013 | 79.69% | 72.12% | 81.62% | 70.06% | 73.92% | 65.78% | 83.67% | -0.00083564 | 0.00064925 | 0.017473 | 0.018678

## Concentrations RSDs
| Sample Id | R | Acquisition Time | QC Status | Dataset File | Method File | Na KED 23/23 Helium KED_Low | Mg KED 24/24 Helium KED_Low | Al KED 27/27 Helium KED_Low | P Oshift 31/47 Oxygen DRC | S (34) -1 34/34 Helium KED_Low | K KED 39/39 Helium KED_Low | Ca KED 43/43 Helium KED_Low | V KED 51/51 Helium KED_Low | Cr (52) 52/52 Helium KED_Low | Cr (53) NH3 53/53 Ammonia DRC | Fe (54) NH3 54/54 Ammonia DRC | Mn Oshift 55/71 Oxygen DRC | Fe (56) KED 56/56 Helium KED_Low | Co KED 59/59 Helium KED_Low | Ni KED 60/60 Helium KED_Low | Cu (63) KED 63/63 Helium KED_Low | Cu (65) NH3 65/65 Ammonia DRC | Zn NH3 66/66 Ammonia DRC | As 75/75 | As KED 75/75 Helium KED_Low | As RH Oshift 75/91 Oxygen DRC | As Oshift 75/91 Oxygen DRC | Se (77) 77/77 | Se (77) Oshift 77/93 Oxygen DRC | Se (78) Oshift 78/94 Oxygen DRC | Se (80) Oshift 80/96 Oxygen DRC | Se (82) 82/82 | Se (82) Oshift 82/98 Oxygen DRC | Mo (95) NH3 95/95 Ammonia DRC | Mo (95) OShift 95/127 Oxygen DRC | Rh (IS1) 103/103 (IS) | Rh (IS2) 103/103 (IS) | Rh (IS3) 103/103 (IS) | Rh (IS4) 103/103 (IS) | Rh (IS5) 103/103 (IS) | Rh (IS6) 103/103 (IS) | Rh OShift (IS1) 103/119 (IS) | Cd (111) NH3 111/111 Ammonia DRC | Cd (114) NH3 114/114 Ammonia DRC | Pb 208/208 | Pb NH3 208/208 Ammonia DRC
1 | CAL BLANK | R | 6/2/2023 3:17:36 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CAL BLANK.004 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | [30 empty] | 3.277% | 0.3095% | 0.7573% | 0.4141% | 0.5558% | 2.017% | 1.979%
2 | LOW 0.01 | R | 6/2/2023 3:23:43 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LOW 0.01.005 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.108% | 1.397% | 10.85% | 14.63% | 186.9% | 2.484% | 3.867% | 15.57% | 37.79% | 16.18% | 0.5481% | 14.97% | 4.738% | 32.85% | 183.5% | 45.75% | 8.939% | 5.397% | 406% | 144.8% | 77.82% | 57.92% | 254.6% | 61.52% | 95.56% | 1670% | 248% | 72.73% | 3.909% | 6.584% | 0.6899% | 1.464% | 2.25% | 1.661% | 1% | 1.345% | 3.114% | 16.18% | 8.973% | 8.344% | 6.245%
3 | LOW 0.05 | R | 6/2/2023 3:29:46 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LOW 0.05.006 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 3.123% | 1.815% | 22.54% | 12.05% | 105.1% | 5.081% | 1.629% | 7.276% | 11.49% | 11.29% | 2.354% | 1.744% | 10.21% | 6.745% | 12.73% | 27.79% | 8.488% | 1.923% | 143.7% | 40.6% | 13.61% | 11.29% | 172.4% | 63.18% | 64.99% | 22.33% | 139.1% | 31.5% | 3.015% | 7.473% | 3.037% | 3.901% | 6.34% | 2.055% | 2.358% | 3.639% | 5.768% | 5.5% | 9.265% | 8.697% | 4.117%
4 | LOW 0.1 | R | 6/2/2023 3:35:48 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LOW 0.1.007 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 5.272% | 8.797% | 6.488% | 3.355% | 86.26% | 6.635% | 8.06% | 14.32% | 10.33% | 6.918% | 2.176% | 5.875% | 2.791% | 4.595% | 15.26% | 10.41% | 9.756% | 8.399% | 25.83% | 14.66% | 2.495% | 5.761% | 119.8% | 13.41% | 15.41% | 6.036% | 112% | 12.19% | 2.56% | 2.265% | 1.229% | 0.8802% | 2.017% | 1.21% | 0.9493% | 0.0713% | 2.965% | 0.9823% | 3.288% | 0.9838% | 1.538%
5 | LOW 0.5 | R | 6/2/2023 3:41:51 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LOW 0.5.008 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 26.03% | 23.4% | 117% | 4.773% | 16.46% | 11.61% | 37.3% | 14.9% | 9.392% | 0.3795% | 7.488% | 11.82% | 30.49% | 3.359% | 4.328% | 4.59% | 2.328% | 1.799% | 4.517% | 16.87% | 10.88% | 10.96% | 25.15% | 7.102% | 10.67% | 11.62% | 22.56% | 32.23% | 4.001% | 11.26% | 0.4323% | 8.264% | 5.283% | 1.121% | 0.4188% | 2.735% | 8.717% | 2.286% | 0.4938% | 0.6198% | 1.529%
6 | STD 1 | R | 6/2/2023 3:47:53 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\STD 1.009 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.53% | 3.744% | 53.38% | 3.714% | 3.357% | 2.553% | 4.274% | 1.256% | 1.602% | 1.7% | 4.483% | 0.6076% | 6.197% | 3.692% | 1.389% | 5.147% | 1.734% | 2.765% | 3.984% | 9.832% | 3.071% | 2.408% | 16.42% | 10.73% | 6.388% | 3.356% | 19.89% | 10.17% | 2.022% | 0.6339% | 0.6733% | 0.3953% | 2.17% | 1.01% | 9.13% | 0.9068% | 2.259% | 1.301% | 1.498% | 0.2846% | 10.91%
7 | STD 5 | R | 6/2/2023 3:53:56 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\STD 5.010 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.597% | 3.042% | 3.72% | 8.024% | 5.596% | 2.544% | 4.515% | 3.916% | 3.647% | 2.918% | 0.6036% | 0.642% | 2.393% | 2.016% | 2.721% | 3.516% | 1.876% | 1.332% | 2.646% | 3.21% | 2.327% | 1.459% | 1.766% | 6.238% | 1.896% | 3.871% | 8.606% | 5.287% | 1.787% | 1.239% | 1.937% | 1.007% | 1.888% | 0.7846% | 0.8302% | 1.916% | 2.833% | 2.304% | 1.396% | 2.228% | 0.8799%
8 | STD 10 | R | 6/2/2023 3:59:58 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\STD 10.011 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.7815% | 1.728% | 1.952% | 1.031% | 5.96% | 0.5903% | 1.79% | 0.67% | 1.042% | 1.088% | 2.691% | 1.835% | 1.057% | 1.088% | 0.6605% | 0.2345% | 1.275% | 1.529% | 2.184% | 2.552% | 2.104% | 1.926% | 4.609% | 3.611% | 3.444% | 1.758% | 1.736% | 5.916% | 1.02% | 1.428% | 0.4851% | 1.248% | 0.6937% | 0.8391% | 1.152% | 1.515% | 3.184% | 0.5492% | 1.025% | 0.5872% | 1.898%
9 | STD 25 | R | 6/2/2023 4:06:00 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\STD 25.012 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 6.564% | 5.268% | 5.847% | 1.049% | 5.589% | 6.523% | 7.08% | 6.001% | 5.527% | 1.229% | 1.165% | 1.613% | 6.907% | 6.813% | 6.779% | 5.294% | 1.079% | 0.2219% | 2.767% | 4.543% | 3.126% | 1.53% | 4.794% | 0.2835% | 2.776% | 2.167% | 1.515% | 3.47% | 0.8789% | 0.9552% | 0.7074% | 1.095% | 4.204% | 0.3375% | 1.562% | 0.3715% | 3.595% | 0.1098% | 0.8436% | 1.016% | 0.9908%
10 | STD 50 | R | 6/2/2023 4:12:02 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\STD 50.013 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.943% | 0.7592% | 4.386% | 0.3753% | 1.68% | 1.004% | 1.001% | 1.347% | 1.022% | 1.506% | 1.08% | 1.44% | 0.8983% | 1.193% | 0.9181% | 0.9986% | 1.027% | 0.58% | 3.197% | 1.582% | 2.128% | 0.731% | 3.456% | 0.2134% | 0.8603% | 1.546% | 4.025% | 0.8253% | 0.7457% | 0.8754% | 3.075% | 0.651% | 0.2639% | 1.039% | 0.9888% | 1.18% | 2.279% | 1.101% | 1.062% | 6.203% | 1.389%
11 | STD 100 | R | 6/2/2023 4:18:04 PM |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\STD 100.014 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.22% | 1.725% | 3.163% | 0.4734% | 2.515% | 2.269% | 1.373% | 1.824% | 2.448% | 0.6476% | 0.8878% | 0.5354% | 1.842% | 1.709% | 2.178% | 1.507% | 0.1297% | 0.3536% | 1.78% | 2.21% | 0.8053% | 0.2924% | 3.052% | 0.9716% | 0.9739% | 0.6654% | 1.798% | 0.4773% | 0.9045% | 0.2908% | 1.824% | 0.3891% | 2.22% | 0.9116% | 0.2953% | 0.2254% | 0.3844% | 0.5678% | 0.3087% | 2.301% | 0.3141%
13 | LBLANK | R | 6/2/2023 4:24:07 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LBLANK.015 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.9233% | 2.521% | 2.406% | 11.13% | 24.14% | 4.481% | 39168% | 12.16% | 74.63% | 69.51% | 1.112% | 17.47% | 3.882% | 67.27% | 20.62% | 1794% | 3.173% | 0.6495% | 419.4% | 310.7% | 34.46% | 35.11% | 268.8% | 1.913e-06% | 15.32% | 58.22% | 258.3% | 409% | 26.64% | 116% | 17.26% | 1.191% | 0.8535% | 1.196% | 0.8029% | 0.3994% | 1.61% | 84.81% | 13.63% | 167.5% | 29.59%
14 | ICV PE | R | 6/2/2023 4:30:09 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\ICV PE.016 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 3.37% | 2.885% | 4.71% | 0.6954% | 1.437% | 1.75% | 4.527% | 2.514% | 4.132% | 1.668% | 1.421% | 0.362% | 3.225% | 1.756% | 1.828% | 0.9877% | 1.695% | 1.302% | 2.892% | 1.305% | 1.207% | 1.215% | 1.743% | 2.005% | 2.48% | 1.512% | 1.61% | 1.216% | 0.4603% | 0.6177% | 0.5684% | 0.7782% | 2.672% | 1.203% | 0.7208% | 1.13% | 0.4998% | 1.822% | 1.62% | 0.745% | 0.3812%
15 | CCV 1ppb TRACE | R | 6/2/2023 4:36:12 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.017 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 3.661% | 5.083% | 28.26% | 6.21% | 6.327% | 5.035% | 1.963% | 1.94% | 1.032% | 2.068% | 1.167% | 0.9468% | 11.4% | 3.894% | 4.11% | 2.609% | 1.695% | 4.102% | 4.003% | 17.86% | 1.162% | 1.24% | 8.868% | 5.437% | 2.516% | 2.285% | 3.765% | 2.641% | 1.04% | 0.5805% | 0.3011% | 0.2138% | 1.258% | 0.8222% | 1.164% | 1.04% | 0.2663% | 1.52% | 2.083% | 1.771% | 1.07%
16 | CCV 1ppb TRACE | R | 6/2/2023 4:42:44 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.018 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.11% | 5.11% | 35.48% | 5.271% | 5.239% | 4.991% | 3.768% | 1.457% | 2.787% | 0.6205% | 0.8254% | 0.5057% | 10.24% | 3.251% | 5.067% | 1.435% | 1.486% | 4.954% | 4.094% | 9.137% | 2.42% | 3.118% | 3.353% | 8.526% | 8.389% | 1.611% | 15.8% | 9.361% | 1.073% | 0.4602% | 1.782% | 0.6125% | 1.518% | 0.4966% | 0.2556% | 0.8769% | 1.613% | 0.8061% | 1.542% | 1.289% | 0.6489%
17 | LFB 10ppb | R | 6/2/2023 4:48:46 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LFB 10ppb.019 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.7392% | 1.685% | 3.844% | 0.7648% | 2.846% | 2.293% | 0.2622% | 2.239% | 2.804% | 0.185% | 1.969% | 0.6288% | 2.272% | 2.108% | 2.449% | 1.816% | 0.722% | 0.9904% | 3.154% | 3.77% | 4.243% | 1.04% | 1.371% | 1.242% | 1.498% | 2.154% | 3.276% | 1.469% | 0.5991% | 2.543% | 1.052% | 1.135% | 1.665% | 0.5565% | 0.1487% | 0.4907% | 3.728% | 1.832% | 0.5359% | 0.8176% | 1.79%
18 | LFB 10ppb | R | 6/2/2023 4:55:16 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LFB 10ppb.020 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.253% | 1.555% | 1.829% | 2.581% | 3.572% | 1.864% | 1.568% | 0.5776% | 1.619% | 0.5383% | 1.584% | 0.825% | 1.659% | 0.949% | 0.0717% | 1.904% | 0.9943% | 0.1316% | 1.15% | 2.004% | 1.937% | 1.137% | 2.95% | 4.151% | 2.6% | 1.173% | 6.293% | 4.086% | 0.9595% | 1.106% | 1.041% | 1.084% | 0.8078% | 0.9704% | 1.585% | 1.534% | 2.448% | 0.2347% | 0.6711% | 0.7907% | 1.882%
19 | LLOQ 0.01 | R | 6/2/2023 5:01:19 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LLOQ 0.01.021 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.573% | 9.924% | 21.39% | 12.67% | 10.44% | 3.596% | 46.03% | 24.96% | 8.872% | 8.708% | 3.047% | 27.52% | 12.08% | 25.66% | 32.3% | 25.33% | 33.07% | 5.813% | 191.4% | 166.6% | 42.94% | 43.87% | 1657% | 31.53% | 44.21% | 221.5% | 167.6% | 303.7% | 7.168% | 9.21% | 0.3595% | 1.23% | 0.9884% | 0.27% | 2.44% | 0.9701% | 2.43% | 9.782% | 12.85% | 17.29% | 7.982%
20 | CCV 10 ppb MAJORS | R | 6/2/2023 5:07:22 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.022 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 4.284% | 4.464% | 3.984% | 2.383% | 5.26% | 4.973% | 7.115% | 3.607% | 5.447% | 1.3% | 1.062% | 1.795% | 5.423% | 4.844% | 4.786% | 5.096% | 1.104% | 1.481% | 1.851% | 6.902% | 1.475% | 1.008% | 3.048% | 3.827% | 1.523% | 1.448% | 1.848% | 5.275% | 0.9162% | 0.8824% | 0.7594% | 0.8657% | 4.005% | 0.9703% | 0.4442% | 1.381% | 1.44% | 0.8733% | 0.3795% | 0.6226% | 2.545%
21 | CCV 10 ppb MAJORS | R | 6/2/2023 5:13:54 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.023 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.156% | 2.298% | 1.295% | 0.9557% | 3.22% | 1.411% | 1.326% | 1.685% | 1.97% | 1.009% | 2.242% | 7.045% | 1.555% | 2.894% | 0.8619% | 2.112% | 0.958% | 1.207% | 2.756% | 1.864% | 6.627% | 6.278% | 3.352% | 11.32% | 5.742% | 7.033% | 6.364% | 5.859% | 0.3593% | 6.58% | 1.208% | 5.284% | 1.377% | 1.12% | 1.755% | 0.5646% | 5.952% | 0.4712% | 1.547% | 0.7882% | 2.225%
22 | LLOQ 0.1 | R | 6/2/2023 5:19:57 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LLOQ 0.1.024 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 81.77% | 15.1% | 20.71% | 17.95% | 14.47% | 42.91% | 21.02% | 15.81% | 8.459% | 14.22% | 15.23% | 4.229% | 1720% | 2.256% | 17.79% | 2.909% | 15.08% | 295.9% | 64.16% | 20.07% | 7.595% | 6.002% | 760.6% | 65.96% | 16.73% | 1.758% | 60.5% | 23.49% | 14.11% | 10.74% | 9.201% | 1.188% | 2.087% | 9.1% | 2.183% | 1.131% | 3.106% | 13.62% | 9.962% | 9.904% | 2.003%
23 | LBLANK | R | 6/2/2023 5:25:59 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LBLANK.025 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.3211% | 0% | 8.923% | 4.171% | 18.53% | 3.205% | 0% | 15.97% | 100.8% | 77.8% | 2.13% | 14.86% | 22.4% | 136.4% | 19.24% | 180% | 5.877% | 0.6843% | 993% | 337.6% | 5.337% | 5.626% | 384.2% | 32.94% | 12.98% | 16.33% | 107.9% | 177.5% | 4.594% | 4.484% | 0.7912% | 0.3726% | 1.346% | 6.44% | 0.4287% | 0.6385% | 1.826% | 16.97% | 266.8% | 27.26% | 20.98%
24 | LFB 6.25ppb & 2500ppb | R | 6/2/2023 5:32:01 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LFB 6.25ppb & 2500ppb.026 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.549% | 0.3865% | 5.811% | 3.43% | 1.007% | 1.807% | 2.386% | 1.967% | 2.034% | 0.6874% | 1.715% | 11.53% | 2.495% | 1.249% | 1.264% | 3.167% | 0.5928% | 1.743% | 1.989% | 2.984% | 12.87% | 10.39% | 5.923% | 14.79% | 9.112% | 8.705% | 2.904% | 10.35% | 1.707% | 11.31% | 1.492% | 8.57% | 1.654% | 0.2609% | 0.5971% | 0.9686% | 11.66% | 2.06% | 1.463% | 1.195% | 1.399%
25 | MT 1 | R | 6/2/2023 5:38:02 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 1.027 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.9883% | 12.24% | 0.7114% | 2.795% | 2.432% | 1.772% | 1.579% | 6.933% | 0.7885% | 1.407% | 0.6656% | 2.394% | 0.8796% | 10.48% | 0.6753% | 1.332% | 1.13% | 0.6094% | 5.394% | 9.042% | 8.331% | 2.113% | 8.12% | 21.28% | 8.056% | 7.17% | 9.414% | 12.29% | 5.247% | 2.409% | 0.9372% | 0.5176% | 1.204% | 0.4976% | 0.2489% | 2.413% | 5.928% | 1.453% | 5.151% | 1.375% | 1.424%
26 | MT 1 LD | R | 6/2/2023 5:44:04 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 1 LD.028 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.6156% | 14.59% | 0.7963% | 15.59% | 4.766% | 2.722% | 15.29% | 3.511% | 1.654% | 5.369% | 4.621% | 0.9773% | 0.9279% | 9.111% | 7.123% | 1.53% | 5.034% | 4.016% | 8.849% | 9.219% | 4.581% | 2.179% | 12.14% | 7.077% | 4.106% | 4.871% | 20.33% | 10.45% | 12.2% | 3.468% | 4.661% | 0.2501% | 0.2011% | 4.336% | 0.3626% | 12.94% | 3.78% | 13.08% | 4.951% | 6.249% | 2.387%
27 | LBLANK | R | 6/2/2023 6:05:56 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LBLANK.029 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.708% | 2.645% | 13.3% | 19.53% | 338.9% | 4.375% | 0% | 18.93% | 51.24% | 58.25% | 1.07% | 26.61% | 73.78% | 12.06% | 8.963% | 751.7% | 15.22% | 1.024% | 131.5% | 0% | 14.87% | 18.23% | 358.5% | 40.44% | 21.79% | 70.03% | 66.64% | 0% | 10.16% | 7.126% | 0.8819% | 3.788% | 9.354% | 1.37% | 0.708% | 0.9342% | 2.821% | 1.895e-06% | 64.63% | 6.07% | 21.19%
28 | STD 1 | R | 6/2/2023 6:11:58 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\STD 1.030 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.227% | 3.916% | 23.38% | 0.8094% | 5.347% | 2.813% | 6.333% | 1.717% | 3.103% | 1.902% | 3.499% | 2.508% | 5.313% | 2.13% | 4.977% | 1.037% | 1.415% | 1.265% | 2.975% | 8.448% | 1.212% | 1.976% | 14.97% | 1.492% | 3.228% | 1.729% | 6.666% | 12.92% | 1.257% | 0.915% | 0.8004% | 0.4601% | 1.337% | 0.4157% | 0.9085% | 1.846% | 2.49% | 2.05% | 0.6991% | 0.9771% | 2.113%
29 | STD 0.5 | R | 6/2/2023 6:18:06 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\STD 0.5.031 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.996% | 5.252% | 33.9% | 0.5496% | 13.38% | 5.834% | 8.385% | 4.333% | 7.648% | 2.022% | 7.084% | 2.198% | 1.818% | 2.03% | 5.065% | 4.51% | 1.706% | 1.56% | 7.17% | 6.575% | 1.877% | 3.323% | 12.83% | 27.9% | 6.102% | 7.818% | 6.461% | 4.295% | 0.9211% | 1.8% | 5.292% | 0.7411% | 1.748% | 0.8597% | 0.9448% | 0.8309% | 3.994% | 5.061% | 0.3473% | 4.943% | 0.7842%
30 | ICV PE | R | 6/2/2023 6:24:08 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\ICV PE.032 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.8838% | 2.574% | 4.373% | 1.345% | 0.8144% | 4.193% | 1.831% | 1.823% | 0.7844% | 1.226% | 2.911% | 0.3676% | 1.559% | 1.4% | 1.432% | 0.1444% | 1.479% | 2.506% | 2.325% | 2.716% | 1.11% | 1.556% | 2.268% | 1.757% | 1.845% | 0.6239% | 1.68% | 7.124% | 2.014% | 0.6545% | 0.4835% | 0.5159% | 0.3361% | 1.553% | 0.3284% | 0.9607% | 2.663% | 2.401% | 2.538% | 0.6703% | 1.605%
31 | ICV PE | R | 6/2/2023 6:30:39 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\ICV PE.033 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.521% | 1.367% | 2.782% | 1.658% | 2.275% | 1.671% | 3.181% | 1.384% | 0.3291% | 2.573% | 1.734% | 1.354% | 0.8064% | 0.9122% | 1.392% | 1.558% | 0.8248% | 1.336% | 2.134% | 6.154% | 2.528% | 1.221% | 2.095% | 2.528% | 1.482% | 0.8713% | 0.7146% | 4.541% | 0.9984% | 0.2326% | 0.9306% | 1.079% | 0.9531% | 1.59% | 1.321% | 1.55% | 2.529% | 1.426% | 0.7498% | 0.4838% | 1.799%
32 | CCV 1ppb TRACE | R | 6/2/2023 6:36:41 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.034 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 5.891% | 0.4015% | 48.45% | 6.495% | 7.703% | 3.638% | 6.581% | 3.823% | 7.057% | 1.462% | 5.064% | 5.874% | 8.257% | 3.664% | 2.151% | 1.784% | 0.7273% | 2.565% | 4.954% | 11.69% | 7.102% | 5.773% | 15.52% | 9.622% | 8.144% | 5.373% | 4.02% | 18.92% | 0.6316% | 4.986% | 1.422% | 4.075% | 3.134% | 0.9271% | 0.1638% | 0.9257% | 5.339% | 0.6606% | 0.8623% | 0.5593% | 0.9829%
33 | CCV 1ppb TRACE | R | 6/2/2023 6:43:13 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.035 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.89% | 4.917% | 13.23% | 3.659% | 1.833% | 1.559% | 6.697% | 3.753% | 1.993% | 1.234% | 1.568% | 1.609% | 4.521% | 1.962% | 1.237% | 1.825% | 1.542% | 2.321% | 3.576% | 6.461% | 3.852% | 1.589% | 6.97% | 12.45% | 5.245% | 2.896% | 6.55% | 8.024% | 0.9768% | 0.7565% | 0.4997% | 1.523% | 0.2567% | 1.24% | 1.106% | 0.315% | 3.524% | 2.595% | 1.273% | 0.5409% | 1.049%
34 | CCV 10 ppb MAJORS | R | 6/2/2023 6:49:16 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.036 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.737% | 3.17% | 7.648% | 0.5549% | 0.7386% | 2.71% | 2.845% | 1.84% | 1.286% | 1.276% | 1.182% | 0.6145% | 0.8802% | 0.9013% | 1.576% | 0.2795% | 0.4121% | 1.457% | 1.95% | 4.528% | 3.449% | 0.4631% | 2.598% | 2.333% | 0.3943% | 1.402% | 0.5427% | 5.189% | 1.467% | 1.708% | 1.213% | 0.4581% | 0.9478% | 1.095% | 2.5% | 1.903% | 3.479% | 0.9399% | 1.432% | 0.2888% | 1.552%
35 | CCV 10 ppb MAJORS | R | 6/2/2023 6:55:47 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.037 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.2999% | 1.367% | 7.392% | 1.373% | 4.006% | 1.527% | 2.864% | 1.311% | 1.8% | 1.107% | 2.928% | 0.9147% | 1.356% | 0.6006% | 0.7316% | 1.242% | 1.541% | 2.81% | 1.074% | 2.834% | 1.372% | 0.9298% | 1.45% | 2.871% | 3.766% | 1.604% | 0.6972% | 1.417% | 0.9582% | 0.3421% | 0.4744% | 0.1205% | 0.8444% | 2.254% | 1.46% | 1.43% | 1.266% | 1.093% | 2.006% | 0.9765% | 0.7457%
36 | MT 1 | R | 6/2/2023 7:01:50 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 1.038 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.673% | 2.569% | 0.8845% | 0.8532% | 6.07% | 1.041% | 13.74% | 3.842% | 8.806% | 3.574% | 1.413% | 12.7% | 0.2749% | 6.175% | 5.916% | 0.628% | 1.93% | 0.6891% | 6.162% | 8.469% | 13.82% | 12.32% | 12.09% | 4.791% | 13.53% | 17.62% | 8.741% | 16.66% | 2.977% | 8.974% | 0.6626% | 10.46% | 0.5676% | 0.6905% | 1.108% | 1.269% | 11.49% | 12.9% | 2.507% | 2.291% | 3.282%
37 | MT 1 LD | R | 6/2/2023 7:07:53 PM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 1 LD.039 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 6.822% | 7.823% | 6.6% | 0.9083% | 14% | 6.69% | 18.29% | 4.484% | 12.41% | 0.3388% | 0.3646% | 2.3% | 7.417% | 6.574% | 7.384% | 7.139% | 1.45% | 1.336% | 3.233% | 9.258% | 1.101% | 3.351% | 15.1% | 7.247% | 15.69% | 4.774% | 2.291% | 17.96% | 5.829% | 3.666% | 0.4137% | 1.447% | 5.661% | 0.9494% | 0.5818% | 1.856% | 1.371% | 5.648% | 2.319% | 1.289% | 0.8339%
38 | MT 2 | R | 6/2/2023 7:13:55 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 2.040 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.721% | 8.3% | 3.278% | 0.9013% | 2.322% | 1.762% | 13.24% | 1.374% | 3.427% | 1.495% | 1.947% | 7.358% | 3.119% | 7.157% | 10.91% | 2.963% | 2.019% | 1.397% | 5.045% | 2.099% | 12.48% | 9.316% | 10.83% | 18.63% | 11.66% | 11.94% | 5.997% | 8.844% | 6.488% | 12.47% | 0.7604% | 6.356% | 1.884% | 0.8882% | 0.8442% | 1.233% | 9.465% | 4.782% | 3.225% | 1.12% | 2.084%
39 | MT 2 LSP | R | 6/2/2023 7:19:58 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 2 LSP.041 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 3.873% | 5.136% | 6.013% | 5.762% | 5.744% | 6.582% | 4.947% | 4.33% | 5.807% | 1.094% | 0.3701% | 0.533% | 5.631% | 6.297% | 6.599% | 6.791% | 0.6409% | 0.6067% | 2.794% | 12.01% | 1.876% | 1.453% | 2.468% | 6.358% | 3.332% | 2.367% | 2.052% | 1.022% | 1.2% | 1.472% | 0.5755% | 0.9017% | 5.028% | 0.8139% | 1.077% | 5.053% | 0.6228% | 1.1% | 0.5802% | 0.3513% | 1.546%
40 | MT 2 LSP | R | 6/2/2023 7:26:09 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 2 LSP.042 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 4.602% | 4.826% | 2.954% | 1.36% | 3.226% | 2.215% | 2.884% | 1.679% | 2.185% | 0.7604% | 0.7922% | 2.082% | 3.052% | 4.248% | 3.645% | 2.768% | 1.089% | 0.7546% | 1.627% | 2.505% | 3.322% | 1.471% | 4.705% | 2.449% | 1.914% | 1.688% | 3.535% | 3.135% | 1.653% | 0.662% | 1.223% | 1.209% | 3.109% | 0.5526% | 0.5652% | 0.4011% | 3.085% | 1.917% | 0.506% | 1.018% | 1.437%
41 | MT 3 | R | 6/2/2023 7:32:11 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 3 .043 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.171% | 9.207% | 3.629% | 1.546% | 7.048% | 4.028% | 6.089% | 4.878% | 4.502% | 4.711% | 0.6689% | 1.527% | 2.228% | 8.187% | 1.753% | 2.224% | 0.4469% | 1.077% | 0.8155% | 7.865% | 2.005% | 1.886% | 17.82% | 4.074% | 3.796% | 2.833% | 11.35% | 3.512% | 4.175% | 1.27% | 0.7661% | 0.9923% | 1.467% | 1.192% | 3.551% | 0.8537% | 3.301% | 7.641% | 7.416% | 0.5022% | 3.663%
42 | MT 4 | R | 6/2/2023 7:38:14 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 4.044 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.677% | 4.705% | 0.3803% | 1.231% | 4.453% | 2.372% | 10.21% | 5.79% | 2.768% | 2.208% | 1.039% | 1.507% | 1.357% | 7.295% | 10.69% | 1.422% | 1.714% | 0.8096% | 4.391% | 6.247% | 1.62% | 0.807% | 4.294% | 8.625% | 8.66% | 1.124% | 4.367% | 6.264% | 0.316% | 4.216% | 1.747% | 1.595% | 0.5574% | 1.393% | 0.9051% | 2.019% | 1.509% | 3.455% | 4.264% | 4.673% | 1.119%
43 | LBLANK | R | 6/2/2023 7:44:17 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LBLANK.045 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.714% | 2.635% | 4.327% | 35.33% | 22.56% | 2.934% | 0% | 16.35% | 24.11% | 154.8% | 3.324% | 20.74% | 13.6% | 53.06% | 170.9% | 20.86% | 13.7% | 2.923% | 52.1% | 1062% | 24.98% | 28.94% | 146.1% | 44.66% | 9.666% | 40.8% | 382.9% | 233.2% | 6.163% | 1.378% | 0.9507% | 0.8704% | 1.308% | 6.585% | 2.3% | 11.33% | 1.138% | 14.74% | 28.67% | 5.706% | 23.97%
44 | CCV 1ppb TRACE | R | 6/2/2023 7:50:20 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.046 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.3585% | 2.912% | 45.27% | 3.469% | 11.34% | 3.154% | 1.743% | 5.451% | 2.116% | 3.577% | 2.778% | 0.5439% | 7.402% | 2.309% | 0.2095% | 1.442% | 1.598% | 0.8381% | 5.996% | 12.35% | 1.371% | 2.079% | 13.47% | 5.013% | 11.95% | 3.203% | 7.582% | 10.9% | 3.494% | 1.635% | 8.804% | 0.1298% | 0.7926% | 1.001% | 1.473% | 0.9488% | 1.764% | 1.337% | 0.527% | 9.521% | 1.804%
45 | CCV 1ppb TRACE | R | 6/2/2023 7:56:52 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.047 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 3.577% | 5.436% | 20.78% | 6.354% | 4.773% | 5.461% | 2.732% | 1.285% | 2.34% | 2.697% | 4.398% | 1.151% | 1.221% | 2.177% | 3.192% | 1.593% | 0.8376% | 4.18% | 2.628% | 7.151% | 3.414% | 0.2311% | 3.699% | 16.3% | 3.168% | 2.484% | 5.644% | 25.73% | 0.91% | 1.211% | 0.5208% | 0.8058% | 0.5806% | 0.6088% | 1.214% | 0.795% | 2.535% | 2.139% | 1.309% | 0.6382% | 1.158%
46 | CCV 10 ppb MAJORS | R | 6/2/2023 8:02:55 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.048 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.302% | 1.351% | 3.989% | 1.226% | 4.691% | 2.161% | 3.511% | 2.01% | 1.835% | 0.8878% | 2.048% | 1.929% | 2.78% | 1.481% | 2.077% | 2.362% | 0.9373% | 0.6617% | 2.662% | 6.689% | 1.358% | 2.889% | 0.9506% | 6.159% | 2.982% | 1.234% | 2.479% | 5.54% | 1.561% | 2.027% | 1.593% | 1.513% | 1.337% | 0.1982% | 1.956% | 2.399% | 0.9019% | 1.083% | 0.9206% | 0.816% | 0.452%
47 | CCV 10 ppb MAJORS | R | 6/2/2023 8:09:27 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.049 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 3.299% | 2.934% | 1.896% | 0.9745% | 2.857% | 2.03% | 4.508% | 2.443% | 2.558% | 2.789% | 2.605% | 1.13% | 3.951% | 2.111% | 2.145% | 3.881% | 1.949% | 2.088% | 3.611% | 5.001% | 1.994% | 0.408% | 3.525% | 3.651% | 1.388% | 1% | 4.417% | 4.983% | 1.667% | 0.7078% | 1.685% | 0.492% | 1.812% | 1.45% | 4.316% | 0.8608% | 1.801% | 0.3463% | 0.8332% | 0.2791% | 5.201%
48 | MT 5 | R | 6/2/2023 8:15:30 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 5.050 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.325% | 9.614% | 2.915% | 1.672% | 2.369% | 1.143% | 3.957% | 3.447% | 2.702% | 0.6876% | 2.859% | 1.4% | 2.553% | 12.58% | 6.705% | 1.405% | 1.865% | 2.291% | 4.013% | 3.204% | 1.687% | 2.212% | 3.294% | 14.98% | 7.055% | 8.706% | 10.14% | 11.87% | 3.641% | 3.449% | 0.09245% | 1.465% | 2.185% | 1.019% | 1.38% | 0.9517% | 1.248% | 2.019% | 2.163% | 1.864% | 0.7342%
49 | MT 6 | R | 6/2/2023 8:21:32 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 6.051 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.792% | 1.468% | 3.426% | 1.208% | 3.699% | 1.669% | 2.289% | 4.549% | 1.637% | 1.1% | 1.151% | 1.779% | 1.792% | 2.732% | 1.064% | 2.077% | 1.647% | 1.111% | 1.982% | 3.564% | 1.55% | 0.1609% | 1.655% | 2.901% | 4.274% | 0.6331% | 3.41% | 1.215% | 5.42% | 3.901% | 0.1585% | 0.8318% | 1.771% | 0.2038% | 3.655% | 0.5005% | 0.8938% | 6.331% | 1.277% | 0.1671% | 6.235%
50 | MT 7 | R | 6/2/2023 8:27:35 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 7.052 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.278% | 13.97% | 1.854% | 2.419% | 1.753% | 0.4474% | 4.398% | 0.975% | 0.9707% | 16.58% | 15.99% | 0.9864% | 1.308% | 2.209% | 2.617% | 1.088% | 16.92% | 15.47% | 2.367% | 1.596% | 5.993% | 1.432% | 14.33% | 3.708% | 5.082% | 1.007% | 4.687% | 3.234% | 18.86% | 1.672% | 1.298% | 0.5251% | 0.5266% | 13.62% | 2.073% | 2.553% | 5.283% | 16.44% | 14.11% | 1.665% | 0.8361%
51 | MT 8 | R | 6/2/2023 8:33:38 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 8.053 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 7.794% | 3.933% | 10.88% | 1.416% | 10.46% | 9.902% | 13.78% | 9.367% | 7.965% | 3.454% | 0.3211% | 0.5287% | 8.455% | 13.29% | 11.14% | 10.33% | 0.8466% | 0.3502% | 1.702% | 10.43% | 2.57% | 1.785% | 10.25% | 5.522% | 2.478% | 1.852% | 3.792% | 4.119% | 2.858% | 3.134% | 1.041% | 0.941% | 8.569% | 0.6776% | 1.889% | 0.8354% | 3.476% | 8.653% | 2.748% | 2.388% | 2.755%
52 | MT 9 | R | 6/2/2023 8:39:40 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 9.054 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.01043% | 7.095% | 0.6402% | 1.861% | 2.69% | 1.774% | 10.03% | 5.304% | 2.845% | 0.9479% | 0.9172% | 2.504% | 0.8696% | 7.223% | 6.214% | 1.859% | 0.6775% | 0.745% | 3.664% | 8.357% | 2.86% | 0.8485% | 4.388% | 7.476% | 5.796% | 7.088% | 7.508% | 8.653% | 1.923% | 3.866% | 2.996% | 2.092% | 1.329% | 0.3894% | 0.5578% | 0.8843% | 1.274% | 3.792% | 3.769% | 4.957% | 1.822%
53 | MT 10 | R | 6/2/2023 8:45:42 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 10.055 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.35% | 10.52% | 4.126% | 1.021% | 2.183% | 2.385% | 6.227% | 3.262% | 1.198% | 2.4% | 1.387% | 0.8303% | 1.623% | 2.518% | 2.305% | 2.335% | 1.904% | 1.348% | 2.222% | 8.493% | 1.668% | 0.3364% | 3.716% | 16.52% | 1.025% | 4.943% | 3.904% | 13.07% | 2.494% | 3.203% | 0.6311% | 0.1809% | 2.343% | 0.8768% | 1.155% | 0.2671% | 1.945% | 4.831% | 2.237% | 1.986% | 1.592%
54 | MT 11 | R | 6/2/2023 8:51:44 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 11 .056 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.668% | 2.111% | 1.275% | 0.6638% | 0.1169% | 3.331% | 5.79% | 3.46% | 0.5172% | 0.5408% | 0.05037% | 0.8465% | 1.359% | 0.9114% | 5.009% | 0.3242% | 0.4379% | 0.8144% | 1.858% | 4.613% | 2.153% | 1.502% | 4.704% | 6.519% | 3.285% | 3.012% | 0.8337% | 1.665% | 1.085% | 3.502% | 0.771% | 1.447% | 0.4026% | 0.4229% | 1.155% | 1.645% | 1.41% | 0.4168% | 1.9% | 3.232% | 1.38%
55 | MT 11 LD | R | 6/2/2023 8:57:47 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 11 LD.057 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.242% | 4.118% | 2.834% | 1.815% | 1.939% | 1.913% | 3.683% | 1.744% | 2.282% | 1.677% | 1.403% | 2.177% | 2.387% | 0.8653% | 3.103% | 1.922% | 2.278% | 1.723% | 2.335% | 2.905% | 3.43% | 0.9462% | 4.411% | 8.436% | 3.237% | 1.839% | 5.632% | 2.317% | 1.398% | 2.407% | 0.9997% | 0.233% | 2.233% | 1.473% | 0.3604% | 1.211% | 2.723% | 4.017% | 2.312% | 0.9368% | 0.5458%
56 | MT 11 LD | R | 6/2/2023 9:04:18 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 11 LD.058 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 11.01% | 13.61% | 14% | 1.038% | 14.07% | 12.81% | 13.44% | 15.13% | 13.3% | 0.7543% | 1.634% | 1.063% | 13.76% | 14.67% | 10.29% | 11.69% | 0.4053% | 1.268% | 1.612% | 14.07% | 3.072% | 0.587% | 2.516% | 2.184% | 3.432% | 0.8912% | 6.93% | 3.805% | 1.82% | 2.6% | 1.546% | 0.8353% | 10.39% | 0.6253% | 1.273% | 1.641% | 2.827% | 1.874% | 1.752% | 3.166% | 2.178%
57 | MT 12 | R | 6/2/2023 9:10:20 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 12.059 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 4.862% | 8.084% | 4.965% | 1.033% | 5.56% | 5.11% | 3.782% | 4.699% | 2.902% | 23.14% | 22.99% | 1.945% | 5.911% | 6.109% | 1.731% | 4.316% | 22.63% | 22.53% | 5.694% | 6.549% | 1.134% | 1.976% | 7.742% | 3.601% | 9.578% | 1.029% | 4.689% | 7.279% | 23.94% | 1.587% | 4.179% | 1.193% | 4.493% | 18.38% | 2.352% | 2.59% | 0.208% | 21.04% | 20.79% | 3.753% | 2.09%
58 | MT 12 LSP | R | 6/2/2023 9:16:23 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 12 LSP.060 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.009% | 2.555% | 1.04% | 0.8617% | 2.491% | 1.008% | 0.3533% | 1.004% | 1.548% | 1.743% | 1.321% | 1.265% | 0.5247% | 1.192% | 2.066% | 0.5379% | 1.379% | 1.927% | 7.502% | 4.492% | 3.675% | 1.117% | 7.825% | 4.583% | 2.188% | 0.3151% | 8.454% | 3.451% | 0.794% | 1.34% | 6.188% | 0.7726% | 0.7554% | 1.518% | 0.2408% | 1.442% | 4.271% | 1.465% | 1.311% | 7.247% | 1.223%
59 | LBLANK | R | 6/2/2023 9:22:26 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LBLANK.061 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.02% | 5.517% | 8.996% | 58.81% | 8.105% | 2.701% | 0% | 19.66% | 41.69% | 128.9% | 0.2755% | 5.577% | 43.47% | 104.7% | 219.4% | 14.84% | 13.63% | 11.7% | 54.63% | 148.8% | 10.49% | 12.6% | 134% | 46.06% | 41.45% | 198.2% | 628.7% | 996.2% | 266.6% | 1.828% | 9.207% | 1.278% | 1.399% | 0.9841% | 1.208% | 0.6588% | 1.331% | 116.2% | 60.6% | 17.39% | 22.83%
60 | CCV 1ppb TRACE | R | 6/2/2023 9:28:29 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.062 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 5.284% | 5.189% | 17.41% | 3.35% | 5.695% | 3.07% | 10.07% | 6.173% | 7.737% | 1.108% | 1.613% | 2.287% | 4.141% | 4.749% | 4.079% | 8.541% | 1.897% | 1.56% | 2.419% | 9.036% | 3.932% | 2.539% | 12.42% | 10.55% | 5.198% | 1.172% | 6.711% | 3.715% | 1.15% | 2.751% | 1.834% | 1.485% | 2.968% | 1.363% | 1.804% | 0.514% | 2.523% | 1.686% | 0.9935% | 1.256% | 0.8751%
61 | CCV 1ppb TRACE | R | 6/2/2023 9:35:00 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.063 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.7007% | 4.037% | 55.07% | 8.822% | 3.271% | 5.848% | 4.011% | 1.834% | 1.856% | 10.44% | 16.65% | 2.145% | 5.875% | 2.379% | 4.366% | 3.234% | 10.67% | 12.94% | 5.685% | 10.4% | 0.9819% | 1.949% | 6.863% | 16.29% | 1.314% | 0.9249% | 12.65% | 10.46% | 9.698% | 2.374% | 1.099% | 1.942% | 2.372% | 8.414% | 0.7711% | 0.6928% | 3.014% | 8.062% | 10.68% | 1.08% | 0.9483%
62 | LFB 10ppb | R | 6/2/2023 9:41:03 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LFB 10ppb.064 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.917% | 2.937% | 5.502% | 0.9751% | 6.072% | 4.236% | 1.3% | 2.58% | 0.9721% | 0.8373% | 4.427% | 16.94% | 2.59% | 2.418% | 3.182% | 1.699% | 1.812% | 2.239% | 6.725% | 2.288% | 16.24% | 15.77% | 7.102% | 13.42% | 14.79% | 15.43% | 6.719% | 13.68% | 1.264% | 17.31% | 4.147% | 13.4% | 2.311% | 1.59% | 1.582% | 0.8024% | 13.85% | 2.455% | 1.38% | 5.791% | 1.497%
63 | LFB 10ppb | R | 6/2/2023 9:47:33 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LFB 10ppb.065 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.166% | 1.934% | 7.782% | 1.051% | 5.551% | 1.371% | 2.201% | 2.322% | 1.208% | 1.293% | 1.731% | 2.424% | 2.272% | 0.73% | 0.1395% | 1.025% | 1.842% | 0.9654% | 0.915% | 3.016% | 2.978% | 1.153% | 5.239% | 2.011% | 3.167% | 1.549% | 2.187% | 0.7447% | 0.4035% | 1.097% | 0.8385% | 0.9454% | 1.375% | 1.43% | 1.04% | 1.192% | 2.932% | 0.8786% | 1.111% | 0.6742% | 2.21%
64 | CCV 10 ppb MAJORS | R | 6/2/2023 9:53:35 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.066 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.6188% | 1.303% | 11.53% | 1.589% | 3.904% | 2.333% | 4.199% | 0.6944% | 2.04% | 15.54% | 16.33% | 1.517% | 3.125% | 1.758% | 0.8815% | 0.2248% | 14.63% | 14.85% | 13.23% | 5.698% | 0.8613% | 1.763% | 15.07% | 6.935% | 3.474% | 2.731% | 11.59% | 4.361% | 16.45% | 1.003% | 9.805% | 1.096% | 0.7243% | 13.02% | 0.8641% | 0.6459% | 0.7573% | 16.09% | 15.38% | 10.94% | 0.7986%
65 | CCV 10 ppb MAJORS | R | 6/2/2023 10:00:06 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.067 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 11.3% | 11.29% | 14.23% | 13.1% | 13.08% | 11.78% | 12.43% | 10.22% | 11.5% | 17.38% | 19.66% | 0.1295% | 12.76% | 12.46% | 13.25% | 12% | 16.31% | 16.63% | 3.626% | 13.17% | 1.069% | 0.3856% | 2.441% | 5.465% | 1.968% | 2.294% | 1.68% | 1.453% | 16.71% | 1.643% | 1.235% | 0.3058% | 8.985% | 13.17% | 0.8413% | 9.628% | 0.7577% | 15.62% | 16.48% | 2.197% | 1.147%
66 | MT 13 | R | 6/2/2023 10:06:09 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 13.068 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.956% | 6.979% | 2.892% | 1.587% | 5.309% | 3.046% | 4.265% | 0.4151% | 1.618% | 1.913% | 2.022% | 2.177% | 1.238% | 2.618% | 3.629% | 0.812% | 1.249% | 1.982% | 13.54% | 1.427% | 4.58% | 0.3389% | 19.9% | 2.654% | 1.751% | 1.274% | 15.17% | 6.436% | 2.713% | 0.5155% | 12.36% | 0.9991% | 1.253% | 1.454% | 1.079% | 1.472% | 4.465% | 1.834% | 1.343% | 15.53% | 1.434%
67 | MT 14 | R | 6/2/2023 10:12:11 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 14.069 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.234% | 3.461% | 1.881% | 0.07655% | 2.423% | 0.889% | 7.072% | 1.412% | 3.078% | 11% | 7.306% | 0.6111% | 1.058% | 4.187% | 2.145% | 1.816% | 0.9116% | 0.9512% | 2.027% | 2.769% | 4.16% | 1.789% | 6.67% | 3.31% | 1.876% | 0.4904% | 1.614% | 9.852% | 3.438% | 5.8% | 2.437% | 1.154% | 0.7913% | 0.5921% | 0.6574% | 0.9968% | 3.947% | 2.117% | 2.921% | 4.018% | 0.428%
68 | MT 15 | R | 6/2/2023 10:18:14 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 15.070 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 3.613% | 7.217% | 3.293% | 3.522% | 2.477% | 1.186% | 1.56% | 3.697% | 2.759% | 0.4389% | 0.8361% | 2.259% | 1.035% | 3.364% | 1.856% | 2.3% | 0.6576% | 1.646% | 0.5683% | 7.534% | 4.193% | 0.7789% | 0.8199% | 10.2% | 2.619% | 0.5279% | 2.475% | 11.18% | 1.429% | 3.255% | 0.8478% | 0.8153% | 1.37% | 0.3174% | 0.3711% | 0.6713% | 3.628% | 3.803% | 2.037% | 3.703% | 1.053%
69 | MT 16 | R | 6/2/2023 10:24:17 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 16.071 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.344% | 1.658% | 6.436% | 0.3291% | 4.954% | 1.665% | 5.368% | 4.47% | 1.417% | 1.816% | 1.12% | 0.904% | 0.3393% | 1.829% | 4.706% | 0.4745% | 0.9314% | 1.23% | 16.45% | 5.98% | 4.705% | 1.313% | 26.4% | 5.59% | 1.69% | 3.815% | 22.51% | 4.501% | 2.519% | 1.387% | 13.4% | 0.8907% | 0.5334% | 0.9951% | 0.3905% | 0.4574% | 3.255% | 2.057% | 2.648% | 15.9% | 0.9175%
70 | MT 17 | R | 6/2/2023 10:30:20 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 17.072 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.158% | 3.036% | 1.568% | 3.091% | 1.863% | 1.319% | 4.047% | 8.373% | 1.345% | 1.817% | 0.9843% | 10.45% | 0.7678% | 0.9783% | 3.884% | 1.048% | 1.616% | 1.018% | 1.474% | 4.437% | 9.783% | 10.26% | 4.657% | 5.031% | 8.216% | 10.53% | 1.298% | 6.49% | 1.406% | 9.888% | 1.278% | 7.663% | 0.6128% | 1.382% | 0.6488% | 2.926% | 7.486% | 0.6166% | 0.8778% | 1.591% | 0.6524%
71 | MT 18 | R | 6/2/2023 10:36:32 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 18.073 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.927% | 6.235% | 3.94% | 1.868% | 4.343% | 3.974% | 14.97% | 1.436% | 1.523% | 1.305% | 1.414% | 2.313% | 1.495% | 4.212% | 3.125% | 2.324% | 1.031% | 1.321% | 1.544% | 5.128% | 0.368% | 0.9555% | 7.371% | 4.514% | 2.545% | 2.362% | 0.8598% | 4.995% | 0.3205% | 1.79% | 0.6094% | 0.8108% | 2.343% | 0.8281% | 0.9725% | 3.367% | 0.6683% | 0.8716% | 0.7466% | 2.133% | 2.252%
72 | MT 19 | R | 6/2/2023 10:42:43 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 19.074 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 3.218% | 5.356% | 2.473% | 0.03002% | 3.965% | 3.621% | 16.34% | 8.23% | 17.38% | 21.83% | 20.06% | 2.729% | 3.941% | 8.51% | 1.886% | 3.94% | 18.3% | 18.8% | 1.054% | 7.567% | 1.895% | 1.74% | 8.559% | 9.774% | 2.513% | 5.508% | 1.792% | 0.9775% | 19.96% | 1.988% | 0.7626% | 0.952% | 2.992% | 16.28% | 0.8677% | 0.7387% | 0.3675% | 18.13% | 19.46% | 0.8672% | 1.548%
73 | MT 20 | R | 6/2/2023 10:48:45 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 20.075 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.7387% | 5.287% | 6.148% | 2.169% | 2.172% | 2.764% | 5.283% | 4.053% | 5.742% | 2.826% | 0.9874% | 1.259% | 0.9009% | 8.308% | 1.717% | 0.8351% | 0.9958% | 0.877% | 0.8% | 4.6% | 3.251% | 1.371% | 5.313% | 4.487% | 1.507% | 4.232% | 3.583% | 0.9645% | 0.6743% | 0.7527% | 0.6466% | 1.382% | 1.375% | 0.476% | 1.106% | 1.98% | 3.274% | 0.2579% | 0.4491% | 2.264% | 1.34%
74 | MT 21 | R | 6/2/2023 10:54:57 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 21.076 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.507% | 5.05% | 4.533% | 0.5244% | 2.28% | 2.197% | 1.873% | 4.271% | 1.865% | 7.809% | 1.307% | 3.111% | 0.8425% | 2.192% | 2.145% | 0.7357% | 1.315% | 1.409% | 3.468% | 3.006% | 5.297% | 3.952% | 7.12% | 2.228% | 1.102% | 6.659% | 1.441% | 4.495% | 1.918% | 2.771% | 4.339% | 3.162% | 0.9335% | 1.814% | 9.275% | 0.7241% | 4.421% | 2.339% | 0.5922% | 6.683% | 8.456%
75 | MT 21 LD | R | 6/2/2023 11:00:59 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 21 LD.077 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 8.691% | 0% | 2.434% | 12.31% | 3.538% | 9.734% | 0% | 21.75% | 90.94% | 53.65% | 1.87% | 3.06% | 373.3% | 48.77% | 22.76% | 173% | 17.68% | 1.18% | 1.1% | 142% | 7.045% | 7.219% | 1.234% | 45.08% | 20.1% | 95.1% | 55.12% | 0% | 47.03% | 3.534% | 4.263% | 1.299% | 3.683% | 1.791% | 2.635% | 5.031% | 1.434% | 61.43% | 345.1% | 9.773% | 3.281%
76 | LBLANK | R | 6/2/2023 11:07:00 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LBLANK.078 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.574% | 0% | 13.98% | 9.38% | 32.92% | 3.834% | 126.2% | 23.83% | 61.17% | 44.15% | 2.359% | 15.98% | 136% | 123.6% | 107.9% | 91.17% | 21.43% | 1.899% | 36.85% | 0% | 8.154% | 8.721% | 46.2% | 38.44% | 18.28% | 177.6% | 18.23% | 0% | 6.64% | 2.555% | 0.7277% | 2.293% | 9.042% | 1.143% | 1.085% | 1.805% | 1.055% | 24.67% | 26.92% | 23.38% | 27.87%
77 | CCV 1ppb TRACE | R | 6/2/2023 11:13:03 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.079 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.258% | 1.369% | 32.37% | 4.023% | 1.246% | 4.581% | 6.803% | 0.649% | 2.397% | 6.601% | 10.14% | 2.426% | 2.838% | 3.198% | 1.277% | 0.121% | 6.062% | 11.11% | 6.901% | 10.97% | 4.464% | 2.37% | 14.57% | 8.395% | 4.915% | 2.755% | 4.009% | 4.847% | 5.066% | 1.514% | 3.006% | 0.9754% | 0.5034% | 6.526% | 2.304% | 0.2884% | 5.331% | 7.011% | 8.415% | 3.187% | 0.6868%
78 | CCV 1ppb TRACE | R | 6/2/2023 11:19:35 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.080 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.378% | 2.033% | 14.51% | 11.57% | 5.61% | 2.813% | 5.129% | 3.38% | 2.033% | 1.034% | 5.599% | 1.271% | 11.39% | 0.781% | 1.5% | 5.476% | 1.362% | 0.7508% | 5.629% | 16.07% | 0.5757% | 1.181% | 6.078% | 23.23% | 7.85% | 3.682% | 8.741% | 5.562% | 0.3095% | 2.803% | 2.323% | 1.079% | 1.707% | 0.8071% | 3.131% | 0.7261% | 0.7534% | 2.16% | 2.728% | 2.958% | 2.938%
79 | CCV 10 ppb MAJORS | R | 6/2/2023 11:25:38 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.081 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.5039% | 2.153% | 3.228% | 1.104% | 3.32% | 1.774% | 1.661% | 1.852% | 1.541% | 1.461% | 0.7219% | 2.159% | 0.8019% | 0.8111% | 0.2968% | 1.295% | 0.9056% | 3.463% | 0.9114% | 1.512% | 1.945% | 1.945% | 4.208% | 4.659% | 1.698% | 0.4512% | 2.583% | 1.664% | 1.805% | 1.909% | 0.8125% | 0.7045% | 0.6476% | 1.288% | 0.5617% | 1.271% | 3.113% | 0.7551% | 0.7673% | 1.002% | 1.455%
80 | CCV 10 ppb MAJORS | R | 6/2/2023 11:32:09 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.082 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.5259% | 1.006% | 10.83% | 3.008% | 1.991% | 0.7668% | 2.644% | 1.565% | 2.106% | 1.345% | 0.4222% | 1.644% | 2.431% | 1.557% | 1.505% | 1.059% | 0.5879% | 0.3183% | 1.517% | 3.251% | 1.055% | 1.853% | 2.096% | 7.284% | 1.296% | 1.91% | 4.276% | 4.817% | 0.3524% | 0.3% | 0.2871% | 1.117% | 2.235% | 0.5438% | 1.721% | 1.967% | 0.5508% | 1.544% | 1.244% | 0.8927% | 0.3412%
81 | MT 22 | R | 6/2/2023 11:38:12 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 22.083 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.471% | 6.384% | 3.195% | 2.379% | 1.222% | 1.547% | 6.474% | 8.328% | 4.861% | 4.979% | 1.846% | 2.388% | 1.253% | 8.669% | 2.405% | 2.335% | 1.228% | 0.7782% | 1.362% | 3.302% | 3.277% | 1.826% | 13.12% | 6.75% | 1.864% | 0.502% | 3.353% | 6.862% | 2.254% | 1.443% | 0.7876% | 1.307% | 1.936% | 0.6126% | 3.071% | 2.294% | 3.163% | 1.531% | 0.895% | 0.8977% | 3.239%
82 | MT 22 LSP | R | 6/2/2023 11:44:15 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 22 LSP.084 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.317% | 1.162% | 6.005% | 4.14% | 1.334% | 2.674% | 2.237% | 1.04% | 6.185% | 1.623% | 1.065% | 0.5054% | 0.8329% | 1.522% | 1.841% | 0.8893% | 1.746% | 2.458% | 0.5422% | 2.031% | 2.795% | 2.212% | 5.016% | 8.69% | 2.852% | 1.945% | 1.556% | 2.123% | 1.909% | 0.7522% | 0.6553% | 0.3218% | 0.9114% | 1.679% | 1.875% | 0.8518% | 1.923% | 2.304% | 2.333% | 0.3176% | 2.223%
83 | MT 22 LSP | R | 6/2/2023 11:50:45 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 22 LSP.085 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 10.99% | 12% | 12.04% | 1.35% | 13.3% | 9.625% | 12.16% | 11.81% | 12.86% | 1.351% | 2.662% | 0.8745% | 12.73% | 11.79% | 12.33% | 11.76% | 1.768% | 1.766% | 3.187% | 9.149% | 2.297% | 0.7867% | 3.871% | 2.58% | 2.537% | 1.054% | 4.805% | 2.8% | 0.359% | 0.6898% | 1.837% | 0.5864% | 10.06% | 1.614% | 0.776% | 1.682% | 2.474% | 0.4868% | 0.9798% | 2.286% | 3.131%
84 | MT 23 | R | 6/2/2023 11:56:47 PM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 23.086 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.139% | 8.176% | 7.676% | 2.205% | 2.583% | 1.703% | 2.773% | 15.09% | 3.953% | 3.137% | 0.7619% | 4.77% | 0.751% | 5.75% | 2.933% | 1.692% | 0.843% | 1.599% | 3.264% | 5.757% | 7.235% | 3.705% | 1.584% | 8.91% | 7.373% | 4.537% | 4.393% | 9.768% | 2.272% | 2.797% | 0.974% | 2.535% | 0.398% | 1.489% | 18.35% | 2.303% | 6.36% | 2.117% | 1.348% | 1.472% | 21.39%
85 | MT 24 | R | 6/3/2023 12:02:49 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 24.087 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 4.201% | 11.86% | 11.57% | 1.326% | 9.347% | 3.075% | 8.43% | 11.3% | 4.961% | 5.011% | 0.9568% | 1.859% | 2.521% | 2.769% | 9.63% | 3.016% | 0.5661% | 0.809% | 1.736% | 19.34% | 2.254% | 1.661% | 14.9% | 11.35% | 1.896% | 5.141% | 4.86% | 11.58% | 0.7362% | 0.7873% | 1.66% | 1.933% | 1.351% | 0.139% | 0.1458% | 1.582% | 1.749% | 1.484% | 0.5749% | 2.53% | 0.3167%
86 | MT 25 | R | 6/3/2023 12:08:52 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 25.088 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.215% | 4.165% | 2.756% | 2.359% | 1.975% | 3.877% | 10.77% | 10.29% | 4.287% | 3.606% | 0.8148% | 2.519% | 2.701% | 4.576% | 6.811% | 1.122% | 1.892% | 1.091% | 4.407% | 7.835% | 3.137% | 1.679% | 11.63% | 10.83% | 7.768% | 3.331% | 3.68% | 9.737% | 1.554% | 2.935% | 5.11% | 1.511% | 2.406% | 1.124% | 1.667% | 0.9929% | 2.834% | 0.59% | 0.745% | 6.13% | 1.564%
87 | MT 26 | R | 6/3/2023 12:14:55 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 26.089 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.7031% | 10.79% | 4.823% | 0.8579% | 4.235% | 1.542% | 17.83% | 13.12% | 10.7% | 0.8781% | 1.015% | 2.133% | 0.6893% | 2.698% | 6.176% | 1.082% | 1.308% | 1.092% | 24.2% | 8.856% | 2.884% | 2.395% | 46.97% | 12.91% | 7.727% | 3.962% | 33.36% | 14.06% | 1.708% | 1.956% | 21.85% | 1.883% | 1.35% | 1.088% | 4.918% | 2.001% | 1.104% | 2.539% | 1.031% | 28.61% | 8.269%
88 | MT 27 | R | 6/3/2023 12:21:03 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 27.090 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.043% | 1.108% | 4.583% | 2.607% | 0.5448% | 1.286% | 5.504% | 3.12% | 1.677% | 1.913% | 0.7257% | 1.515% | 1.346% | 3.758% | 4.158% | 1.618% | 0.9869% | 1.014% | 3.993% | 4.545% | 4.862% | 0.448% | 9.357% | 4.399% | 5.274% | 2.699% | 8.413% | 6.543% | 0.6589% | 0.2412% | 2.21% | 0.2049% | 1.669% | 0.7935% | 1.777% | 0.1771% | 4.994% | 1.774% | 0.8392% | 3.835% | 0.2609%
89 | MT 28 | R | 6/3/2023 12:27:06 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 28.091 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 3.866% | 6.313% | 2.693% | 0.8543% | 6.312% | 4.65% | 10.96% | 9.451% | 5.521% | 1.111% | 1.083% | 15.39% | 5.883% | 5.313% | 6.422% | 5.057% | 1.695% | 1.146% | 5.546% | 11.3% | 15.81% | 15.09% | 2.687% | 23.65% | 13.55% | 16.35% | 6.987% | 13.71% | 2.118% | 15.25% | 4.661% | 13.76% | 4.393% | 1.624% | 1.204% | 1.882% | 14.63% | 0.5278% | 0.8633% | 7.68% | 1.308%
90 | MT 29 | R | 6/3/2023 12:33:08 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 29.092 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.106% | 3.027% | 3.002% | 0.6309% | 1.438% | 1.397% | 13.94% | 8.096% | 3.78% | 4.424% | 3.251% | 2.08% | 3.354% | 1.728% | 0.9347% | 2.718% | 2.868% | 3.206% | 2.57% | 3.82% | 6.424% | 1.977% | 4.226% | 3.577% | 2.352% | 0.532% | 1.318% | 1.767% | 2.741% | 1.798% | 1.462% | 1.066% | 2.094% | 2.753% | 16.85% | 1.757% | 5.311% | 3.484% | 3.874% | 1.861% | 20.24%
91 | LBLANK | R | 6/3/2023 12:39:20 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LBLANK.093 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.37% | 0% | 6.649% | 5.276% | 12.82% | 3.926% | 509.1% | 15.93% | 73.45% | 85.52% | 4.994% | 18.63% | 157.3% | 96.52% | 135.8% | 41.38% | 30.47% | 1.225% | 53.41% | 53.49% | 5.734% | 6.007% | 21.84% | 1.913e-06% | 14.96% | 61.07% | 129.6% | 10869% | 32.64% | 4.501% | 0.7584% | 0.9775% | 1.084% | 2.596% | 3.789% | 0.669% | 0.5209% | 1853% | 132.7% | 10.51% | 14.29%
92 | CCV 1ppb TRACE | R | 6/3/2023 12:45:23 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.094 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 5.948% | 3.501% | 54.42% | 3.939% | 6.618% | 4.002% | 1.136% | 4.296% | 1.003% | 3.804% | 7.579% | 9.006% | 4.403% | 3.025% | 4.557% | 1.887% | 2.686% | 4.815% | 4.829% | 9.42% | 10.86% | 9.438% | 9.622% | 2.905% | 12.69% | 9.804% | 15.23% | 14.02% | 5.871% | 8.455% | 1.268% | 8.323% | 1.44% | 4.163% | 4.235% | 1.407% | 10.07% | 4.553% | 2.589% | 2.2% | 5.234%
93 | CCV 1ppb TRACE | R | 6/3/2023 12:51:55 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.095 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.75% | 2.87% | 23.42% | 5.958% | 3.814% | 3.214% | 3.675% | 2.902% | 4.847% | 2.204% | 3.891% | 1.656% | 5.663% | 1.162% | 2.027% | 5.078% | 2.98% | 5.546% | 7.185% | 8.832% | 4.362% | 1.988% | 19.85% | 5.485% | 6.725% | 3.624% | 1.032% | 8.339% | 3.167% | 1.826% | 2.273% | 0.9341% | 1.173% | 1.371% | 0.5272% | 1.209% | 3.074% | 1.878% | 2.029% | 0.6825% | 0.5742%
94 | LFB 10ppb | R | 6/3/2023 12:57:57 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LFB 10ppb.096 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 6.479% | 6.773% | 10.44% | 0.6184% | 8.324% | 5.229% | 5.563% | 6.562% | 6.595% | 2.646% | 0.734% | 16.26% | 9.819% | 6.173% | 5.83% | 6.387% | 1.221% | 0.7609% | 1.977% | 10.19% | 15.19% | 17.44% | 0.3818% | 14.59% | 15.4% | 16.71% | 6.627% | 14.32% | 1.044% | 17.36% | 2.397% | 15.87% | 5.813% | 0.4013% | 0.8652% | 1.072% | 13.21% | 1.089% | 0.4799% | 0.9497% | 0.6716%
95 | LFB 10ppb | R | 6/3/2023 1:04:27 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LFB 10ppb.097 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 3.226% | 4.909% | 5.834% | 1.612% | 4.988% | 2.604% | 4.783% | 4.938% | 4.362% | 1.094% | 2.212% | 0.7559% | 2.498% | 3.242% | 1.443% | 4.107% | 1.085% | 1.055% | 1.501% | 1.933% | 1.978% | 1.98% | 3.624% | 6.41% | 3.781% | 0.6603% | 1.488% | 6.422% | 1.074% | 1.513% | 1.132% | 0.8321% | 2.777% | 0.9989% | 0.9717% | 1.255% | 1.482% | 1.332% | 1.483% | 1.531% | 1.077%
96 | CCV 10 ppb MAJORS | R | 6/3/2023 1:10:29 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.098 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.624% | 3.234% | 3.879% | 3.1% | 3.534% | 5.123% | 5.181% | 3.604% | 2.152% | 2.706% | 5.683% | 0.2545% | 2.86% | 1.923% | 2.331% | 2.025% | 3.592% | 4.034% | 2.886% | 2.293% | 6.948% | 1.066% | 2.286% | 3.462% | 3.019% | 1.253% | 2.62% | 0.9666% | 1.44% | 0.8506% | 0.3343% | 0.3827% | 1.653% | 2.011% | 1.12% | 2.201% | 5.925% | 2.72% | 2.504% | 1.391% | 0.5987%
97 | CCV 10 ppb MAJORS | R | 6/3/2023 1:17:01 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.099 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.098% | 3.487% | 5.53% | 0.5864% | 1.797% | 2.294% | 1.805% | 3.304% | 2.421% | 2.26% | 1.377% | 1.18% | 2.519% | 0.4378% | 0.3904% | 1.041% | 2.024% | 1.82% | 8.918% | 0.9244% | 3.257% | 1.484% | 9.366% | 3.707% | 4.173% | 2.946% | 7.817% | 5.04% | 2.602% | 0.4093% | 6.855% | 0.9804% | 0.7228% | 1.684% | 0.4096% | 0.9582% | 4.193% | 1.87% | 3.021% | 7.997% | 1.179%
98 | MT 30 | R | 6/3/2023 1:23:04 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 30.100 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 5.106% | 9.317% | 3.099% | 4.249% | 1.862% | 43.69% | 7.28% | 1.814% | 2.209% | 2.047% | 7.813% | 3.373% | 2.749% | 5.812% | 4.025% | 2.346% | 5.245% | 4.769% | 0.9263% | 8.015% | 3.952% | 4.201% | 16.94% | 15.67% | 3.221% | 12.63% | 8.191% | 23.22% | 6.608% | 5.629% | 0.4139% | 1.587% | 1.933% | 4.209% | 0.3621% | 5.189% | 3.401% | 6.179% | 4.421% | 0.9682% | 0.5173%
99 | MT 31 | R | 6/3/2023 1:29:15 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 31.101 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 22.37% | 4.343% | 3.392% | 2.847% | 9.62% | 6.21% | 7.881% | 3.222% | 9.721% | 5.678% | 5.811% | 3.127% | 1.493% | 1.272% | 3.517% | 2.507% | 6.198% | 5.607% | 2.905% | 7.175% | 5.801% | 2.287% | 35.75% | 7.471% | 8.088% | 11.1% | 24.87% | 20.15% | 3.699% | 6.893% | 1.469% | 2.868% | 1.476% | 5.268% | 1.575% | 1.621% | 5.433% | 5.362% | 4.596% | 1.479% | 1.366%
100 | MT 31 LD | R | 6/3/2023 1:35:29 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 31 LD.102 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 34.02% | 7.602% | 4.386% | 1.819% | 15.27% | 1.439% | 4.965% | 1.628% | 11.66% | 0.5166% | 1.179% | 0.2318% | 5.053% | 5.599% | 6.682% | 4.972% | 2.199% | 1.947% | 5.867% | 13.16% | 2.651% | 0.5154% | 7.415% | 0.3956% | 7.853% | 8.603% | 18.08% | 20.98% | 1.492% | 4.096% | 1.459% | 0.4163% | 3.757% | 1.122% | 0.5867% | 0.6212% | 1.875% | 2.1% | 1.481% | 1.579% | 0.3002%
101 | MT 31 LD | R | 6/3/2023 1:42:00 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 31 LD.103 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 27.67% | 4.83% | 5.93% | 0.08368% | 7.076% | 4.173% | 12.83% | 7.687% | 5.77% | 1.206% | 1.511% | 1.088% | 4.336% | 7.065% | 12.47% | 5.135% | 1.652% | 0.9098% | 4.842% | 15.1% | 4.597% | 4.381% | 20.3% | 14.9% | 2.27% | 7.898% | 35.47% | 28.36% | 2.115% | 0.146% | 1.708% | 0.6476% | 4.019% | 0.7374% | 0.9438% | 0.5114% | 3.925% | 1.828% | 1.262% | 1.643% | 3.397%
102 | MT 32 | R | 6/3/2023 1:48:02 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 32.104 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.71% | 3.35% | 34.27% | 2.673% | 4.767% | 9.346% | 63.28% | 13.06% | 21.42% | 11.1% | 1.825% | 6.28% | 81.44% | 110.1% | 39.65% | 4.391% | 2.826% | 3.815% | 61.09% | 0% | 12.15% | 12.25% | 50.84% | 40.02% | 20.93% | 20.62% | 1984% | 25070% | 96.36% | 34.76% | 0.8806% | 1.86% | 1.304% | 3.723% | 7.835% | 2.144% | 1.085% | 62.56% | 51.03% | 2.494% | 12.68%
103 | MT 32 LSP | R | 6/3/2023 1:54:03 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 32 LSP.105 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.917% | 1.295% | 9.692% | 3.626% | 1.855% | 1.628% | 1.25% | 2.219% | 3.069% | 1.455% | 3.866% | 1.855% | 0.9288% | 1.584% | 2.217% | 0.4509% | 1.751% | 1.04% | 3.079% | 1.014% | 1.844% | 0.6079% | 0.5495% | 7.547% | 3.097% | 3.101% | 5.471% | 4.332% | 1.029% | 1.727% | 1.062% | 1.271% | 0.8434% | 0.7635% | 2.126% | 4.725% | 2.145% | 0.5837% | 0.7866% | 1.877% | 3.979%
104 | MT 32 LSP | R | 6/3/2023 2:00:34 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 32 LSP.106 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 3.175% | 0.813% | 5.883% | 2.136% | 2.74% | 1.72% | 2.152% | 0.369% | 2.416% | 7.759% | 8.407% | 1.056% | 2.208% | 3.157% | 1.165% | 2.662% | 8.969% | 7.508% | 14% | 3.551% | 4.408% | 2.746% | 14.76% | 5.567% | 3.231% | 4.298% | 17.32% | 2.207% | 8.441% | 1.999% | 10.81% | 1.89% | 1.409% | 6.059% | 1.08% | 1.811% | 4.381% | 6.393% | 8.135% | 13.04% | 1.272%
105 | MT 33 | R | 6/3/2023 2:06:37 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 33.107 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 3.617% | 3.185% | 10.04% | 1.712% | 13.5% | 2.766% | 63.7% | 19.07% | 36.73% | 24.96% | 11.4% | 31.12% | 115.8% | 45.09% | 65.21% | 5.507% | 25.86% | 7.733% | 26.58% | 0% | 14.72% | 17.43% | 9.773% | 1.913e-06% | 11.62% | 52.79% | 136.9% | 360.1% | 87.75% | 5.252% | 0.4069% | 0.386% | 1.525% | 18.74% | 1.845% | 0.9518% | 0.1119% | 79.48% | 24.28% | 6.943% | 3.776%
106 | MT 34 | R | 6/3/2023 2:12:39 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 34.108 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.679% | 0% | 14.96% | 1.122% | 4.277% | 3.703% | 127.7% | 18.68% | 12.21% | 17.89% | 0.9594% | 6.57% | 5.074% | 72.44% | 25.35% | 34.97% | 26% | 6.102% | 48.56% | 543.8% | 22.42% | 23.39% | 76.31% | 51.65% | 23.28% | 20.13% | 2633% | 6629% | 10.09% | 5.759% | 12.01% | 0.2628% | 2.413% | 0.8015% | 2.602% | 1.5% | 2.47% | 58.34% | 9.322% | 19.07% | 5.522%
107 | MT 35 | R | 6/3/2023 2:18:42 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 35.109 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.8654% | 3.191% | 23.2% | 2.244% | 10.92% | 5.154% | 59.72% | 11.68% | 30.04% | 30.19% | 38.96% | 96.32% | 3.408% | 36.06% | 118.6% | 13.75% | 31.86% | 5.084% | 43% | 545.6% | 2.669% | 2.52% | 24.63% | 56.5% | 19.65% | 77.5% | 989.6% | 0% | 12.41% | 3.683% | 7.479% | 4.168% | 1.385% | 7.165% | 2.089% | 2.639% | 4.96% | 41.4% | 84.64% | 11.49% | 7.22%
108 | LBLANK | R | 6/3/2023 2:24:45 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LBLANK.110 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.554% | 0% | 9.879% | 5.998% | 2.511% | 2.243% | 516.3% | 19.37% | 27.34% | 77.62% | 2.989% | 55.72% | 580.5% | 1% | 108% | 335.7% | 7.564% | 0.9124% | 17.97% | 0% | 7.962% | 8.612% | 24.04% | 1.913e-06% | 30.95% | 15.38% | 125.7% | 0% | 10.93% | 0.9506% | 0.5983% | 1.21% | 0.8168% | 0.6004% | 7.613% | 1.681% | 4.429% | 16.41% | 121.9% | 16.26% | 23.58%
109 | CCV 1ppb TRACE | R | 6/3/2023 2:30:48 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.111 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 3.442% | 6.171% | 38.12% | 3.446% | 4.616% | 3.734% | 5.519% | 3.343% | 4.769% | 13.35% | 17.5% | 4.487% | 8.635% | 2.523% | 3.298% | 3.274% | 12.67% | 15.5% | 4.159% | 21.59% | 0.6092% | 2.786% | 9.757% | 0.5536% | 6.978% | 5.144% | 3.783% | 7.349% | 10.93% | 1.445% | 0.601% | 2.968% | 2.31% | 8.74% | 5.73% | 1.739% | 0.9405% | 11.36% | 11.91% | 0.327% | 6.637%
110 | CCV 1ppb TRACE | R | 6/3/2023 2:37:20 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.112 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.023% | 5.322% | 25.63% | 38.95% | 12.8% | 8.904% | 5.019% | 9.531% | 3.68% | 3.555% | 7.277% | 2.887% | 1.833% | 2.06% | 2.503% | 5.291% | 1.141% | 1.809% | 10.22% | 7.592% | 1.208% | 1.308% | 7.961% | 6.522% | 1.532% | 3.655% | 7.48% | 2.498% | 3.383% | 2.064% | 2.934% | 1.181% | 3.025% | 1.702% | 3.526% | 23.81% | 1.665% | 3.322% | 1.461% | 4.874% | 2.72%
111 | CCV 10 ppb MAJORS | R | 6/3/2023 2:43:23 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.113 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 3.521% | 1.558% | 12.72% | 0.31% | 2.552% | 1.782% | 4.088% | 4.162% | 3.817% | 0.8657% | 1.042% | 3.068% | 1.401% | 4.096% | 2.311% | 3.634% | 0.7136% | 1.003% | 2.384% | 5.94% | 6.486% | 4.228% | 4.947% | 4.066% | 1.923% | 0.736% | 2.108% | 1.743% | 0.2951% | 3.179% | 1.635% | 2.953% | 3.071% | 0.2421% | 1.407% | 0.7519% | 5.327% | 0.8994% | 0.4301% | 0.8668% | 0.7729%
112 | CCV 10 ppb MAJORS | R | 6/3/2023 2:49:56 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.114 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 4.187% | 5.734% | 3.78% | 1.775% | 3.484% | 2.948% | 5.583% | 5.234% | 4.433% | 2.651% | 3.604% | 2.718% | 5.387% | 3.912% | 3.249% | 3.922% | 3.29% | 3.495% | 10.33% | 1.573% | 1.812% | 1.554% | 10.43% | 2.432% | 2.5% | 3.137% | 5.798% | 8% | 2.784% | 1.282% | 8.621% | 1.044% | 3.708% | 2.385% | 0.6965% | 1.192% | 1.563% | 4.619% | 2.197% | 10.49% | 0.8466%
113 | MT 36 | R | 6/3/2023 2:55:59 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 36.115 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 3.731% | 5.811% | 23.23% | 2.289% | 4.474% | 3.681% | 96.23% | 14.9% | 14.51% | 4.098% | 41.04% | 19.01% | 8.632% | 70.03% | 96.32% | 33.45% | 13% | 6.195% | 19.56% | 629.8% | 2.789% | 3.012% | 54.39% | 1.913e-06% | 10.49% | 15.16% | 633.6% | 0% | 33.94% | 8.795% | 0.8226% | 0.5796% | 3.575% | 0.7782% | 1.6% | 1.374% | 2.5% | 35.99% | 26.19% | 8.647% | 6.124%
114 | MT 37 | R | 6/3/2023 3:02:01 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 37.116 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.245% | 3.739% | 26.24% | 2.652% | 15.85% | 8.207% | 19.76% | 12.63% | 27.83% | 48.91% | 3.296% | 27.99% | 49.69% | 123.3% | 26.29% | 9.857% | 4.98% | 8.335% | 68.1% | 459.1% | 25.96% | 31.84% | 61.95% | 51.04% | 9.853% | 101% | 60.44% | 0% | 5.621% | 3.525% | 4.949% | 1.962% | 5.157% | 1.024% | 0.7732% | 2.17% | 2.238% | 25.4% | 40.24% | 11.57% | 9.194%
115 | MT 38 | R | 6/3/2023 3:08:04 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 38.117 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.8208% | 3.302% | 33.41% | 8.5% | 2.598% | 0.8989% | 176.6% | 10.85% | 20.44% | 21.09% | 3.219% | 117.2% | 40.88% | 57.02% | 1055% | 14.02% | 13.28% | 4.475% | 9.276% | 536.3% | 9.607% | 9.214% | 28.85% | 1.913e-06% | 40.3% | 16.63% | 366% | 0% | 7.194% | 5.311% | 0.8482% | 1.332% | 2.295% | 0.8309% | 5.724% | 4.168% | 4.579% | 62.32% | 72.38% | 7.691% | 8.309%
116 | MT 39 | R | 6/3/2023 3:14:07 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 39.118 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.7144% | 3.373% | 10.76% | 4.033% | 9.543% | 1.177% | 56.14% | 20.93% | 38.85% | 9.028% | 5.792% | 29.71% | 8.809% | 104.9% | 287.5% | 12.01% | 10.56% | 7.671% | 66.94% | 578.2% | 3.718% | 4.862% | 115.1% | 52.7% | 33.16% | 161.3% | 197.6% | 147.2% | 5.816% | 0.6795% | 19.11% | 3.826% | 2.82% | 0.4397% | 0.3909% | 1.465% | 3.923% | 18.66% | 28.8% | 37.77% | 7.187%
117 | MT 40 | R | 6/3/2023 3:20:10 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 40.119 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.466% | 23.65% | 14.17% | 2.797% | 7.649% | 4.032% | 16.9% | 23.13% | 44.91% | 38.22% | 9.609% | 3.964% | 39.83% | 48.8% | 67.84% | 9.978% | 6.72% | 6.331% | 199.9% | 133.6% | 167.9% | 110.7% | 189.3% | 1.913e-06% | 49.61% | 49.08% | 75.88% | 6914% | 20.02% | 59.93% | 10.63% | 1.623% | 2.306% | 1.998% | 0.6156% | 1.489% | 0.5008% | 38.1% | 29.02% | 8.73% | 4.482%
118 | MT 31 | R | 6/3/2023 3:26:12 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 31 .120 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 17.83% | 13.18% | 2.139% | 1.99% | 4.064% | 9.164% | 3.943% | 4.525% | 4.884% | 17.74% | 15.91% | 1.564% | 1.023% | 4.959% | 3.793% | 1.854% | 15.16% | 14.15% | 1.32% | 1.576% | 1.051% | 2.538% | 38.74% | 26.83% | 13.14% | 17.31% | 17.44% | 23.44% | 20.31% | 5.563% | 2.015% | 1.575% | 0.7747% | 12.16% | 0.8676% | 1.007% | 0.9116% | 15.24% | 14.59% | 2.643% | 0.9184%
119 | MT 31 LD | R | 6/3/2023 3:32:23 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 31 LD.121 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 21.14% | 30.66% | 11.95% | 1.656% | 14.42% | 17.45% | 15.44% | 9.072% | 12.26% | 5.465% | 8.428% | 3.007% | 11.54% | 9.294% | 7.667% | 10.53% | 8.143% | 7.744% | 2.342% | 14.73% | 4.46% | 4.599% | 24.68% | 38.01% | 19.13% | 14.58% | 43.13% | 12.58% | 7.2% | 4.137% | 1.263% | 2.565% | 8.462% | 5.954% | 6.666% | 0.5211% | 2.805% | 7.886% | 8.1% | 1.254% | 10.11%
120 | MT 31 LD | R | 6/3/2023 3:38:54 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 31 LD.122 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 28.83% | 8.065% | 3.892% | 1.411% | 9.259% | 5.563% | 10.47% | 9.684% | 0.6946% | 14.18% | 9.583% | 2.72% | 1.885% | 0.947% | 8.411% | 2.519% | 10.09% | 10.44% | 3.486% | 16.2% | 3.931% | 2.506% | 57.26% | 30.78% | 10.73% | 11% | 8.011% | 22% | 9.121% | 2.648% | 0.8365% | 0.5805% | 1.543% | 7.989% | 6.517% | 0.6336% | 2.476% | 11.87% | 10.63% | 3.039% | 8.731%
121 | MT 32 | R | 6/3/2023 3:44:56 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 32.123 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.8656% | 0% | 15.71% | 3.894% | 8.323% | 2.396% | 98.84% | 13.58% | 29.11% | 1.681% | 2.879% | 40.42% | 6.886% | 101.7% | 100.6% | 15.87% | 3.643% | 1.479% | 40.4% | 0% | 27.08% | 30.47% | 10.58% | 54.25% | 35.63% | 22.38% | 173.2% | 0% | 259.4% | 120.8% | 1.818% | 0.7765% | 1.688% | 1.42% | 0.739% | 0.6329% | 3.134% | 38.36% | 109.9% | 6.084% | 4.469%
122 | MT 32 LSP | R | 6/3/2023 3:50:58 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 32 LSP.124 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 3.166% | 1.131% | 6.6% | 14.91% | 2.02% | 2.269% | 1.652% | 2.565% | 2.021% | 1.788% | 0.2836% | 1.338% | 1.662% | 2.782% | 1.58% | 2.013% | 0.4357% | 1.117% | 1.375% | 7.527% | 3.726% | 2.196% | 1.321% | 9.731% | 2.853% | 2.953% | 5.781% | 9.786% | 0.2524% | 1.711% | 0.2819% | 2.234% | 1.325% | 1.038% | 7.669% | 12.56% | 4.573% | 0.6668% | 1.67% | 0.4211% | 7.915%
123 | MT 32 LSP | R | 6/3/2023 3:57:29 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 32 LSP.125 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.153% | 1.386% | 5.68% | 3.198% | 1.044% | 2.244% | 1.516% | 1.398% | 1.913% | 2.12% | 2.576% | 1.428% | 0.7183% | 1.277% | 1.566% | 1.989% | 1.555% | 1.976% | 2.746% | 6.983% | 4.155% | 1.783% | 6.168% | 8.162% | 5.033% | 3.166% | 2.152% | 4.272% | 0.6917% | 1.93% | 0.3078% | 1.06% | 0.3866% | 1.909% | 7.714% | 1.227% | 5.496% | 2.247% | 1.994% | 0.8457% | 10.81%
124 | LBLANK | R | 6/3/2023 4:03:32 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LBLANK.126 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.06% | 3.069% | 12.06% | 26.5% | 2.724% | 6.055% | 396.1% | 23.2% | 26.36% | 17.41% | 2.059% | 27.36% | 1426% | 25.64% | 116.3% | 188.6% | 11.93% | 1.452% | 24.68% | 270.3% | 6.718% | 11.97% | 49.03% | 1.913e-06% | 10.34% | 159.4% | 78.17% | 0% | 54.05% | 3.984% | 5.275% | 1.506% | 1.192% | 0.8449% | 0.8595% | 0.4695% | 4.999% | 17.92% | 108.9% | 25.08% | 27.36%
125 | CCV 1ppb TRACE | R | 6/3/2023 4:09:35 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.127 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.82% | 3.482% | 19.57% | 3.575% | 2.652% | 1.283% | 6.773% | 6.288% | 3.292% | 2.5% | 1.034% | 2.452% | 2.793% | 5.893% | 0.8813% | 0.715% | 2.603% | 2.989% | 5.178% | 2.059% | 1.766% | 0.4995% | 12.49% | 10.69% | 6.703% | 4.41% | 1.572% | 9.213% | 5.312% | 0.6206% | 1.071% | 1.119% | 2.577% | 0.4788% | 1.127% | 0.8431% | 1.445% | 2.708% | 1.174% | 1.693% | 0.8479%
126 | CCV 1ppb TRACE | R | 6/3/2023 4:16:07 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.128 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 11.46% | 10.72% | 14.03% | 1.845% | 19.84% | 11.46% | 17.48% | 13.92% | 12.7% | 3.525% | 12.56% | 1.424% | 21.12% | 13.49% | 9% | 11.95% | 4.876% | 4.437% | 1.643% | 25.8% | 3.439% | 4.757% | 23.84% | 6.899% | 0.4799% | 11.22% | 5.006% | 4.875% | 3.343% | 2.117% | 5.431% | 0.7013% | 10.07% | 3.118% | 1.005% | 1.87% | 1.716% | 3.606% | 4.251% | 5.683% | 1.473%
127 | LFB 10ppb | R | 6/3/2023 4:22:09 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LFB 10ppb.129 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.532% | 2.704% | 5.621% | 1.622% | 0.7128% | 3.042% | 1.268% | 1.101% | 2.181% | 0.8263% | 0.69% | 1.434% | 1.483% | 2.019% | 2.898% | 0.5611% | 0.9446% | 1.948% | 1.518% | 5.708% | 1.903% | 1.63% | 2.445% | 5.341% | 3.224% | 0.7633% | 2.007% | 1.244% | 1.879% | 2.143% | 1.054% | 0.6958% | 0.4979% | 0.5974% | 7.309% | 0.6011% | 1.086% | 1.547% | 1.357% | 0.4232% | 5.574%
128 | LFB 10ppb | R | 6/3/2023 4:28:39 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LFB 10ppb.130 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.41% | 1.113% | 5.777% | 0.9092% | 1.288% | 0.7453% | 5.049% | 1.052% | 1.281% | 3.475% | 5.484% | 2.945% | 0.4864% | 1.09% | 2.489% | 1.393% | 3.904% | 3.691% | 4.12% | 4.296% | 2.968% | 3.157% | 6.041% | 6.106% | 2.817% | 4.777% | 3.111% | 5.204% | 3.312% | 4.524% | 1.593% | 2.194% | 0.86% | 4.347% | 1.464% | 0.8585% | 2.151% | 3.61% | 3.13% | 1.838% | 1.404%
129 | CCV 10 ppb MAJORS | R | 6/3/2023 4:34:42 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.131 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 17.06% | 17.78% | 12.01% | 0.6695% | 19.36% | 15.46% | 17.94% | 16.05% | 16.74% | 3.114% | 1.364% | 14.14% | 18.74% | 17.84% | 17.48% | 15.59% | 2.199% | 2.163% | 3.628% | 14.9% | 16.24% | 13.97% | 2.727% | 10.53% | 12.14% | 11.74% | 4.725% | 11.87% | 1.189% | 13.28% | 1.699% | 12% | 15.21% | 0.6787% | 1.577% | 0.1978% | 14.5% | 1.788% | 1.813% | 0.8238% | 0.2824%
130 | CCV 10 ppb MAJORS | R | 6/3/2023 4:41:13 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.132 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.886% | 3.729% | 6.583% | 1.519% | 3.408% | 2.21% | 4.438% | 3.545% | 2.756% | 2.549% | 4.905% | 1.891% | 2.408% | 1.387% | 2.797% | 1.824% | 4.024% | 3.939% | 1.24% | 5.405% | 2.658% | 1.699% | 2.477% | 1.585% | 3.474% | 1.607% | 4.392% | 0.08698% | 2.917% | 1.429% | 0.9461% | 0.08694% | 1.411% | 2.23% | 13.62% | 1.327% | 1.117% | 0.8087% | 2.53% | 0.7924% | 17.07%
131 | MT 33 | R | 6/3/2023 4:47:16 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 33.133 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.19% | 3.167% | 17.6% | 1.402% | 4.107% | 3.218% | 74.55% | 12.33% | 48.03% | 35.24% | 5.567% | 33.35% | 73.46% | 45.78% | 105.7% | 8.417% | 17.83% | 1.256% | 36.72% | 268.2% | 10.61% | 11.74% | 37.53% | 1.913e-06% | 23.81% | 21.04% | 3609% | 0% | 24.01% | 2.382% | 1.426% | 0.847% | 1.316% | 4.714% | 0.5451% | 1.164% | 0.8964% | 1.705% | 116.2% | 6.219% | 0.9917%
132 | MT 34 | R | 6/3/2023 4:53:19 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 34.134 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.566% | 3.123% | 44.72% | 2.02% | 9.622% | 8.05% | 94.63% | 10.55% | 33.19% | 7.335% | 5.98% | 38.94% | 8.145% | 105.6% | 28.05% | 21.89% | 75.63% | 2.009% | 55.38% | 545% | 6.522% | 8.03% | 28.85% | 80.48% | 10.47% | 20.67% | 828.2% | 0% | 11.18% | 2.764% | 4.029% | 0.6439% | 4.583% | 4.222% | 0.4155% | 0.8662% | 2.321% | 38.52% | 109.3% | 4.381% | 3.107%
133 | MT 35 | R | 6/3/2023 4:59:21 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 35.135 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.26% | 0% | 30.76% | 0.9326% | 8.333% | 3.415% | 65.3% | 13.22% | 12.11% | 33.05% | 12.89% | 34.3% | 12.59% | 123% | 39.36% | 25.75% | 45.16% | 6.499% | 71.59% | 129.7% | 14.44% | 15.51% | 44.15% | 1.913e-06% | 10.77% | 96.71% | 1517% | 5740% | 2.398% | 6.951% | 1.775% | 0.8806% | 1.538% | 4.688% | 1.014% | 1.391% | 2.195% | 38.76% | 67.03% | 1.515% | 8.488%
134 | MT 36 | R | 6/3/2023 5:05:25 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 36.136 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.07978% | 3.206% | 3.206% | 2.589% | 7.925% | 9.367% | 142.2% | 13.1% | 24.65% | 25.05% | 134.2% | 48.65% | 83.38% | 53.04% | 197.5% | 19.51% | 23.34% | 8.741% | 128.3% | 130.9% | 13.6% | 16.12% | 99.77% | 52.96% | 50.2% | 21.5% | 39.59% | 169.5% | 5.281% | 9.279% | 10.71% | 1.222% | 1.98% | 5.821% | 0.4248% | 1.06% | 1.019% | 19.77% | 125% | 27.17% | 6.886%
135 | MT 37 | R | 6/3/2023 5:11:28 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 37.137 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 3.003% | 0% | 12.34% | 1.529% | 10.11% | 5.594% | 29.21% | 24.17% | 32.13% | 24.75% | 10.64% | 30.28% | 16.64% | 391.3% | 39.2% | 7.381% | 22.04% | 58.63% | 3.989% | 523.9% | 6.758% | 8.337% | 40.8% | 58.51% | 10.8% | 61.82% | 58.06% | 0% | 13.43% | 11.19% | 4.519% | 2.484% | 2.159% | 18.06% | 1.801% | 1.231% | 3.304% | 17.34% | 92.12% | 7.04% | 2.894%
136 | MT 38 | R | 6/3/2023 5:17:31 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 38.138 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.4582% | 10.05% | 23.87% | 1.18% | 8.072% | 6.814% | 456.8% | 23.64% | 1.936% | 8.091% | 6.019% | 63.89% | 33.79% | 26.82% | 64.12% | 18.32% | 14.16% | 5.825% | 20.01% | 131.4% | 9.411% | 10.04% | 59.14% | 1.913e-06% | 0.05029% | 32.45% | 29% | 39063% | 14.65% | 4.888% | 1.347% | 0.2564% | 0.7302% | 0.8453% | 1.715% | 0.7716% | 1.678% | 25.95% | 57.05% | 6.854% | 6.305%
137 | MT 39 | R | 6/3/2023 5:23:33 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 39.139 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.873% | 3.463% | 22.48% | 1.557% | 8.827% | 12.71% | 96.17% | 12.39% | 19.49% | 7.85% | 6.899% | 29.85% | 7.382% | 48.66% | 66.12% | 19.11% | 25.45% | 9.484% | 32.78% | 498% | 13.91% | 16.29% | 76.49% | 62.12% | 22.4% | 45.74% | 145.1% | 0% | 9.316% | 1.569% | 1.948% | 9.875% | 0.3066% | 2.624% | 0.9186% | 2.377% | 10.02% | 25.03% | 72.74% | 4.352% | 5.333%
138 | MT 40 | R | 6/3/2023 5:29:36 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 40.140 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.048% | 6.226% | 17.96% | 1.096% | 15.15% | 3.185% | 35.79% | 25.11% | 16.56% | 47.66% | 2.689% | 8.254% | 37.76% | 146.9% | 119.8% | 10.27% | 16.81% | 3.037% | 22.53% | 125.2% | 129.7% | 73.36% | 31.15% | 52.65% | 0% | 20.52% | 317% | 169% | 10.8% | 42.09% | 1.063% | 1.543% | 6.92% | 1.427% | 1.69% | 0.3667% | 0.7927% | 69.21% | 119.4% | 4.117% | 8.831%
139 | MT 32 B | R | 6/3/2023 5:35:38 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 32 B.141 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.248% | 3.227% | 44.58% | 1.286% | 7.227% | 6.965% | 134.4% | 12.35% | 12.42% | 9.332% | 2.597% | 116.2% | 45.93% | 77.21% | 56.56% | 14.77% | 5.802% | 7.295% | 41.17% | 548.4% | 1.713% | 3.483% | 9.395% | 1.913e-06% | 0% | 50.33% | 1972% | 0% | 10.44% | 6.284% | 0.6845% | 7.508% | 3.367% | 6.031% | 0.6136% | 0.6488% | 4.381% | 38.13% | 55.16% | 2.221% | 2.057%
140 | MT 32 B LD | R | 6/3/2023 5:41:40 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 32 B LD.142 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 3.596% | 3.4% | 22.23% | 0.9923% | 0.7548% | 8.769% | 79.43% | 13.35% | 12.21% | 10.44% | 1.908% | 82.96% | 38.89% | 120.4% | 126.8% | 14.09% | 8.058% | 4.481% | 65.61% | 481.2% | 16.48% | 18.45% | 49.1% | 1.913e-06% | 10.55% | 34.17% | 867.1% | 3139% | 9.086% | 5.606% | 0.3621% | 1.906% | 1.226% | 1.253% | 0.9606% | 0.4011% | 3.324% | 48.15% | 54.99% | 8.808% | 1.298%
141 | LBLANK | R | 6/3/2023 5:47:43 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LBLANK.143 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.4898% | 0% | 3.346% | 4.157% | 8.847% | 2.587% | 147.3% | 20.61% | 23.33% | 35.53% | 5.356% | 27.26% | 46.57% | 117.5% | 282.4% | 143.9% | 16.13% | 4.173% | 117.6% | 0% | 9.837% | 10.63% | 139.6% | 41.71% | 11.32% | 16.19% | 444.6% | 2502% | 4.165% | 2.904% | 6.251% | 1.114% | 1.838% | 7.988% | 2.69% | 1.065% | 2.821% | 38.75% | 256.2% | 10.01% | 8.619%
142 | CCV 1ppb TRACE | R | 6/3/2023 5:53:46 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.144 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 3.549% | 6.082% | 35.65% | 2.878% | 17.12% | 1.63% | 5.382% | 4.88% | 2.566% | 3.759% | 9.931% | 0.8264% | 1.093% | 2.84% | 3.528% | 4.12% | 6.474% | 3.446% | 3.651% | 23.42% | 4.97% | 3.216% | 11.97% | 11.57% | 8.034% | 7.816% | 8.758% | 11.88% | 6.21% | 1.298% | 1.078% | 0.3289% | 2.55% | 3.953% | 4.668% | 0.1957% | 2.008% | 2.573% | 5.656% | 1.998% | 4.935%
143 | CCV 1ppb TRACE | R | 6/3/2023 6:00:18 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.145 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 5.798% | 3.995% | 6.357% | 4.009% | 7.222% | 4.101% | 8.477% | 5.29% | 2.46% | 1.034% | 5.74% | 5.78% | 5.3% | 1.839% | 4.952% | 2.527% | 0.7814% | 2.838% | 7.915% | 7.688% | 5.425% | 4.996% | 24.8% | 7.429% | 2.635% | 2.889% | 4.758% | 6.808% | 0.8545% | 3.136% | 8.766% | 3.463% | 0.7633% | 0.6012% | 1.045% | 1.7% | 2.406% | 0.8395% | 0.8334% | 8.722% | 0.594%
144 | CCV 10 ppb MAJORS | R | 6/3/2023 6:06:21 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.146 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.753% | 3.469% | 3.909% | 0.4774% | 2.865% | 1.867% | 2.547% | 0.6524% | 1.112% | 1.267% | 1.062% | 1.464% | 1.067% | 2.983% | 2.914% | 1.602% | 0.531% | 1.444% | 2.364% | 3.024% | 3.875% | 1.96% | 5.14% | 6.383% | 1.951% | 1.729% | 2.465% | 4.278% | 0.821% | 2.333% | 2.134% | 0.7508% | 1.599% | 0.7586% | 1.771% | 1.899% | 2.694% | 1.577% | 0.6331% | 4.737% | 0.3781%
145 | CCV 10 ppb MAJORS | R | 6/3/2023 6:12:53 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.147 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 3.199% | 2.121% | 6.239% | 4.136% | 4.862% | 0.5411% | 0.9494% | 2.337% | 1.839% | 2.444% | 1.926% | 2.813% | 1.297% | 2.308% | 2.161% | 1.688% | 0.8851% | 0.2873% | 2.187% | 1.007% | 1.572% | 1.955% | 2.708% | 7.177% | 1.1% | 1.8% | 0.3137% | 3.387% | 0.8533% | 1.453% | 0.2414% | 1.628% | 1.252% | 0.6619% | 2.003% | 1.375% | 2.372% | 1.003% | 1.227% | 1.18% | 0.5548%
146 | MT 33 B | R | 6/3/2023 6:18:55 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 33 B.148 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.935% | 6.016% | 40.13% | 2.181% | 9.794% | 3.232% | 212.8% | 35.29% | 13.77% | 28.67% | 4.842% | 331.6% | 95.31% | 3.383% | 76.35% | 12.97% | 6.321% | 3.838% | 93.19% | 129.4% | 11.03% | 11.88% | 44.36% | 53.99% | 35.25% | 0% | 120.7% | 0% | 27.84% | 8.503% | 2.552% | 1.458% | 3.217% | 0.1093% | 1.454% | 2.021% | 3.766% | 158.5% | 26.74% | 13.92% | 6.482%
147 | MT 33 B LSP | R | 6/3/2023 6:24:56 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 33 B LSP.149 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.304% | 3.168% | 5.181% | 0.6773% | 3.592% | 2.638% | 2.155% | 3.986% | 2.861% | 17.68% | 17.93% | 1.785% | 2.042% | 0.7066% | 2.497% | 2.187% | 16.58% | 16.84% | 3.439% | 2.6% | 2.2% | 1.161% | 3.664% | 3.21% | 0.4546% | 1.902% | 0.7594% | 2.752% | 16.53% | 1.853% | 1.071% | 0.9567% | 2.891% | 13.69% | 1.78% | 0.411% | 2.661% | 16.45% | 16.34% | 3.08% | 0.9972%
148 | MT 34 B | R | 6/3/2023 6:30:59 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 34 B.150 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.789% | 0.1085% | 41.51% | 1.791% | 7.258% | 3.684% | 80.88% | 8.384% | 13.96% | 15.58% | 9.025% | 75.4% | 60.31% | 72.06% | 15.08% | 27.46% | 4.461% | 3.407% | 34.48% | 78.83% | 71.9% | 103.8% | 20.7% | 81.36% | 32.4% | 23.95% | 327.4% | 0% | 9.757% | 35.42% | 1.729% | 1.099% | 1.847% | 1.225% | 20.37% | 1.395% | 2.369% | 1.45% | 17.1% | 6.106% | 20.37%
149 | MT 35 B | R | 6/3/2023 6:37:01 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 35 B.151 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.2919% | 3.056% | 32.74% | 0.4262% | 14.12% | 4.574% | 80.62% | 26.05% | 5.211% | 21.44% | 1.813% | 96.7% | 87.4% | 144.1% | 69.93% | 21% | 4.788% | 1.184% | 29.29% | 629.8% | 16.17% | 21.55% | 48.19% | 53.15% | 32.39% | 97.69% | 69131% | 0% | 25.48% | 5.505% | 2.23% | 0.8355% | 5.945% | 0.5605% | 1.321% | 1.406% | 4.894% | 27.8% | 72.14% | 3.75% | 4.824%
150 | MT 36 B | R | 6/3/2023 6:43:03 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 36 B.152 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.186% | 0% | 12.22% | 1.612% | 4.51% | 9.325% | 1783% | 11.58% | 15.88% | 4.209% | 13.84% | 65.66% | 17.99% | 81.11% | 36.23% | 20.33% | 34.88% | 0.9452% | 24.54% | 128.5% | 24.61% | 29.26% | 34.39% | 1.913e-06% | 0% | 56.37% | 151% | 6296% | 4.475% | 2.502% | 2.537% | 0.8917% | 3.069% | 1.455% | 2.594% | 2.508% | 2.122% | 38.25% | 75.74% | 6.363% | 6.249%
151 | MT 37 B | R | 6/3/2023 6:49:06 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 37 B.153 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 3.585% | 3.347% | 18.98% | 1.09% | 17.96% | 3.381% | 214.6% | 19.31% | 30.39% | 23.13% | 4.685% | 11848% | 884.9% | 95.23% | 40.05% | 17.8% | 11.81% | 6.287% | 47.22% | 509.3% | 14.69% | 16.19% | 59.46% | 53.72% | 65.87% | 15.84% | 23.23% | 11459% | 3.433% | 9.742% | 7.634% | 1.203% | 1.467% | 4.744% | 0.639% | 0.4963% | 2.766% | 35.04% | 50.53% | 11.69% | 7.617%
152 | MT 38 B | R | 6/3/2023 6:55:09 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 38 B.154 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.005% | 3.358% | 44% | 1.439% | 17.66% | 9.429% | 173.2% | 18.86% | 4.615% | 26.05% | 2.373% | 36.68% | 1382% | 36.48% | 83.65% | 11.59% | 40.59% | 5.524% | 55.41% | 0% | 2.35% | 3.29% | 40.75% | 53.82% | 11.18% | 16.54% | 383.9% | 164.4% | 21.31% | 28.74% | 0.8405% | 1.932% | 3.406% | 5.569% | 1.434% | 2.02% | 2.232% | 58% | 105.8% | 9.576% | 7.88%
153 | MT 39 B | R | 6/3/2023 7:01:12 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 39 B.155 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.504% | 3.515% | 70.97% | 0.1078% | 11.31% | 9.404% | 29.05% | 10.69% | 13.56% | 25.91% | 6.541% | 32.36% | 87.4% | 117.5% | 61.86% | 8.698% | 26.37% | 6.66% | 63.36% | 128.7% | 7.382% | 8.397% | 70.59% | 53.29% | 19.9% | 57.92% | 129.7% | 6381% | 111.7% | 91.29% | 1.212% | 0.4638% | 2.127% | 15.68% | 0.6322% | 1.901% | 1.441% | 171.5% | 104.1% | 0.6844% | 5.3%
154 | MT 40 B | R | 6/3/2023 7:07:15 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 40 B.156 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.8179% | 3.614% | 35.58% | 2.17% | 12.2% | 3.044% | 82.06% | 20.67% | 23.47% | 42.96% | 6.994% | 70.2% | 157.4% | 42.83% | 78.37% | 7.057% | 4.467% | 2.853% | 57.83% | 461.9% | 18.91% | 21.31% | 35.32% | 54.86% | 20.93% | 93.09% | 80.47% | 2944% | 14.44% | 2.278% | 1.492% | 0.04178% | 5.454% | 2.629% | 1.283% | 0.2662% | 2.496% | 49.5% | 26.09% | 3.42% | 2.001%
155 | MT 32 C | R | 6/3/2023 7:13:17 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 32 C.157 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.8096% | 3.363% | 56.36% | 1.271% | 9.661% | 2.956% | 386.9% | 28.72% | 7.178% | 20.25% | 1.761% | 707.1% | 259.2% | 50.33% | 14.53% | 12.47% | 8.271% | 2.596% | 27.43% | 256.4% | 8.213% | 9.366% | 97.91% | 1.913e-06% | 21.76% | 42.21% | 135.4% | 0% | 6.278% | 9.953% | 1.581% | 4.394% | 0.7933% | 0.516% | 2.115% | 2.579% | 2.367% | 32.12% | 24.89% | 7.477% | 3.843%
156 | LBLANK | R | 6/3/2023 7:19:20 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LBLANK.158 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.03082% | 0% | 17.2% | 4.553% | 12.47% | 6.607% | 0% | 29.38% | 24.75% | 35.88% | 0.9258% | 52.46% | 124.7% | 4.655% | 109.5% | 923% | 21.51% | 1.935% | 6.987% | 553.4% | 5.978% | 7.441% | 74.85% | 1.913e-06% | 21.09% | 1.117% | 57.56% | 159.8% | 6.927% | 1.863% | 2.229% | 3.111% | 3.885% | 1.139% | 1.64% | 1.283% | 6.081% | 18.16% | 102% | 21.06% | 17.54%
157 | CCV 1ppb TRACE | R | 6/3/2023 7:25:23 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.159 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 6.661% | 4.622% | 17.55% | 2.424% | 15.74% | 3.409% | 8.538% | 7.455% | 1.862% | 0.5696% | 6.957% | 1.409% | 7.348% | 3.99% | 6.512% | 5.407% | 0.6342% | 2.083% | 3.042% | 11.28% | 1.861% | 2.106% | 11.55% | 22.06% | 16.14% | 5.461% | 1.799% | 13.5% | 1.538% | 3.377% | 2.265% | 1.394% | 4.541% | 0.2398% | 1.555% | 0.7395% | 0.3333% | 0.8452% | 2.203% | 0.6902% | 2.269%
158 | CCV 1ppb TRACE | R | 6/3/2023 7:31:56 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.160 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.293% | 4.351% | 14.63% | 3.113% | 5.675% | 3.013% | 5.161% | 0.851% | 2.701% | 16.17% | 25.65% | 14.09% | 5.031% | 2.098% | 0.08961% | 4.342% | 16.13% | 20.06% | 3.756% | 14.52% | 17.2% | 17.55% | 8.201% | 31.28% | 12.63% | 22.52% | 10.22% | 22.82% | 14.87% | 15.21% | 0.2753% | 10.9% | 1.271% | 12.3% | 1.138% | 1.043% | 10.2% | 17.6% | 16.73% | 1.3% | 0.9839%
159 | LFB 10ppb | R | 6/3/2023 7:37:58 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LFB 10ppb.161 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 3.325% | 2.422% | 5.135% | 1.87% | 4.712% | 1.915% | 3.104% | 1.558% | 0.9391% | 1.066% | 1.81% | 17.08% | 2.209% | 0.7619% | 1.082% | 0.8892% | 0.7268% | 0.4717% | 2.463% | 4.256% | 20.53% | 18.94% | 1.76% | 23.01% | 14.35% | 16.81% | 3.82% | 18.48% | 2.256% | 18.8% | 0.3911% | 15.61% | 0.6502% | 0.5084% | 1.486% | 1.66% | 17.85% | 1.076% | 0.3317% | 1.202% | 1.687%
160 | LFB 10ppb | R | 6/3/2023 7:44:29 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LFB 10ppb.162 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 9.066% | 10.97% | 10.06% | 2.493% | 14.43% | 9.54% | 8.343% | 9.884% | 10.9% | 150.3% | 514.7% | 6.759% | 9.486% | 8.815% | 8.168% | 8.993% | 155% | 220.7% | 5.447% | 10.14% | 5.021% | 5.78% | 3.541% | 11.58% | 6.319% | 5.554% | 50.08% | 4.488% | 151.5% | 6.978% | 7.329% | 4.752% | 6.464% | 11.35% | 4.381% | 1.244% | 3.787% | 151.3% | 152.6% | 132.2% | 176.8%
161 | CCV 10 ppb MAJORS | R | 6/3/2023 7:50:31 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.163 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.728% | 2.435% | 2.871% | 0.5512% | 1.203% | 0.8872% | 2.329% | 2.323% | 1.701% | 3.105% | 2.051% | 1.468% | 3.434% | 1.156% | 3.434% | 0.6154% | 3.211% | 1.705% | 14.42% | 2.932% | 4.248% | 1.094% | 18.05% | 4.241% | 1.327% | 0.6735% | 14.53% | 2.254% | 2.59% | 2.01% | 11.22% | 0.9423% | 1.508% | 1.205% | 1.575% | 0.5462% | 5.632% | 2.311% | 0.1801% | 15.23% | 1.535%
162 | CCV 10 ppb MAJORS | R | 6/3/2023 7:57:03 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.164 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 5.536% | 7.793% | 9.761% | 2.522% | 7.306% | 5.908% | 5.716% | 5.82% | 5.847% | 0.9588% | 0.8408% | 9.362% | 4.207% | 3.824% | 4.363% | 3.21% | 0.2883% | 0.4082% | 6.592% | 1.38% | 8.013% | 10.3% | 7.777% | 14.4% | 6.294% | 8.512% | 10.22% | 8.431% | 1.174% | 9.647% | 5.403% | 7.973% | 3.675% | 0.9733% | 16.26% | 0.7824% | 6.295% | 1.159% | 1.336% | 7.418% | 18.6%
163 | MT 32 C LD | R | 6/3/2023 8:03:06 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 32 C LD.165 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1.523% | 3.543% | 32.91% | 2.045% | 9.724% | 1.096% | 185.3% | 22.44% | 7.987% | 6.473% | 4.397% | 143.4% | 50.51% | 73.39% | 34.02% | 11.81% | 22.37% | 0.9479% | 154.2% | 460.9% | 36.58% | 43.83% | 96.07% | 1.913e-06% | 11.56% | 51.88% | 165.4% | 3240% | 43.5% | 9.714% | 8.649% | 7.805% | 3.406% | 0.9663% | 1.024% | 1.317% | 6.748% | 51.46% | 40.12% | 4.875% | 4.574%
164 | MT 33 C | R | 6/3/2023 8:09:08 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 33 C.166 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.506% | 0% | 41.93% | 1.196% | 9.478% | 9.377% | 97.24% | 12.97% | 37.96% | 4.101% | 2.245% | 555% | 41.56% | 57.26% | 25.25% | 18.36% | 11.9% | 4.171% | 58.05% | 0% | 4.123% | 5.049% | 89.19% | 54.54% | 10.51% | 34.63% | 162.2% | 0% | 11% | 1.096% | 12.8% | 0.8302% | 4.486% | 0.7645% | 2.292% | 1.733% | 2.175% | 134.9% | 61.07% | 24.63% | 1.649%
165 | MT 33 C LSP | R | 6/3/2023 8:15:09 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 33 C LSP.167 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 4.166% | 6.76% | 5.324% | 1.02% | 5.363% | 2.256% | 5.508% | 3.835% | 4.29% | 2.925% | 5.458% | 5.558% | 4.307% | 4.39% | 4.271% | 4.137% | 3.4% | 3.249% | 13.17% | 2.788% | 7.605% | 4.36% | 12.68% | 7.734% | 3.342% | 4.714% | 13.15% | 5.476% | 9.879% | 4.662% | 11.4% | 3.145% | 3.735% | 2.278% | 1.312% | 0.9037% | 6.544% | 3.536% | 3.125% | 14.01% | 1.622%
166 | MT 33 C LSP | R | 6/3/2023 8:21:40 AM | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 33 C LSP.168 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 5.891% | 4.427% | 1.863% | 2.763% | 2.03% | 4.6% | 5.495% | 4.722% | 4.965% | 1.809% | 1.369% | 1.817% | 4.415% | 5.213% | 5.669% | 4.898% | 1.937% | 1.245% | 3.41% | 5.81% | 4.53% | 0.5199% | 6.883% | 2.631% | 1.177% | 1.509% | 3.199% | 1.816% | 0.8775% | 0.7923% | 2.904% | 0.81% | 3.376% | 1.277% | 1.147% | 0.9664% | 4.506% | 0.5425% | 0.5% | 1.98% | 0.2707%
167 | MT 34 C | R | 6/3/2023 8:27:42 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 34 C.169 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 0.4974% | 3.436% | 36.72% | 1.269% | 1.613% | 3.673% | 690.8% | 43.24% | 28.21% | 15.16% | 4.439% | 1069% | 31.82% | 55.73% | 11.7% | 27.46% | 10.41% | 1.049% | 27.35% | 128.3% | 49.98% | 63.78% | 18.78% | 55.27% | 0.7004% | 33.82% | 236.9% | 0% | 8.549% | 14.95% | 8.495% | 3.462% | 0.4912% | 0.3255% | 1.149% | 1.153% | 2.364% | 19.36% | 101% | 2.525% | 8.578%
168 | MT 35 C | R | 6/3/2023 8:33:45 AM | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 35 C.170 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2.186% | 3.395% | 69.67% | 1.926% | 12.55% | 6.948% | 255.7% | 0.7164% | 12.42% | 14.3% | 1.374% | 61.26% | 47.56% | 95.37% | 46.78% | 12.85% | 16.41% | 5.384% | 40.41% | 0% | 3.963% | 4.932% | 22.78% | 1.913e-06% | 13.96% | 0% | 203.5% | 2345% | 92.28% | 2.427% | 2.723% | 6.665% | 3.68% | 3.089% | 3.027% | 2.378% | 4.968% | 51.26% | 12.91% | 14.61% | 3.396%

## Unfactored Concentrations
| Sample Id | R | Acquisition Time | QC Status | Dataset File | Method File | Na KED 23/23 Helium KED_Low (ppb) | Mg KED 24/24 Helium KED_Low (ppb) | Al KED 27/27 Helium KED_Low (ppb) | P Oshift 31/47 Oxygen DRC (ppb) | S (34) -1 34/34 Helium KED_Low (ppb) | K KED 39/39 Helium KED_Low (ppb) | Ca KED 43/43 Helium KED_Low (ppb) | V KED 51/51 Helium KED_Low (ppb) | Cr (52) 52/52 Helium KED_Low (ppb) | Cr (53) NH3 53/53 Ammonia DRC (ppb) | Fe (54) NH3 54/54 Ammonia DRC (ppb) | Mn Oshift 55/71 Oxygen DRC (ppb) | Fe (56) KED 56/56 Helium KED_Low (ppb) | Co KED 59/59 Helium KED_Low (ppb) | Ni KED 60/60 Helium KED_Low (ppb) | Cu (63) KED 63/63 Helium KED_Low (ppb) | Cu (65) NH3 65/65 Ammonia DRC (ppb) | Zn NH3 66/66 Ammonia DRC (ppb) | As 75/75 (ppb) | As KED 75/75 Helium KED_Low (ppb) | As RH Oshift 75/91 Oxygen DRC (ppb) | As Oshift 75/91 Oxygen DRC (ppb) | Se (77) 77/77 (ppb) | Se (77) Oshift 77/93 Oxygen DRC (ppb) | Se (78) Oshift 78/94 Oxygen DRC (ppb) | Se (80) Oshift 80/96 Oxygen DRC (ppb) | Se (82) 82/82 (ppb) | Se (82) Oshift 82/98 Oxygen DRC (ppb) | Mo (95) NH3 95/95 Ammonia DRC (ppb) | Mo (95) OShift 95/127 Oxygen DRC (ppb) | Rh (IS1) 103/103 (IS) | Rh (IS2) 103/103 (IS) Oxygen DRC | Rh (IS3) 103/103 (IS) Helium KED_Low | Rh (IS4) 103/103 (IS) Ammonia DRC | Rh (IS5) 103/103 (IS) Ammonia DRC | Rh (IS6) 103/103 (IS) Oxygen DRC | Rh OShift (IS1) 103/119 (IS) Oxygen DRC | Cd (111) NH3 111/111 Ammonia DRC (ppb) | Cd (114) NH3 114/114 Ammonia DRC (ppb) | Pb 208/208 (ppb) | Pb NH3 208/208 Ammonia DRC (ppb)
1 | CAL BLANK | R | 06/02/2023 15:17:36 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CAL BLANK.004 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth
2 | LOW 0.01 | R | 06/02/2023 15:23:43 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LOW 0.01.005 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | -422.3993 | -327.5683 | -0.93583 | 0.38721 | 154.9971 | -303.9575 | -295.0181 | 0.017673 | 0.0098479 | 0.0097913 | -0.69414 | 0.0087781 | -1.0252 | 0.010013 | 0.0047174 | 0.015504 | 0.0047898 | -0.18306 | 0.0060177 | 0.0083017 | 0.0039041 | 0.0057828 | -0.041074 | 0.024948 | -0.016212 | -0.00031501 | -0.026953 | 0.022713 | 0.010451 | 0.010344 | [7 empty] | 0.0099663 | 0.0099165 | 0.0096244 | 0.009868
3 | LOW 0.05 | R | 06/02/2023 15:29:46 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LOW 0.05.006 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | -382.0795 | -279.2759 | -0.59744 | 0.97096 | 207.3922 | -255.2057 | -249.3962 | 0.050063 | 0.055303 | 0.055384 | -0.63168 | 0.048923 | -0.90336 | 0.050103 | 0.05348 | 0.058267 | 0.048548 | -0.2017 | 0.049533 | 0.052476 | 0.05064 | 0.051525 | -0.036241 | 0.04765 | 0.024099 | 0.051803 | 0.029596 | 0.06202 | 0.040989 | 0.042998 | [7 empty] | 0.049938 | 0.05099 | 0.057051 | 0.053985
4 | LOW 0.1 | R | 06/02/2023 15:35:48 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LOW 0.1.007 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | -316.1199 | -225.8575 | -0.87139 | 0.75218 | 97.956 | -219.0009 | -225.018 | 0.097748 | 0.091177 | 0.096914 | -0.44375 | 0.094897 | -0.74407 | 0.095425 | 0.082266 | 0.096822 | 0.1015 | -0.081046 | 0.10107 | 0.089531 | 0.098214 | 0.09125 | 0.10263 | 0.096584 | 0.097405 | 0.097935 | 0.10217 | 0.096353 | 0.089224 | 0.091731 | [7 empty] | 0.10342 | 0.10428 | 0.11144 | 0.09848
5 | LOW 0.5 | R | 06/02/2023 15:41:51 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LOW 0.5.008 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 136.2129 | 193.5249 | 0.1294 | 1.8636 | 597.907 | 257.6033 | 192.3505 | 0.48108 | 0.47365 | 0.48243 | 0.50147 | 0.47199 | 0.50256 | 0.46838 | 0.47691 | 0.53363 | 0.52159 | 0.44981 | 0.48482 | 0.55327 | 0.49777 | 0.48646 | 0.46062 | 0.47137 | 0.43518 | 0.47763 | 0.46859 | 0.42841 | 0.4904 | 0.47815 | [7 empty] | 0.5109 | 0.51221 | 0.53569 | 0.5287
6 | STD 1 | R | 06/02/2023 15:47:53 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\STD 1.009 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 990.5164 | 999.0291 | 1.0046 | 1.9976 | 1300.8923 | 993.6008 | 1002.1128 | 0.9878 | 0.98469 | 1.0152 | 1.6265 | 1.0029 | 1.6127 | 0.97921 | 1.0292 | 1.0428 | 1.0734 | 1.1343 | 1.0225 | 0.94708 | 1.0103 | 1.0098 | 0.95732 | 1.0784 | 0.98316 | 0.98738 | 1.0685 | 1.0665 | 1.022 | 1.0373 | [7 empty] | 1.0056 | 0.99349 | 1.049 | 1.2687
7 | STD 5 | R | 06/02/2023 15:53:56 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\STD 5.010 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 5157.6412 | 5044.1152 | 4.8911 | 6.3929 | 4949.6683 | 5129.6499 | 4932.2957 | 4.8816 | 4.8161 | 4.9794 | 4.8582 | 4.9955 | 4.7704 | 4.8589 | 4.9788 | 4.9116 | 5.0183 | 5.0507 | 4.9976 | 4.8557 | 5.0325 | 4.9796 | 5.2796 | 4.6996 | 5.08 | 5.0466 | 5.1723 | 4.9666 | 5.1069 | 5.1578 | [7 empty] | 4.9739 | 4.9009 | 4.9029 | 4.8232
8 | STD 10 | R | 06/02/2023 15:59:58 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\STD 10.011 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 10252.7747 | 9920.0692 | 9.5999 | 11.0451 | 10079.4311 | 10070.508 | 9927.3956 | 9.6599 | 9.7247 | 9.7897 | 9.9137 | 10.1422 | 9.6373 | 9.6803 | 10.1272 | 9.5721 | 9.7527 | 10.0025 | 10.153 | 9.8883 | 10.1937 | 9.9592 | 10.2911 | 9.7273 | 9.9298 | 9.6592 | 10.1021 | 9.7844 | 10.321 | 10.436 | [7 empty] | 9.7515 | 9.8295 | 9.368 | 9.7051
9 | STD 25 | R | 06/02/2023 16:06:00 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\STD 25.012 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 26477.3189 | 25003.5244 | 28.1438 | 26.4679 | 26352.1901 | 26299.7528 | 26184.4979 | 26.9715 | 26.5589 | 24.2227 | 25.2453 | 24.9675 | 26.7371 | 26.5526 | 27.2977 | 25.7008 | 23.8781 | 25.1448 | 25.5185 | 27.26 | 25.7851 | 25.7309 | 25.7657 | 25.1569 | 26.178 | 25.6105 | 25.4178 | 25.563 | 26.9742 | 26.9903 | [7 empty] | 24.851 | 24.7974 | 21.1033 | 21.8212
10 | STD 50 | R | 06/02/2023 16:12:02 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\STD 50.013 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 46753.0659 | 43797.3242 | 46.9741 | 50.028 | 47875.4863 | 46071.4173 | 48565.4265 | 49.7824 | 48.5088 | 47.8411 | 50.0775 | 50.1112 | 48.5368 | 47.9347 | 48.6293 | 46.2521 | 46.3961 | 48.6774 | 47.544 | 48.0934 | 48.1739 | 51.0684 | 48.2345 | 51.687 | 52.4448 | 51.5451 | 46.5481 | 51.9872 | 54.0085 | 55.4237 | [7 empty] | 49.1032 | 48.8427 | 29.915 | 48.2553
11 | STD 100 | R | 06/02/2023 16:18:04 |  | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\STD 100.014 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 95852.3792 | 86684.2364 | 99.1938 | 99.4349 | 99052.5906 | 91223.4826 | 104446.7551 | 106.8424 | 102.4972 | 98.0373 | 102.2693 | 102.3935 | 103.685 | 103.8368 | 101.9906 | 93.9563 | 90.6201 | 94.3714 | 110.1467 | 99.0107 | 93.374 | 100.9361 | 115.2894 | 106.0795 | 106.1982 | 104.5918 | 111.5283 | 105.2095 | 112.948 | 115.2779 | [7 empty] | 96.9352 | 96.6369 | 49.4726 | 96.4856
12 | Calibration Curves | R | 06/02/2023 16:18:04
13 | LBLANK | R | 06/02/2023 16:24:07 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LBLANK.015 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | -145.773 | -48.7752 | -0.92607 | -0.32009 | 467.0084 | -167.5555 | -0.01244 | 0.0096825 | 0.0050838 | 0.00097923 | -0.56933 | -0.0036322 | -0.53599 | 0.00289 | -0.0094525 | 0.00025038 | -0.010745 | -0.23801 | -0.0071944 | 0.0039143 | -0.0036233 | -0.0038431 | 0.10924 | 0.014903 | -0.029849 | -0.0036589 | 0.020548 | -0.0016216 | -0.0039439 | -0.0015604 | [7 empty] | -0.00031202 | 0.00075458 | -0.00075595 | -0.00039769
14 | ICV PE | R | 06/02/2023 16:30:09 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\ICV PE.016 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 9823.9524 | 9944.6776 | 11.1131 | 20.0669 | 10681.7877 | 9422.7641 | 10726.211 | 10.0739 | 10.0157 | 9.7619 | 10.0029 | 10.2034 | 10.5905 | 10.3283 | 10.5191 | 9.9341 | 9.7802 | 9.9155 | 9.885 | 8.7886 | 9.0956 | 10.0139 | 7.2921 | 9.9484 | 10.457 | 10.2681 | 10.0259 | 10.2691 | 10.1446 | 10.3165 | [7 empty] | 9.8959 | 9.8193 | 8.3318 | 9.1537
15 | CCV 1ppb TRACE | R | 06/02/2023 16:36:12 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.017 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 917.1242 | 1032.8229 | 0.67866 | 0.82457 | 1812.5312 | 818.0184 | 1092.2421 | 1.0775 | 1.0519 | 1.0305 | 0.5955 | 1.0632 | 0.79028 | 1.1017 | 1.119 | 1.0813 | 1.0815 | 0.87659 | 0.97221 | 1.0031 | 0.86087 | 0.98823 | 0.98123 | 1.0375 | 0.95556 | 0.98496 | 1.0295 | 0.97738 | 1.0469 | 1.0568 | [7 empty] | 1.0145 | 1.0178 | 0.92704 | 0.92218
16 | CCV 1ppb TRACE | R | 06/02/2023 16:42:44 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.018 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 988.4925 | 1014.6011 | 0.53485 | 0.52706 | 2076.2749 | 907.2816 | 1058.5224 | 1.0715 | 1.0503 | 1.021 | 0.55656 | 1.1252 | 0.76502 | 1.107 | 1.1572 | 1.0602 | 1.0954 | 1.4936 | 0.98863 | 0.93905 | 0.88344 | 1.0413 | 1.0361 | 1.0852 | 1.0643 | 1.0471 | 0.95568 | 1.1535 | 1.0515 | 1.0617 | [7 empty] | 1.0101 | 0.99231 | 0.93319 | 0.91035
17 | LFB 10ppb | R | 06/02/2023 16:48:46 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LFB 10ppb.019 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 11220.7527 | 10584.4469 | 10.9725 | 11.6099 | 11646.8755 | 10407.1031 | 11307.4482 | 10.388 | 10.1814 | 10.2906 | 11.2461 | 11.0973 | 11.538 | 10.9081 | 10.9852 | 9.9463 | 10.1897 | 13.9928 | 10.5608 | 9.573 | 9.0555 | 10.448 | 10.8077 | 11.0004 | 10.7592 | 10.5684 | 10.5406 | 10.893 | 10.6068 | 10.5799 | [7 empty] | 9.825 | 9.718 | 8.1616 | 8.3769
18 | LFB 10ppb | R | 06/02/2023 16:55:16 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LFB 10ppb.020 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 11164.1162 | 10847.4002 | 12.22 | 12.0531 | 11713.2675 | 10386.4177 | 11380.2089 | 10.5083 | 10.3136 | 10.1861 | 11.2148 | 11.2524 | 11.5015 | 10.9732 | 11.1528 | 10.0312 | 10.2523 | 14.0507 | 10.3909 | 10.2885 | 9.1395 | 10.3933 | 11.1021 | 11.0533 | 10.7623 | 10.7371 | 10.3819 | 10.6638 | 10.5372 | 10.6573 | [7 empty] | 9.7462 | 9.649 | 8.0232 | 8.3515
19 | LLOQ 0.01 | R | 06/02/2023 17:01:19 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LLOQ 0.01.021 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | -138.0924 | -39.7596 | -0.58716 | 0.43129 | 941.3472 | -158.7711 | 14.3678 | 0.018883 | 0.016598 | 0.010949 | -0.46865 | 0.0075582 | -0.24595 | 0.012491 | 0.003325 | 0.014066 | 0.006141 | -0.13024 | -0.021182 | 0.0033634 | 0.0041725 | 0.0058062 | 0.0070255 | 0.041888 | -0.025776 | 0.0028406 | -0.017262 | 0.0021742 | 0.012904 | 0.014265 | [7 empty] | 0.0087397 | 0.010384 | 0.0079997 | 0.0087649
20 | CCV 10 ppb MAJORS | R | 06/02/2023 17:07:22 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.022 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 12218.3718 | 11786.596 | 11.7543 | 11.4738 | 12395.218 | 11100.4799 | 11922.9999 | 10.785 | 10.7097 | 10.5238 | 10.7419 | 11.6461 | 11.6607 | 11.419 | 11.4898 | 10.3056 | 10.3968 | 10.5241 | 10.2955 | 9.9808 | 9.2524 | 10.5532 | 11.0721 | 10.5724 | 10.8825 | 10.68 | 10.3157 | 10.6049 | 10.592 | 10.774 | [7 empty] | 9.7957 | 9.7371 | 8.1524 | 8.203
21 | CCV 10 ppb MAJORS | R | 06/02/2023 17:13:54 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.023 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 12155.8073 | 11534.2219 | 11.5892 | 10.6853 | 12260.8312 | 11190.5038 | 11962.7689 | 10.818 | 10.652 | 10.5538 | 10.8271 | 12.4547 | 11.3666 | 11.3338 | 11.4317 | 10.2051 | 10.4441 | 10.5546 | 10.4084 | 9.8099 | 9.3471 | 11.0214 | 11.0873 | 11.2832 | 11.4539 | 11.2855 | 10.6633 | 11.103 | 10.6018 | 11.2965 | [7 empty] | 9.762 | 9.6951 | 8.2167 | 8.5509
22 | LLOQ 0.1 | R | 06/02/2023 17:19:57 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LLOQ 0.1.024 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | -21.9855 | 50.9662 | -0.41297 | 0.28369 | 1215.184 | -46.7358 | 125.3788 | 0.10957 | 0.11069 | 0.098192 | -0.2392 | 0.11878 | -0.0053462 | 0.11058 | 0.11615 | 0.10939 | 0.10227 | -0.010861 | 0.078753 | 0.095414 | 0.081962 | 0.09817 | 0.02152 | 0.13285 | 0.071288 | 0.10055 | 0.084762 | 0.078256 | 0.090107 | 0.098494 | [7 empty] | 0.090996 | 0.0939 | 0.084129 | 0.087
23 | LBLANK | R | 06/02/2023 17:25:59 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LBLANK.025 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | -147.691 | -49.4851 | -0.86327 | -0.29428 | 1195.2686 | -163.5813 | -4.861 | 0.0076166 | 0.0026875 | 0.00079703 | -0.55016 | -0.0032707 | -0.30372 | 0.00028526 | -0.0115 | -0.0017169 | -0.0098416 | -0.21784 | -0.0036543 | 0.003273 | -0.0075308 | -0.0080116 | 0.030613 | 0.02405 | -0.034409 | -0.0071849 | -0.055017 | 0.0098666 | -0.011061 | -0.012106 | [7 empty] | -0.0005172 | 8.349e-05 | -0.0017076 | -0.00056047
24 | LFB 6.25ppb & 2500ppb | R | 06/02/2023 17:32:01 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LFB 6.25ppb & 2500ppb.026 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 32284.453 | 30307.3697 | 24.6718 | 10.165 | 30868.3027 | 30109.4028 | 31301.7661 | 7.1933 | 7.052 | 7.0597 | 7.7449 | 7.5174 | 8.4385 | 7.5935 | 7.5147 | 6.5514 | 6.7455 | 7.1095 | 6.8655 | 6.395 | 5.6646 | 6.7761 | 7.4238 | 7.1031 | 7.0332 | 6.863 | 7.1388 | 6.994 | 7.0511 | 6.7685 | [7 empty] | 6.4283 | 6.3118 | 5.4246 | 5.385
25 | MT 1 | R | 06/02/2023 17:38:02 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 1.027 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 3938.0267 | 228.4614 | 422.4351 | 2517.3629 | 3274.1776 | 3290.2344 | 673.5447 | 0.45867 | 0.55734 | 0.53116 | 91.6986 | 0.8713 | 98.3796 | 0.072481 | 0.37763 | 4.0977 | 4.1035 | 12.9328 | 1.5666 | 1.3648 | 1.3534 | 1.6312 | 0.7488 | 0.92616 | 0.84418 | 0.86674 | 0.97459 | 0.93372 | 0.067212 | 0.062977 | [7 empty] | 0.068459 | 0.069387 | 0.15119 | 0.14031
26 | MT 1 LD | R | 06/02/2023 17:44:04 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 1 LD.028 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 4213.4507 | 238.2359 | 450.616 | 2626.5344 | 3456.5058 | 3568.3228 | 717.1952 | 0.45768 | 0.56047 | 0.57715 | 95.2097 | 0.90562 | 96.1427 | 0.078058 | 0.37349 | 4.0827 | 4.3734 | 14.2192 | 1.7187 | 1.3844 | 1.4645 | 1.709 | 0.93357 | 1.0354 | 0.89558 | 0.96807 | 1.1174 | 1.0082 | 0.072039 | 0.065195 | [7 empty] | 0.078664 | 0.074799 | 0.16115 | 0.14655
27 | LBLANK | R | 06/02/2023 18:05:56 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LBLANK.029 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | -146.2195 | -48.7408 | -0.87234 | -0.15189 | 356.4664 | -165.3765 | -4.861 | 0.0069751 | 0.006241 | 0.0010477 | -0.53254 | -0.0033132 | -0.42668 | 0.00031495 | -0.0077825 | -0.00017509 | -0.009862 | -0.23012 | -0.021261 | -0.0031069 | -0.006368 | -0.0065477 | 0.050564 | 0.019441 | -0.031308 | -0.0054913 | -0.04187 | -0.0054505 | -0.011445 | -0.010879 | [7 empty] | -0.0013303 | 0.00038503 | -0.0018674 | -0.00075614
28 | STD 1 | R | 06/02/2023 18:11:58 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\STD 1.030 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1256.5884 | 1252.5356 | 1.3331 | 1.9354 | 2533.1423 | 1136.5104 | 1275.4261 | 1.1409 | 1.1777 | 1.1119 | 1.9447 | 1.2679 | 2.2161 | 1.2071 | 1.1837 | 1.1593 | 1.1592 | 1.2236 | 1.0991 | 0.92583 | 0.90431 | 1.0711 | 1.1759 | 1.279 | 1.1325 | 1.1418 | 1.0397 | 1.2834 | 1.0595 | 1.0761 | [7 empty] | 1.0072 | 0.99053 | 0.9298 | 0.89562
29 | STD 0.5 | R | 06/02/2023 18:18:06 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\STD 0.5.031 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 570.7201 | 558.1965 | 0.98988 | 1.9026 | 2058.4941 | 453.9212 | 578.0099 | 0.56019 | 0.57207 | 0.5629 | 0.68882 | 0.64012 | 1.0056 | 0.57829 | 0.59989 | 0.56297 | 0.60095 | 0.54788 | 0.50923 | 0.5244 | 0.49099 | 0.57946 | 0.39351 | 0.56899 | 0.58736 | 0.55312 | 0.42028 | 0.59043 | 0.54074 | 0.54631 | [7 empty] | 0.50244 | 0.50494 | 0.43321 | 0.44793
30 | ICV PE | R | 06/02/2023 18:24:08 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\ICV PE.032 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 13897.8579 | 13749.6052 | 13.9589 | 28.1108 | 13951.4949 | 12791.5569 | 13327.2364 | 11.5708 | 11.3416 | 11.1409 | 11.7665 | 13.0434 | 12.4062 | 11.8958 | 11.9854 | 10.4681 | 10.6167 | 10.7588 | 10.7083 | 8.7767 | 9.2977 | 11.1401 | 8.1861 | 11.2813 | 11.719 | 11.5341 | 11.1233 | 11.3428 | 10.2514 | 10.8331 | [7 empty] | 9.7436 | 9.6854 | 8.1639 | 8.5864
31 | ICV PE | R | 06/02/2023 18:30:39 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\ICV PE.033 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 13661.1271 | 13405.2844 | 13.5666 | 25.9672 | 13345.0393 | 12236.5207 | 13307.5124 | 11.4975 | 11.1834 | 11.1563 | 11.9394 | 12.86 | 12.1774 | 11.8808 | 11.9216 | 10.4523 | 10.6991 | 10.8552 | 10.5169 | 8.9232 | 9.0987 | 10.8854 | 8.0598 | 11.6457 | 11.4559 | 11.268 | 10.8422 | 11.3762 | 10.3886 | 10.7405 | [7 empty] | 9.9262 | 9.8307 | 8.0438 | 8.4798
32 | CCV 1ppb TRACE | R | 06/02/2023 18:36:41 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.034 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1197.1876 | 1248.5498 | 0.39246 | 1.2004 | 2197.7058 | 1089.3757 | 1249.0218 | 1.1299 | 1.1132 | 1.1117 | 0.69584 | 1.3335 | 0.95012 | 1.1826 | 1.215 | 1.0725 | 1.1209 | 0.89552 | 1.0664 | 1.0087 | 0.93385 | 1.1345 | 1.2297 | 1.2676 | 1.1228 | 1.1986 | 1.0296 | 1.1019 | 1.0679 | 1.1433 | [7 empty] | 0.99734 | 0.99206 | 0.90466 | 0.89042
33 | CCV 1ppb TRACE | R | 06/02/2023 18:43:13 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.035 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1280.3979 | 1330.1937 | 1.023 | 0.83521 | 2768.7563 | 1081.081 | 1381.8236 | 1.1716 | 1.187 | 1.1511 | 0.69211 | 1.2998 | 1.1162 | 1.2419 | 1.261 | 1.1349 | 1.1422 | 0.90637 | 1.0183 | 0.88583 | 0.91925 | 1.0956 | 1.1404 | 1.0998 | 1.0763 | 1.1457 | 1.0603 | 1.1365 | 1.0646 | 1.0782 | [7 empty] | 1.0174 | 0.99391 | 0.91094 | 0.89151
34 | CCV 10 ppb MAJORS | R | 06/02/2023 18:49:16 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.036 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 14116.8842 | 13633.8579 | 13.5979 | 13.1568 | 14424.6605 | 12415.4571 | 13460.7718 | 11.6073 | 11.3983 | 11.4099 | 11.7938 | 12.9671 | 12.2666 | 12.157 | 12.3103 | 10.5781 | 10.9908 | 11.1114 | 10.7506 | 9.7871 | 9.1837 | 10.9285 | 11.6208 | 11.2458 | 11.6391 | 11.1064 | 10.9319 | 10.9671 | 10.873 | 10.9608 | [7 empty] | 9.9046 | 9.7928 | 8.0981 | 8.4022
35 | CCV 10 ppb MAJORS | R | 06/02/2023 18:55:47 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.037 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 13859.7015 | 13002.0386 | 13.0026 | 13.0325 | 13944.7964 | 12276.3185 | 12850.3773 | 11.6348 | 11.3822 | 11.3969 | 11.6512 | 12.8255 | 12.1949 | 12.0396 | 12.1705 | 10.6666 | 10.8653 | 11.14 | 10.8657 | 10.1504 | 9.2303 | 11.0926 | 11.7398 | 11.6659 | 11.5002 | 11.1047 | 10.9412 | 11.0898 | 10.877 | 11.1551 | [7 empty] | 9.8475 | 9.6894 | 8.1076 | 8.7508
36 | MT 1 | R | 06/02/2023 19:01:50 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 1.038 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 4390.7358 | 251.0427 | 480.1772 | 2566.2268 | 3789.2436 | 3596.5431 | 806.9309 | 0.48063 | 0.58234 | 0.55929 | 97.0559 | 1.0589 | 103.677 | 0.082321 | 0.37773 | 4.1209 | 4.1489 | 13.3781 | 1.5977 | 1.3501 | 1.6209 | 1.9139 | 0.81317 | 1.0787 | 1.028 | 1.0591 | 0.97694 | 0.97873 | 0.070888 | 0.078951 | [7 empty] | 0.072153 | 0.072524 | 0.14317 | 0.1374
37 | MT 1 LD | R | 06/02/2023 19:07:53 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 1 LD.039 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 4737.7152 | 274.1657 | 522.9345 | 2687.2307 | 3935.6455 | 3968.5109 | 879.3574 | 0.48204 | 0.5954 | 0.56471 | 96.0594 | 0.97603 | 106.4012 | 0.084645 | 0.44127 | 4.3902 | 4.3454 | 14.174 | 1.6925 | 1.4061 | 1.4484 | 1.7711 | 0.79004 | 1.0325 | 0.98088 | 1.0047 | 1.0746 | 1.063 | 0.066619 | 0.065967 | [7 empty] | 0.079197 | 0.074548 | 0.14617 | 0.14074
38 | MT 2 | R | 06/02/2023 19:13:55 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 2.040 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 5158.7022 | 252.262 | 560.1251 | 2891.4774 | 4116.3693 | 4139.4865 | 851.201 | 0.50451 | 0.62417 | 0.61414 | 99.9173 | 0.95328 | 104.0469 | 0.084898 | 0.40195 | 4.5614 | 4.7273 | 14.9376 | 1.7748 | 1.6273 | 1.6088 | 1.9035 | 0.87084 | 0.95784 | 1.0829 | 1.0914 | 1.1047 | 1.0527 | 0.098176 | 0.098976 | [7 empty] | 0.076314 | 0.076029 | 0.14735 | 0.14372
39 | MT 2 LSP | R | 06/02/2023 19:19:58 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 2 LSP.041 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 43041.2575 | 35051.0008 | 610.1414 | 3362.0097 | 37498.7812 | 38471.4225 | 35018.4305 | 8.3381 | 8.2533 | 7.9847 | 115.5293 | 9.2264 | 121.2611 | 7.9125 | 8.2952 | 11.4312 | 11.3947 | 22.9213 | 8.7648 | 8.4286 | 7.7299 | 8.9702 | 8.4342 | 8.3947 | 8.4973 | 8.2237 | 8.2555 | 8.6432 | 7.037 | 7.0037 | [7 empty] | 6.2989 | 6.2167 | 5.3218 | 5.4114
40 | MT 2 LSP | R | 06/02/2023 19:26:09 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 2 LSP.042 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 40167.212 | 32459.88 | 574.1295 | 3227.6575 | 35669.1004 | 35523.6786 | 32959.6604 | 7.93 | 7.6956 | 7.9018 | 113.666 | 9.2359 | 114.8888 | 7.634 | 8.0017 | 11.0315 | 11.3609 | 22.6737 | 8.8028 | 8.1391 | 7.6772 | 9.1221 | 8.3266 | 8.6324 | 8.444 | 8.397 | 8.1963 | 8.4878 | 7.03 | 7.2615 | [7 empty] | 6.2396 | 6.1679 | 5.3555 | 5.3794
41 | MT 3 | R | 06/02/2023 19:32:11 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 3 .043 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 4004.6025 | 212.2501 | 435.5836 | 2351.5468 | 3706.9383 | 3294.5661 | 619.4182 | 0.42402 | 0.67014 | 0.66024 | 83.7203 | 0.79953 | 84.5065 | 0.073066 | 0.42675 | 3.759 | 3.9661 | 12.4638 | 1.581 | 1.4894 | 1.3547 | 1.6234 | 0.8665 | 1.0681 | 0.90249 | 0.90039 | 1.0432 | 0.93624 | 0.071804 | 0.065333 | [7 empty] | 0.065051 | 0.068259 | 0.14473 | 0.13871
42 | MT 4 | R | 06/02/2023 19:38:14 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 4.044 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 5342.3849 | 259.1695 | 587.2254 | 3022.489 | 4880.4377 | 4463.1677 | 866.4666 | 0.55984 | 0.66918 | 0.63449 | 134.5167 | 1.1584 | 145.0237 | 0.097116 | 0.43049 | 4.8579 | 4.8848 | 15.7073 | 1.9539 | 1.7358 | 1.6878 | 2.0419 | 0.92327 | 1.1444 | 1.0885 | 1.1468 | 1.1902 | 1.0775 | 0.083875 | 0.080313 | [7 empty] | 0.086431 | 0.085481 | 0.15863 | 0.15228
43 | LBLANK | R | 06/02/2023 19:44:17 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LBLANK.045 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | -148.4385 | -48.7435 | -0.94902 | 0.47215 | 1503.6932 | -154.0031 | -4.861 | 0.0074153 | 0.0039509 | 0.000477 | -0.52315 | -0.0032491 | -0.2402 | 0.0004931 | -0.0061775 | -0.0020108 | -0.0087663 | -0.21823 | -0.047512 | 0.00060526 | -0.0056794 | -0.005765 | 0.055805 | 0.02008 | -0.030431 | -0.0034118 | -0.0091382 | 0.0032203 | -0.010332 | -0.012062 | [7 empty] | -0.0012259 | 0.00081252 | -0.0020589 | -0.00072911
44 | CCV 1ppb TRACE | R | 06/02/2023 19:50:20 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.046 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1350.6062 | 1376.7676 | 0.98963 | 0.99616 | 2948.8417 | 1202.9902 | 1431.5703 | 1.2657 | 1.1917 | 1.219 | 0.80702 | 1.35 | 1.1813 | 1.2799 | 1.326 | 1.151 | 1.1641 | 0.96691 | 1.0057 | 1.0306 | 0.90977 | 1.1105 | 1.4257 | 1.1527 | 1.1236 | 1.1697 | 0.98577 | 1.1579 | 1.0792 | 1.1104 | [7 empty] | 1.0115 | 1.0049 | 0.90438 | 0.87798
45 | CCV 1ppb TRACE | R | 06/02/2023 19:56:52 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.047 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1341.2632 | 1409.3016 | 1.1294 | 0.80754 | 3029.6237 | 1128.3677 | 1350.0194 | 1.2141 | 1.1959 | 1.1948 | 0.75335 | 1.3215 | 1.1268 | 1.2477 | 1.2858 | 1.162 | 1.1752 | 0.9289 | 1.0808 | 1.021 | 0.90009 | 1.0863 | 1.2791 | 1.121 | 1.1189 | 1.142 | 1.0566 | 1.2841 | 1.0706 | 1.0766 | [7 empty] | 0.99114 | 0.99077 | 0.87733 | 0.86759
46 | CCV 10 ppb MAJORS | R | 06/02/2023 20:02:55 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.048 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 15136.6738 | 14616.0052 | 14.6761 | 14.3822 | 15742.228 | 13354.1633 | 13892.798 | 12.3561 | 11.9936 | 11.8475 | 12.4123 | 13.3335 | 12.9346 | 12.8082 | 12.7879 | 10.9096 | 11.0973 | 11.41 | 10.9634 | 10.3172 | 9.3036 | 11.1423 | 11.956 | 11.5902 | 11.4282 | 11.2403 | 11.051 | 11.2373 | 10.8506 | 11.0599 | [7 empty] | 9.8588 | 9.6914 | 7.7179 | 8.2375
47 | CCV 10 ppb MAJORS | R | 06/02/2023 20:09:27 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.049 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 14892.7634 | 14174.5086 | 14.4885 | 14.0731 | 14651.0299 | 13072.0905 | 13669.1522 | 12.2021 | 11.7414 | 12.0275 | 12.452 | 13.5502 | 12.469 | 12.641 | 12.5974 | 10.7941 | 11.3803 | 11.4364 | 10.8255 | 10.2835 | 9.3139 | 11.2431 | 12.1584 | 11.5709 | 11.4959 | 11.5765 | 11.0262 | 11.3981 | 11.0708 | 11.2378 | [7 empty] | 9.9257 | 9.7906 | 7.7776 | 8.7688
48 | MT 5 | R | 06/02/2023 20:15:30 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 5.050 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 5444.3989 | 303.142 | 572.47 | 3141.2391 | 4613.6915 | 4290.7112 | 924.0429 | 0.53388 | 0.64217 | 0.61705 | 110.827 | 1.055 | 119.3673 | 0.097861 | 0.43402 | 4.6243 | 4.6274 | 15.2062 | 1.8163 | 1.514 | 1.4764 | 1.8379 | 1.0096 | 1.0629 | 1.0186 | 1.0114 | 1.1309 | 0.93376 | 0.075855 | 0.06983 | [7 empty] | 0.08105 | 0.079407 | 0.14204 | 0.14285
49 | MT 6 | R | 06/02/2023 20:21:32 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 6.051 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 13254.0143 | 719.7689 | 1332.6762 | 7172.5836 | 9196.9916 | 10665.9048 | 2172.6291 | 1.2201 | 1.3237 | 1.2586 | 243.5425 | 2.3934 | 261.7472 | 0.20286 | 0.97031 | 10.8734 | 10.7793 | 35.549 | 4.2744 | 4.0632 | 3.7545 | 4.5278 | 2.0552 | 2.3986 | 2.5983 | 2.4247 | 2.5625 | 2.5676 | 0.18769 | 0.17022 | [7 empty] | 0.1924 | 0.18337 | 0.2689 | 0.27379
50 | MT 7 | R | 06/02/2023 20:27:35 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 7.052 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 13093.7643 | 758.1468 | 1388.7267 | 7189.1611 | 9166.6726 | 10793.4359 | 1935.6723 | 1.1856 | 1.3271 | 1.3165 | 238.1719 | 2.3275 | 255.2242 | 0.20901 | 1.0616 | 10.8057 | 10.8248 | 35.487 | 4.3018 | 3.8956 | 3.7332 | 4.5399 | 2.1609 | 2.3122 | 2.6021 | 2.4548 | 2.7629 | 2.3709 | 0.2028 | 0.17851 | [7 empty] | 0.17877 | 0.18277 | 0.31886 | 0.31049
51 | MT 8 | R | 06/02/2023 20:33:38 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 8.053 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 12795.8848 | 718.3922 | 1342.167 | 6781.8018 | 9112.9461 | 10496.7611 | 2058.432 | 1.1633 | 1.3687 | 1.2594 | 217.3419 | 2.2428 | 241.7167 | 0.20255 | 1.0324 | 10.4651 | 10.0661 | 32.6747 | 3.9864 | 3.7723 | 3.6098 | 4.2297 | 1.992 | 2.2354 | 2.3749 | 2.3077 | 2.4231 | 2.3432 | 0.17724 | 0.17091 | [7 empty] | 0.17447 | 0.1666 | 0.26395 | 0.26116
52 | MT 9 | R | 06/02/2023 20:39:40 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 9.054 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 12208.9519 | 653.6601 | 1285.0152 | 6982.8798 | 8663.2891 | 9719.0445 | 1911.2318 | 1.1438 | 1.2449 | 1.2333 | 224.1746 | 2.2126 | 233.4657 | 0.19984 | 0.93338 | 11.7302 | 11.9763 | 34.085 | 4.27 | 3.6746 | 3.5425 | 4.2531 | 2.2316 | 2.5109 | 2.3457 | 2.3332 | 2.7059 | 2.4277 | 0.18493 | 0.16324 | [7 empty] | 0.1752 | 0.1756 | 0.27053 | 0.25304
53 | MT 10 | R | 06/02/2023 20:45:42 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 10.055 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 12729.0397 | 760.0704 | 1375.8636 | 7377.1544 | 8982.8171 | 10488.9016 | 1993.7665 | 1.1939 | 1.5046 | 1.4337 | 246.053 | 2.4661 | 260.7206 | 0.20506 | 1.0248 | 11.9928 | 11.9431 | 36.2005 | 4.3419 | 3.758 | 3.8499 | 4.572 | 2.0949 | 2.3888 | 2.5134 | 2.543 | 2.593 | 2.4176 | 0.19157 | 0.18912 | [7 empty] | 0.18258 | 0.17426 | 0.26714 | 0.25831
54 | MT 11 | R | 06/02/2023 20:51:44 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 11 .056 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 25247.4894 | 1464.7179 | 2644.5768 | 14676.4934 | 16195.242 | 20678.5465 | 4332.207 | 2.4004 | 2.5326 | 2.5679 | 478.5436 | 4.9298 | 506.2743 | 0.40944 | 2.0135 | 20.152 | 20.4554 | 71.2293 | 8.4965 | 7.3843 | 7.5744 | 8.9291 | 4.1273 | 5.0453 | 4.9622 | 5.0111 | 5.5323 | 5.092 | 0.41223 | 0.35751 | [7 empty] | 0.36724 | 0.353 | 0.46897 | 0.47136
55 | MT 11 LD | R | 06/02/2023 20:57:47 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 11 LD.057 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 25625.0538 | 1403.6442 | 2638.7866 | 14003.7903 | 15954.7379 | 20235.4264 | 4416.3157 | 2.3315 | 2.5971 | 2.5805 | 469.8813 | 4.7747 | 497.1705 | 1.0024 | 1.9664 | 20.1502 | 20.3807 | 70.4693 | 8.5636 | 7.6082 | 7.4979 | 8.6688 | 3.985 | 4.785 | 4.8997 | 4.9278 | 5.3562 | 4.9587 | 0.38904 | 0.34505 | [7 empty] | 0.35397 | 0.34548 | 0.48258 | 0.46516
56 | MT 11 LD | R | 06/02/2023 21:04:18 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 11 LD.058 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 26893.5705 | 1560.6703 | 2833.4009 | 14906.4788 | 17127.6831 | 22054.5885 | 4690.7287 | 2.5068 | 2.8262 | 2.5254 | 473.5868 | 4.8903 | 538.092 | 1.0642 | 2.1366 | 21.7048 | 20.1428 | 69.8057 | 8.4572 | 8.1214 | 7.456 | 8.8817 | 3.7901 | 4.6811 | 4.98 | 4.856 | 5.295 | 4.7465 | 0.37081 | 0.35376 | [7 empty] | 0.34739 | 0.33726 | 0.46585 | 0.44984
57 | MT 12 | R | 06/02/2023 21:10:20 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 12.059 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 26342.5624 | 1560.1194 | 2680.7375 | 15291.7572 | 16365.7034 | 21281.8487 | 4253.8227 | 2.389 | 2.5949 | 2.7041 | 499.7453 | 4.7999 | 498.9965 | 0.43527 | 1.9143 | 20.7387 | 22.1678 | 75.2042 | 8.8558 | 7.7155 | 7.4343 | 9.0999 | 4.2069 | 5.3851 | 5.364 | 5.1638 | 5.537 | 5.2368 | 0.40404 | 0.35362 | [7 empty] | 0.39152 | 0.36931 | 0.48594 | 0.47577
58 | MT 12 LSP | R | 06/02/2023 21:16:23 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 12 LSP.060 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 68607.0345 | 38472.7417 | 2748.4494 | 15852.8372 | 53419.8926 | 59320.7928 | 41018.542 | 10.8119 | 10.4666 | 10.7756 | 511.793 | 13.6767 | 516.1271 | 8.6999 | 10.3288 | 27.6434 | 28.2797 | 82.0636 | 16.8457 | 14.8397 | 13.6994 | 16.7925 | 13.1244 | 13.2042 | 13.3986 | 12.8997 | 13.7037 | 13.2825 | 7.6431 | 7.6106 | [7 empty] | 6.6157 | 6.5279 | 5.5967 | 5.6225
59 | LBLANK | R | 06/02/2023 21:22:26 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LBLANK.061 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | -143.7588 | -46.5218 | -0.78556 | 0.15823 | 1683.4894 | -160.4684 | -4.861 | 0.0094303 | 0.005387 | 0.00051521 | -0.50395 | -0.0029249 | -0.14883 | 0.00049207 | 0.0021397 | 0.0025971 | -0.0069244 | -0.18386 | -0.047651 | 0.0043281 | -0.0070991 | -0.0074371 | 0.14709 | 0.020302 | -0.026509 | -0.0045254 | 0.016283 | -0.00080727 | 0.0011055 | -0.011562 | [7 empty] | -0.0006411 | 0.00059854 | -0.0018058 | -0.00074913
60 | CCV 1ppb TRACE | R | 06/02/2023 21:28:29 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.062 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1517.2125 | 1520.3576 | 0.68273 | 1.2064 | 3447.1305 | 1288.0561 | 1473.8406 | 1.3139 | 1.2404 | 1.2524 | 0.82199 | 1.3582 | 1.3041 | 1.3271 | 1.3551 | 1.2355 | 1.2313 | 1.0177 | 1.074 | 1.0126 | 0.91882 | 1.1237 | 1.3492 | 1.0093 | 1.1465 | 1.1391 | 1.1309 | 1.0883 | 1.0961 | 1.0759 | [7 empty] | 1.0191 | 1.0049 | 0.8405 | 0.83944
61 | CCV 1ppb TRACE | R | 06/02/2023 21:35:00 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.063 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1538.606 | 1596.5673 | 0.98671 | 1.1278 | 3499.2738 | 1288.1703 | 1500.661 | 1.2957 | 1.2455 | 1.2826 | 0.89831 | 1.3918 | 1.292 | 1.3285 | 1.3127 | 1.2022 | 1.2776 | 1.0317 | 1.0332 | 0.96423 | 0.91999 | 1.1448 | 1.3521 | 1.1303 | 1.2286 | 1.1622 | 1.0164 | 1.1517 | 1.1242 | 1.1045 | [7 empty] | 1.0321 | 1.0195 | 0.84105 | 0.83925
62 | LFB 10ppb | R | 06/02/2023 21:41:03 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LFB 10ppb.064 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 16963.3012 | 15798.9625 | 15.1889 | 16.611 | 16494.2433 | 14309.8878 | 14470.7536 | 12.6252 | 12.1419 | 12.6794 | 13.6782 | 12.2563 | 13.6614 | 12.7325 | 12.9815 | 11.0887 | 11.6226 | 15.8761 | 10.6027 | 10.3309 | 8.3066 | 10.2467 | 11.8793 | 10.5443 | 10.9292 | 10.454 | 10.3203 | 10.8775 | 10.9785 | 10.1044 | [7 empty] | 9.7768 | 9.6408 | 7.1442 | 7.7084
63 | LFB 10ppb | R | 06/02/2023 21:47:33 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LFB 10ppb.065 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 16616.0665 | 15582.4334 | 15.9918 | 17.8106 | 16644.221 | 14030.1672 | 14342.6585 | 12.6007 | 12.1688 | 12.405 | 13.4389 | 13.7796 | 13.7973 | 12.9494 | 13.1534 | 11.1804 | 11.3402 | 15.4052 | 11.2858 | 10.2763 | 9.6147 | 11.5632 | 12.5099 | 12.0041 | 11.6658 | 11.3745 | 11.5239 | 11.5838 | 10.7815 | 11.2615 | [7 empty] | 9.6619 | 9.5592 | 7.4358 | 7.6839
64 | CCV 10 ppb MAJORS | R | 06/02/2023 21:53:35 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.066 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 16818.1446 | 15942.7348 | 15.6709 | 16.672 | 17044.9012 | 14246.7741 | 14415.2684 | 12.6443 | 12.306 | 12.3118 | 12.7164 | 13.9061 | 13.1644 | 13.0024 | 13.286 | 11.1987 | 11.2807 | 11.6984 | 13.0962 | 10.9236 | 9.4231 | 11.477 | 14.5036 | 11.2376 | 11.6545 | 11.1099 | 13.1031 | 11.4886 | 10.7319 | 11.2272 | [7 empty] | 9.5234 | 9.3744 | 8.6335 | 7.8975
65 | CCV 10 ppb MAJORS | R | 06/02/2023 22:00:06 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.067 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 16930.2435 | 15671.7282 | 15.1116 | 13.8925 | 16418.3915 | 14348.0151 | 14649.3146 | 12.7903 | 12.4661 | 13.4725 | 13.8084 | 13.8865 | 13.2379 | 13.1027 | 13.0618 | 11.236 | 12.3435 | 12.5048 | 11.359 | 10.6404 | 9.2898 | 11.5157 | 12.5503 | 11.7664 | 11.6988 | 11.6708 | 11.4076 | 11.2852 | 11.6897 | 11.2583 | [7 empty] | 10.358 | 10.1935 | 7.5404 | 7.9021
66 | MT 13 | R | 06/02/2023 22:06:09 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 13.068 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 30737.5572 | 1820.8405 | 3100.6238 | 17961.8263 | 19195.8508 | 24300.9309 | 4801.868 | 2.7657 | 2.9328 | 2.8446 | 531.2763 | 5.3806 | 557.2625 | 0.41973 | 2.1147 | 23.0437 | 22.8116 | 77.5063 | 10.0831 | 8.3222 | 8.201 | 9.8608 | 4.8127 | 5.5166 | 5.3644 | 5.3412 | 5.9779 | 5.0301 | 0.45202 | 0.36972 | [7 empty] | 0.39763 | 0.37004 | 0.51607 | 0.47666
67 | MT 14 | R | 06/02/2023 22:12:11 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 14.069 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 30703.9049 | 1770.0111 | 3009.2567 | 17140.4128 | 18615.777 | 24097.4289 | 4432.4352 | 2.6847 | 2.8559 | 3.0333 | 583.4065 | 5.853 | 549.5117 | 0.44979 | 2.19 | 22.6285 | 22.5197 | 75.6795 | 9.4708 | 8.3281 | 7.9865 | 9.6979 | 4.3 | 5.5278 | 5.431 | 5.4105 | 5.6735 | 5.4869 | 0.46434 | 0.37948 | [7 empty] | 0.38501 | 0.36444 | 0.49709 | 0.49228
68 | MT 15 | R | 06/02/2023 22:18:14 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 15.070 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 30329.0572 | 1775.8103 | 2986.0743 | 17146.9435 | 18819.1582 | 24220.4616 | 4441.4852 | 2.6714 | 2.9493 | 2.9503 | 517.054 | 9.7159 | 545.2228 | 0.43221 | 2.7244 | 22.5383 | 22.1493 | 76.1901 | 9.4913 | 7.9056 | 8.0273 | 9.7551 | 4.4638 | 5.4427 | 5.5721 | 5.4362 | 5.7642 | 5.4677 | 0.47372 | 0.39614 | [7 empty] | 0.39311 | 0.37653 | 0.49425 | 0.48125
69 | MT 16 | R | 06/02/2023 22:24:17 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 16.071 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 8952.299 | 1172.257 | 24.6276 | 13630.4333 | 13152.7928 | 9335.3707 | 502.9978 | 0.25554 | 2.7674 | 2.7594 | 1249.2524 | 8.199 | 1139.5852 | 0.2229 | 2.1867 | 22.6325 | 22.9954 | 66.5366 | 7.0708 | 5.1974 | 5.2779 | 6.4969 | 4.7078 | 4.9034 | 5.0627 | 5.0403 | 5.7142 | 4.8317 | 2.231 | 2.2188 | [7 empty] | 11.8326 | 11.2179 | 0.21164 | 0.17442
70 | MT 17 | R | 06/02/2023 22:30:20 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 17.072 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 21229.9516 | 2684.8274 | 52.6549 | 30446.6843 | 28066.5028 | 22237.4235 | 1132.122 | 0.57723 | 6.0914 | 5.8882 | 2749.1227 | 19.764 | 2738.849 | 0.51094 | 4.9551 | 50.6334 | 49.3095 | 144.4342 | 14.2194 | 12.3232 | 13.2062 | 15.7811 | 9.0058 | 12.269 | 12.6082 | 12.1684 | 11.3891 | 11.9798 | 4.7945 | 5.2846 | [7 empty] | 25.856 | 24.7805 | 0.37231 | 0.35988
71 | MT 18 | R | 06/02/2023 22:36:32 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 18.073 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 18623.5557 | 2281.0516 | 46.3191 | 27615.8101 | 24636.082 | 19477.6248 | 1004.3962 | 0.51498 | 4.7032 | 4.7108 | 2418.7587 | 15.5559 | 2185.5614 | 0.44386 | 4.4388 | 43.0445 | 43.2757 | 124.9763 | 12.5026 | 10.3976 | 10.5607 | 12.5271 | 7.7149 | 9.6203 | 9.9076 | 9.6502 | 9.9133 | 9.835 | 4.2291 | 4.2799 | [7 empty] | 22.396 | 21.2256 | 0.29733 | 0.28141
72 | MT 19 | R | 06/02/2023 22:42:43 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 19.074 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 9214.5213 | 1187.8096 | 24.0498 | 13890.2763 | 13076.5655 | 9523.619 | 505.9479 | 0.25464 | 2.286 | 2.2286 | 1234.8446 | 7.8282 | 1084.3919 | 0.22943 | 3.4478 | 24.9884 | 25.5039 | 63.8711 | 5.9391 | 5.1523 | 5.0124 | 6.293 | 3.8344 | 4.5419 | 4.9302 | 4.797 | 4.6841 | 4.6831 | 2.0523 | 2.0806 | [7 empty] | 11.254 | 10.6989 | 0.18799 | 0.17659
73 | MT 20 | R | 06/02/2023 22:48:45 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 20.075 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 20144.4177 | 2380.2906 | 48.5462 | 28884.8142 | 25867.958 | 20396.1114 | 995.6703 | 0.50021 | 4.7815 | 4.8015 | 2499.7317 | 16.1044 | 2265.957 | 0.44006 | 4.0102 | 45.2108 | 45.3714 | 130.9351 | 12.2058 | 10.8696 | 10.7466 | 12.8257 | 7.7118 | 9.8549 | 10.3885 | 9.8474 | 9.8735 | 9.8703 | 4.1843 | 4.2133 | [7 empty] | 23.357 | 22.0995 | 0.30253 | 0.28984
74 | MT 21 | R | 06/02/2023 22:54:57 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 21.076 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 10368.6275 | 1389.894 | 26.1933 | 16791.9613 | 13708.9997 | 10586.5301 | 573.1915 | 0.26084 | 2.8347 | 2.8774 | 1356.5094 | 8.6074 | 1209.7909 | 0.23841 | 2.4675 | 23.6617 | 24.0391 | 72.0528 | 6.2961 | 5.5014 | 5.3945 | 6.5977 | 4.0134 | 5.1139 | 5.1603 | 5.0197 | 5.1573 | 4.7603 | 2.2103 | 2.2105 | [7 empty] | 12.2157 | 11.7715 | 0.21137 | 0.19643
75 | MT 21 LD | R | 06/02/2023 23:00:59 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 21 LD.077 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | -139.1487 | -49.4851 | -1.0164 | 0.38984 | -1914.8967 | -161.295 | -4.861 | 0.0071048 | 0.008229 | 0.00097882 | -0.50997 | -0.004217 | -0.33164 | 0.00052829 | -0.0099214 | 0.023057 | -0.010154 | -0.25227 | 0.62232 | 0.0048715 | -0.0075785 | -0.0082167 | -1.6368 | 0.020145 | -0.030202 | -0.0063845 | -0.09456 | -0.0054505 | 0.0016234 | -0.012515 | [7 empty] | -0.00077791 | 0.00021411 | -0.0025041 | -0.0012344
76 | LBLANK | R | 06/02/2023 23:07:00 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LBLANK.078 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | -145.4882 | -49.4851 | -0.80611 | 0.55563 | 1671.8764 | -154.3215 | -2.8124 | 0.0079497 | 0.0071477 | 0.0043221 | -0.4533 | -0.0033659 | -0.087474 | 0.00043961 | -0.0024758 | 0.0016345 | -0.007604 | -0.224 | -0.10021 | -0.0031069 | -0.0067848 | -0.0071568 | 0.2154 | 0.026794 | -0.032915 | -0.0039661 | -0.059389 | -0.0054505 | -0.0090693 | -0.012204 | [7 empty] | -0.00078141 | 0.00033854 | -0.0020194 | -0.00071883
77 | CCV 1ppb TRACE | R | 06/02/2023 23:13:03 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.079 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1921.9981 | 1865.4923 | 1.028 | 2.1966 | 3700.4663 | 1516.1001 | 1591.8276 | 1.398 | 1.2745 | 1.2502 | 0.82723 | 1.4553 | 1.2864 | 1.3102 | 1.31 | 1.227 | 1.1453 | 0.93786 | 1.1259 | 1.0218 | 0.96293 | 1.144 | 1.2541 | 1.1468 | 1.2247 | 1.1432 | 1.1947 | 1.2084 | 1.0331 | 1.1041 | [7 empty] | 0.94444 | 0.93824 | 0.82146 | 0.82402
78 | CCV 1ppb TRACE | R | 06/02/2023 23:19:35 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.080 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 1937.1723 | 1884.1061 | 1.4867 | 1.9533 | 4118.1488 | 1574.5214 | 1515.4807 | 1.3645 | 1.3246 | 1.3493 | 0.93922 | 1.4504 | 1.4205 | 1.3774 | 1.3459 | 1.2303 | 1.2412 | 1.0183 | 1.0005 | 1.1337 | 0.9604 | 1.1573 | 1.2777 | 1.2424 | 1.2232 | 1.1422 | 1.0542 | 1.1886 | 1.0934 | 1.1163 | [7 empty] | 1.0072 | 0.98545 | 0.82437 | 0.81035
79 | CCV 10 ppb MAJORS | R | 06/02/2023 23:25:38 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.081 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 21067.7647 | 18684.2381 | 17.5281 | 18.0676 | 18822.0334 | 17038.6265 | 16145.4684 | 13.2759 | 12.979 | 13.0011 | 13.6353 | 14.4695 | 13.6609 | 13.3696 | 13.6205 | 11.5073 | 11.6265 | 11.9894 | 11.4145 | 11.0874 | 9.5235 | 11.6226 | 12.2196 | 11.8034 | 12.3136 | 11.4598 | 11.5536 | 11.3066 | 10.8811 | 11.3095 | [7 empty] | 9.7467 | 9.6583 | 7.311 | 7.7987
80 | CCV 10 ppb MAJORS | R | 06/02/2023 23:32:09 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.082 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 21074.6872 | 18835.8147 | 18.6471 | 17.3845 | 18858.5816 | 16776.0468 | 15757.751 | 13.4103 | 12.9431 | 13.2497 | 13.6733 | 14.7788 | 13.6744 | 13.4601 | 13.5352 | 11.6628 | 11.7513 | 12.0642 | 11.4824 | 11.0534 | 9.6273 | 11.7135 | 12.2649 | 11.513 | 11.8556 | 11.7302 | 12.0614 | 11.7759 | 11.1743 | 11.3448 | [7 empty] | 9.8131 | 9.8278 | 7.3186 | 7.5647
81 | MT 22 | R | 06/02/2023 23:38:12 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 22.083 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 7350.3081 | 1011.1728 | 19.7143 | 10961.948 | 10871.874 | 7490.4282 | 374.7182 | 0.22444 | 1.6051 | 1.6466 | 941.3146 | 6.019 | 822.8003 | 0.16519 | 1.4798 | 16.2837 | 16.3145 | 46.6807 | 4.2972 | 3.6982 | 3.867 | 4.5236 | 2.8171 | 3.5258 | 3.6479 | 3.486 | 3.4347 | 3.3286 | 1.4823 | 1.4479 | [7 empty] | 8.1357 | 7.6787 | 0.17965 | 0.17139
82 | MT 22 LSP | R | 06/02/2023 23:44:15 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 22 LSP.084 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 61874.9849 | 48592.044 | 32.2961 | 2.8966 | 51960.4154 | 51552.8568 | 41753.7778 | 9.1641 | 10.5983 | 10.1595 | 993.632 | 15.6714 | 849.2299 | 8.7239 | 10.0689 | 23.5674 | 23.2071 | 55.4439 | 11.9743 | 10.6146 | 10.0857 | 12.2178 | 10.8796 | 11.4073 | 11.8509 | 11.2157 | 11.3957 | 10.9367 | 8.586 | 8.7843 | [7 empty] | 14.2626 | 13.8916 | 4.864 | 4.9946
83 | MT 22 LSP | R | 06/02/2023 23:50:45 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 22 LSP.085 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 56021.5458 | 42869.2132 | 26.5192 | 12429.388 | 45997.6668 | 46870.1974 | 37071.8348 | 8.026 | 9.1533 | 10.1918 | 992.1416 | 15.7762 | 757.7792 | 7.8074 | 8.9608 | 21.1906 | 23.1063 | 55.2389 | 11.9942 | 9.6926 | 10.1601 | 12.1738 | 10.9098 | 11.6752 | 11.7744 | 11.2124 | 11.2857 | 10.8573 | 8.6205 | 8.7875 | [7 empty] | 14.4183 | 13.8689 | 4.9356 | 4.8327
84 | MT 23 | R | 06/02/2023 23:56:47 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 23.086 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 4167.6084 | 531.5582 | 12.1815 | 6379.9953 | 7387.8538 | 4361.5988 | 237.885 | 0.11409 | 1.1276 | 1.0868 | 504.7472 | 3.5213 | 508.0426 | 0.10454 | 0.94733 | 9.5858 | 9.6734 | 28.5944 | 2.6441 | 2.2552 | 2.1525 | 2.647 | 2.0474 | 2.193 | 2.0502 | 1.9592 | 2.189 | 2.1011 | 0.96868 | 0.92851 | [7 empty] | 4.8808 | 4.6181 | 0.10234 | 0.083308
85 | MT 24 | R | 06/03/2023 00:02:49 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 24.087 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 3953.5362 | 461.5362 | 10.051 | 5575.9935 | 6759.4653 | 3941.4553 | 190.6957 | 0.11138 | 1.0913 | 1.1048 | 452.6408 | 3.3004 | 451.843 | 0.093842 | 0.91476 | 9.0554 | 9.0861 | 26.0721 | 2.4435 | 1.9862 | 2.0786 | 2.518 | 1.6085 | 2.2221 | 2.0217 | 1.8397 | 1.9844 | 1.9986 | 0.82636 | 0.84657 | [7 empty] | 4.4974 | 4.2525 | 0.074568 | 0.071418
86 | MT 25 | R | 06/03/2023 00:08:52 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 25.088 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 4647.6853 | 593.1143 | 13.2612 | 6963.9529 | 7198.9212 | 4765.6722 | 296.9423 | 0.13343 | 0.79615 | 0.85817 | 523.3037 | 3.8095 | 514.8923 | 0.10586 | 0.68377 | 10.5038 | 10.8708 | 30.6891 | 2.894 | 2.3309 | 2.4426 | 2.9203 | 2.3797 | 2.4727 | 2.1586 | 2.2642 | 2.383 | 2.0209 | 0.87431 | 0.8895 | [7 empty] | 5.1671 | 4.909 | 0.10707 | 0.095725
87 | MT 26 | R | 06/03/2023 00:14:55 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 26.089 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 4284.898 | 525.8116 | 14.2656 | 6463.3928 | 7101.0936 | 4202.6383 | 206.2797 | 0.12336 | 1.2065 | 1.1962 | 486.3739 | 3.5335 | 473.6568 | 0.097103 | 0.9879 | 9.6872 | 9.9788 | 28.048 | 3.0733 | 2.2087 | 2.2832 | 2.6812 | 2.696 | 2.1316 | 2.017 | 2.0221 | 2.6589 | 1.9497 | 0.89117 | 0.91208 | [7 empty] | 4.7494 | 4.5586 | 0.14609 | 0.08439
88 | MT 27 | R | 06/03/2023 00:21:03 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 27.090 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 12290.216 | 1704.7488 | 29.9034 | 17410.0987 | 16171.658 | 12068.5986 | 733.8392 | 0.30382 | 2.7257 | 2.6016 | 1472.115 | 9.683 | 1274.2666 | 0.25533 | 2.361 | 24.8398 | 24.9139 | 72.3287 | 6.6712 | 5.7115 | 6.0217 | 7.1672 | 4.1722 | 5.7849 | 5.4283 | 5.4354 | 5.4076 | 5.2683 | 2.3074 | 2.3514 | [7 empty] | 12.4902 | 11.8986 | 0.1834 | 0.17704
89 | MT 28 | R | 06/03/2023 00:27:06 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 28.091 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 13748.68 | 1689.5292 | 32.522 | 18895.0879 | 17998.3442 | 13362.4792 | 689.1893 | 0.32406 | 2.5894 | 2.6384 | 1608.5639 | 10.0255 | 1429.693 | 0.28189 | 2.3779 | 28.2912 | 27.2659 | 80.1424 | 7.7783 | 6.8868 | 6.5432 | 7.6258 | 5.2128 | 6.0895 | 6.0755 | 5.9121 | 6.3238 | 5.9183 | 2.4737 | 2.4263 | [7 empty] | 13.7015 | 13.2527 | 0.19538 | 0.17245
90 | MT 29 | R | 06/03/2023 00:33:08 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 29.092 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 25930.7988 | 3315.7679 | 59.2385 | 35007.4459 | 30986.87 | 25936.1238 | 1380.7219 | 0.60663 | 5.4587 | 5.6832 | 3144.7884 | 19.8541 | 2626.8061 | 0.52079 | 5.0038 | 51.8753 | 52.5969 | 154.718 | 14.9214 | 12.7158 | 13.0098 | 15.0609 | 9.4267 | 12.3204 | 11.871 | 11.6849 | 12.2219 | 11.6484 | 4.9995 | 5.0585 | [7 empty] | 26.9175 | 25.6433 | 0.33965 | 0.29001
91 | LBLANK | R | 06/03/2023 00:39:20 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LBLANK.093 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | -145.5986 | -49.4851 | -0.84614 | 1.9448 | 2279.0338 | -149.0433 | -0.70675 | 0.0091747 | 0.0054108 | 0.0031163 | -0.42463 | -0.0024772 | 0.023472 | 0.001005 | -0.0033192 | 0.0028118 | -0.0076276 | -0.21115 | -0.1254 | 0.015162 | -0.0071337 | -0.0075301 | 0.43478 | 0.014903 | -0.02441 | -0.0071842 | -0.02621 | -8.5498e-05 | -0.0024399 | -0.012072 | [7 empty] | 1.9841e-05 | 0.00049455 | -0.0018493 | -0.00072881
92 | CCV 1ppb TRACE | R | 06/03/2023 00:45:23 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.094 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2047.6878 | 1948.7255 | 1.3815 | 2.2013 | 3767.647 | 1589.153 | 1598.6672 | 1.3709 | 1.2941 | 1.3767 | 1.0321 | 1.5752 | 1.2912 | 1.329 | 1.3189 | 1.2415 | 1.2681 | 1.042 | 1.0363 | 1.0971 | 1.0197 | 1.2447 | 1.5493 | 1.3729 | 1.24 | 1.196 | 1.0755 | 1.1923 | 1.1168 | 1.2005 | [7 empty] | 1.0115 | 1.0069 | 0.78832 | 0.81479
93 | CCV 1ppb TRACE | R | 06/03/2023 00:51:55 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.095 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2146.4917 | 2009.1125 | 1.5307 | 1.4234 | 4216.6121 | 1634.9721 | 1633.7414 | 1.3741 | 1.3396 | 1.3082 | 0.91974 | 1.5133 | 1.4226 | 1.3757 | 1.3973 | 1.2453 | 1.2151 | 0.98749 | 1.0144 | 1.0212 | 0.97742 | 1.1631 | 1.4628 | 1.1878 | 1.1896 | 1.1879 | 1.057 | 0.98405 | 1.0508 | 1.1257 | [7 empty] | 0.96441 | 0.96082 | 0.79491 | 0.80002
94 | LFB 10ppb | R | 06/03/2023 00:57:57 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LFB 10ppb.096 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 22743.1014 | 20507.6626 | 19.6136 | 20.8277 | 20007.7152 | 17912.0399 | 17261.8613 | 14.2062 | 14.006 | 13.1488 | 14.7685 | 13.5946 | 15.4391 | 14.1674 | 14.0358 | 12.2584 | 11.6318 | 15.9831 | 11.2699 | 11.6136 | 9.2992 | 11.0607 | 12.7432 | 11.166 | 11.5117 | 11.0525 | 11.8479 | 11.436 | 10.8598 | 10.589 | [7 empty] | 9.7247 | 9.5267 | 6.8686 | 7.3993
95 | LFB 10ppb | R | 06/03/2023 01:04:27 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LFB 10ppb.097 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 22121.7484 | 19529.2105 | 17.9986 | 20.2756 | 19349.5194 | 17860.4334 | 16677.1287 | 13.5251 | 13.2079 | 12.7902 | 14.0186 | 14.3222 | 14.4788 | 13.5183 | 13.5772 | 11.7374 | 11.265 | 15.6674 | 11.4809 | 10.6821 | 10.2316 | 11.5979 | 12.5989 | 11.5237 | 11.9089 | 11.5022 | 11.783 | 11.4808 | 10.5915 | 11.0736 | [7 empty] | 9.3144 | 9.2147 | 6.9899 | 7.3386
96 | CCV 10 ppb MAJORS | R | 06/03/2023 01:10:29 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.098 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 22589.6947 | 20393.5272 | 17.3366 | 19.5679 | 19425.7924 | 17763.5482 | 16839.7548 | 14.0308 | 13.6793 | 13.7998 | 14.2817 | 15.1421 | 13.7263 | 13.6078 | 13.8226 | 11.7902 | 12.0951 | 12.5491 | 11.2899 | 10.8352 | 9.8758 | 11.9667 | 12.4697 | 12.7502 | 12.2478 | 11.7033 | 11.4765 | 12.1988 | 11.2995 | 11.581 | [7 empty] | 9.9651 | 9.826 | 7.0548 | 7.5393
97 | CCV 10 ppb MAJORS | R | 06/03/2023 01:17:01 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.099 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 21768.7225 | 19295.4051 | 17.1835 | 19.0631 | 18877.5652 | 17566.8589 | 16012.6827 | 13.6025 | 13.0867 | 13.4579 | 13.9046 | 14.8392 | 13.4556 | 13.3476 | 13.6075 | 11.8217 | 11.7167 | 12.0015 | 12.1097 | 10.6211 | 10.2884 | 11.8442 | 13.8678 | 11.8764 | 11.8966 | 11.4785 | 11.9035 | 11.4277 | 10.8188 | 11.2882 | [7 empty] | 9.6592 | 9.5858 | 7.5939 | 7.5223
98 | MT 30 | R | 06/03/2023 01:23:04 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 30.100 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | -110.7365 | 221.1253 | 728.2531 | 2886.1545 | 4884.4676 | 23.8532 | 2048.7213 | 0.91961 | 0.6147 | 0.64576 | 768.3711 | 121.3975 | 585.3216 | 0.57087 | 1.006 | 70.9182 | 80.1046 | 154.8468 | 1.4167 | 1.4386 | 1.3133 | 1.5507 | 0.44731 | 0.42512 | 0.37839 | 0.44905 | 0.50002 | 0.44832 | 0.35519 | 0.30881 | [7 empty] | 1.8206 | 1.7278 | 2.2988 | 2.4146
99 | MT 31 | R | 06/03/2023 01:29:15 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 31.101 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 67.1371 | 422.5248 | 548.3727 | 2843.9193 | 3847.9022 | 742.1862 | 1148.704 | 0.66841 | 0.46758 | 0.45224 | 429.6306 | 87.6994 | 434.4449 | 0.56428 | 0.64272 | 52.3942 | 51.031 | 160.4809 | 1.273 | 1.2283 | 1.1539 | 1.3629 | 0.4243 | 0.25557 | 0.3279 | 0.29876 | 0.35966 | 0.25233 | 0.17954 | 0.19105 | [7 empty] | 1.7808 | 1.6966 | 1.6865 | 1.7046
100 | MT 31 LD | R | 06/03/2023 01:35:29 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 31 LD.102 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 79.2036 | 397.7084 | 567.784 | 2921.2983 | 4121.6521 | 781.8006 | 1187.5135 | 0.71018 | 0.46475 | 0.47577 | 442.7182 | 89.6347 | 442.5217 | 0.56217 | 0.64128 | 53.4535 | 51.9474 | 168.1468 | 1.3057 | 1.1802 | 1.1779 | 1.4009 | 0.45437 | 0.28631 | 0.27571 | 0.30633 | 0.40555 | 0.35221 | 0.18775 | 0.19691 | [7 empty] | 1.8267 | 1.7531 | 1.6913 | 1.7004
101 | MT 31 LD | R | 06/03/2023 01:42:00 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 31 LD.103 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 78.5828 | 372.4812 | 550.5283 | 2913.8358 | 3936.957 | 713.0268 | 1122.1404 | 0.66631 | 0.48509 | 0.46498 | 441.1994 | 89.8688 | 426.6917 | 0.57672 | 0.60685 | 51.0588 | 51.8841 | 165.7256 | 1.3194 | 1.0225 | 1.1705 | 1.3978 | 0.39925 | 0.27937 | 0.30642 | 0.30079 | 0.32646 | 0.25441 | 0.18376 | 0.18537 | [7 empty] | 1.8206 | 1.7338 | 1.6575 | 1.813
102 | MT 32 | R | 06/03/2023 01:48:02 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 32.104 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | -144.7487 | -47.6428 | 1.3164 | 15.509 | 2360.0329 | -141.8897 | 5.6901 | 0.0067194 | 0.0193 | 0.010786 | -0.37785 | 0.0031724 | 0.062826 | 0.0010276 | 0.0081659 | 0.050337 | 0.04171 | -0.17729 | -0.10322 | -0.0031069 | -0.0062945 | -0.0066437 | 0.22868 | 0.027702 | -0.034678 | -0.0081024 | -0.0027766 | -3.7398e-05 | 0.0019113 | 0.003399 | [7 empty] | -0.00097726 | 0.00076153 | 0.028341 | 0.027134
103 | MT 32 LSP | R | 06/03/2023 01:54:03 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 32 LSP.105 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 59887.2643 | 50644.7811 | 15.872 | 32.7045 | 45903.8404 | 47309.0756 | 43655.7692 | 9.3293 | 8.7804 | 9.1232 | 11.2615 | 10.2735 | 11.0396 | 8.8125 | 12.7901 | 7.548 | 7.6739 | 30.6556 | 7.5337 | 7.017 | 6.3827 | 7.6321 | 8.4141 | 7.4448 | 7.897 | 7.7676 | 7.9748 | 7.5708 | 7.1118 | 7.299 | [7 empty] | 6.1736 | 6.0817 | 4.5774 | 4.6939
104 | MT 32 LSP | R | 06/03/2023 02:00:34 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 32 LSP.106 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 58372.3945 | 49490.2435 | 16.4419 | 27.2185 | 44894.4671 | 45555.581 | 42167.8164 | 9.1397 | 8.7203 | 9.0414 | 11.1499 | 9.8701 | 11.038 | 8.8242 | 12.7897 | 7.6911 | 7.7355 | 30.9724 | 7.2768 | 7.2956 | 6.5845 | 7.5691 | 7.9927 | 7.786 | 7.8443 | 7.4632 | 7.3677 | 7.3296 | 7.181 | 7.1225 | [7 empty] | 6.1848 | 6.1477 | 4.4876 | 4.6328
105 | MT 33 | R | 06/03/2023 02:06:37 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 33.107 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | -142.7512 | -47.7302 | 1.8161 | 13.5315 | 2534.3942 | -154.9333 | 11.3607 | 0.007671 | 0.017918 | 0.012695 | -0.36244 | 0.0020194 | 0.078175 | 0.0031046 | 0.012863 | 0.05066 | 0.054577 | 0.7638 | -0.1182 | -0.0031069 | -0.0066087 | -0.006872 | 0.37803 | 0.014903 | -0.030705 | -0.0082007 | -0.044485 | 0.0050517 | 0.0029657 | -0.0099611 | [7 empty] | -0.00074733 | 0.00081791 | 0.021308 | 0.024079
106 | MT 34 | R | 06/03/2023 02:12:39 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 34.108 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | -142.178 | -49.4851 | 2.1566 | 15.1166 | 2761.4014 | -155.1564 | 5.323 | 0.0076619 | 0.028469 | 0.024526 | 3.6869 | 0.020232 | 4.0946 | 0.001539 | 0.045571 | 0.018414 | 0.006774 | -0.19065 | -0.16402 | 0.001452 | -0.0066964 | -0.0071322 | 0.4468 | 0.021236 | -0.026554 | -0.0081853 | -0.00071836 | -0.00013879 | -0.0071368 | -0.011437 | [7 empty] | -0.00088286 | 0.00091382 | 0.01857 | 0.019687
107 | MT 35 | R | 06/03/2023 02:18:42 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 35.109 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | -146.0364 | -48.5898 | 1.4573 | 14.9942 | 2873.0389 | -154.5021 | 11.4599 | 0.0071812 | 0.018914 | 0.011039 | 0.16023 | 0.0017853 | 0.59007 | 0.0015291 | 0.0079942 | 0.010703 | 0.0047741 | -0.1682 | -0.20451 | 0.0014451 | -0.0060964 | -0.0064258 | 0.42152 | 0.034891 | -0.032558 | -0.0061885 | -0.0036452 | -0.0054505 | -0.008408 | -0.011446 | [7 empty] | -0.00046169 | 0.00061959 | 0.020539 | 0.018983
108 | LBLANK | R | 06/03/2023 02:24:45 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LBLANK.110 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | -146.3403 | -49.4851 | -0.76116 | 1.5027 | 2955.6642 | -144.5824 | -0.69835 | 0.0076988 | 0.0086455 | 0.0011485 | -0.46784 | -0.0028974 | 0.0065413 | 0.00045 | 0.0070436 | 0.0020642 | -0.0095033 | -0.21796 | -0.10159 | -0.0031069 | -0.0076273 | -0.0081989 | 0.33343 | 0.014903 | -0.03055 | -0.010994 | -0.024012 | -0.0054505 | -0.00885 | -0.012493 | [7 empty] | -0.0012151 | 0.00025078 | -0.0019169 | -0.00072586
109 | CCV 1ppb TRACE | R | 06/03/2023 02:30:48 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.111 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2114.1886 | 1935.993 | 1.3111 | 1.8749 | 4368.4181 | 1685.1854 | 1692.7601 | 1.429 | 1.3469 | 1.4085 | 1.0811 | 1.6008 | 1.4819 | 1.3633 | 1.4149 | 1.3133 | 1.2914 | 1.1048 | 1.0234 | 1.1213 | 1.004 | 1.2204 | 1.2548 | 1.3467 | 1.1259 | 1.1653 | 1.1173 | 1.0488 | 1.1063 | 1.1574 | [7 empty] | 1.0032 | 1.0051 | 0.77883 | 0.72678
110 | CCV 1ppb TRACE | R | 06/03/2023 02:37:20 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.112 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2231.5756 | 2061.9625 | 1.7574 | 1.6109 | 4848.718 | 1729.6854 | 1729.0406 | 1.5125 | 1.432 | 1.4234 | 1.0527 | 1.5528 | 1.5316 | 1.4302 | 1.4569 | 1.3318 | 1.2927 | 1.0688 | 1.1164 | 1.2217 | 0.9724 | 1.1521 | 1.4405 | 1.2776 | 1.1988 | 1.1568 | 1.2388 | 1.2345 | 1.0896 | 1.1011 | [7 empty] | 1.0047 | 1.0047 | 0.77226 | 0.78767
111 | CCV 10 ppb MAJORS | R | 06/03/2023 02:43:23 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.113 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 24000.1796 | 21306.2792 | 19.2387 | 19.6868 | 21033.2194 | 19175.1195 | 18411.3094 | 14.7288 | 14.2458 | 13.7585 | 13.992 | 15.0344 | 14.5708 | 14.5704 | 14.6214 | 12.557 | 11.8956 | 12.138 | 11.3734 | 11.395 | 9.6021 | 11.6359 | 12.9356 | 11.9372 | 11.7948 | 11.4913 | 11.2424 | 11.6494 | 10.9499 | 11.233 | [7 empty] | 9.5604 | 9.4731 | 6.7632 | 7.5041
112 | CCV 10 ppb MAJORS | R | 06/03/2023 02:49:56 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.114 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 22860.7575 | 20634.5933 | 18.2955 | 19.3139 | 19326.5307 | 17904.7891 | 16952.1775 | 14.0346 | 13.5491 | 14.3133 | 14.6135 | 15.1383 | 13.9184 | 13.9297 | 13.828 | 11.9315 | 12.3968 | 12.5668 | 10.7862 | 10.649 | 10.0675 | 11.8767 | 11.6157 | 12.0555 | 12.0109 | 11.6614 | 10.6018 | 11.4443 | 11.4496 | 11.3598 | [7 empty] | 10.0635 | 9.9163 | 6.5032 | 7.4578
113 | MT 36 | R | 06/03/2023 02:55:59 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 36.115 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | -138.8261 | -46.8185 | 1.5042 | 12.4404 | 2471.587 | -152.5087 | 21.1422 | 0.0090932 | 0.036496 | 0.028221 | -0.038993 | 0.0047601 | 0.35193 | 0.001851 | 0.0075976 | 0.018681 | 0.0080063 | -0.15337 | -0.093735 | 0.0011785 | -0.0071259 | -0.0076165 | 0.2949 | 0.014903 | -0.034669 | -0.011007 | 0.013328 | -0.0054505 | -0.0048643 | -0.0096318 | [7 empty] | -0.00097943 | 0.00093873 | 0.018862 | 0.019869
114 | MT 37 | R | 06/03/2023 03:02:01 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 37.116 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | -140.2416 | -44.7635 | 1.5199 | 17.0547 | 3011.2968 | -133.9243 | 25.4643 | 0.0067013 | 0.018068 | 0.011265 | -0.34804 | 0.0081076 | 0.23325 | 0.00048963 | 0.020798 | 0.10779 | 0.093351 | -0.051341 | -0.073072 | 0.0018824 | -0.0066633 | -0.0069524 | 0.1898 | 0.021128 | -0.036776 | -0.0043526 | -0.080683 | -0.0054505 | -0.0097029 | -0.01114 | [7 empty] | -0.00076081 | 0.0011405 | 0.019808 | 0.021373
115 | MT 38 | R | 06/03/2023 03:08:04 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 38.117 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | -144.4965 | -47.6679 | 1.0819 | 14.2231 | 2846.886 | -147.2849 | 5.3362 | 0.0081775 | 0.020448 | 0.014544 | -0.32728 | 0.001601 | 0.11684 | 0.00099439 | 0.0005993 | 0.027401 | 0.01953 | -0.17024 | -0.11902 | 0.0014821 | -0.0066717 | -0.0070585 | 0.4243 | 0.014903 | -0.024163 | -0.010042 | 0.024851 | -0.0054505 | -0.0094473 | -0.011318 | [7 empty] | -0.00097828 | 0.00020832 | 0.019512 | 0.02359
116 | MT 39 | R | 06/03/2023 03:14:07 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 39.118 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | -134.7712 | -47.6299 | 1.6257 | 15.6344 | 2650.2677 | -157.7532 | 15.8529 | 0.0081898 | 0.01626 | 0.01702 | 0.64455 | 0.0051882 | 1.0537 | 0.0011666 | -0.0018028 | 0.02747 | 0.018787 | -0.091264 | -0.17617 | 0.0013288 | -0.0073028 | -0.0077536 | 0.45038 | 0.021419 | -0.028402 | -0.0041884 | -0.053277 | 0.01114 | -0.010167 | -0.011165 | [7 empty] | -0.0010944 | 0.00060832 | 0.024216 | 0.020295
117 | MT 40 | R | 06/03/2023 03:20:10 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 40.119 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | -130.1622 | -33.4273 | 1.6128 | 15.8166 | 2505.0298 | -145.3452 | 82.7629 | 0.0071706 | 0.011932 | 0.0088556 | -0.2576 | 0.036936 | 0.099912 | 0.00061681 | 0.0031326 | 0.043908 | 0.031512 | -0.18111 | -0.020935 | 0.0057309 | 0.0015706 | 0.0031237 | 0.081057 | 0.014903 | -0.030214 | -0.0070302 | -0.03282 | 0.00014006 | -0.0024888 | -0.0023504 | [7 empty] | -0.0010904 | 0.00034364 | 0.016941 | 0.019652
118 | MT 31 | R | 06/03/2023 03:26:12 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 31 .120 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 83.5412 | 466.9858 | 602.5057 | 3024.0423 | 4750.0796 | 754.125 | 1249.7838 | 0.73216 | 0.47644 | 0.43272 | 433.9076 | 90.3855 | 454.3924 | 0.58535 | 0.64747 | 54.4451 | 50.5312 | 162.0076 | 1.2794 | 1.1534 | 1.2055 | 1.4148 | 0.37912 | 0.3112 | 0.30343 | 0.3073 | 0.32009 | 0.26047 | 0.17255 | 0.18702 | [7 empty] | 1.7646 | 1.677 | 1.6107 | 1.6969
119 | MT 31 LD | R | 06/03/2023 03:32:23 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 31 LD.121 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 73.1438 | 388.7412 | 570.0477 | 3129.5429 | 4051.3574 | 715.5443 | 1194.2733 | 0.65451 | 0.44693 | 0.47966 | 445.8888 | 94.9684 | 436.3192 | 0.55864 | 0.65944 | 51.6342 | 51.9486 | 166.9469 | 1.1992 | 1.172 | 1.2218 | 1.4658 | 0.39034 | 0.34419 | 0.30823 | 0.31203 | 0.25357 | 0.38872 | 0.18854 | 0.19889 | [7 empty] | 1.7986 | 1.7487 | 1.6464 | 1.4671
120 | MT 31 LD | R | 06/03/2023 03:38:54 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 31 LD.122 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 67.038 | 375.452 | 562.7079 | 3054.865 | 4198.8439 | 765.9647 | 1181.7597 | 0.68464 | 0.47282 | 0.41395 | 389.7312 | 91.0084 | 449.3221 | 0.57222 | 0.61579 | 53.1815 | 45.4802 | 146.8858 | 1.2322 | 1.2036 | 1.1556 | 1.3834 | 0.21561 | 0.36417 | 0.25757 | 0.28355 | 0.20481 | 0.33233 | 0.15981 | 0.18743 | [7 empty] | 1.59 | 1.5282 | 1.4726 | 1.5959
121 | MT 32 | R | 06/03/2023 03:44:56 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 32.123 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | -140.5046 | -49.4851 | 1.2009 | 14.5063 | 1944.0591 | -160.3118 | 3.9832 | 0.0067412 | 0.021352 | 0.012143 | -0.35325 | 0.0035975 | 0.06037 | 0.00086552 | 0.0086911 | 0.054733 | 0.048041 | -0.1888 | -0.089208 | -0.0031069 | -0.0064072 | -0.0067431 | 0.27347 | 0.021699 | -0.032235 | -0.0079112 | 0.015592 | -0.0054505 | 0.00081964 | 0.00024197 | [7 empty] | -0.00095934 | 0.00077851 | 0.027026 | 0.025532
122 | MT 32 LSP | R | 06/03/2023 03:50:58 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 32 LSP.124 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 59198.2805 | 50868.7813 | 16.8765 | 20.9067 | 46657.696 | 46604.017 | 44772.8246 | 9.2949 | 8.9604 | 9.0711 | 11.2355 | 10.314 | 11.0842 | 8.797 | 12.9905 | 7.6546 | 7.6865 | 30.7115 | 7.4744 | 6.8332 | 6.3044 | 7.7194 | 8.3895 | 7.9213 | 7.9794 | 7.628 | 7.9449 | 7.5412 | 7.1191 | 7.3003 | [7 empty] | 6.2121 | 6.0604 | 4.5079 | 5.2764
123 | MT 32 LSP | R | 06/03/2023 03:57:29 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 32 LSP.125 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 56997.9224 | 49732.8893 | 15.7411 | 28.3522 | 46033.9925 | 46333.5436 | 42233.8606 | 9.1036 | 8.8745 | 8.2322 | 10.3291 | 10.0136 | 10.9836 | 8.749 | 12.7293 | 7.6139 | 7.1273 | 28.649 | 6.6981 | 6.9968 | 6.2945 | 7.5056 | 7.4153 | 7.6028 | 7.5414 | 7.4016 | 6.727 | 7.2523 | 6.5845 | 7.1769 | [7 empty] | 5.695 | 5.6849 | 4.0294 | 4.9599
124 | LBLANK | R | 06/03/2023 04:03:32 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LBLANK.126 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | -144.6105 | -48.6235 | -0.82122 | 0.42084 | 2439.7664 | -161.6167 | -0.8722 | 0.0071738 | 0.010563 | 0.0047714 | -0.48381 | -0.0025897 | 0.002076 | 0.0011311 | 0.0039121 | 0.00098345 | -0.0075572 | -0.22481 | -0.15217 | 0.0055406 | -0.0058034 | -0.0059339 | 0.42485 | 0.014903 | -0.034722 | -0.0042137 | -0.04304 | -0.0054505 | -0.0037191 | -0.011543 | [7 empty] | -0.0011022 | 0.00048136 | -0.0019962 | -0.00075634
125 | CCV 1ppb TRACE | R | 06/03/2023 04:09:35 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.127 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2169.5302 | 2077.1071 | 1.8108 | 1.6366 | 4667.7849 | 1628.9883 | 1680.7396 | 1.4254 | 1.4072 | 1.3982 | 0.99482 | 1.5712 | 1.4668 | 1.4546 | 1.4535 | 1.2922 | 1.2846 | 1.0523 | 1.0545 | 1.1446 | 0.95661 | 1.1729 | 1.4275 | 1.1772 | 1.1383 | 1.1683 | 1.1016 | 0.86819 | 1.0974 | 1.132 | [7 empty] | 0.96462 | 0.97355 | 0.75007 | 0.77849
126 | CCV 1ppb TRACE | R | 06/03/2023 04:16:07 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.128 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2136.6886 | 1976.8154 | 1.7489 | 1.5025 | 4491.1948 | 1681.6968 | 1675.864 | 1.4986 | 1.4337 | 1.4723 | 1.1102 | 1.5598 | 1.4883 | 1.4067 | 1.3914 | 1.2773 | 1.3201 | 1.1024 | 1.0608 | 1.102 | 0.97129 | 1.1798 | 1.4626 | 1.045 | 1.139 | 1.1908 | 1.0564 | 1.2602 | 1.1131 | 1.1286 | [7 empty] | 0.99318 | 1.0264 | 0.75077 | 0.79293
127 | LFB 10ppb | R | 06/03/2023 04:22:09 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LFB 10ppb.129 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 22092.3553 | 19429.9035 | 18.5431 | 21.0107 | 19544.9564 | 17317.922 | 16680.2598 | 14.0329 | 13.5186 | 13.3624 | 14.7285 | 14.641 | 14.6787 | 13.5133 | 13.7731 | 11.6876 | 11.7126 | 16.2393 | 11.5546 | 10.5749 | 9.6165 | 11.6384 | 12.4886 | 11.9852 | 11.7756 | 11.5678 | 11.6554 | 11.2988 | 10.8439 | 11.0488 | [7 empty] | 9.4908 | 9.4459 | 6.5651 | 7.3977
128 | LFB 10ppb | R | 06/03/2023 04:28:39 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LFB 10ppb.130 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 21552.9202 | 18945.8218 | 19.5101 | 21.6908 | 19307.5939 | 16896.9813 | 16158.8223 | 13.8152 | 13.0991 | 13.9098 | 15.1718 | 15.1977 | 14.3208 | 13.4938 | 13.7069 | 11.6415 | 11.8928 | 16.344 | 11.5375 | 11.1034 | 10.0911 | 12.0182 | 12.8541 | 11.374 | 12.2237 | 12.0905 | 11.7723 | 11.6705 | 11.1487 | 11.5015 | [7 empty] | 9.7282 | 9.5247 | 6.7343 | 7.2543
129 | CCV 10 ppb MAJORS | R | 06/03/2023 04:34:42 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.131 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 21093.3047 | 18897.1161 | 18.4399 | 20.276 | 19070.8214 | 16932.2362 | 15843.9077 | 13.5792 | 12.7624 | 14.4052 | 14.519 | 20.7953 | 13.133 | 13.0339 | 13.3477 | 11.3365 | 12.2772 | 12.5252 | 11.4074 | 10.3838 | 13.2541 | 16.3241 | 12.7961 | 16.5329 | 16.4322 | 15.7281 | 11.567 | 16.0732 | 11.3664 | 15.6033 | [7 empty] | 9.8015 | 9.7809 | 6.6986 | 7.4315
130 | CCV 10 ppb MAJORS | R | 06/03/2023 04:41:13 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.132 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 21761.6891 | 19380.299 | 18.1671 | 19.6343 | 19339.3999 | 17668.7693 | 16603.32 | 14.1108 | 13.3492 | 14.2487 | 14.9199 | 15.1752 | 13.8782 | 13.6736 | 13.8828 | 11.8883 | 12.3379 | 12.3305 | 11.4732 | 10.8783 | 9.7348 | 11.7961 | 12.6374 | 12.3362 | 11.7149 | 11.8037 | 11.4107 | 11.5964 | 11.3804 | 11.4001 | [7 empty] | 9.77 | 9.817 | 6.5307 | 7.9157
131 | MT 33 | R | 06/03/2023 04:47:16 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 33.133 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | -144.1848 | -48.5966 | 1.6614 | 12.8542 | 2663.8684 | -157.0125 | 9.2859 | 0.0066889 | 0.018093 | 0.011816 | -0.38767 | 0.001292 | 0.090013 | 0.0029399 | 0.021779 | 0.05714 | 0.055856 | -0.21232 | -0.13866 | 0.0056633 | -0.0071523 | -0.0076247 | 0.35949 | 0.014903 | -0.026322 | -0.0081015 | -0.0015029 | -0.0054505 | 0.0023444 | -0.0087058 | [7 empty] | -0.00097609 | 0.00076346 | 0.020306 | 0.020036
132 | MT 34 | R | 06/03/2023 04:53:19 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 34.134 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | -139.8359 | -47.7653 | 1.2166 | 15.18 | 2413.4705 | -153.9784 | 8.8892 | 0.0081279 | 0.027228 | 0.022024 | 3.547 | 0.023692 | 3.9087 | 0.0016985 | 0.029443 | 0.016106 | 0.0036195 | -0.19201 | -0.10293 | 0.0014472 | -0.0075884 | -0.0080568 | 0.34761 | 0.027839 | -0.034675 | -0.0081149 | -0.0068125 | -0.0054505 | -0.0067535 | -0.011695 | [7 empty] | -0.0010882 | 0.00092831 | 0.018388 | 0.018767
133 | MT 35 | R | 06/03/2023 04:59:21 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 35.135 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | -146.2957 | -49.4851 | 1.6412 | 15.3975 | 2722.4985 | -152.9544 | 5.7575 | 0.0067262 | 0.020743 | 0.015156 | 0.21904 | 0.0017489 | 0.7646 | 0.00046485 | 0.0082274 | 0.018128 | 0.0064002 | -0.15915 | -0.067794 | 0.010943 | -0.0068913 | -0.0072599 | 0.25266 | 0.014903 | -0.034567 | -0.0030236 | 0.0036586 | 0.00016959 | -0.0078448 | -0.011495 | [7 empty] | -0.00095241 | 0.00073003 | 0.018863 | 0.018729
134 | MT 36 | R | 06/03/2023 05:05:25 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 36.136 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | -143.2209 | -48.5857 | 1.3364 | 14.9673 | 2686.8262 | -152.2694 | 7.8563 | 0.0067211 | 0.029478 | 0.028491 | -0.016615 | 0.0038097 | 1.037 | 0.0011945 | 0.0038847 | 0.015518 | 0.0094586 | -0.16028 | -0.084038 | 0.0060886 | -0.006711 | -0.0069625 | 0.34936 | 0.021467 | -0.030134 | -0.0079939 | -0.045816 | 0.0056974 | -0.0082046 | -0.01081 | [7 empty] | -0.001084 | 0.00052797 | 0.019232 | 0.01992
135 | MT 37 | R | 06/03/2023 05:11:28 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 37.137 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | -142.3234 | -49.4851 | 1.688 | 18.6929 | 2664.7185 | -147.7208 | 24.8616 | 0.0072205 | 0.015679 | 0.0098111 | -0.36234 | 0.0088126 | 0.18426 | 8.152e-05 | 0.011156 | 0.10984 | 0.083711 | -0.065297 | -0.11469 | 0.0015343 | -0.0073583 | -0.0077437 | 0.36634 | 0.035954 | -0.036586 | -0.0088214 | -0.032594 | -0.0054505 | -0.0096882 | -0.011173 | [7 empty] | -0.0012092 | 0.001004 | 0.023374 | 0.020876
136 | MT 38 | R | 06/03/2023 05:17:31 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 38.138 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | -145.9849 | -46.7705 | 1.5713 | 16.214 | 2981.6137 | -151.3022 | 1.3542 | 0.0071949 | 0.022751 | 0.014046 | -0.32481 | 0.0020243 | 0.19475 | 0.0021002 | 0.0097205 | 0.02356 | 0.017619 | -0.17052 | -0.17556 | 0.0060034 | -0.0079559 | -0.008591 | 0.31306 | 0.014903 | -0.032497 | -0.009034 | -0.069323 | 2.4275e-05 | -0.0097384 | -0.011372 | [7 empty] | -0.00083264 | 0.00067931 | 0.018568 | 0.019137
137 | MT 39 | R | 06/03/2023 05:23:33 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 39.139 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | -134.5654 | -47.5824 | 1.8089 | 15.5881 | 2895.9019 | -150.5773 | 10.404 | 0.0072502 | 0.021796 | 0.018083 | 0.75999 | 0.0059778 | 1.1598 | 0.0018257 | 0.019226 | 0.030593 | 0.020001 | -0.088074 | -0.1151 | 0.0016566 | -0.0065536 | -0.0068192 | 0.28764 | 0.023236 | -0.031889 | -0.004136 | 0.066347 | -0.0054505 | -0.0098128 | -0.011514 | [7 empty] | -0.00082737 | 0.00049034 | 0.019165 | 0.01946
138 | MT 40 | R | 06/03/2023 05:29:36 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 40.140 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | -131.7014 | -30.8969 | 1.7648 | 16.0031 | 3040.3577 | -152.8283 | 68.5884 | 0.0072661 | 0.023501 | 0.013608 | -0.23487 | 0.035183 | 0.35471 | 0.0010606 | 0.0039844 | 0.042751 | 0.032599 | -0.18283 | -0.11061 | 0.0070286 | 0.0011609 | 0.0022731 | 0.21148 | 0.021412 | -0.038868 | -0.0080261 | -0.024686 | 0.0057302 | -0.0015225 | -0.0034334 | [7 empty] | -0.00095047 | 0.00036892 | 0.018976 | 0.019604
139 | MT 32 B | R | 06/03/2023 05:35:38 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 32 B.141 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | -145.0978 | -48.5802 | 0.86127 | 13.7504 | 2911.3754 | -150.9518 | -2.7367 | 0.0072459 | 0.02172 | 0.014899 | -0.32135 | 0.00077306 | 0.17932 | 0.00083963 | 0.025931 | 0.049526 | 0.040111 | -0.18887 | -0.091773 | 0.0014342 | -0.0068012 | -0.0070993 | 0.26294 | 0.014903 | -0.038868 | -0.0058054 | -0.0027996 | -0.0054505 | -0.007516 | -0.0098047 | [7 empty] | -0.00096132 | 0.00089346 | 0.024929 | 0.026179
140 | MT 32 B LD | R | 06/03/2023 05:41:40 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 32 B LD.142 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | -143.7237 | -48.5325 | 0.81155 | 15.4575 | 2910.3949 | -139.9721 | 8.4222 | 0.0067318 | 0.017968 | 0.013469 | -0.40689 | -0.0012858 | 0.076383 | 0.00048013 | 0.0036544 | 0.068218 | 0.049683 | -0.11571 | -0.063588 | 0.0017472 | -0.007181 | -0.0075861 | 0.18421 | 0.014903 | -0.036635 | -0.008893 | -0.016383 | 0.00031832 | -0.0093333 | -0.0092452 | [7 empty] | -0.00044264 | 0.00060491 | 0.029564 | 0.02875
141 | LBLANK | R | 06/03/2023 05:47:43 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LBLANK.143 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | -147.4082 | -49.4851 | -0.88719 | 1.3695 | 3009.3505 | -153.8548 | -2.6267 | 0.0078023 | 0.011657 | 0.003746 | -0.47825 | -0.0019548 | 0.066039 | 0.00028853 | 0.0022234 | -0.0017249 | -0.0092517 | -0.22657 | -0.043734 | -0.0031069 | -0.0079101 | -0.0084986 | 0.079404 | 0.028747 | -0.034374 | -0.010947 | 0.005303 | 0.00040536 | -0.011249 | -0.012048 | [7 empty] | -0.00095466 | 0.00014374 | -0.0022795 | -0.00079824
142 | CCV 1ppb TRACE | R | 06/03/2023 05:53:46 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.144 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2205.8349 | 2087.0744 | 1.5678 | 2.1631 | 4744.8351 | 1625.2449 | 1926.3046 | 1.4815 | 1.4635 | 1.4636 | 1.1267 | 1.6091 | 1.5448 | 1.4574 | 1.5002 | 1.325 | 1.3135 | 1.0794 | 1.115 | 0.99725 | 1.0021 | 1.2258 | 1.5632 | 1.2048 | 1.1649 | 1.1892 | 1.0693 | 1.1787 | 1.0959 | 1.1221 | [7 empty] | 0.99635 | 0.98862 | 0.77162 | 0.73371
143 | CCV 1ppb TRACE | R | 06/03/2023 06:00:18 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.145 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2164.876 | 2043.9683 | 1.7483 | 1.7014 | 4810.6832 | 1730.5175 | 1740.5768 | 1.4799 | 1.4132 | 1.4158 | 1.0982 | 1.6222 | 1.5524 | 1.4099 | 1.4921 | 1.2996 | 1.2976 | 1.0626 | 1.0805 | 1.1286 | 1.0301 | 1.2456 | 1.5338 | 1.2461 | 1.22 | 1.1741 | 1.0908 | 1.1735 | 1.0968 | 1.1693 | [7 empty] | 0.9788 | 0.96602 | 0.74461 | 0.78011
144 | CCV 10 ppb MAJORS | R | 06/03/2023 06:06:21 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.146 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 23421.8549 | 21366.2416 | 19.7732 | 20.4635 | 20739.4614 | 18765.838 | 17788.2699 | 14.8099 | 14.3279 | 14.0774 | 14.3105 | 15.2628 | 14.539 | 14.6014 | 14.6472 | 12.3025 | 11.9855 | 12.3486 | 11.8731 | 11.5123 | 9.8349 | 12.0184 | 12.9632 | 11.7748 | 12.1936 | 11.8886 | 11.966 | 11.3574 | 10.997 | 11.4155 | [7 empty] | 9.6105 | 9.5282 | 6.9007 | 7.4512
145 | CCV 10 ppb MAJORS | R | 06/03/2023 06:12:53 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.147 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 22335.7108 | 20007.3012 | 19.9551 | 19.6122 | 20925.1471 | 17941.292 | 17143.2219 | 14.2424 | 13.8579 | 14.0292 | 14.0306 | 15.5376 | 14.3482 | 13.928 | 14.1099 | 12.1434 | 12.1387 | 12.3662 | 11.3409 | 11.0035 | 10.1228 | 11.8766 | 12.6111 | 12.3797 | 12.3669 | 11.7063 | 11.0055 | 11.4094 | 11.055 | 11.3472 | [7 empty] | 9.736 | 9.5107 | 6.6961 | 7.4202
146 | MT 33 B | R | 06/03/2023 06:18:55 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 33 B.148 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | -145.8824 | -46.6451 | 1.0636 | 13.0248 | 2594.679 | -160.7379 | 8.3126 | 0.0087229 | 0.019625 | 0.012539 | -0.4581 | 0.00046016 | 0.12024 | 0.0021744 | 0.015497 | 0.043786 | 0.042829 | -0.2049 | -0.054838 | 0.0063233 | -0.0078587 | -0.0084222 | 0.20159 | 0.021652 | -0.032296 | -0.01197 | -0.021156 | -0.0054505 | -0.0065505 | -0.0081916 | [7 empty] | -0.00047903 | 0.00065078 | 0.019249 | 0.020937
147 | MT 33 B LSP | R | 06/03/2023 06:24:56 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 33 B LSP.149 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 60037.5792 | 52232.7355 | 50.6883 | 64.7614 | 48452.8172 | 48010.6355 | 46679.3948 | 38.0646 | 35.8324 | 32.8088 | 34.5058 | 38.9653 | 36.9867 | 36.3855 | 35.9194 | 30.4122 | 26.7644 | 26.9735 | 28.0795 | 27.3384 | 24.0109 | 28.8613 | 29.9328 | 28.8019 | 28.5682 | 27.7929 | 27.2441 | 27.3911 | 25.922 | 29.6321 | [7 empty] | 21.218 | 20.9679 | 14.6263 | 17.9222
148 | MT 34 B | R | 06/03/2023 06:30:59 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 34 B.150 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | -144.091 | -46.7646 | 0.49718 | 12.6368 | 2407.141 | -141.8676 | 7.526 | 0.0096642 | 0.021654 | 0.019091 | -0.33959 | -0.0015364 | 0.14969 | 0.0017216 | 0.048942 | 0.024742 | 0.022368 | -0.18031 | -0.091465 | 0.019522 | -0.0038214 | -0.0034514 | 0.42889 | 0.028103 | -0.030256 | -0.0070095 | -0.010672 | -0.0054505 | 0.023746 | 0.011118 | [7 empty] | -0.00060798 | 0.0014914 | 0.0201 | 0.02244
149 | MT 35 B | R | 06/03/2023 06:37:01 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 35 B.151 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | -142.3015 | -48.6271 | 1.7461 | 13.149 | 2527.6116 | -154.5867 | 11.4399 | 0.0081263 | 0.024197 | 0.013701 | -0.44083 | -0.0016059 | 0.070743 | 0.00083997 | 0.010506 | 0.014452 | 0.011512 | -0.21554 | -0.12167 | 0.0011785 | -0.0066022 | -0.0068769 | 0.27316 | 0.0215 | -0.030337 | -0.0060652 | -9.4583e-05 | -0.0054505 | -0.0053959 | -0.010558 | [7 empty] | -0.00073469 | 0.00042543 | 0.018076 | 0.018867
150 | MT 36 B | R | 06/03/2023 06:43:03 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 36 B.152 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | -143.1065 | -49.4851 | 1.645 | 12.3528 | 2812.1803 | -150.0478 | -0.4303 | 0.0082724 | 0.036789 | 0.029102 | -0.089439 | 0.0016045 | 0.43029 | 0.0013992 | 0.012072 | 0.022717 | 0.01095 | 4.1816 | -0.12225 | 0.0064338 | -0.0072011 | -0.0075538 | 0.25412 | 0.014903 | -0.038868 | -0.0080107 | 0.028903 | 0.00015418 | 0.049177 | 0.048569 | [7 empty] | -0.0009633 | 0.00054837 | 0.018622 | 0.019849
151 | MT 37 B | R | 06/03/2023 06:49:06 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 37 B.153 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | -129.8997 | -48.5469 | 0.90672 | 13.896 | 2647.189 | -158.1526 | 7.9558 | 0.0077171 | 0.017862 | 0.016604 | -0.44311 | 1.0653e-05 | -0.004652 | 0.0010303 | 0.011807 | 0.025949 | 0.017099 | -0.19869 | -0.11583 | 0.0016011 | -0.0081821 | -0.0088374 | 0.29819 | 0.021603 | -0.028159 | -0.010967 | -0.043493 | 8.3653e-05 | -0.0087965 | -0.010645 | [7 empty] | -0.00068102 | 0.00040154 | 0.018762 | 0.01959
152 | MT 38 B | R | 06/03/2023 06:55:09 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 38 B.154 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | -115.919 | -48.5439 | 0.90711 | 13.416 | 2514.8662 | -149.9235 | 5.6562 | 0.0076907 | 0.020634 | 0.015132 | -0.43008 | -0.0019892 | 0.0045858 | 0.001577 | 0.010993 | 0.014237 | 0.0080447 | -0.19643 | -0.082601 | -0.0031069 | -0.0082372 | -0.0088143 | 0.23274 | 0.021621 | -0.034424 | -0.010927 | -0.015497 | 0.0060749 | -0.0050157 | -0.0061113 | [7 empty] | -0.00087741 | 0.0006348 | 0.019958 | 0.019269
153 | MT 39 B | R | 06/03/2023 07:01:12 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 39 B.155 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | -144.1875 | -47.5551 | 0.79419 | 13.132 | 2808.6213 | -149.7511 | 12.5645 | 0.0088225 | 0.024779 | 0.015327 | -0.35789 | -0.0015255 | 0.15413 | 0.00028775 | 0.013561 | 0.019459 | 0.007667 | -0.19099 | -0.092568 | 0.0063991 | -0.0083597 | -0.0089997 | 0.2994 | 0.021525 | -0.0324 | -0.0089705 | -0.036218 | 0.00015208 | 0.0026576 | 0.0018925 | [7 empty] | -0.00058069 | 0.00033827 | 0.019918 | 0.019003
154 | MT 40 B | R | 06/03/2023 07:07:15 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 40 B.156 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | -145.849 | -47.5025 | 1.1167 | 12.9709 | 2811.8769 | -157.9235 | 8.345 | 0.0078325 | 0.02096 | 0.011695 | -0.40188 | -0.0017854 | 0.07357 | 0.0033634 | 0.010359 | 0.049049 | 0.03284 | -0.19278 | -0.088147 | 0.0018642 | -0.0065638 | -0.0068346 | 0.19634 | 0.021811 | -0.032141 | -0.0057686 | -0.051613 | 0.00034068 | -0.0099003 | -0.010252 | [7 empty] | -0.00084652 | 0.00035425 | 0.020678 | 0.018126
155 | MT 32 C | R | 06/03/2023 07:13:17 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 32 C.157 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | -146.6446 | -48.5426 | 0.8692 | 13.2677 | 2576.878 | -154.4551 | 1.6925 | 0.0082932 | 0.020278 | 0.014372 | -0.43667 | 8.9921e-05 | 0.01308 | 0.00066745 | 0.0097843 | 0.047896 | 0.041455 | -0.1611 | -0.10618 | 0.0064685 | -0.0076678 | -0.0081116 | 0.15614 | 0.014903 | -0.034529 | -0.0080491 | -0.015383 | -0.0054505 | -0.0084536 | -0.010598 | [7 empty] | -0.00068688 | 0.00062077 | 0.025313 | 0.025226
156 | LBLANK | R | 06/03/2023 07:19:20 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LBLANK.158 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | -147.857 | -49.4851 | -0.82108 | 1.2593 | 2577.0718 | -148.8617 | -4.861 | 0.0082894 | 0.0083397 | 0.0031559 | -0.49735 | -0.0030791 | -0.066008 | 0.00046874 | 0.0023377 | 0.00014746 | -0.0088015 | -0.22328 | -0.12393 | 0.0014155 | -0.0081964 | -0.0087918 | 0.18021 | 0.014903 | -0.03214 | -0.0088262 | -0.06201 | 0.0064591 | -0.010277 | -0.012816 | [7 empty] | -0.001204 | 0.00039887 | -0.0021029 | -0.00087167
157 | CCV 1ppb TRACE | R | 06/03/2023 07:25:23 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.159 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2228.141 | 2157.1104 | 1.5024 | 2.0579 | 4498.6333 | 1689.8757 | 1771.6199 | 1.4935 | 1.431 | 1.4461 | 1.0944 | 1.5717 | 1.5615 | 1.4555 | 1.4413 | 1.3101 | 1.307 | 1.0783 | 1.068 | 1.1073 | 1.0091 | 1.202 | 1.561 | 1.1841 | 1.2241 | 1.156 | 1.0996 | 1.1393 | 1.0875 | 1.1248 | [7 empty] | 0.97166 | 0.98536 | 0.77275 | 0.79166
158 | CCV 1ppb TRACE | R | 06/03/2023 07:31:56 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 1ppb TRACE.160 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 2259.4002 | 2060.9943 | 1.9346 | 1.6836 | 5002.6423 | 1707.8261 | 1797.7708 | 1.5021 | 1.4091 | 1.6226 | 1.287 | 1.5681 | 1.533 | 1.4472 | 1.5517 | 1.3739 | 1.4609 | 1.2397 | 1.0986 | 1.1559 | 0.9536 | 1.1959 | 1.2941 | 1.2704 | 1.1499 | 1.1814 | 1.1382 | 1.226 | 1.2247 | 1.1593 | [7 empty] | 1.1178 | 1.1136 | 0.75428 | 0.78438
159 | LFB 10ppb | R | 06/03/2023 07:37:58 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LFB 10ppb.161 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 22008.8046 | 19739.8604 | 20.3066 | 21.6092 | 19930.9088 | 17005.7439 | 16892.2286 | 14.1463 | 13.5856 | 13.7366 | 14.7745 | 14.3862 | 14.7431 | 13.7783 | 14.0832 | 11.8005 | 11.7028 | 16.2186 | 11.447 | 10.8957 | 8.6279 | 11.0052 | 12.7434 | 10.786 | 11.4947 | 11.0227 | 11.1616 | 11.435 | 10.8083 | 10.5312 | [7 empty] | 9.5094 | 9.2732 | 6.4994 | 7.2632
160 | LFB 10ppb | R | 06/03/2023 07:44:29 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\LFB 10ppb.162 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 19528.3549 | 17890.3419 | 16.5575 | 20.7182 | 18187.6028 | 15780.3209 | 15180.9207 | 12.8424 | 12.4285 | 0.67159 | 0.23309 | 14.8657 | 13.4224 | 12.3185 | 12.5409 | 10.5825 | 0.67244 | 0.54305 | 0.62992 | 10.071 | 9.4619 | 11.8757 | -1.5593 | 11.5906 | 12.5208 | 12.0519 | -0.14536 | 11.4872 | 0.59682 | 11.1602 | [7 empty] | 0.48101 | 0.46492 | -0.00044911 | -5.4767e-05
161 | CCV 10 ppb MAJORS | R | 06/03/2023 07:50:31 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.163 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 22358.4488 | 20023.5587 | 18.7649 | 21.1372 | 20285.393 | 17789.0904 | 16721.3055 | 14.3841 | 13.8884 | 14.4481 | 14.6342 | 15.4213 | 14.215 | 14.0465 | 13.8967 | 12.0496 | 12.2425 | 12.5482 | 13.6506 | 10.808 | 9.6335 | 12.0002 | 15.1006 | 12.1458 | 12.0296 | 11.7464 | 13.5256 | 11.2244 | 11.2821 | 11.4594 | [7 empty] | 9.7471 | 9.6685 | 7.8487 | 8.1409
162 | CCV 10 ppb MAJORS | R | 06/03/2023 07:57:03 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\CCV 10 ppb MAJORS.164 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 22680.5735 | 20310.9018 | 18.1759 | 20.0575 | 20410.5381 | 18361.8529 | 17484.8905 | 14.8212 | 13.8826 | 14.2275 | 14.4131 | 13.3342 | 14.4153 | 14.1999 | 14.2842 | 12.3249 | 12.066 | 12.2722 | 11.0053 | 11.796 | 8.5796 | 10.4726 | 11.8505 | 10.6998 | 10.5519 | 10.2262 | 10.8794 | 10.3456 | 11.1253 | 10.0612 | [7 empty] | 9.6075 | 9.4285 | 6.2012 | 8.1981
163 | MT 32 C LD | R | 06/03/2023 08:03:06 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 32 C LD.165 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | -140.7397 | -48.4932 | 0.74952 | 11.7177 | 2728.4124 | -141.8514 | 4.2651 | 0.0083861 | 0.021652 | 0.012689 | -0.42061 | -0.001051 | 0.16325 | 0.00087204 | 0.012975 | 0.05466 | 0.040552 | -0.15442 | -0.043464 | 0.0018705 | -0.0061107 | -0.0062816 | 0.2069 | 0.014903 | -0.029407 | -0.0075113 | -0.027017 | 0.00030781 | -0.0036483 | -0.0086778 | [7 empty] | -0.00083449 | 0.0011359 | 0.025057 | 0.025618
164 | MT 33 C | R | 06/03/2023 08:09:08 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 33 C.166 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | -145.7086 | -49.4851 | 0.88039 | 12.4185 | 2886.3951 | -156.478 | 4.1413 | 0.0073074 | 0.020773 | 0.012549 | -0.40958 | 0.0003478 | 0.1611 | 0.0010916 | 0.023486 | 0.050381 | 0.040755 | -0.17665 | -0.17692 | -0.0031069 | -0.0083293 | -0.0089575 | 0.37803 | 0.021753 | -0.036644 | -0.0088826 | -0.043529 | -0.0054505 | -0.0077281 | -0.0098835 | [7 empty] | -0.00056455 | 0.00056094 | 0.022383 | 0.019597
165 | MT 33 C LSP | R | 06/03/2023 08:15:09 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 33 C LSP.167 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 59050.3916 | 51087.8021 | 51.9107 | 66.4428 | 46445.6396 | 46122.6891 | 45574.0222 | 38.0361 | 35.6948 | 28.0196 | 27.3774 | 38.1977 | 36.282 | 35.5592 | 35.5266 | 29.7314 | 22.5045 | 25.8482 | 31.8323 | 27.5463 | 23.1578 | 27.9778 | 34.0757 | 28.6826 | 28.2372 | 26.906 | 30.253 | 26.7324 | 0.039944 | 28.9852 | [7 empty] | 26.1806 | 26.3261 | 12.2191 | 31.4701
166 | MT 33 C LSP | R | 06/03/2023 08:21:40 | Failed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 33 C LSP.168 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | 58565.1544 | 51486.5813 | 49.8975 | 63.9704 | 46749.5255 | 47082.3799 | 45081.2455 | 38.2164 | 35.7307 | 36.7121 | 38.9269 | 38.9595 | 36.6258 | 36.289 | 35.8593 | 30.2663 | 30.2764 | 30.1235 | 28.8949 | 28.3867 | 23.0697 | 28.5605 | 31.1265 | 28.7723 | 28.654 | 28.0977 | 28.1884 | 26.8109 | 28.9871 | 29.5395 | [7 empty] | 23.4326 | 23.3502 | 15.0327 | 17.9537
167 | MT 34 C | R | 06/03/2023 08:27:42 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 34 C.169 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | -143.464 | -47.5969 | 1.1387 | 12.1541 | 2681.8217 | -154.6356 | -0.54156 | 0.0082801 | 0.024895 | 0.019783 | -0.32762 | -0.00016841 | 0.22723 | 0.0031461 | 0.066185 | 0.015852 | 0.0072428 | -0.20826 | -0.18878 | 0.011085 | -0.0045896 | -0.0044872 | 0.53756 | 0.021887 | -0.032257 | -0.010015 | -0.039947 | -0.0054505 | 0.023349 | 0.012708 | [7 empty] | -0.0010872 | 0.00034451 | 0.020497 | 0.017352
168 | MT 35 C | R | 06/03/2023 08:33:45 | Passed | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\DataSet\230601_USGS_MethodTest\MT 35 C.170 | C:\Users\Public\Documents\PerkinElmer Syngistix\ICPMS\Method\USGS Methods\230601_USGS_MTEST_5%.mth | -142.6634 | -47.621 | 0.84783 | 12.634 | 2618.4619 | -149.882 | 3.9574 | 0.0077539 | 0.014176 | 0.012988 | -0.43659 | -0.0016743 | 0.10955 | 0.0010313 | 0.014336 | 0.016579 | 0.011307 | -0.19544 | -0.13235 | -0.0031069 | -0.0074028 | -0.0078824 | 0.358 | 0.014903 | -0.030298 | -0.01197 | -0.023896 | -0.00037493 | -0.0031493 | -0.011013 | [7 empty] | -0.00083564 | 0.00064925 | 0.017473 | 0.018678

## Internal Standards
Row Index Number | Rh (IS1) 103/103 | Rh (IS2) 103/103 | Rh (IS3) 103/103 | Rh (IS4) 103/103 | Rh (IS5) 103/103 | Rh (IS6) 103/103 | Rh OShift (IS1) 103/119
1 | 100% | 100% | 100% | 100% | 100% | 100% | 100%
2 | 99.88% | 98.17% | 99.42% | 98.02% | 99.55% | 99.39% | 98.22%
3 | 100.3% | 95.71% | 95.83% | 97.53% | 95.26% | 96.53% | 95.79%
4 | 100.5% | 98.82% | 100.7% | 97.75% | 103% | 97.8% | 92.3%
5 | 105.4% | 105% | 100% | 99.91% | 98.85% | 88.43% | 101.9%
6 | 110% | 102.4% | 101.2% | 102.2% | 85.22% | 103.1% | 99.1%
7 | 114.1% | 99.78% | 101.7% | 103.8% | 102.6% | 103.9% | 95.52%
8 | 111% | 94.15% | 100.7% | 99.1% | 99.7% | 97.83% | 88.97%
9 | 102.1% | 79.75% | 83.24% | 86.73% | 97.87% | 83.98% | 77%
10 | 106.2% | 70.89% | 79.92% | 80.11% | 82.27% | 74.38% | 72.69%
11 | 92.58% | 69.83% | 78.64% | 79.66% | 82.73% | 73.46% | 72.99%
13 | 111.9% | 101.1% | 108.8% | 117.2% | 118.5% | 98.82% | 109.1%
14 | 123.4% | 99.04% | 110.5% | 107.5% | 110.8% | 99.44% | 105.4%
15 | 123.2% | 109.7% | 120.9% | 114.8% | 115.8% | 97.9% | 121.7%
16 | 117.4% | 105.2% | 122.8% | 109.6% | 111% | 104.8% | 119.8%
17 | 113.1% | 96.53% | 112.7% | 100.5% | 102.6% | 97.42% | 107.8%
18 | 112.1% | 96.18% | 111.4% | 98.79% | 100.8% | 96.04% | 105.8%
19 | 111.5% | 103.6% | 118.7% | 106.5% | 102.8% | 101.6% | 118.4%
20 | 111.4% | 94.39% | 108.3% | 96.93% | 102.5% | 94.97% | 104.1%
21 | 111.1% | 91.07% | 111.4% | 97.56% | 98.76% | 99.55% | 103.9%
22 | 120.2% | 103.2% | 117.8% | 110.5% | 108.8% | 100.8% | 118.5%
23 | 110.8% | 102.3% | 119.6% | 104.2% | 106.6% | 102% | 120.6%
24 | 101.5% | 90.7% | 101.1% | 88.82% | 91.62% | 87.03% | 105.3%
25 | 106.4% | 98.46% | 111.5% | 99.27% | 102.9% | 87.68% | 115%
26 | 98.44% | 97.08% | 109.9% | 93.43% | 100.6% | 88.32% | 109.6%
27 | 104.4% | 98.65% | 112.4% | 97.73% | 101.1% | 95.94% | 117.3%
28 | 102.2% | 94.28% | 108.4% | 94.49% | 98.58% | 93.97% | 107.9%
29 | 105.9% | 92.76% | 110.5% | 93.25% | 97.34% | 91.86% | 105.7%
30 | 102.1% | 84.96% | 101.2% | 88.66% | 92.57% | 84.94% | 98.43%
31 | 103% | 88.1% | 103.2% | 88.6% | 93.47% | 91.63% | 101.9%
32 | 103% | 91.24% | 113.1% | 95.77% | 98.08% | 94.47% | 107.2%
33 | 102.5% | 93.99% | 110.3% | 93.56% | 97.03% | 91.93% | 108.3%
34 | 102.3% | 88.44% | 105.9% | 87.96% | 92.77% | 88.08% | 101.8%
35 | 101.1% | 87.41% | 104.4% | 88.41% | 88.42% | 85.66% | 101.6%
36 | 100.1% | 83.62% | 107.2% | 92.09% | 95% | 90.82% | 95.62%
37 | 97.82% | 90.03% | 101.9% | 90.08% | 94.59% | 88.04% | 106.3%
38 | 97.04% | 88.5% | 104.4% | 88.15% | 93.28% | 87.34% | 101.6%
39 | 91.73% | 80.31% | 91.17% | 80.69% | 83.07% | 78.77% | 90.11%
40 | 94.12% | 80.04% | 96.3% | 83.2% | 86.26% | 81.27% | 92%
41 | 99.83% | 94.09% | 110% | 93.25% | 94.5% | 90.87% | 109%
42 | 98.19% | 89.72% | 103.4% | 90.12% | 93.81% | 88.11% | 104.9%
43 | 99.33% | 89.97% | 104.6% | 88.3% | 93.57% | 95.7% | 105.9%
44 | 94.2% | 86.26% | 103.7% | 87.38% | 90.89% | 87.36% | 101.7%
45 | 94.19% | 87.53% | 104.8% | 87.61% | 91.75% | 87.12% | 102.1%
46 | 94.32% | 80.62% | 96.31% | 81.6% | 84.73% | 79.77% | 93.33%
47 | 95.17% | 81.16% | 99.65% | 81.42% | 81.38% | 81.55% | 94.74%
48 | 96.07% | 89.31% | 102.8% | 87.6% | 90.9% | 84.45% | 107.4%
49 | 93.37% | 84.1% | 98.96% | 84.57% | 86.13% | 84.71% | 98.05%
50 | 94.93% | 83.8% | 99.07% | 85.33% | 88.67% | 82.6% | 98.71%
51 | 94.14% | 84.6% | 97.74% | 85.84% | 89.38% | 82.91% | 95.88%
52 | 88.84% | 83.4% | 99.46% | 82.93% | 87.81% | 80.55% | 96.8%
53 | 90.74% | 80.72% | 97.23% | 83.04% | 86.91% | 80.64% | 92.68%
54 | 89.06% | 79.53% | 95.37% | 80.34% | 83.98% | 79.08% | 90.64%
55 | 88.67% | 80.54% | 94.54% | 80.64% | 84.36% | 81.56% | 90.09%
56 | 90.73% | 80.42% | 90.99% | 83.04% | 87.94% | 79.26% | 92.66%
57 | 88.36% | 79.89% | 94.62% | 78.96% | 84.64% | 78.34% | 94.53%
58 | 84.6% | 74.1% | 87.53% | 74.45% | 78.26% | 76.26% | 87.9%
59 | 98.95% | 85.36% | 104.6% | 87.03% | 93.37% | 83.76% | 101.9%
60 | 93.52% | 84.19% | 99.78% | 83.6% | 88.21% | 83.04% | 99.49%
61 | 90.5% | 81.75% | 101.3% | 80.57% | 86.68% | 80.87% | 98.24%
62 | 91.41% | 82.7% | 92.07% | 75.91% | 80.66% | 76.02% | 98.71%
63 | 87.21% | 75.19% | 92.66% | 77.19% | 81.13% | 75.16% | 87.47%
64 | 78.69% | 76.51% | 93.71% | 79.29% | 80.66% | 74.59% | 90.09%
65 | 87.44% | 75.8% | 95.11% | 74.16% | 80.37% | 82.95% | 90.84%
66 | 82.49% | 77.12% | 90.35% | 77.39% | 82.22% | 73.12% | 89.77%
67 | 85.79% | 75.97% | 90.6% | 77.42% | 80.63% | 73.74% | 89.27%
68 | 84.13% | 75.77% | 89.33% | 76.92% | 80.78% | 74.88% | 89.1%
69 | 78.27% | 78.76% | 94.92% | 79.14% | 84.23% | 79.1% | 93.82%
70 | 82.19% | 72.19% | 88.36% | 76.58% | 80.47% | 77.42% | 83.37%
71 | 81.35% | 75.75% | 87.51% | 75.93% | 79.75% | 73.85% | 86.87%
72 | 83.77% | 74.99% | 88.16% | 76.25% | 82.36% | 73.95% | 91.01%
73 | 82.37% | 75.26% | 85.07% | 74.03% | 79.43% | 72.74% | 86.89%
74 | 86.08% | 77.21% | 86.72% | 75.52% | 84.58% | 68.91% | 91.35%
75 | 98.13% | 87.83% | 98.07% | 92.42% | 95.67% | 86.51% | 96.61%
76 | 83.05% | 77.62% | 90.89% | 77.65% | 82.47% | 74.71% | 89.92%
77 | 83.82% | 74.29% | 85.81% | 78.61% | 81.3% | 68.94% | 85.39%
78 | 85.2% | 73.52% | 86.02% | 75.57% | 82.58% | 71.63% | 85.58%
79 | 81.94% | 69.55% | 80.87% | 71.19% | 74.5% | 68.15% | 82.08%
80 | 82.91% | 69.55% | 82.23% | 70.88% | 76.78% | 70.71% | 81.8%
81 | 84.62% | 74.16% | 86.67% | 75.85% | 81.5% | 71.95% | 83.91%
82 | 77.2% | 66.06% | 76.72% | 68.93% | 72.14% | 28.74% | 77.38%
83 | 76.92% | 67.06% | 85.95% | 69.39% | 73.95% | 66.36% | 77.71%
84 | 81.67% | 76.92% | 86.97% | 75.88% | 90.86% | 73.65% | 91.63%
85 | 84.56% | 75% | 88.1% | 77.43% | 84.13% | 77.18% | 87.81%
86 | 78.45% | 74.42% | 87% | 76.11% | 82.48% | 72.87% | 86.03%
87 | 70.99% | 74.05% | 87.36% | 75.42% | 96.73% | 72.43% | 84.06%
88 | 85.13% | 74.03% | 85.04% | 75.55% | 82.16% | 72.67% | 85.32%
89 | 77.61% | 75.05% | 81.58% | 74.1% | 80.03% | 72.15% | 84.7%
90 | 78.6% | 72.28% | 81.11% | 71.42% | 84.91% | 74.36% | 81.08%
91 | 83.34% | 73.5% | 84.73% | 75.99% | 85.38% | 73.36% | 86.5%
92 | 83.65% | 70.54% | 84.57% | 73.28% | 80.18% | 70.99% | 83.35%
93 | 83.2% | 72.13% | 84.04% | 75.98% | 79.78% | 71.57% | 82.98%
94 | 82.05% | 72.71% | 74.97% | 68.56% | 73.14% | 66.51% | 83.18%
95 | 81.02% | 68.65% | 78.53% | 71.46% | 73.5% | 68.01% | 75.24%
96 | 82.47% | 67.23% | 78.38% | 69.01% | 74.26% | 68.27% | 78.97%
97 | 77.69% | 69.18% | 81% | 70.67% | 74.46% | 67.89% | 77.06%
98 | 87.88% | 74.18% | 83.86% | 68.42% | 79.2% | 74.7% | 84.66%
99 | 85.8% | 75.76% | 84.55% | 76.69% | 79.83% | 71.05% | 86.61%
100 | 84.7% | 72.32% | 82.51% | 73.06% | 78.81% | 70.96% | 83.14%
101 | 85.63% | 72.49% | 84.83% | 74.25% | 73.35% | 69.86% | 83.69%
102 | 81.74% | 72.34% | 83.23% | 74.57% | 74.51% | 67.67% | 82.3%
103 | 73.8% | 63.4% | 73.33% | 63.93% | 68.35% | 58.53% | 73.31%
104 | 77.56% | 64.43% | 74.04% | 64.85% | 69.45% | 62.64% | 71.68%
105 | 85.08% | 74.42% | 87.06% | 74.99% | 69.21% | 71.93% | 87.88%
106 | 82.42% | 73.98% | 85.47% | 76.28% | 81.22% | 72.55% | 84.23%
107 | 81.79% | 71.18% | 85.94% | 78.51% | 81.26% | 71.47% | 80.43%
108 | 84.41% | 72.5% | 84.44% | 73.52% | 74.9% | 71.65% | 82.62%
109 | 83.08% | 70.54% | 84% | 70.95% | 81.32% | 69.88% | 82.79%
110 | 83.58% | 71.46% | 81.96% | 70.11% | 76.8% | 71.53% | 81.8%
111 | 82.56% | 68.5% | 75.92% | 67.74% | 71.86% | 66.85% | 80.37%
112 | 86.8% | 68.61% | 80.76% | 66.92% | 73.18% | 67.76% | 78.26%
113 | 85.82% | 72.12% | 86.48% | 72.84% | 79.71% | 72.6% | 81.85%
114 | 87.26% | 73.41% | 81.65% | 74.56% | 80.52% | 73.46% | 85.32%
115 | 83.44% | 72.34% | 85.73% | 72.99% | 69.01% | 75.08% | 83.16%
116 | 77.39% | 71.92% | 84.27% | 72.33% | 77.39% | 70.46% | 83.74%
117 | 90.77% | 70.99% | 86.18% | 70.25% | 76.87% | 68.6% | 83.63%
118 | 86.63% | 70.82% | 80.75% | 72.78% | 74.91% | 68.26% | 80.3%
119 | 84.19% | 68.59% | 83.37% | 70.24% | 85.75% | 67.3% | 79.51%
120 | 79.53% | 69.61% | 81.02% | 75.45% | 73.92% | 68.14% | 80.56%
121 | 79.85% | 69% | 79.82% | 68.88% | 75.21% | 68.75% | 78.48%
122 | 73.87% | 60.74% | 72.33% | 61.81% | 59.59% | 81.98% | 71.96%
123 | 77.05% | 62.96% | 73.88% | 65.87% | 61.67% | 61.54% | 72.69%
124 | 79.72% | 72.56% | 88.09% | 74.29% | 80.24% | 70.79% | 84.53%
125 | 86.15% | 73.55% | 83.36% | 72.87% | 78.33% | 73.93% | 87.1%
126 | 84.19% | 71.22% | 85.55% | 68.65% | 74.48% | 69.67% | 83.54%
127 | 81.1% | 66.1% | 77.4% | 65.89% | 68.95% | 64.48% | 77.34%
128 | 80.12% | 65.05% | 78.89% | 65.56% | 69.78% | 66.03% | 74.9%
129 | 80.85% | 51.42% | 83.32% | 65.69% | 70.42% | 65.61% | 61.51%
130 | 82.84% | 66.19% | 79.11% | 66.44% | 68.24% | 64.95% | 77.56%
131 | 85.08% | 72.48% | 86.76% | 72.19% | 80.12% | 69.94% | 82.92%
132 | 84.82% | 72.62% | 87.93% | 73.78% | 77.99% | 69.88% | 86.12%
133 | 83.32% | 70.29% | 83.04% | 68.09% | 75.94% | 69.57% | 82.43%
134 | 77.68% | 70.36% | 83.61% | 69.55% | 76.59% | 69.71% | 84.08%
135 | 78.43% | 67.54% | 82.84% | 80.26% | 76.51% | 68.74% | 81.08%
136 | 80.92% | 71.43% | 84.81% | 68.43% | 74.67% | 69.15% | 81.4%
137 | 81.46% | 62.72% | 80.73% | 67.35% | 74.45% | 69.28% | 73.8%
138 | 80.06% | 70.68% | 78.81% | 68.35% | 72.56% | 67.85% | 81.33%
139 | 80.6% | 67.15% | 82.42% | 66.7% | 74.77% | 68.53% | 79.36%
140 | 81.07% | 67.85% | 79.67% | 67.05% | 72.47% | 61.56% | 80.17%
141 | 87.32% | 67.8% | 79.28% | 70.62% | 74.86% | 67.34% | 79.17%
142 | 80.03% | 67.55% | 79% | 65.49% | 76.95% | 66.62% | 79.85%
143 | 80.28% | 66.07% | 80.44% | 65.78% | 71.02% | 66.11% | 77.17%
144 | 76.69% | 63.08% | 72.91% | 62.65% | 67.06% | 61.8% | 74.57%
145 | 79.46% | 64.21% | 75.62% | 63.64% | 67.53% | 65.41% | 72.84%
146 | 83.29% | 70.37% | 82.03% | 69.96% | 74.17% | 61.67% | 82.9%
147 | 81.27% | 62.63% | 72.48% | 68.6% | 65.79% | 61.05% | 72.81%
148 | 81.94% | 70.34% | 84.82% | 70.7% | 68.21% | 66.71% | 82.45%
149 | 83.51% | 71.48% | 86.86% | 71.37% | 76.69% | 70.46% | 83.16%
150 | 78.97% | 70.2% | 82.08% | 70.26% | 72.13% | 71.1% | 83.8%
151 | 79.15% | 69.92% | 83.04% | 66.02% | 74.06% | 67.58% | 80.5%
152 | 81.65% | 68.54% | 83.43% | 72.18% | 76.69% | 67.59% | 83.44%
153 | 80.2% | 70.2% | 80.67% | 75.37% | 73.82% | 70.58% | 84.1%
154 | 80.77% | 67.68% | 80.1% | 69.49% | 73.99% | 66.78% | 80.15%
155 | 79.89% | 69.84% | 80.86% | 66.07% | 72.96% | 66.02% | 84.83%
156 | 79.46% | 66.8% | 81.75% | 66.59% | 72.01% | 67.23% | 79.44%
157 | 78.79% | 67.83% | 78.94% | 65.96% | 71.04% | 64.86% | 78.05%
158 | 78.56% | 66.67% | 78.3% | 59.14% | 68.93% | 65.16% | 80.67%
159 | 77.49% | 67.22% | 73.82% | 62.67% | 66.01% | 61.62% | 83.29%
160 | 78.76% | 60.66% | 78.98% | 71.23% | 81.31% | 62.98% | 73.57%
161 | 67.95% | 63.43% | 73.75% | 62.01% | 61.54% | 61.52% | 76.51%
162 | 82.69% | 71.94% | 73.64% | 63.98% | 61.26% | 62.79% | 84.7%
163 | 83.54% | 63.18% | 78.63% | 68.73% | 73.69% | 67.44% | 73.93%
164 | 75.53% | 68.1% | 78.53% | 66.52% | 74.27% | 66.56% | 81.18%
165 | 76.36% | 63.03% | 71.99% | 36.71% | 28.83% | 60.04% | 73.8%
166 | 78.76% | 61.33% | 72.2% | 61.68% | 65.38% | 60.45% | 73.5%
167 | 72.12% | 68.9% | 81.45% | 70.08% | 75.91% | 67.55% | 79.81%
168 | 79.69% | 72.12% | 81.62% | 70.06% | 73.92% | 65.78% | 83.67%

## QC
QC Standards
QC STD 1
|  | Sample Id | R | Acquisition Time | Na KED 23/23 | Mg KED 24/24 | Al KED 27/27 | P Oshift 31/47 | S (34) -1 34/34 | K KED 39/39 | Ca KED 43/43 | V KED 51/51 | Cr (52) 52/52 | Cr (53) NH3 53/53 | Fe (54) NH3 54/54 | Mn Oshift 55/71 | Fe (56) KED 56/56 | Co KED 59/59 | Ni KED 60/60 | Cu (63) KED 63/63 | Cu (65) NH3 65/65 | Zn NH3 66/66 | As 75/75 | As KED 75/75 | As RH Oshift 75/91 | As Oshift 75/91 | Se (77) 77/77 | Se (77) Oshift 77/93 | Se (78) Oshift 78/94 | Se (80) Oshift 80/96 | Se (82) 82/82 | Se (82) Oshift 82/98 | Mo (95) NH3 95/95 | Mo (95) OShift 95/127 | Rh (IS1) 103/103 (IS) | Rh (IS2) 103/103 (IS) | Rh (IS3) 103/103 (IS) | Rh (IS4) 103/103 (IS) | Rh (IS5) 103/103 (IS) | Rh (IS6) 103/103 (IS) | Rh OShift (IS1) 103/119 (IS) | Cd (111) NH3 111/111 | Cd (114) NH3 114/114 | Pb 208/208 | Pb NH3 208/208
| 13 | LBLANK | R | 06/02/2023 16:24:07 | -145.773 | -48.7752 | -0.92607 | -0.32009 | 467.0084 | -167.5555 | -0.01244 | 0.0096825 | 0.0050838 | 0.00097923 | -0.56933 | -0.0036322 | -0.53599 | 0.00289 | -0.0094525 | 0.00025038 | -0.010745 | -0.23801 | -0.0071944 | 0.0039143 | -0.0036233 | -0.0038431 | 0.10924 | 0.014903 | -0.029849 | -0.0036589 | 0.020548 | -0.0016216 | -0.0039439 | -0.0015604 | [7 empty] | -0.00031202 | 0.00075458 | -0.00075595 | -0.00039769
|  | RSD |  |  | 0.9233% | 2.521% | 2.406% | 11.13% | 24.14% | 4.481% | 39168% | 12.16% | 74.63% | 69.51% | 1.112% | 17.47% | 3.882% | 67.27% | 20.62% | 1794% | 3.173% | 0.6495% | 419.4% | 310.7% | 34.46% | 35.11% | 268.8% | 1.913e-06% | 15.32% | 58.22% | 258.3% | 409% | 26.64% | 116% | [7 empty] | 84.81% | 13.63% | 167.5% | 29.59%
|  | % Recovery | [32 empty] | 111.9% | 101.1% | 108.8% | 117.2% | 118.5% | 98.82% | 109.1%
| 43 | LBLANK | R | 06/02/2023 19:44:17 | -148.4385 | -48.7435 | -0.94902 | 0.47215 | 1503.6932 | -154.0031 | -4.861 | 0.0074153 | 0.0039509 | 0.000477 | -0.52315 | -0.0032491 | -0.2402 | 0.0004931 | -0.0061775 | -0.0020108 | -0.0087663 | -0.21823 | -0.047512 | 0.00060526 | -0.0056794 | -0.005765 | 0.055805 | 0.02008 | -0.030431 | -0.0034118 | -0.0091382 | 0.0032203 | -0.010332 | -0.012062 | [7 empty] | -0.0012259 | 0.00081252 | -0.0020589 | -0.00072911
|  | RSD |  |  | 0.714% | 2.635% | 4.327% | 35.33% | 22.56% | 2.934% | 0% | 16.35% | 24.11% | 154.8% | 3.324% | 20.74% | 13.6% | 53.06% | 170.9% | 20.86% | 13.7% | 2.923% | 52.1% | 1062% | 24.98% | 28.94% | 146.1% | 44.66% | 9.666% | 40.8% | 382.9% | 233.2% | 6.163% | 1.378% | [7 empty] | 14.74% | 28.67% | 5.706% | 23.97%
|  | % Recovery | [32 empty] | 99.33% | 89.97% | 104.6% | 88.3% | 93.57% | 95.7% | 105.9%
| 59 | LBLANK | R | 06/02/2023 21:22:26 | -143.7588 | -46.5218 | -0.78556 | 0.15823 | 1683.4894 | -160.4684 | -4.861 | 0.0094303 | 0.005387 | 0.00051521 | -0.50395 | -0.0029249 | -0.14883 | 0.00049207 | 0.0021397 | 0.0025971 | -0.0069244 | -0.18386 | -0.047651 | 0.0043281 | -0.0070991 | -0.0074371 | 0.14709 | 0.020302 | -0.026509 | -0.0045254 | 0.016283 | -0.00080727 | 0.0011055 | -0.011562 | [7 empty] | -0.0006411 | 0.00059854 | -0.0018058 | -0.00074913
|  | RSD |  |  | 1.02% | 5.517% | 8.996% | 58.81% | 8.105% | 2.701% | 0% | 19.66% | 41.69% | 128.9% | 0.2755% | 5.577% | 43.47% | 104.7% | 219.4% | 14.84% | 13.63% | 11.7% | 54.63% | 148.8% | 10.49% | 12.6% | 134% | 46.06% | 41.45% | 198.2% | 628.7% | 996.2% | 266.6% | 1.828% | [7 empty] | 116.2% | 60.6% | 17.39% | 22.83%
|  | % Recovery | [32 empty] | 98.95% | 85.36% | 104.6% | 87.03% | 93.37% | 83.76% | 101.9%
| 76 | LBLANK | R | 06/02/2023 23:07:00 | -145.4882 | -49.4851 | -0.80611 | 0.55563 | 1671.8764 | -154.3215 | -2.8124 | 0.0079497 | 0.0071477 | 0.0043221 | -0.4533 | -0.0033659 | -0.087474 | 0.00043961 | -0.0024758 | 0.0016345 | -0.007604 | -0.224 | -0.10021 | -0.0031069 | -0.0067848 | -0.0071568 | 0.2154 | 0.026794 | -0.032915 | -0.0039661 | -0.059389 | -0.0054505 | -0.0090693 | -0.012204 | [7 empty] | -0.00078141 | 0.00033854 | -0.0020194 | -0.00071883
|  | RSD |  |  | 1.574% | 0% | 13.98% | 9.38% | 32.92% | 3.834% | 126.2% | 23.83% | 61.17% | 44.15% | 2.359% | 15.98% | 136% | 123.6% | 107.9% | 91.17% | 21.43% | 1.899% | 36.85% | 0% | 8.154% | 8.721% | 46.2% | 38.44% | 18.28% | 177.6% | 18.23% | 0% | 6.64% | 2.555% | [7 empty] | 24.67% | 26.92% | 23.38% | 27.87%
|  | % Recovery | [32 empty] | 83.05% | 77.62% | 90.89% | 77.65% | 82.47% | 74.71% | 89.92%
| 91 | LBLANK | R | 06/03/2023 00:39:20 | -145.5986 | -49.4851 | -0.84614 | 1.9448 | 2279.0338 | -149.0433 | -0.70675 | 0.0091747 | 0.0054108 | 0.0031163 | -0.42463 | -0.0024772 | 0.023472 | 0.001005 | -0.0033192 | 0.0028118 | -0.0076276 | -0.21115 | -0.1254 | 0.015162 | -0.0071337 | -0.0075301 | 0.43478 | 0.014903 | -0.02441 | -0.0071842 | -0.02621 | -8.5498e-05 | -0.0024399 | -0.012072 | [7 empty] | 1.9841e-05 | 0.00049455 | -0.0018493 | -0.00072881
|  | RSD |  |  | 1.37% | 0% | 6.649% | 5.276% | 12.82% | 3.926% | 509.1% | 15.93% | 73.45% | 85.52% | 4.994% | 18.63% | 157.3% | 96.52% | 135.8% | 41.38% | 30.47% | 1.225% | 53.41% | 53.49% | 5.734% | 6.007% | 21.84% | 1.913e-06% | 14.96% | 61.07% | 129.6% | 10869% | 32.64% | 4.501% | [7 empty] | 1853% | 132.7% | 10.51% | 14.29%
|  | % Recovery | [32 empty] | 83.34% | 73.5% | 84.73% | 75.99% | 85.38% | 73.36% | 86.5%
| 108 | LBLANK | R | 06/03/2023 02:24:45 | -146.3403 | -49.4851 | -0.76116 | 1.5027 | 2955.6642 | -144.5824 | -0.69835 | 0.0076988 | 0.0086455 | 0.0011485 | -0.46784 | -0.0028974 | 0.0065413 | 0.00045 | 0.0070436 | 0.0020642 | -0.0095033 | -0.21796 | -0.10159 | -0.0031069 | -0.0076273 | -0.0081989 | 0.33343 | 0.014903 | -0.03055 | -0.010994 | -0.024012 | -0.0054505 | -0.00885 | -0.012493 | [7 empty] | -0.0012151 | 0.00025078 | -0.0019169 | -0.00072586
|  | RSD |  |  | 2.554% | 0% | 9.879% | 5.998% | 2.511% | 2.243% | 516.3% | 19.37% | 27.34% | 77.62% | 2.989% | 55.72% | 580.5% | 1% | 108% | 335.7% | 7.564% | 0.9124% | 17.97% | 0% | 7.962% | 8.612% | 24.04% | 1.913e-06% | 30.95% | 15.38% | 125.7% | 0% | 10.93% | 0.9506% | [7 empty] | 16.41% | 121.9% | 16.26% | 23.58%
|  | % Recovery | [32 empty] | 84.41% | 72.5% | 84.44% | 73.52% | 74.9% | 71.65% | 82.62%
| 124 | LBLANK | R | 06/03/2023 04:03:32 | -144.6105 | -48.6235 | -0.82122 | 0.42084 | 2439.7664 | -161.6167 | -0.8722 | 0.0071738 | 0.010563 | 0.0047714 | -0.48381 | -0.0025897 | 0.002076 | 0.0011311 | 0.0039121 | 0.00098345 | -0.0075572 | -0.22481 | -0.15217 | 0.0055406 | -0.0058034 | -0.0059339 | 0.42485 | 0.014903 | -0.034722 | -0.0042137 | -0.04304 | -0.0054505 | -0.0037191 | -0.011543 | [7 empty] | -0.0011022 | 0.00048136 | -0.0019962 | -0.00075634
|  | RSD |  |  | 2.06% | 3.069% | 12.06% | 26.5% | 2.724% | 6.055% | 396.1% | 23.2% | 26.36% | 17.41% | 2.059% | 27.36% | 1426% | 25.64% | 116.3% | 188.6% | 11.93% | 1.452% | 24.68% | 270.3% | 6.718% | 11.97% | 49.03% | 1.913e-06% | 10.34% | 159.4% | 78.17% | 0% | 54.05% | 3.984% | [7 empty] | 17.92% | 108.9% | 25.08% | 27.36%
|  | % Recovery | [32 empty] | 79.72% | 72.56% | 88.09% | 74.29% | 80.24% | 70.79% | 84.53%
| 141 | LBLANK | R | 06/03/2023 05:47:43 | -147.4082 | -49.4851 | -0.88719 | 1.3695 | 3009.3505 | -153.8548 | -2.6267 | 0.0078023 | 0.011657 | 0.003746 | -0.47825 | -0.0019548 | 0.066039 | 0.00028853 | 0.0022234 | -0.0017249 | -0.0092517 | -0.22657 | -0.043734 | -0.0031069 | -0.0079101 | -0.0084986 | 0.079404 | 0.028747 | -0.034374 | -0.010947 | 0.005303 | 0.00040536 | -0.011249 | -0.012048 | [7 empty] | -0.00095466 | 0.00014374 | -0.0022795 | -0.00079824
|  | RSD |  |  | 0.4898% | 0% | 3.346% | 4.157% | 8.847% | 2.587% | 147.3% | 20.61% | 23.33% | 35.53% | 5.356% | 27.26% | 46.57% | 117.5% | 282.4% | 143.9% | 16.13% | 4.173% | 117.6% | 0% | 9.837% | 10.63% | 139.6% | 41.71% | 11.32% | 16.19% | 444.6% | 2502% | 4.165% | 2.904% | [7 empty] | 38.75% | 256.2% | 10.01% | 8.619%
|  | % Recovery | [32 empty] | 87.32% | 67.8% | 79.28% | 70.62% | 74.86% | 67.34% | 79.17%
| 156 | LBLANK | R | 06/03/2023 07:19:20 | -147.857 | -49.4851 | -0.82108 | 1.2593 | 2577.0718 | -148.8617 | -4.861 | 0.0082894 | 0.0083397 | 0.0031559 | -0.49735 | -0.0030791 | -0.066008 | 0.00046874 | 0.0023377 | 0.00014746 | -0.0088015 | -0.22328 | -0.12393 | 0.0014155 | -0.0081964 | -0.0087918 | 0.18021 | 0.014903 | -0.03214 | -0.0088262 | -0.06201 | 0.0064591 | -0.010277 | -0.012816 | [7 empty] | -0.001204 | 0.00039887 | -0.0021029 | -0.00087167
|  | RSD |  |  | 0.03082% | 0% | 17.2% | 4.553% | 12.47% | 6.607% | 0% | 29.38% | 24.75% | 35.88% | 0.9258% | 52.46% | 124.7% | 4.655% | 109.5% | 923% | 21.51% | 1.935% | 6.987% | 553.4% | 5.978% | 7.441% | 74.85% | 1.913e-06% | 21.09% | 1.117% | 57.56% | 159.8% | 6.927% | 1.863% | [7 empty] | 18.16% | 102% | 21.06% | 17.54%
|  | % Recovery | [32 empty] | 79.46% | 66.8% | 81.75% | 66.59% | 72.01% | 67.23% | 79.44%
QC STD 2
|  | Sample Id | R | Acquisition Time | Na KED 23/23 | Mg KED 24/24 | Al KED 27/27 | P Oshift 31/47 | S (34) -1 34/34 | K KED 39/39 | Ca KED 43/43 | V KED 51/51 | Cr (52) 52/52 | Cr (53) NH3 53/53 | Fe (54) NH3 54/54 | Mn Oshift 55/71 | Fe (56) KED 56/56 | Co KED 59/59 | Ni KED 60/60 | Cu (63) KED 63/63 | Cu (65) NH3 65/65 | Zn NH3 66/66 | As 75/75 | As KED 75/75 | As RH Oshift 75/91 | As Oshift 75/91 | Se (77) 77/77 | Se (77) Oshift 77/93 | Se (78) Oshift 78/94 | Se (80) Oshift 80/96 | Se (82) 82/82 | Se (82) Oshift 82/98 | Mo (95) NH3 95/95 | Mo (95) OShift 95/127 | Rh (IS1) 103/103 (IS) | Rh (IS2) 103/103 (IS) | Rh (IS3) 103/103 (IS) | Rh (IS4) 103/103 (IS) | Rh (IS5) 103/103 (IS) | Rh (IS6) 103/103 (IS) | Rh OShift (IS1) 103/119 (IS) | Cd (111) NH3 111/111 | Cd (114) NH3 114/114 | Pb 208/208 | Pb NH3 208/208
| 14 | ICV PE | R | 06/02/2023 16:30:09 | 9823.9524 | 9944.6776 | 11.1131 | 20.0669 | 10681.7877 | 9422.7641 | 10726.211 | 10.0739 | 10.0157 | 9.7619 | 10.0029 | 10.2034 | 10.5905 | 10.3283 | 10.5191 | 9.9341 | 9.7802 | 9.9155 | 9.885 | 8.7886 | 9.0956 | 10.0139 | 7.2921 | 9.9484 | 10.457 | 10.2681 | 10.0259 | 10.2691 | 10.1446 | 10.3165 | [7 empty] | 9.8959 | 9.8193 | 8.3318 | 9.1537
|  | RSD |  |  | 3.37% | 2.885% | 4.71% | 0.6954% | 1.437% | 1.75% | 4.527% | 2.514% | 4.132% | 1.668% | 1.421% | 0.362% | 3.225% | 1.756% | 1.828% | 0.9877% | 1.695% | 1.302% | 2.892% | 1.305% | 1.207% | 1.215% | 1.743% | 2.005% | 2.48% | 1.512% | 1.61% | 1.216% | 0.4603% | 0.6177% | [7 empty] | 1.822% | 1.62% | 0.745% | 0.3812%
|  | % Recovery |  |  | 98.24% | 99.45% | 111.1% | 100.3% | 106.8% | 94.23% | 107.3% | 100.7% | 100.2% | 97.62% | 100% | 102% | 105.9% | 103.3% | 105.2% | 99.34% | 97.8% | 99.16% | 98.85% | 87.89% |  | 100.1% | 72.92% | 99.48% | 104.6% | 102.7% | 100.3% | 102.7% | 101.4% | 103.2% | 123.4% | 99.04% | 110.5% | 107.5% | 110.8% | 99.44% | 105.4% | 98.96% | 98.19% | 83.32% | 91.54%
| 30 | ICV PE | R | 06/02/2023 18:24:08 | 13897.8579 | 13749.6052 | 13.9589 | 28.1108 | 13951.4949 | 12791.5569 | 13327.2364 | 11.5708 | 11.3416 | 11.1409 | 11.7665 | 13.0434 | 12.4062 | 11.8958 | 11.9854 | 10.4681 | 10.6167 | 10.7588 | 10.7083 | 8.7767 | 9.2977 | 11.1401 | 8.1861 | 11.2813 | 11.719 | 11.5341 | 11.1233 | 11.3428 | 10.2514 | 10.8331 | [7 empty] | 9.7436 | 9.6854 | 8.1639 | 8.5864
|  | RSD |  |  | 0.8838% | 2.574% | 4.373% | 1.345% | 0.8144% | 4.193% | 1.831% | 1.823% | 0.7844% | 1.226% | 2.911% | 0.3676% | 1.559% | 1.4% | 1.432% | 0.1444% | 1.479% | 2.506% | 2.325% | 2.716% | 1.11% | 1.556% | 2.268% | 1.757% | 1.845% | 0.6239% | 1.68% | 7.124% | 2.014% | 0.6545% | [7 empty] | 2.401% | 2.538% | 0.6703% | 1.605%
|  | % Recovery |  |  | 139% | 137.5% | 139.6% | 140.6% | 139.5% | 127.9% | 133.3% | 115.7% | 113.4% | 111.4% | 117.7% | 130.4% | 124.1% | 119% | 119.9% | 104.7% | 106.2% | 107.6% | 107.1% | 87.77% |  | 111.4% | 81.86% | 112.8% | 117.2% | 115.3% | 111.2% | 113.4% | 102.5% | 108.3% | 102.1% | 84.96% | 101.2% | 88.66% | 92.57% | 84.94% | 98.43% | 97.44% | 96.85% | 81.64% | 85.86%
| 31 | ICV PE | R | 06/02/2023 18:30:39 | 13661.1271 | 13405.2844 | 13.5666 | 25.9672 | 13345.0393 | 12236.5207 | 13307.5124 | 11.4975 | 11.1834 | 11.1563 | 11.9394 | 12.86 | 12.1774 | 11.8808 | 11.9216 | 10.4523 | 10.6991 | 10.8552 | 10.5169 | 8.9232 | 9.0987 | 10.8854 | 8.0598 | 11.6457 | 11.4559 | 11.268 | 10.8422 | 11.3762 | 10.3886 | 10.7405 | [7 empty] | 9.9262 | 9.8307 | 8.0438 | 8.4798
|  | RSD |  |  | 1.521% | 1.367% | 2.782% | 1.658% | 2.275% | 1.671% | 3.181% | 1.384% | 0.3291% | 2.573% | 1.734% | 1.354% | 0.8064% | 0.9122% | 1.392% | 1.558% | 0.8248% | 1.336% | 2.134% | 6.154% | 2.528% | 1.221% | 2.095% | 2.528% | 1.482% | 0.8713% | 0.7146% | 4.541% | 0.9984% | 0.2326% | [7 empty] | 1.426% | 0.7498% | 0.4838% | 1.799%
|  | % Recovery |  |  | 136.6% | 134.1% | 135.7% | 129.8% | 133.5% | 122.4% | 133.1% | 115% | 111.8% | 111.6% | 119.4% | 128.6% | 121.8% | 118.8% | 119.2% | 104.5% | 107% | 108.6% | 105.2% | 89.23% |  | 108.9% | 80.6% | 116.5% | 114.6% | 112.7% | 108.4% | 113.8% | 103.9% | 107.4% | 103% | 88.1% | 103.2% | 88.6% | 93.47% | 91.63% | 101.9% | 99.26% | 98.31% | 80.44% | 84.8%
QC STD 3
|  | Sample Id | R | Acquisition Time | Na KED 23/23 | Mg KED 24/24 | Al KED 27/27 | P Oshift 31/47 | S (34) -1 34/34 | K KED 39/39 | Ca KED 43/43 | V KED 51/51 | Cr (52) 52/52 | Cr (53) NH3 53/53 | Fe (54) NH3 54/54 | Mn Oshift 55/71 | Fe (56) KED 56/56 | Co KED 59/59 | Ni KED 60/60 | Cu (63) KED 63/63 | Cu (65) NH3 65/65 | Zn NH3 66/66 | As 75/75 | As KED 75/75 | As RH Oshift 75/91 | As Oshift 75/91 | Se (77) 77/77 | Se (77) Oshift 77/93 | Se (78) Oshift 78/94 | Se (80) Oshift 80/96 | Se (82) 82/82 | Se (82) Oshift 82/98 | Mo (95) NH3 95/95 | Mo (95) OShift 95/127 | Rh (IS1) 103/103 (IS) | Rh (IS2) 103/103 (IS) | Rh (IS3) 103/103 (IS) | Rh (IS4) 103/103 (IS) | Rh (IS5) 103/103 (IS) | Rh (IS6) 103/103 (IS) | Rh OShift (IS1) 103/119 (IS) | Cd (111) NH3 111/111 | Cd (114) NH3 114/114 | Pb 208/208 | Pb NH3 208/208
| 15 | CCV 1ppb TRACE | R | 06/02/2023 16:36:12 | 917.1242 | 1032.8229 | 0.67866 | 0.82457 | 1812.5312 | 818.0184 | 1092.2421 | 1.0775 | 1.0519 | 1.0305 | 0.5955 | 1.0632 | 0.79028 | 1.1017 | 1.119 | 1.0813 | 1.0815 | 0.87659 | 0.97221 | 1.0031 | 0.86087 | 0.98823 | 0.98123 | 1.0375 | 0.95556 | 0.98496 | 1.0295 | 0.97738 | 1.0469 | 1.0568 | [7 empty] | 1.0145 | 1.0178 | 0.92704 | 0.92218
|  | RSD |  |  | 3.661% | 5.083% | 28.26% | 6.21% | 6.327% | 5.035% | 1.963% | 1.94% | 1.032% | 2.068% | 1.167% | 0.9468% | 11.4% | 3.894% | 4.11% | 2.609% | 1.695% | 4.102% | 4.003% | 17.86% | 1.162% | 1.24% | 8.868% | 5.437% | 2.516% | 2.285% | 3.765% | 2.641% | 1.04% | 0.5805% | [7 empty] | 1.52% | 2.083% | 1.771% | 1.07%
|  | % Recovery |  |  | 91.71% | 103.3% | 67.87% | 82.46% | 181.3% | 81.8% | 109.2% | 107.8% | 105.2% | 103% | 59.55% | 106.3% | 79.03% | 110.2% | 111.9% | 108.1% | 108.1% | 87.66% | 97.22% | 100.3% | 86.09% | 98.82% | 98.12% | 103.7% | 95.56% | 98.5% | 102.9% | 97.74% | 104.7% | 105.7% | 123.2% | 109.7% | 120.9% | 114.8% | 115.8% | 97.9% | 121.7% | 101.5% | 101.8% | 92.7% | 92.22%
| 16 | CCV 1ppb TRACE | R | 06/02/2023 16:42:44 | 988.4925 | 1014.6011 | 0.53485 | 0.52706 | 2076.2749 | 907.2816 | 1058.5224 | 1.0715 | 1.0503 | 1.021 | 0.55656 | 1.1252 | 0.76502 | 1.107 | 1.1572 | 1.0602 | 1.0954 | 1.4936 | 0.98863 | 0.93905 | 0.88344 | 1.0413 | 1.0361 | 1.0852 | 1.0643 | 1.0471 | 0.95568 | 1.1535 | 1.0515 | 1.0617 | [7 empty] | 1.0101 | 0.99231 | 0.93319 | 0.91035
|  | RSD |  |  | 2.11% | 5.11% | 35.48% | 5.271% | 5.239% | 4.991% | 3.768% | 1.457% | 2.787% | 0.6205% | 0.8254% | 0.5057% | 10.24% | 3.251% | 5.067% | 1.435% | 1.486% | 4.954% | 4.094% | 9.137% | 2.42% | 3.118% | 3.353% | 8.526% | 8.389% | 1.611% | 15.8% | 9.361% | 1.073% | 0.4602% | [7 empty] | 0.8061% | 1.542% | 1.289% | 0.6489%
|  | % Recovery |  |  | 98.85% | 101.5% | 53.49% | 52.71% | 207.6% | 90.73% | 105.9% | 107.2% | 105% | 102.1% | 55.66% | 112.5% | 76.5% | 110.7% | 115.7% | 106% | 109.5% | 149.4% | 98.86% | 93.91% | 88.34% | 104.1% | 103.6% | 108.5% | 106.4% | 104.7% | 95.57% | 115.3% | 105.2% | 106.2% | 117.4% | 105.2% | 122.8% | 109.6% | 111% | 104.8% | 119.8% | 101% | 99.23% | 93.32% | 91.03%
| 32 | CCV 1ppb TRACE | R | 06/02/2023 18:36:41 | 1197.1876 | 1248.5498 | 0.39246 | 1.2004 | 2197.7058 | 1089.3757 | 1249.0218 | 1.1299 | 1.1132 | 1.1117 | 0.69584 | 1.3335 | 0.95012 | 1.1826 | 1.215 | 1.0725 | 1.1209 | 0.89552 | 1.0664 | 1.0087 | 0.93385 | 1.1345 | 1.2297 | 1.2676 | 1.1228 | 1.1986 | 1.0296 | 1.1019 | 1.0679 | 1.1433 | [7 empty] | 0.99734 | 0.99206 | 0.90466 | 0.89042
|  | RSD |  |  | 5.891% | 0.4015% | 48.45% | 6.495% | 7.703% | 3.638% | 6.581% | 3.823% | 7.057% | 1.462% | 5.064% | 5.874% | 8.257% | 3.664% | 2.151% | 1.784% | 0.7273% | 2.565% | 4.954% | 11.69% | 7.102% | 5.773% | 15.52% | 9.622% | 8.144% | 5.373% | 4.02% | 18.92% | 0.6316% | 4.986% | [7 empty] | 0.6606% | 0.8623% | 0.5593% | 0.9829%
|  | % Recovery |  |  | 119.7% | 124.9% | 39.25% | 120% | 219.8% | 108.9% | 124.9% | 113% | 111.3% | 111.2% | 69.58% | 133.4% | 95.01% | 118.3% | 121.5% | 107.3% | 112.1% | 89.55% | 106.6% | 100.9% | 93.39% | 113.5% | 123% | 126.8% | 112.3% | 119.9% | 103% | 110.2% | 106.8% | 114.3% | 103% | 91.24% | 113.1% | 95.77% | 98.08% | 94.47% | 107.2% | 99.73% | 99.21% | 90.47% | 89.04%
| 33 | CCV 1ppb TRACE | R | 06/02/2023 18:43:13 | 1280.3979 | 1330.1937 | 1.023 | 0.83521 | 2768.7563 | 1081.081 | 1381.8236 | 1.1716 | 1.187 | 1.1511 | 0.69211 | 1.2998 | 1.1162 | 1.2419 | 1.261 | 1.1349 | 1.1422 | 0.90637 | 1.0183 | 0.88583 | 0.91925 | 1.0956 | 1.1404 | 1.0998 | 1.0763 | 1.1457 | 1.0603 | 1.1365 | 1.0646 | 1.0782 | [7 empty] | 1.0174 | 0.99391 | 0.91094 | 0.89151
|  | RSD |  |  | 1.89% | 4.917% | 13.23% | 3.659% | 1.833% | 1.559% | 6.697% | 3.753% | 1.993% | 1.234% | 1.568% | 1.609% | 4.521% | 1.962% | 1.237% | 1.825% | 1.542% | 2.321% | 3.576% | 6.461% | 3.852% | 1.589% | 6.97% | 12.45% | 5.245% | 2.896% | 6.55% | 8.024% | 0.9768% | 0.7565% | [7 empty] | 2.595% | 1.273% | 0.5409% | 1.049%
|  | % Recovery |  |  | 128% | 133% | 102.3% | 83.52% | 276.9% | 108.1% | 138.2% | 117.2% | 118.7% | 115.1% | 69.21% | 130% | 111.6% | 124.2% | 126.1% | 113.5% | 114.2% | 90.64% | 101.8% | 88.58% | 91.93% | 109.6% | 114% | 110% | 107.6% | 114.6% | 106% | 113.6% | 106.5% | 107.8% | 102.5% | 93.99% | 110.3% | 93.56% | 97.03% | 91.93% | 108.3% | 101.7% | 99.39% | 91.09% | 89.15%
| 44 | CCV 1ppb TRACE | R | 06/02/2023 19:50:20 | 1350.6062 | 1376.7676 | 0.98963 | 0.99616 | 2948.8417 | 1202.9902 | 1431.5703 | 1.2657 | 1.1917 | 1.219 | 0.80702 | 1.35 | 1.1813 | 1.2799 | 1.326 | 1.151 | 1.1641 | 0.96691 | 1.0057 | 1.0306 | 0.90977 | 1.1105 | 1.4257 | 1.1527 | 1.1236 | 1.1697 | 0.98577 | 1.1579 | 1.0792 | 1.1104 | [7 empty] | 1.0115 | 1.0049 | 0.90438 | 0.87798
|  | RSD |  |  | 0.3585% | 2.912% | 45.27% | 3.469% | 11.34% | 3.154% | 1.743% | 5.451% | 2.116% | 3.577% | 2.778% | 0.5439% | 7.402% | 2.309% | 0.2095% | 1.442% | 1.598% | 0.8381% | 5.996% | 12.35% | 1.371% | 2.079% | 13.47% | 5.013% | 11.95% | 3.203% | 7.582% | 10.9% | 3.494% | 1.635% | [7 empty] | 1.337% | 0.527% | 9.521% | 1.804%
|  | % Recovery |  |  | 135.1% | 137.7% | 98.96% | 99.62% | 294.9% | 120.3% | 143.2% | 126.6% | 119.2% | 121.9% | 80.7% | 135% | 118.1% | 128% | 132.6% | 115.1% | 116.4% | 96.69% | 100.6% | 103.1% | 90.98% | 111.1% | 142.6% | 115.3% | 112.4% | 117% | 98.58% | 115.8% | 107.9% | 111% | 94.2% | 86.26% | 103.7% | 87.38% | 90.89% | 87.36% | 101.7% | 101.2% | 100.5% | 90.44% | 87.8%
| 45 | CCV 1ppb TRACE | R | 06/02/2023 19:56:52 | 1341.2632 | 1409.3016 | 1.1294 | 0.80754 | 3029.6237 | 1128.3677 | 1350.0194 | 1.2141 | 1.1959 | 1.1948 | 0.75335 | 1.3215 | 1.1268 | 1.2477 | 1.2858 | 1.162 | 1.1752 | 0.9289 | 1.0808 | 1.021 | 0.90009 | 1.0863 | 1.2791 | 1.121 | 1.1189 | 1.142 | 1.0566 | 1.2841 | 1.0706 | 1.0766 | [7 empty] | 0.99114 | 0.99077 | 0.87733 | 0.86759
|  | RSD |  |  | 3.577% | 5.436% | 20.78% | 6.354% | 4.773% | 5.461% | 2.732% | 1.285% | 2.34% | 2.697% | 4.398% | 1.151% | 1.221% | 2.177% | 3.192% | 1.593% | 0.8376% | 4.18% | 2.628% | 7.151% | 3.414% | 0.2311% | 3.699% | 16.3% | 3.168% | 2.484% | 5.644% | 25.73% | 0.91% | 1.211% | [7 empty] | 2.139% | 1.309% | 0.6382% | 1.158%
|  | % Recovery |  |  | 134.1% | 140.9% | 112.9% | 80.75% | 303% | 112.8% | 135% | 121.4% | 119.6% | 119.5% | 75.34% | 132.2% | 112.7% | 124.8% | 128.6% | 116.2% | 117.5% | 92.89% | 108.1% | 102.1% | 90.01% | 108.6% | 127.9% | 112.1% | 111.9% | 114.2% | 105.7% | 128.4% | 107.1% | 107.7% | 94.19% | 87.53% | 104.8% | 87.61% | 91.75% | 87.12% | 102.1% | 99.11% | 99.08% | 87.73% | 86.76%
| 60 | CCV 1ppb TRACE | R | 06/02/2023 21:28:29 | 1517.2125 | 1520.3576 | 0.68273 | 1.2064 | 3447.1305 | 1288.0561 | 1473.8406 | 1.3139 | 1.2404 | 1.2524 | 0.82199 | 1.3582 | 1.3041 | 1.3271 | 1.3551 | 1.2355 | 1.2313 | 1.0177 | 1.074 | 1.0126 | 0.91882 | 1.1237 | 1.3492 | 1.0093 | 1.1465 | 1.1391 | 1.1309 | 1.0883 | 1.0961 | 1.0759 | [7 empty] | 1.0191 | 1.0049 | 0.8405 | 0.83944
|  | RSD |  |  | 5.284% | 5.189% | 17.41% | 3.35% | 5.695% | 3.07% | 10.07% | 6.173% | 7.737% | 1.108% | 1.613% | 2.287% | 4.141% | 4.749% | 4.079% | 8.541% | 1.897% | 1.56% | 2.419% | 9.036% | 3.932% | 2.539% | 12.42% | 10.55% | 5.198% | 1.172% | 6.711% | 3.715% | 1.15% | 2.751% | [7 empty] | 1.686% | 0.9935% | 1.256% | 0.8751%
|  | % Recovery |  |  | 151.7% | 152% | 68.27% | 120.6% | 344.7% | 128.8% | 147.4% | 131.4% | 124% | 125.2% | 82.2% | 135.8% | 130.4% | 132.7% | 135.5% | 123.5% | 123.1% | 101.8% | 107.4% | 101.3% | 91.88% | 112.4% | 134.9% | 100.9% | 114.7% | 113.9% | 113.1% | 108.8% | 109.6% | 107.6% | 93.52% | 84.19% | 99.78% | 83.6% | 88.21% | 83.04% | 99.49% | 101.9% | 100.5% | 84.05% | 83.94%
| 61 | CCV 1ppb TRACE | R | 06/02/2023 21:35:00 | 1538.606 | 1596.5673 | 0.98671 | 1.1278 | 3499.2738 | 1288.1703 | 1500.661 | 1.2957 | 1.2455 | 1.2826 | 0.89831 | 1.3918 | 1.292 | 1.3285 | 1.3127 | 1.2022 | 1.2776 | 1.0317 | 1.0332 | 0.96423 | 0.91999 | 1.1448 | 1.3521 | 1.1303 | 1.2286 | 1.1622 | 1.0164 | 1.1517 | 1.1242 | 1.1045 | [7 empty] | 1.0321 | 1.0195 | 0.84105 | 0.83925
|  | RSD |  |  | 0.7007% | 4.037% | 55.07% | 8.822% | 3.271% | 5.848% | 4.011% | 1.834% | 1.856% | 10.44% | 16.65% | 2.145% | 5.875% | 2.379% | 4.366% | 3.234% | 10.67% | 12.94% | 5.685% | 10.4% | 0.9819% | 1.949% | 6.863% | 16.29% | 1.314% | 0.9249% | 12.65% | 10.46% | 9.698% | 2.374% | [7 empty] | 8.062% | 10.68% | 1.08% | 0.9483%
|  | % Recovery |  |  | 153.9% | 159.7% | 98.67% | 112.8% | 349.9% | 128.8% | 150.1% | 129.6% | 124.5% | 128.3% | 89.83% | 139.2% | 129.2% | 132.8% | 131.3% | 120.2% | 127.8% | 103.2% | 103.3% | 96.42% | 92% | 114.5% | 135.2% | 113% | 122.9% | 116.2% | 101.6% | 115.2% | 112.4% | 110.5% | 90.5% | 81.75% | 101.3% | 80.57% | 86.68% | 80.87% | 98.24% | 103.2% | 101.9% | 84.1% | 83.92%
| 77 | CCV 1ppb TRACE | R | 06/02/2023 23:13:03 | 1921.9981 | 1865.4923 | 1.028 | 2.1966 | 3700.4663 | 1516.1001 | 1591.8276 | 1.398 | 1.2745 | 1.2502 | 0.82723 | 1.4553 | 1.2864 | 1.3102 | 1.31 | 1.227 | 1.1453 | 0.93786 | 1.1259 | 1.0218 | 0.96293 | 1.144 | 1.2541 | 1.1468 | 1.2247 | 1.1432 | 1.1947 | 1.2084 | 1.0331 | 1.1041 | [7 empty] | 0.94444 | 0.93824 | 0.82146 | 0.82402
|  | RSD |  |  | 2.258% | 1.369% | 32.37% | 4.023% | 1.246% | 4.581% | 6.803% | 0.649% | 2.397% | 6.601% | 10.14% | 2.426% | 2.838% | 3.198% | 1.277% | 0.121% | 6.062% | 11.11% | 6.901% | 10.97% | 4.464% | 2.37% | 14.57% | 8.395% | 4.915% | 2.755% | 4.009% | 4.847% | 5.066% | 1.514% | [7 empty] | 7.011% | 8.415% | 3.187% | 0.6868%
|  | % Recovery |  |  | 192.2% | 186.5% | 102.8% | 219.7% | 370% | 151.6% | 159.2% | 139.8% | 127.5% | 125% | 82.72% | 145.5% | 128.6% | 131% | 131% | 122.7% | 114.5% | 93.79% | 112.6% | 102.2% | 96.29% | 114.4% | 125.4% | 114.7% | 122.5% | 114.3% | 119.5% | 120.8% | 103.3% | 110.4% | 83.82% | 74.29% | 85.81% | 78.61% | 81.3% | 68.94% | 85.39% | 94.44% | 93.82% | 82.15% | 82.4%
| 78 | CCV 1ppb TRACE | R | 06/02/2023 23:19:35 | 1937.1723 | 1884.1061 | 1.4867 | 1.9533 | 4118.1488 | 1574.5214 | 1515.4807 | 1.3645 | 1.3246 | 1.3493 | 0.93922 | 1.4504 | 1.4205 | 1.3774 | 1.3459 | 1.2303 | 1.2412 | 1.0183 | 1.0005 | 1.1337 | 0.9604 | 1.1573 | 1.2777 | 1.2424 | 1.2232 | 1.1422 | 1.0542 | 1.1886 | 1.0934 | 1.1163 | [7 empty] | 1.0072 | 0.98545 | 0.82437 | 0.81035
|  | RSD |  |  | 1.378% | 2.033% | 14.51% | 11.57% | 5.61% | 2.813% | 5.129% | 3.38% | 2.033% | 1.034% | 5.599% | 1.271% | 11.39% | 0.781% | 1.5% | 5.476% | 1.362% | 0.7508% | 5.629% | 16.07% | 0.5757% | 1.181% | 6.078% | 23.23% | 7.85% | 3.682% | 8.741% | 5.562% | 0.3095% | 2.803% | [7 empty] | 2.16% | 2.728% | 2.958% | 2.938%
|  | % Recovery |  |  | 193.7% | 188.4% | 148.7% | 195.3% | 411.8% | 157.5% | 151.5% | 136.5% | 132.5% | 134.9% | 93.92% | 145% | 142.1% | 137.7% | 134.6% | 123% | 124.1% | 101.8% | 100.1% | 113.4% | 96.04% | 115.7% | 127.8% | 124.2% | 122.3% | 114.2% | 105.4% | 118.9% | 109.3% | 111.6% | 85.2% | 73.52% | 86.02% | 75.57% | 82.58% | 71.63% | 85.58% | 100.7% | 98.54% | 82.44% | 81.04%
| 92 | CCV 1ppb TRACE | R | 06/03/2023 00:45:23 | 2047.6878 | 1948.7255 | 1.3815 | 2.2013 | 3767.647 | 1589.153 | 1598.6672 | 1.3709 | 1.2941 | 1.3767 | 1.0321 | 1.5752 | 1.2912 | 1.329 | 1.3189 | 1.2415 | 1.2681 | 1.042 | 1.0363 | 1.0971 | 1.0197 | 1.2447 | 1.5493 | 1.3729 | 1.24 | 1.196 | 1.0755 | 1.1923 | 1.1168 | 1.2005 | [7 empty] | 1.0115 | 1.0069 | 0.78832 | 0.81479
|  | RSD |  |  | 5.948% | 3.501% | 54.42% | 3.939% | 6.618% | 4.002% | 1.136% | 4.296% | 1.003% | 3.804% | 7.579% | 9.006% | 4.403% | 3.025% | 4.557% | 1.887% | 2.686% | 4.815% | 4.829% | 9.42% | 10.86% | 9.438% | 9.622% | 2.905% | 12.69% | 9.804% | 15.23% | 14.02% | 5.871% | 8.455% | [7 empty] | 4.553% | 2.589% | 2.2% | 5.234%
|  | % Recovery |  |  | 204.8% | 194.9% | 138.1% | 220.1% | 376.8% | 158.9% | 159.9% | 137.1% | 129.4% | 137.7% | 103.2% | 157.5% | 129.1% | 132.9% | 131.9% | 124.2% | 126.8% | 104.2% | 103.6% | 109.7% | 102% | 124.5% | 154.9% | 137.3% | 124% | 119.6% | 107.5% | 119.2% | 111.7% | 120% | 83.65% | 70.54% | 84.57% | 73.28% | 80.18% | 70.99% | 83.35% | 101.2% | 100.7% | 78.83% | 81.48%
| 93 | CCV 1ppb TRACE | R | 06/03/2023 00:51:55 | 2146.4917 | 2009.1125 | 1.5307 | 1.4234 | 4216.6121 | 1634.9721 | 1633.7414 | 1.3741 | 1.3396 | 1.3082 | 0.91974 | 1.5133 | 1.4226 | 1.3757 | 1.3973 | 1.2453 | 1.2151 | 0.98749 | 1.0144 | 1.0212 | 0.97742 | 1.1631 | 1.4628 | 1.1878 | 1.1896 | 1.1879 | 1.057 | 0.98405 | 1.0508 | 1.1257 | [7 empty] | 0.96441 | 0.96082 | 0.79491 | 0.80002
|  | RSD |  |  | 2.75% | 2.87% | 23.42% | 5.958% | 3.814% | 3.214% | 3.675% | 2.902% | 4.847% | 2.204% | 3.891% | 1.656% | 5.663% | 1.162% | 2.027% | 5.078% | 2.98% | 5.546% | 7.185% | 8.832% | 4.362% | 1.988% | 19.85% | 5.485% | 6.725% | 3.624% | 1.032% | 8.339% | 3.167% | 1.826% | [7 empty] | 1.878% | 2.029% | 0.6825% | 0.5742%
|  | % Recovery |  |  | 214.6% | 200.9% | 153.1% | 142.3% | 421.7% | 163.5% | 163.4% | 137.4% | 134% | 130.8% | 91.97% | 151.3% | 142.3% | 137.6% | 139.7% | 124.5% | 121.5% | 98.75% | 101.4% | 102.1% | 97.74% | 116.3% | 146.3% | 118.8% | 119% | 118.8% | 105.7% | 98.4% | 105.1% | 112.6% | 83.2% | 72.13% | 84.04% | 75.98% | 79.78% | 71.57% | 82.98% | 96.44% | 96.08% | 79.49% | 80%
| 109 | CCV 1ppb TRACE | R | 06/03/2023 02:30:48 | 2114.1886 | 1935.993 | 1.3111 | 1.8749 | 4368.4181 | 1685.1854 | 1692.7601 | 1.429 | 1.3469 | 1.4085 | 1.0811 | 1.6008 | 1.4819 | 1.3633 | 1.4149 | 1.3133 | 1.2914 | 1.1048 | 1.0234 | 1.1213 | 1.004 | 1.2204 | 1.2548 | 1.3467 | 1.1259 | 1.1653 | 1.1173 | 1.0488 | 1.1063 | 1.1574 | [7 empty] | 1.0032 | 1.0051 | 0.77883 | 0.72678
|  | RSD |  |  | 3.442% | 6.171% | 38.12% | 3.446% | 4.616% | 3.734% | 5.519% | 3.343% | 4.769% | 13.35% | 17.5% | 4.487% | 8.635% | 2.523% | 3.298% | 3.274% | 12.67% | 15.5% | 4.159% | 21.59% | 0.6092% | 2.786% | 9.757% | 0.5536% | 6.978% | 5.144% | 3.783% | 7.349% | 10.93% | 1.445% | [7 empty] | 11.36% | 11.91% | 0.327% | 6.637%
|  | % Recovery |  |  | 211.4% | 193.6% | 131.1% | 187.5% | 436.8% | 168.5% | 169.3% | 142.9% | 134.7% | 140.8% | 108.1% | 160.1% | 148.2% | 136.3% | 141.5% | 131.3% | 129.1% | 110.5% | 102.3% | 112.1% | 100.4% | 122% | 125.5% | 134.7% | 112.6% | 116.5% | 111.7% | 104.9% | 110.6% | 115.7% | 83.08% | 70.54% | 84% | 70.95% | 81.32% | 69.88% | 82.79% | 100.3% | 100.5% | 77.88% | 72.68%
| 110 | CCV 1ppb TRACE | R | 06/03/2023 02:37:20 | 2231.5756 | 2061.9625 | 1.7574 | 1.6109 | 4848.718 | 1729.6854 | 1729.0406 | 1.5125 | 1.432 | 1.4234 | 1.0527 | 1.5528 | 1.5316 | 1.4302 | 1.4569 | 1.3318 | 1.2927 | 1.0688 | 1.1164 | 1.2217 | 0.9724 | 1.1521 | 1.4405 | 1.2776 | 1.1988 | 1.1568 | 1.2388 | 1.2345 | 1.0896 | 1.1011 | [7 empty] | 1.0047 | 1.0047 | 0.77226 | 0.78767
|  | RSD |  |  | 1.023% | 5.322% | 25.63% | 38.95% | 12.8% | 8.904% | 5.019% | 9.531% | 3.68% | 3.555% | 7.277% | 2.887% | 1.833% | 2.06% | 2.503% | 5.291% | 1.141% | 1.809% | 10.22% | 7.592% | 1.208% | 1.308% | 7.961% | 6.522% | 1.532% | 3.655% | 7.48% | 2.498% | 3.383% | 2.064% | [7 empty] | 3.322% | 1.461% | 4.874% | 2.72%
|  | % Recovery |  |  | 223.2% | 206.2% | 175.7% | 161.1% | 484.9% | 173% | 172.9% | 151.2% | 143.2% | 142.3% | 105.3% | 155.3% | 153.2% | 143% | 145.7% | 133.2% | 129.3% | 106.9% | 111.6% | 122.2% | 97.24% | 115.2% | 144% | 127.8% | 119.9% | 115.7% | 123.9% | 123.4% | 109% | 110.1% | 83.58% | 71.46% | 81.96% | 70.11% | 76.8% | 71.53% | 81.8% | 100.5% | 100.5% | 77.23% | 78.77%
| 125 | CCV 1ppb TRACE | R | 06/03/2023 04:09:35 | 2169.5302 | 2077.1071 | 1.8108 | 1.6366 | 4667.7849 | 1628.9883 | 1680.7396 | 1.4254 | 1.4072 | 1.3982 | 0.99482 | 1.5712 | 1.4668 | 1.4546 | 1.4535 | 1.2922 | 1.2846 | 1.0523 | 1.0545 | 1.1446 | 0.95661 | 1.1729 | 1.4275 | 1.1772 | 1.1383 | 1.1683 | 1.1016 | 0.86819 | 1.0974 | 1.132 | [7 empty] | 0.96462 | 0.97355 | 0.75007 | 0.77849
|  | RSD |  |  | 0.82% | 3.482% | 19.57% | 3.575% | 2.652% | 1.283% | 6.773% | 6.288% | 3.292% | 2.5% | 1.034% | 2.452% | 2.793% | 5.893% | 0.8813% | 0.715% | 2.603% | 2.989% | 5.178% | 2.059% | 1.766% | 0.4995% | 12.49% | 10.69% | 6.703% | 4.41% | 1.572% | 9.213% | 5.312% | 0.6206% | [7 empty] | 2.708% | 1.174% | 1.693% | 0.8479%
|  | % Recovery |  |  | 217% | 207.7% | 181.1% | 163.7% | 466.8% | 162.9% | 168.1% | 142.5% | 140.7% | 139.8% | 99.48% | 157.1% | 146.7% | 145.5% | 145.3% | 129.2% | 128.5% | 105.2% | 105.5% | 114.5% | 95.66% | 117.3% | 142.8% | 117.7% | 113.8% | 116.8% | 110.2% | 86.82% | 109.7% | 113.2% | 86.15% | 73.55% | 83.36% | 72.87% | 78.33% | 73.93% | 87.1% | 96.46% | 97.36% | 75.01% | 77.85%
| 126 | CCV 1ppb TRACE | R | 06/03/2023 04:16:07 | 2136.6886 | 1976.8154 | 1.7489 | 1.5025 | 4491.1948 | 1681.6968 | 1675.864 | 1.4986 | 1.4337 | 1.4723 | 1.1102 | 1.5598 | 1.4883 | 1.4067 | 1.3914 | 1.2773 | 1.3201 | 1.1024 | 1.0608 | 1.102 | 0.97129 | 1.1798 | 1.4626 | 1.045 | 1.139 | 1.1908 | 1.0564 | 1.2602 | 1.1131 | 1.1286 | [7 empty] | 0.99318 | 1.0264 | 0.75077 | 0.79293
|  | RSD |  |  | 11.46% | 10.72% | 14.03% | 1.845% | 19.84% | 11.46% | 17.48% | 13.92% | 12.7% | 3.525% | 12.56% | 1.424% | 21.12% | 13.49% | 9% | 11.95% | 4.876% | 4.437% | 1.643% | 25.8% | 3.439% | 4.757% | 23.84% | 6.899% | 0.4799% | 11.22% | 5.006% | 4.875% | 3.343% | 2.117% | [7 empty] | 3.606% | 4.251% | 5.683% | 1.473%
|  | % Recovery |  |  | 213.7% | 197.7% | 174.9% | 150.2% | 449.1% | 168.2% | 167.6% | 149.9% | 143.4% | 147.2% | 111% | 156% | 148.8% | 140.7% | 139.1% | 127.7% | 132% | 110.2% | 106.1% | 110.2% | 97.13% | 118% | 146.3% | 104.5% | 113.9% | 119.1% | 105.6% | 126% | 111.3% | 112.9% | 84.19% | 71.22% | 85.55% | 68.65% | 74.48% | 69.67% | 83.54% | 99.32% | 102.6% | 75.08% | 79.29%
| 142 | CCV 1ppb TRACE | R | 06/03/2023 05:53:46 | 2205.8349 | 2087.0744 | 1.5678 | 2.1631 | 4744.8351 | 1625.2449 | 1926.3046 | 1.4815 | 1.4635 | 1.4636 | 1.1267 | 1.6091 | 1.5448 | 1.4574 | 1.5002 | 1.325 | 1.3135 | 1.0794 | 1.115 | 0.99725 | 1.0021 | 1.2258 | 1.5632 | 1.2048 | 1.1649 | 1.1892 | 1.0693 | 1.1787 | 1.0959 | 1.1221 | [7 empty] | 0.99635 | 0.98862 | 0.77162 | 0.73371
|  | RSD |  |  | 3.549% | 6.082% | 35.65% | 2.878% | 17.12% | 1.63% | 5.382% | 4.88% | 2.566% | 3.759% | 9.931% | 0.8264% | 1.093% | 2.84% | 3.528% | 4.12% | 6.474% | 3.446% | 3.651% | 23.42% | 4.97% | 3.216% | 11.97% | 11.57% | 8.034% | 7.816% | 8.758% | 11.88% | 6.21% | 1.298% | [7 empty] | 2.573% | 5.656% | 1.998% | 4.935%
|  | % Recovery |  |  | 220.6% | 208.7% | 156.8% | 216.3% | 474.5% | 162.5% | 192.6% | 148.2% | 146.3% | 146.4% | 112.7% | 160.9% | 154.5% | 145.7% | 150% | 132.5% | 131.3% | 107.9% | 111.5% | 99.72% | 100.2% | 122.6% | 156.3% | 120.5% | 116.5% | 118.9% | 106.9% | 117.9% | 109.6% | 112.2% | 80.03% | 67.55% | 79% | 65.49% | 76.95% | 66.62% | 79.85% | 99.64% | 98.86% | 77.16% | 73.37%
| 143 | CCV 1ppb TRACE | R | 06/03/2023 06:00:18 | 2164.876 | 2043.9683 | 1.7483 | 1.7014 | 4810.6832 | 1730.5175 | 1740.5768 | 1.4799 | 1.4132 | 1.4158 | 1.0982 | 1.6222 | 1.5524 | 1.4099 | 1.4921 | 1.2996 | 1.2976 | 1.0626 | 1.0805 | 1.1286 | 1.0301 | 1.2456 | 1.5338 | 1.2461 | 1.22 | 1.1741 | 1.0908 | 1.1735 | 1.0968 | 1.1693 | [7 empty] | 0.9788 | 0.96602 | 0.74461 | 0.78011
|  | RSD |  |  | 5.798% | 3.995% | 6.357% | 4.009% | 7.222% | 4.101% | 8.477% | 5.29% | 2.46% | 1.034% | 5.74% | 5.78% | 5.3% | 1.839% | 4.952% | 2.527% | 0.7814% | 2.838% | 7.915% | 7.688% | 5.425% | 4.996% | 24.8% | 7.429% | 2.635% | 2.889% | 4.758% | 6.808% | 0.8545% | 3.136% | [7 empty] | 0.8395% | 0.8334% | 8.722% | 0.594%
|  | % Recovery |  |  | 216.5% | 204.4% | 174.8% | 170.1% | 481.1% | 173.1% | 174.1% | 148% | 141.3% | 141.6% | 109.8% | 162.2% | 155.2% | 141% | 149.2% | 130% | 129.8% | 106.3% | 108% | 112.9% | 103% | 124.6% | 153.4% | 124.6% | 122% | 117.4% | 109.1% | 117.4% | 109.7% | 116.9% | 80.28% | 66.07% | 80.44% | 65.78% | 71.02% | 66.11% | 77.17% | 97.88% | 96.6% | 74.46% | 78.01%
| 157 | CCV 1ppb TRACE | R | 06/03/2023 07:25:23 | 2228.141 | 2157.1104 | 1.5024 | 2.0579 | 4498.6333 | 1689.8757 | 1771.6199 | 1.4935 | 1.431 | 1.4461 | 1.0944 | 1.5717 | 1.5615 | 1.4555 | 1.4413 | 1.3101 | 1.307 | 1.0783 | 1.068 | 1.1073 | 1.0091 | 1.202 | 1.561 | 1.1841 | 1.2241 | 1.156 | 1.0996 | 1.1393 | 1.0875 | 1.1248 | [7 empty] | 0.97166 | 0.98536 | 0.77275 | 0.79166
|  | RSD |  |  | 6.661% | 4.622% | 17.55% | 2.424% | 15.74% | 3.409% | 8.538% | 7.455% | 1.862% | 0.5696% | 6.957% | 1.409% | 7.348% | 3.99% | 6.512% | 5.407% | 0.6342% | 2.083% | 3.042% | 11.28% | 1.861% | 2.106% | 11.55% | 22.06% | 16.14% | 5.461% | 1.799% | 13.5% | 1.538% | 3.377% | [7 empty] | 0.8452% | 2.203% | 0.6902% | 2.269%
|  | % Recovery |  |  | 222.8% | 215.7% | 150.2% | 205.8% | 449.9% | 169% | 177.2% | 149.3% | 143.1% | 144.6% | 109.4% | 157.2% | 156.2% | 145.5% | 144.1% | 131% | 130.7% | 107.8% | 106.8% | 110.7% | 100.9% | 120.2% | 156.1% | 118.4% | 122.4% | 115.6% | 110% | 113.9% | 108.8% | 112.5% | 78.79% | 67.83% | 78.94% | 65.96% | 71.04% | 64.86% | 78.05% | 97.17% | 98.54% | 77.28% | 79.17%
| 158 | CCV 1ppb TRACE | R | 06/03/2023 07:31:56 | 2259.4002 | 2060.9943 | 1.9346 | 1.6836 | 5002.6423 | 1707.8261 | 1797.7708 | 1.5021 | 1.4091 | 1.6226 | 1.287 | 1.5681 | 1.533 | 1.4472 | 1.5517 | 1.3739 | 1.4609 | 1.2397 | 1.0986 | 1.1559 | 0.9536 | 1.1959 | 1.2941 | 1.2704 | 1.1499 | 1.1814 | 1.1382 | 1.226 | 1.2247 | 1.1593 | [7 empty] | 1.1178 | 1.1136 | 0.75428 | 0.78438
|  | RSD |  |  | 1.293% | 4.351% | 14.63% | 3.113% | 5.675% | 3.013% | 5.161% | 0.851% | 2.701% | 16.17% | 25.65% | 14.09% | 5.031% | 2.098% | 0.08961% | 4.342% | 16.13% | 20.06% | 3.756% | 14.52% | 17.2% | 17.55% | 8.201% | 31.28% | 12.63% | 22.52% | 10.22% | 22.82% | 14.87% | 15.21% | [7 empty] | 17.6% | 16.73% | 1.3% | 0.9839%
|  | % Recovery |  |  | 225.9% | 206.1% | 193.5% | 168.4% | 500.3% | 170.8% | 179.8% | 150.2% | 140.9% | 162.3% | 128.7% | 156.8% | 153.3% | 144.7% | 155.2% | 137.4% | 146.1% | 124% | 109.9% | 115.6% | 95.36% | 119.6% | 129.4% | 127% | 115% | 118.1% | 113.8% | 122.6% | 122.5% | 115.9% | 78.56% | 66.67% | 78.3% | 59.14% | 68.93% | 65.16% | 80.67% | 111.8% | 111.4% | 75.43% | 78.44%
QC STD 4
|  | Sample Id | R | Acquisition Time | Na KED 23/23 | Mg KED 24/24 | Al KED 27/27 | P Oshift 31/47 | S (34) -1 34/34 | K KED 39/39 | Ca KED 43/43 | V KED 51/51 | Cr (52) 52/52 | Cr (53) NH3 53/53 | Fe (54) NH3 54/54 | Mn Oshift 55/71 | Fe (56) KED 56/56 | Co KED 59/59 | Ni KED 60/60 | Cu (63) KED 63/63 | Cu (65) NH3 65/65 | Zn NH3 66/66 | As 75/75 | As KED 75/75 | As RH Oshift 75/91 | As Oshift 75/91 | Se (77) 77/77 | Se (77) Oshift 77/93 | Se (78) Oshift 78/94 | Se (80) Oshift 80/96 | Se (82) 82/82 | Se (82) Oshift 82/98 | Mo (95) NH3 95/95 | Mo (95) OShift 95/127 | Rh (IS1) 103/103 (IS) | Rh (IS2) 103/103 (IS) | Rh (IS3) 103/103 (IS) | Rh (IS4) 103/103 (IS) | Rh (IS5) 103/103 (IS) | Rh (IS6) 103/103 (IS) | Rh OShift (IS1) 103/119 (IS) | Cd (111) NH3 111/111 | Cd (114) NH3 114/114 | Pb 208/208 | Pb NH3 208/208
| 17 | LFB 10ppb | R | 06/02/2023 16:48:46 | 11220.7527 | 10584.4469 | 10.9725 | 11.6099 | 11646.8755 | 10407.1031 | 11307.4482 | 10.388 | 10.1814 | 10.2906 | 11.2461 | 11.0973 | 11.538 | 10.9081 | 10.9852 | 9.9463 | 10.1897 | 13.9928 | 10.5608 | 9.573 | 9.0555 | 10.448 | 10.8077 | 11.0004 | 10.7592 | 10.5684 | 10.5406 | 10.893 | 10.6068 | 10.5799 | [7 empty] | 9.825 | 9.718 | 8.1616 | 8.3769
|  | RSD |  |  | 0.7392% | 1.685% | 3.844% | 0.7648% | 2.846% | 2.293% | 0.2622% | 2.239% | 2.804% | 0.185% | 1.969% | 0.6288% | 2.272% | 2.108% | 2.449% | 1.816% | 0.722% | 0.9904% | 3.154% | 3.77% | 4.243% | 1.04% | 1.371% | 1.242% | 1.498% | 2.154% | 3.276% | 1.469% | 0.5991% | 2.543% | [7 empty] | 1.832% | 0.5359% | 0.8176% | 1.79%
|  | % Recovery |  |  | 112.2% | 105.8% | 109.7% | 116.1% | 116.5% | 104.1% | 113.1% | 103.9% | 101.8% | 102.9% | 112.5% | 111% | 115.4% | 109.1% | 109.9% | 99.46% | 101.9% | 139.9% | 105.6% | 95.73% | 90.55% | 104.5% | 108.1% | 110% | 107.6% | 105.7% | 105.4% | 108.9% | 106.1% | 105.8% | 113.1% | 96.53% | 112.7% | 100.5% | 102.6% | 97.42% | 107.8% | 98.25% | 97.18% | 81.62% | 83.77%
| 18 | LFB 10ppb | R | 06/02/2023 16:55:16 | 11164.1162 | 10847.4002 | 12.22 | 12.0531 | 11713.2675 | 10386.4177 | 11380.2089 | 10.5083 | 10.3136 | 10.1861 | 11.2148 | 11.2524 | 11.5015 | 10.9732 | 11.1528 | 10.0312 | 10.2523 | 14.0507 | 10.3909 | 10.2885 | 9.1395 | 10.3933 | 11.1021 | 11.0533 | 10.7623 | 10.7371 | 10.3819 | 10.6638 | 10.5372 | 10.6573 | [7 empty] | 9.7462 | 9.649 | 8.0232 | 8.3515
|  | RSD |  |  | 2.253% | 1.555% | 1.829% | 2.581% | 3.572% | 1.864% | 1.568% | 0.5776% | 1.619% | 0.5383% | 1.584% | 0.825% | 1.659% | 0.949% | 0.0717% | 1.904% | 0.9943% | 0.1316% | 1.15% | 2.004% | 1.937% | 1.137% | 2.95% | 4.151% | 2.6% | 1.173% | 6.293% | 4.086% | 0.9595% | 1.106% | [7 empty] | 0.2347% | 0.6711% | 0.7907% | 1.882%
|  | % Recovery |  |  | 111.6% | 108.5% | 122.2% | 120.5% | 117.1% | 103.9% | 113.8% | 105.1% | 103.1% | 101.9% | 112.1% | 112.5% | 115% | 109.7% | 111.5% | 100.3% | 102.5% | 140.5% | 103.9% | 102.9% | 91.4% | 103.9% | 111% | 110.5% | 107.6% | 107.4% | 103.8% | 106.6% | 105.4% | 106.6% | 112.1% | 96.18% | 111.4% | 98.79% | 100.8% | 96.04% | 105.8% | 97.46% | 96.49% | 80.23% | 83.51%
| 62 | LFB 10ppb | R | 06/02/2023 21:41:03 | 16963.3012 | 15798.9625 | 15.1889 | 16.611 | 16494.2433 | 14309.8878 | 14470.7536 | 12.6252 | 12.1419 | 12.6794 | 13.6782 | 12.2563 | 13.6614 | 12.7325 | 12.9815 | 11.0887 | 11.6226 | 15.8761 | 10.6027 | 10.3309 | 8.3066 | 10.2467 | 11.8793 | 10.5443 | 10.9292 | 10.454 | 10.3203 | 10.8775 | 10.9785 | 10.1044 | [7 empty] | 9.7768 | 9.6408 | 7.1442 | 7.7084
|  | RSD |  |  | 1.917% | 2.937% | 5.502% | 0.9751% | 6.072% | 4.236% | 1.3% | 2.58% | 0.9721% | 0.8373% | 4.427% | 16.94% | 2.59% | 2.418% | 3.182% | 1.699% | 1.812% | 2.239% | 6.725% | 2.288% | 16.24% | 15.77% | 7.102% | 13.42% | 14.79% | 15.43% | 6.719% | 13.68% | 1.264% | 17.31% | [7 empty] | 2.455% | 1.38% | 5.791% | 1.497%
|  | % Recovery |  |  | 169.6% | 158% | 151.9% | 166.1% | 164.9% | 143.1% | 144.7% | 126.3% | 121.4% | 126.8% | 136.8% | 122.6% | 136.6% | 127.3% | 129.8% | 110.9% | 116.2% | 158.8% | 106% | 103.3% | 83.07% | 102.5% | 118.8% | 105.4% | 109.3% | 104.5% | 103.2% | 108.8% | 109.8% | 101% | 91.41% | 82.7% | 92.07% | 75.91% | 80.66% | 76.02% | 98.71% | 97.77% | 96.41% | 71.44% | 77.08%
| 63 | LFB 10ppb | R | 06/02/2023 21:47:33 | 16616.0665 | 15582.4334 | 15.9918 | 17.8106 | 16644.221 | 14030.1672 | 14342.6585 | 12.6007 | 12.1688 | 12.405 | 13.4389 | 13.7796 | 13.7973 | 12.9494 | 13.1534 | 11.1804 | 11.3402 | 15.4052 | 11.2858 | 10.2763 | 9.6147 | 11.5632 | 12.5099 | 12.0041 | 11.6658 | 11.3745 | 11.5239 | 11.5838 | 10.7815 | 11.2615 | [7 empty] | 9.6619 | 9.5592 | 7.4358 | 7.6839
|  | RSD |  |  | 1.166% | 1.934% | 7.782% | 1.051% | 5.551% | 1.371% | 2.201% | 2.322% | 1.208% | 1.293% | 1.731% | 2.424% | 2.272% | 0.73% | 0.1395% | 1.025% | 1.842% | 0.9654% | 0.915% | 3.016% | 2.978% | 1.153% | 5.239% | 2.011% | 3.167% | 1.549% | 2.187% | 0.7447% | 0.4035% | 1.097% | [7 empty] | 0.8786% | 1.111% | 0.6742% | 2.21%
|  | % Recovery |  |  | 166.2% | 155.8% | 159.9% | 178.1% | 166.4% | 140.3% | 143.4% | 126% | 121.7% | 124% | 134.4% | 137.8% | 138% | 129.5% | 131.5% | 111.8% | 113.4% | 154.1% | 112.9% | 102.8% | 96.15% | 115.6% | 125.1% | 120% | 116.7% | 113.7% | 115.2% | 115.8% | 107.8% | 112.6% | 87.21% | 75.19% | 92.66% | 77.19% | 81.13% | 75.16% | 87.47% | 96.62% | 95.59% | 74.36% | 76.84%
| 94 | LFB 10ppb | R | 06/03/2023 00:57:57 | 22743.1014 | 20507.6626 | 19.6136 | 20.8277 | 20007.7152 | 17912.0399 | 17261.8613 | 14.2062 | 14.006 | 13.1488 | 14.7685 | 13.5946 | 15.4391 | 14.1674 | 14.0358 | 12.2584 | 11.6318 | 15.9831 | 11.2699 | 11.6136 | 9.2992 | 11.0607 | 12.7432 | 11.166 | 11.5117 | 11.0525 | 11.8479 | 11.436 | 10.8598 | 10.589 | [7 empty] | 9.7247 | 9.5267 | 6.8686 | 7.3993
|  | RSD |  |  | 6.479% | 6.773% | 10.44% | 0.6184% | 8.324% | 5.229% | 5.563% | 6.562% | 6.595% | 2.646% | 0.734% | 16.26% | 9.819% | 6.173% | 5.83% | 6.387% | 1.221% | 0.7609% | 1.977% | 10.19% | 15.19% | 17.44% | 0.3818% | 14.59% | 15.4% | 16.71% | 6.627% | 14.32% | 1.044% | 17.36% | [7 empty] | 1.089% | 0.4799% | 0.9497% | 0.6716%
|  | % Recovery |  |  | 227.4% | 205.1% | 196.1% | 208.3% | 200.1% | 179.1% | 172.6% | 142.1% | 140.1% | 131.5% | 147.7% | 135.9% | 154.4% | 141.7% | 140.4% | 122.6% | 116.3% | 159.8% | 112.7% | 116.1% | 92.99% | 110.6% | 127.4% | 111.7% | 115.1% | 110.5% | 118.5% | 114.4% | 108.6% | 105.9% | 82.05% | 72.71% | 74.97% | 68.56% | 73.14% | 66.51% | 83.18% | 97.25% | 95.27% | 68.69% | 73.99%
| 95 | LFB 10ppb | R | 06/03/2023 01:04:27 | 22121.7484 | 19529.2105 | 17.9986 | 20.2756 | 19349.5194 | 17860.4334 | 16677.1287 | 13.5251 | 13.2079 | 12.7902 | 14.0186 | 14.3222 | 14.4788 | 13.5183 | 13.5772 | 11.7374 | 11.265 | 15.6674 | 11.4809 | 10.6821 | 10.2316 | 11.5979 | 12.5989 | 11.5237 | 11.9089 | 11.5022 | 11.783 | 11.4808 | 10.5915 | 11.0736 | [7 empty] | 9.3144 | 9.2147 | 6.9899 | 7.3386
|  | RSD |  |  | 3.226% | 4.909% | 5.834% | 1.612% | 4.988% | 2.604% | 4.783% | 4.938% | 4.362% | 1.094% | 2.212% | 0.7559% | 2.498% | 3.242% | 1.443% | 4.107% | 1.085% | 1.055% | 1.501% | 1.933% | 1.978% | 1.98% | 3.624% | 6.41% | 3.781% | 0.6603% | 1.488% | 6.422% | 1.074% | 1.513% | [7 empty] | 1.332% | 1.483% | 1.531% | 1.077%
|  | % Recovery |  |  | 221.2% | 195.3% | 180% | 202.8% | 193.5% | 178.6% | 166.8% | 135.3% | 132.1% | 127.9% | 140.2% | 143.2% | 144.8% | 135.2% | 135.8% | 117.4% | 112.7% | 156.7% | 114.8% | 106.8% | 102.3% | 116% | 126% | 115.2% | 119.1% | 115% | 117.8% | 114.8% | 105.9% | 110.7% | 81.02% | 68.65% | 78.53% | 71.46% | 73.5% | 68.01% | 75.24% | 93.14% | 92.15% | 69.9% | 73.39%
| 127 | LFB 10ppb | R | 06/03/2023 04:22:09 | 22092.3553 | 19429.9035 | 18.5431 | 21.0107 | 19544.9564 | 17317.922 | 16680.2598 | 14.0329 | 13.5186 | 13.3624 | 14.7285 | 14.641 | 14.6787 | 13.5133 | 13.7731 | 11.6876 | 11.7126 | 16.2393 | 11.5546 | 10.5749 | 9.6165 | 11.6384 | 12.4886 | 11.9852 | 11.7756 | 11.5678 | 11.6554 | 11.2988 | 10.8439 | 11.0488 | [7 empty] | 9.4908 | 9.4459 | 6.5651 | 7.3977
|  | RSD |  |  | 1.532% | 2.704% | 5.621% | 1.622% | 0.7128% | 3.042% | 1.268% | 1.101% | 2.181% | 0.8263% | 0.69% | 1.434% | 1.483% | 2.019% | 2.898% | 0.5611% | 0.9446% | 1.948% | 1.518% | 5.708% | 1.903% | 1.63% | 2.445% | 5.341% | 3.224% | 0.7633% | 2.007% | 1.244% | 1.879% | 2.143% | [7 empty] | 1.547% | 1.357% | 0.4232% | 5.574%
|  | % Recovery |  |  | 220.9% | 194.3% | 185.4% | 210.1% | 195.4% | 173.2% | 166.8% | 140.3% | 135.2% | 133.6% | 147.3% | 146.4% | 146.8% | 135.1% | 137.7% | 116.9% | 117.1% | 162.4% | 115.5% | 105.7% | 96.16% | 116.4% | 124.9% | 119.9% | 117.8% | 115.7% | 116.6% | 113% | 108.4% | 110.5% | 81.1% | 66.1% | 77.4% | 65.89% | 68.95% | 64.48% | 77.34% | 94.91% | 94.46% | 65.65% | 73.98%
| 128 | LFB 10ppb | R | 06/03/2023 04:28:39 | 21552.9202 | 18945.8218 | 19.5101 | 21.6908 | 19307.5939 | 16896.9813 | 16158.8223 | 13.8152 | 13.0991 | 13.9098 | 15.1718 | 15.1977 | 14.3208 | 13.4938 | 13.7069 | 11.6415 | 11.8928 | 16.344 | 11.5375 | 11.1034 | 10.0911 | 12.0182 | 12.8541 | 11.374 | 12.2237 | 12.0905 | 11.7723 | 11.6705 | 11.1487 | 11.5015 | [7 empty] | 9.7282 | 9.5247 | 6.7343 | 7.2543
|  | RSD |  |  | 2.41% | 1.113% | 5.777% | 0.9092% | 1.288% | 0.7453% | 5.049% | 1.052% | 1.281% | 3.475% | 5.484% | 2.945% | 0.4864% | 1.09% | 2.489% | 1.393% | 3.904% | 3.691% | 4.12% | 4.296% | 2.968% | 3.157% | 6.041% | 6.106% | 2.817% | 4.777% | 3.111% | 5.204% | 3.312% | 4.524% | [7 empty] | 3.61% | 3.13% | 1.838% | 1.404%
|  | % Recovery |  |  | 215.5% | 189.5% | 195.1% | 216.9% | 193.1% | 169% | 161.6% | 138.2% | 131% | 139.1% | 151.7% | 152% | 143.2% | 134.9% | 137.1% | 116.4% | 118.9% | 163.4% | 115.4% | 111% | 100.9% | 120.2% | 128.5% | 113.7% | 122.2% | 120.9% | 117.7% | 116.7% | 111.5% | 115% | 80.12% | 65.05% | 78.89% | 65.56% | 69.78% | 66.03% | 74.9% | 97.28% | 95.25% | 67.34% | 72.54%
| 159 | LFB 10ppb | R | 06/03/2023 07:37:58 | 22008.8046 | 19739.8604 | 20.3066 | 21.6092 | 19930.9088 | 17005.7439 | 16892.2286 | 14.1463 | 13.5856 | 13.7366 | 14.7745 | 14.3862 | 14.7431 | 13.7783 | 14.0832 | 11.8005 | 11.7028 | 16.2186 | 11.447 | 10.8957 | 8.6279 | 11.0052 | 12.7434 | 10.786 | 11.4947 | 11.0227 | 11.1616 | 11.435 | 10.8083 | 10.5312 | [7 empty] | 9.5094 | 9.2732 | 6.4994 | 7.2632
|  | RSD |  |  | 3.325% | 2.422% | 5.135% | 1.87% | 4.712% | 1.915% | 3.104% | 1.558% | 0.9391% | 1.066% | 1.81% | 17.08% | 2.209% | 0.7619% | 1.082% | 0.8892% | 0.7268% | 0.4717% | 2.463% | 4.256% | 20.53% | 18.94% | 1.76% | 23.01% | 14.35% | 16.81% | 3.82% | 18.48% | 2.256% | 18.8% | [7 empty] | 1.076% | 0.3317% | 1.202% | 1.687%
|  | % Recovery |  |  | 220.1% | 197.4% | 203.1% | 216.1% | 199.3% | 170.1% | 168.9% | 141.5% | 135.9% | 137.4% | 147.7% | 143.9% | 147.4% | 137.8% | 140.8% | 118% | 117% | 162.2% | 114.5% | 109% | 86.28% | 110.1% | 127.4% | 107.9% | 114.9% | 110.2% | 111.6% | 114.3% | 108.1% | 105.3% | 77.49% | 67.22% | 73.82% | 62.67% | 66.01% | 61.62% | 83.29% | 95.09% | 92.73% | 64.99% | 72.63%
| 160 | LFB 10ppb | R | 06/03/2023 07:44:29 | 19528.3549 | 17890.3419 | 16.5575 | 20.7182 | 18187.6028 | 15780.3209 | 15180.9207 | 12.8424 | 12.4285 | 0.67159 | 0.23309 | 14.8657 | 13.4224 | 12.3185 | 12.5409 | 10.5825 | 0.67244 | 0.54305 | 0.62992 | 10.071 | 9.4619 | 11.8757 | -1.5593 | 11.5906 | 12.5208 | 12.0519 | -0.14536 | 11.4872 | 0.59682 | 11.1602 | [7 empty] | 0.48101 | 0.46492 | -0.00044911 | -5.4767e-05
|  | RSD |  |  | 9.066% | 10.97% | 10.06% | 2.493% | 14.43% | 9.54% | 8.343% | 9.884% | 10.9% | 150.3% | 514.7% | 6.759% | 9.486% | 8.815% | 8.168% | 8.993% | 155% | 220.7% | 5.447% | 10.14% | 5.021% | 5.78% | 3.541% | 11.58% | 6.319% | 5.554% | 50.08% | 4.488% | 151.5% | 6.978% | [7 empty] | 151.3% | 152.6% | 132.2% | 176.8%
|  | % Recovery |  |  | 195.3% | 178.9% | 165.6% | 207.2% | 181.9% | 157.8% | 151.8% | 128.4% | 124.3% | 6.716% | 2.331% | 148.7% | 134.2% | 123.2% | 125.4% | 105.8% | 6.724% | 5.43% | 6.299% | 100.7% | 94.62% | 118.8% | -15.59% | 115.9% | 125.2% | 120.5% | -1.454% | 114.9% | 5.968% | 111.6% | 78.76% | 60.66% | 78.98% | 71.23% | 81.31% | 62.98% | 73.57% | 4.81% | 4.649% | -0.004491% | -0.0005477%
QC STD 5
|  | Sample Id | R | Acquisition Time | Na KED 23/23 | Mg KED 24/24 | Al KED 27/27 | P Oshift 31/47 | S (34) -1 34/34 | K KED 39/39 | Ca KED 43/43 | V KED 51/51 | Cr (52) 52/52 | Cr (53) NH3 53/53 | Fe (54) NH3 54/54 | Mn Oshift 55/71 | Fe (56) KED 56/56 | Co KED 59/59 | Ni KED 60/60 | Cu (63) KED 63/63 | Cu (65) NH3 65/65 | Zn NH3 66/66 | As 75/75 | As KED 75/75 | As RH Oshift 75/91 | As Oshift 75/91 | Se (77) 77/77 | Se (77) Oshift 77/93 | Se (78) Oshift 78/94 | Se (80) Oshift 80/96 | Se (82) 82/82 | Se (82) Oshift 82/98 | Mo (95) NH3 95/95 | Mo (95) OShift 95/127 | Rh (IS1) 103/103 (IS) | Rh (IS2) 103/103 (IS) | Rh (IS3) 103/103 (IS) | Rh (IS4) 103/103 (IS) | Rh (IS5) 103/103 (IS) | Rh (IS6) 103/103 (IS) | Rh OShift (IS1) 103/119 (IS) | Cd (111) NH3 111/111 | Cd (114) NH3 114/114 | Pb 208/208 | Pb NH3 208/208
| 19 | LLOQ 0.01 | R | 06/02/2023 17:01:19 | -138.0924 | -39.7596 | -0.58716 | 0.43129 | 941.3472 | -158.7711 | 14.3678 | 0.018883 | 0.016598 | 0.010949 | -0.46865 | 0.0075582 | -0.24595 | 0.012491 | 0.003325 | 0.014066 | 0.006141 | -0.13024 | -0.021182 | 0.0033634 | 0.0041725 | 0.0058062 | 0.0070255 | 0.041888 | -0.025776 | 0.0028406 | -0.017262 | 0.0021742 | 0.012904 | 0.014265 | [7 empty] | 0.0087397 | 0.010384 | 0.0079997 | 0.0087649
|  | RSD |  |  | 1.573% | 9.924% | 21.39% | 12.67% | 10.44% | 3.596% | 46.03% | 24.96% | 8.872% | 8.708% | 3.047% | 27.52% | 12.08% | 25.66% | 32.3% | 25.33% | 33.07% | 5.813% | 191.4% | 166.6% | 42.94% | 43.87% | 1657% | 31.53% | 44.21% | 221.5% | 167.6% | 303.7% | 7.168% | 9.21% | [7 empty] | 9.782% | 12.85% | 17.29% | 7.982%
|  | % Recovery |  |  | -138.1% | -39.76% | -5872% | 4313% | 941.3% | -158.8% | 14.37% | 188.8% | 166% | 109.5% | -4687% | 75.58% | -2460% | 124.9% | 33.25% | 140.7% | 61.41% | -1302% | -211.8% | 33.63% |  | 58.06% | 70.25% | 418.9% | -257.8% | 28.41% | -172.6% | 21.74% | 129% | 142.7% | 111.5% | 103.6% | 118.7% | 106.5% | 102.8% | 101.6% | 118.4% | 87.4% | 103.8% | 80% | 87.65%
QC STD 6
|  | Sample Id | R | Acquisition Time | Na KED 23/23 | Mg KED 24/24 | Al KED 27/27 | P Oshift 31/47 | S (34) -1 34/34 | K KED 39/39 | Ca KED 43/43 | V KED 51/51 | Cr (52) 52/52 | Cr (53) NH3 53/53 | Fe (54) NH3 54/54 | Mn Oshift 55/71 | Fe (56) KED 56/56 | Co KED 59/59 | Ni KED 60/60 | Cu (63) KED 63/63 | Cu (65) NH3 65/65 | Zn NH3 66/66 | As 75/75 | As KED 75/75 | As RH Oshift 75/91 | As Oshift 75/91 | Se (77) 77/77 | Se (77) Oshift 77/93 | Se (78) Oshift 78/94 | Se (80) Oshift 80/96 | Se (82) 82/82 | Se (82) Oshift 82/98 | Mo (95) NH3 95/95 | Mo (95) OShift 95/127 | Rh (IS1) 103/103 (IS) | Rh (IS2) 103/103 (IS) | Rh (IS3) 103/103 (IS) | Rh (IS4) 103/103 (IS) | Rh (IS5) 103/103 (IS) | Rh (IS6) 103/103 (IS) | Rh OShift (IS1) 103/119 (IS) | Cd (111) NH3 111/111 | Cd (114) NH3 114/114 | Pb 208/208 | Pb NH3 208/208
| 20 | CCV 10 ppb MAJORS | R | 06/02/2023 17:07:22 | 12218.3718 | 11786.596 | 11.7543 | 11.4738 | 12395.218 | 11100.4799 | 11922.9999 | 10.785 | 10.7097 | 10.5238 | 10.7419 | 11.6461 | 11.6607 | 11.419 | 11.4898 | 10.3056 | 10.3968 | 10.5241 | 10.2955 | 9.9808 | 9.2524 | 10.5532 | 11.0721 | 10.5724 | 10.8825 | 10.68 | 10.3157 | 10.6049 | 10.592 | 10.774 | [7 empty] | 9.7957 | 9.7371 | 8.1524 | 8.203
|  | RSD |  |  | 4.284% | 4.464% | 3.984% | 2.383% | 5.26% | 4.973% | 7.115% | 3.607% | 5.447% | 1.3% | 1.062% | 1.795% | 5.423% | 4.844% | 4.786% | 5.096% | 1.104% | 1.481% | 1.851% | 6.902% | 1.475% | 1.008% | 3.048% | 3.827% | 1.523% | 1.448% | 1.848% | 5.275% | 0.9162% | 0.8824% | [7 empty] | 0.8733% | 0.3795% | 0.6226% | 2.545%
|  | % Recovery |  |  | 122.2% | 117.9% | 117.5% | 114.7% | 124% | 111% | 119.2% | 107.9% | 107.1% | 105.2% | 107.4% | 116.5% | 116.6% | 114.2% | 114.9% | 103.1% | 104% | 105.2% | 103% | 99.81% | 92.52% | 105.5% | 110.7% | 105.7% | 108.8% | 106.8% | 103.2% | 106% | 105.9% | 107.7% | 111.4% | 94.39% | 108.3% | 96.93% | 102.5% | 94.97% | 104.1% | 97.96% | 97.37% | 81.52% | 82.03%
| 21 | CCV 10 ppb MAJORS | R | 06/02/2023 17:13:54 | 12155.8073 | 11534.2219 | 11.5892 | 10.6853 | 12260.8312 | 11190.5038 | 11962.7689 | 10.818 | 10.652 | 10.5538 | 10.8271 | 12.4547 | 11.3666 | 11.3338 | 11.4317 | 10.2051 | 10.4441 | 10.5546 | 10.4084 | 9.8099 | 9.3471 | 11.0214 | 11.0873 | 11.2832 | 11.4539 | 11.2855 | 10.6633 | 11.103 | 10.6018 | 11.2965 | [7 empty] | 9.762 | 9.6951 | 8.2167 | 8.5509
|  | RSD |  |  | 2.156% | 2.298% | 1.295% | 0.9557% | 3.22% | 1.411% | 1.326% | 1.685% | 1.97% | 1.009% | 2.242% | 7.045% | 1.555% | 2.894% | 0.8619% | 2.112% | 0.958% | 1.207% | 2.756% | 1.864% | 6.627% | 6.278% | 3.352% | 11.32% | 5.742% | 7.033% | 6.364% | 5.859% | 0.3593% | 6.58% | [7 empty] | 0.4712% | 1.547% | 0.7882% | 2.225%
|  | % Recovery |  |  | 121.6% | 115.3% | 115.9% | 106.9% | 122.6% | 111.9% | 119.6% | 108.2% | 106.5% | 105.5% | 108.3% | 124.5% | 113.7% | 113.3% | 114.3% | 102.1% | 104.4% | 105.5% | 104.1% | 98.1% | 93.47% | 110.2% | 110.9% | 112.8% | 114.5% | 112.9% | 106.6% | 111% | 106% | 113% | 111.1% | 91.07% | 111.4% | 97.56% | 98.76% | 99.55% | 103.9% | 97.62% | 96.95% | 82.17% | 85.51%
| 34 | CCV 10 ppb MAJORS | R | 06/02/2023 18:49:16 | 14116.8842 | 13633.8579 | 13.5979 | 13.1568 | 14424.6605 | 12415.4571 | 13460.7718 | 11.6073 | 11.3983 | 11.4099 | 11.7938 | 12.9671 | 12.2666 | 12.157 | 12.3103 | 10.5781 | 10.9908 | 11.1114 | 10.7506 | 9.7871 | 9.1837 | 10.9285 | 11.6208 | 11.2458 | 11.6391 | 11.1064 | 10.9319 | 10.9671 | 10.873 | 10.9608 | [7 empty] | 9.9046 | 9.7928 | 8.0981 | 8.4022
|  | RSD |  |  | 1.737% | 3.17% | 7.648% | 0.5549% | 0.7386% | 2.71% | 2.845% | 1.84% | 1.286% | 1.276% | 1.182% | 0.6145% | 0.8802% | 0.9013% | 1.576% | 0.2795% | 0.4121% | 1.457% | 1.95% | 4.528% | 3.449% | 0.4631% | 2.598% | 2.333% | 0.3943% | 1.402% | 0.5427% | 5.189% | 1.467% | 1.708% | [7 empty] | 0.9399% | 1.432% | 0.2888% | 1.552%
|  | % Recovery |  |  | 141.2% | 136.3% | 136% | 131.6% | 144.2% | 124.2% | 134.6% | 116.1% | 114% | 114.1% | 117.9% | 129.7% | 122.7% | 121.6% | 123.1% | 105.8% | 109.9% | 111.1% | 107.5% | 97.87% | 91.84% | 109.3% | 116.2% | 112.5% | 116.4% | 111.1% | 109.3% | 109.7% | 108.7% | 109.6% | 102.3% | 88.44% | 105.9% | 87.96% | 92.77% | 88.08% | 101.8% | 99.05% | 97.93% | 80.98% | 84.02%
| 35 | CCV 10 ppb MAJORS | R | 06/02/2023 18:55:47 | 13859.7015 | 13002.0386 | 13.0026 | 13.0325 | 13944.7964 | 12276.3185 | 12850.3773 | 11.6348 | 11.3822 | 11.3969 | 11.6512 | 12.8255 | 12.1949 | 12.0396 | 12.1705 | 10.6666 | 10.8653 | 11.14 | 10.8657 | 10.1504 | 9.2303 | 11.0926 | 11.7398 | 11.6659 | 11.5002 | 11.1047 | 10.9412 | 11.0898 | 10.877 | 11.1551 | [7 empty] | 9.8475 | 9.6894 | 8.1076 | 8.7508
|  | RSD |  |  | 0.2999% | 1.367% | 7.392% | 1.373% | 4.006% | 1.527% | 2.864% | 1.311% | 1.8% | 1.107% | 2.928% | 0.9147% | 1.356% | 0.6006% | 0.7316% | 1.242% | 1.541% | 2.81% | 1.074% | 2.834% | 1.372% | 0.9298% | 1.45% | 2.871% | 3.766% | 1.604% | 0.6972% | 1.417% | 0.9582% | 0.3421% | [7 empty] | 1.093% | 2.006% | 0.9765% | 0.7457%
|  | % Recovery |  |  | 138.6% | 130% | 130% | 130.3% | 139.4% | 122.8% | 128.5% | 116.3% | 113.8% | 114% | 116.5% | 128.3% | 121.9% | 120.4% | 121.7% | 106.7% | 108.7% | 111.4% | 108.7% | 101.5% | 92.3% | 110.9% | 117.4% | 116.7% | 115% | 111% | 109.4% | 110.9% | 108.8% | 111.6% | 101.1% | 87.41% | 104.4% | 88.41% | 88.42% | 85.66% | 101.6% | 98.48% | 96.89% | 81.08% | 87.51%
| 46 | CCV 10 ppb MAJORS | R | 06/02/2023 20:02:55 | 15136.6738 | 14616.0052 | 14.6761 | 14.3822 | 15742.228 | 13354.1633 | 13892.798 | 12.3561 | 11.9936 | 11.8475 | 12.4123 | 13.3335 | 12.9346 | 12.8082 | 12.7879 | 10.9096 | 11.0973 | 11.41 | 10.9634 | 10.3172 | 9.3036 | 11.1423 | 11.956 | 11.5902 | 11.4282 | 11.2403 | 11.051 | 11.2373 | 10.8506 | 11.0599 | [7 empty] | 9.8588 | 9.6914 | 7.7179 | 8.2375
|  | RSD |  |  | 1.302% | 1.351% | 3.989% | 1.226% | 4.691% | 2.161% | 3.511% | 2.01% | 1.835% | 0.8878% | 2.048% | 1.929% | 2.78% | 1.481% | 2.077% | 2.362% | 0.9373% | 0.6617% | 2.662% | 6.689% | 1.358% | 2.889% | 0.9506% | 6.159% | 2.982% | 1.234% | 2.479% | 5.54% | 1.561% | 2.027% | [7 empty] | 1.083% | 0.9206% | 0.816% | 0.452%
|  | % Recovery |  |  | 151.4% | 146.2% | 146.8% | 143.8% | 157.4% | 133.5% | 138.9% | 123.6% | 119.9% | 118.5% | 124.1% | 133.3% | 129.3% | 128.1% | 127.9% | 109.1% | 111% | 114.1% | 109.6% | 103.2% | 93.04% | 111.4% | 119.6% | 115.9% | 114.3% | 112.4% | 110.5% | 112.4% | 108.5% | 110.6% | 94.32% | 80.62% | 96.31% | 81.6% | 84.73% | 79.77% | 93.33% | 98.59% | 96.91% | 77.18% | 82.37%
| 47 | CCV 10 ppb MAJORS | R | 06/02/2023 20:09:27 | 14892.7634 | 14174.5086 | 14.4885 | 14.0731 | 14651.0299 | 13072.0905 | 13669.1522 | 12.2021 | 11.7414 | 12.0275 | 12.452 | 13.5502 | 12.469 | 12.641 | 12.5974 | 10.7941 | 11.3803 | 11.4364 | 10.8255 | 10.2835 | 9.3139 | 11.2431 | 12.1584 | 11.5709 | 11.4959 | 11.5765 | 11.0262 | 11.3981 | 11.0708 | 11.2378 | [7 empty] | 9.9257 | 9.7906 | 7.7776 | 8.7688
|  | RSD |  |  | 3.299% | 2.934% | 1.896% | 0.9745% | 2.857% | 2.03% | 4.508% | 2.443% | 2.558% | 2.789% | 2.605% | 1.13% | 3.951% | 2.111% | 2.145% | 3.881% | 1.949% | 2.088% | 3.611% | 5.001% | 1.994% | 0.408% | 3.525% | 3.651% | 1.388% | 1% | 4.417% | 4.983% | 1.667% | 0.7078% | [7 empty] | 0.3463% | 0.8332% | 0.2791% | 5.201%
|  | % Recovery |  |  | 148.9% | 141.7% | 144.9% | 140.7% | 146.5% | 130.7% | 136.7% | 122% | 117.4% | 120.3% | 124.5% | 135.5% | 124.7% | 126.4% | 126% | 107.9% | 113.8% | 114.4% | 108.3% | 102.8% | 93.14% | 112.4% | 121.6% | 115.7% | 115% | 115.8% | 110.3% | 114% | 110.7% | 112.4% | 95.17% | 81.16% | 99.65% | 81.42% | 81.38% | 81.55% | 94.74% | 99.26% | 97.91% | 77.78% | 87.69%
| 64 | CCV 10 ppb MAJORS | R | 06/02/2023 21:53:35 | 16818.1446 | 15942.7348 | 15.6709 | 16.672 | 17044.9012 | 14246.7741 | 14415.2684 | 12.6443 | 12.306 | 12.3118 | 12.7164 | 13.9061 | 13.1644 | 13.0024 | 13.286 | 11.1987 | 11.2807 | 11.6984 | 13.0962 | 10.9236 | 9.4231 | 11.477 | 14.5036 | 11.2376 | 11.6545 | 11.1099 | 13.1031 | 11.4886 | 10.7319 | 11.2272 | [7 empty] | 9.5234 | 9.3744 | 8.6335 | 7.8975
|  | RSD |  |  | 0.6188% | 1.303% | 11.53% | 1.589% | 3.904% | 2.333% | 4.199% | 0.6944% | 2.04% | 15.54% | 16.33% | 1.517% | 3.125% | 1.758% | 0.8815% | 0.2248% | 14.63% | 14.85% | 13.23% | 5.698% | 0.8613% | 1.763% | 15.07% | 6.935% | 3.474% | 2.731% | 11.59% | 4.361% | 16.45% | 1.003% | [7 empty] | 16.09% | 15.38% | 10.94% | 0.7986%
|  | % Recovery |  |  | 168.2% | 159.4% | 156.7% | 166.7% | 170.4% | 142.5% | 144.2% | 126.4% | 123.1% | 123.1% | 127.2% | 139.1% | 131.6% | 130% | 132.9% | 112% | 112.8% | 117% | 131% | 109.2% | 94.23% | 114.8% | 145% | 112.4% | 116.5% | 111.1% | 131% | 114.9% | 107.3% | 112.3% | 78.69% | 76.51% | 93.71% | 79.29% | 80.66% | 74.59% | 90.09% | 95.23% | 93.74% | 86.33% | 78.98%
| 65 | CCV 10 ppb MAJORS | R | 06/02/2023 22:00:06 | 16930.2435 | 15671.7282 | 15.1116 | 13.8925 | 16418.3915 | 14348.0151 | 14649.3146 | 12.7903 | 12.4661 | 13.4725 | 13.8084 | 13.8865 | 13.2379 | 13.1027 | 13.0618 | 11.236 | 12.3435 | 12.5048 | 11.359 | 10.6404 | 9.2898 | 11.5157 | 12.5503 | 11.7664 | 11.6988 | 11.6708 | 11.4076 | 11.2852 | 11.6897 | 11.2583 | [7 empty] | 10.358 | 10.1935 | 7.5404 | 7.9021
|  | RSD |  |  | 11.3% | 11.29% | 14.23% | 13.1% | 13.08% | 11.78% | 12.43% | 10.22% | 11.5% | 17.38% | 19.66% | 0.1295% | 12.76% | 12.46% | 13.25% | 12% | 16.31% | 16.63% | 3.626% | 13.17% | 1.069% | 0.3856% | 2.441% | 5.465% | 1.968% | 2.294% | 1.68% | 1.453% | 16.71% | 1.643% | [7 empty] | 15.62% | 16.48% | 2.197% | 1.147%
|  | % Recovery |  |  | 169.3% | 156.7% | 151.1% | 138.9% | 164.2% | 143.5% | 146.5% | 127.9% | 124.7% | 134.7% | 138.1% | 138.9% | 132.4% | 131% | 130.6% | 112.4% | 123.4% | 125% | 113.6% | 106.4% | 92.9% | 115.2% | 125.5% | 117.7% | 117% | 116.7% | 114.1% | 112.9% | 116.9% | 112.6% | 87.44% | 75.8% | 95.11% | 74.16% | 80.37% | 82.95% | 90.84% | 103.6% | 101.9% | 75.4% | 79.02%
| 79 | CCV 10 ppb MAJORS | R | 06/02/2023 23:25:38 | 21067.7647 | 18684.2381 | 17.5281 | 18.0676 | 18822.0334 | 17038.6265 | 16145.4684 | 13.2759 | 12.979 | 13.0011 | 13.6353 | 14.4695 | 13.6609 | 13.3696 | 13.6205 | 11.5073 | 11.6265 | 11.9894 | 11.4145 | 11.0874 | 9.5235 | 11.6226 | 12.2196 | 11.8034 | 12.3136 | 11.4598 | 11.5536 | 11.3066 | 10.8811 | 11.3095 | [7 empty] | 9.7467 | 9.6583 | 7.311 | 7.7987
|  | RSD |  |  | 0.5039% | 2.153% | 3.228% | 1.104% | 3.32% | 1.774% | 1.661% | 1.852% | 1.541% | 1.461% | 0.7219% | 2.159% | 0.8019% | 0.8111% | 0.2968% | 1.295% | 0.9056% | 3.463% | 0.9114% | 1.512% | 1.945% | 1.945% | 4.208% | 4.659% | 1.698% | 0.4512% | 2.583% | 1.664% | 1.805% | 1.909% | [7 empty] | 0.7551% | 0.7673% | 1.002% | 1.455%
|  | % Recovery |  |  | 210.7% | 186.8% | 175.3% | 180.7% | 188.2% | 170.4% | 161.5% | 132.8% | 129.8% | 130% | 136.4% | 144.7% | 136.6% | 133.7% | 136.2% | 115.1% | 116.3% | 119.9% | 114.1% | 110.9% | 95.23% | 116.2% | 122.2% | 118% | 123.1% | 114.6% | 115.5% | 113.1% | 108.8% | 113.1% | 81.94% | 69.55% | 80.87% | 71.19% | 74.5% | 68.15% | 82.08% | 97.47% | 96.58% | 73.11% | 77.99%
| 80 | CCV 10 ppb MAJORS | R | 06/02/2023 23:32:09 | 21074.6872 | 18835.8147 | 18.6471 | 17.3845 | 18858.5816 | 16776.0468 | 15757.751 | 13.4103 | 12.9431 | 13.2497 | 13.6733 | 14.7788 | 13.6744 | 13.4601 | 13.5352 | 11.6628 | 11.7513 | 12.0642 | 11.4824 | 11.0534 | 9.6273 | 11.7135 | 12.2649 | 11.513 | 11.8556 | 11.7302 | 12.0614 | 11.7759 | 11.1743 | 11.3448 | [7 empty] | 9.8131 | 9.8278 | 7.3186 | 7.5647
|  | RSD |  |  | 0.5259% | 1.006% | 10.83% | 3.008% | 1.991% | 0.7668% | 2.644% | 1.565% | 2.106% | 1.345% | 0.4222% | 1.644% | 2.431% | 1.557% | 1.505% | 1.059% | 0.5879% | 0.3183% | 1.517% | 3.251% | 1.055% | 1.853% | 2.096% | 7.284% | 1.296% | 1.91% | 4.276% | 4.817% | 0.3524% | 0.3% | [7 empty] | 1.544% | 1.244% | 0.8927% | 0.3412%
|  | % Recovery |  |  | 210.7% | 188.4% | 186.5% | 173.8% | 188.6% | 167.8% | 157.6% | 134.1% | 129.4% | 132.5% | 136.7% | 147.8% | 136.7% | 134.6% | 135.4% | 116.6% | 117.5% | 120.6% | 114.8% | 110.5% | 96.27% | 117.1% | 122.6% | 115.1% | 118.6% | 117.3% | 120.6% | 117.8% | 111.7% | 113.4% | 82.91% | 69.55% | 82.23% | 70.88% | 76.78% | 70.71% | 81.8% | 98.13% | 98.28% | 73.19% | 75.65%
| 96 | CCV 10 ppb MAJORS | R | 06/03/2023 01:10:29 | 22589.6947 | 20393.5272 | 17.3366 | 19.5679 | 19425.7924 | 17763.5482 | 16839.7548 | 14.0308 | 13.6793 | 13.7998 | 14.2817 | 15.1421 | 13.7263 | 13.6078 | 13.8226 | 11.7902 | 12.0951 | 12.5491 | 11.2899 | 10.8352 | 9.8758 | 11.9667 | 12.4697 | 12.7502 | 12.2478 | 11.7033 | 11.4765 | 12.1988 | 11.2995 | 11.581 | [7 empty] | 9.9651 | 9.826 | 7.0548 | 7.5393
|  | RSD |  |  | 2.624% | 3.234% | 3.879% | 3.1% | 3.534% | 5.123% | 5.181% | 3.604% | 2.152% | 2.706% | 5.683% | 0.2545% | 2.86% | 1.923% | 2.331% | 2.025% | 3.592% | 4.034% | 2.886% | 2.293% | 6.948% | 1.066% | 2.286% | 3.462% | 3.019% | 1.253% | 2.62% | 0.9666% | 1.44% | 0.8506% | [7 empty] | 2.72% | 2.504% | 1.391% | 0.5987%
|  | % Recovery |  |  | 225.9% | 203.9% | 173.4% | 195.7% | 194.3% | 177.6% | 168.4% | 140.3% | 136.8% | 138% | 142.8% | 151.4% | 137.3% | 136.1% | 138.2% | 117.9% | 121% | 125.5% | 112.9% | 108.4% | 98.76% | 119.7% | 124.7% | 127.5% | 122.5% | 117% | 114.8% | 122% | 113% | 115.8% | 82.47% | 67.23% | 78.38% | 69.01% | 74.26% | 68.27% | 78.97% | 99.65% | 98.26% | 70.55% | 75.39%
| 97 | CCV 10 ppb MAJORS | R | 06/03/2023 01:17:01 | 21768.7225 | 19295.4051 | 17.1835 | 19.0631 | 18877.5652 | 17566.8589 | 16012.6827 | 13.6025 | 13.0867 | 13.4579 | 13.9046 | 14.8392 | 13.4556 | 13.3476 | 13.6075 | 11.8217 | 11.7167 | 12.0015 | 12.1097 | 10.6211 | 10.2884 | 11.8442 | 13.8678 | 11.8764 | 11.8966 | 11.4785 | 11.9035 | 11.4277 | 10.8188 | 11.2882 | [7 empty] | 9.6592 | 9.5858 | 7.5939 | 7.5223
|  | RSD |  |  | 1.098% | 3.487% | 5.53% | 0.5864% | 1.797% | 2.294% | 1.805% | 3.304% | 2.421% | 2.26% | 1.377% | 1.18% | 2.519% | 0.4378% | 0.3904% | 1.041% | 2.024% | 1.82% | 8.918% | 0.9244% | 3.257% | 1.484% | 9.366% | 3.707% | 4.173% | 2.946% | 7.817% | 5.04% | 2.602% | 0.4093% | [7 empty] | 1.87% | 3.021% | 7.997% | 1.179%
|  | % Recovery |  |  | 217.7% | 193% | 171.8% | 190.6% | 188.8% | 175.7% | 160.1% | 136% | 130.9% | 134.6% | 139% | 148.4% | 134.6% | 133.5% | 136.1% | 118.2% | 117.2% | 120% | 121.1% | 106.2% | 102.9% | 118.4% | 138.7% | 118.8% | 119% | 114.8% | 119% | 114.3% | 108.2% | 112.9% | 77.69% | 69.18% | 81% | 70.67% | 74.46% | 67.89% | 77.06% | 96.59% | 95.86% | 75.94% | 75.22%
| 111 | CCV 10 ppb MAJORS | R | 06/03/2023 02:43:23 | 24000.1796 | 21306.2792 | 19.2387 | 19.6868 | 21033.2194 | 19175.1195 | 18411.3094 | 14.7288 | 14.2458 | 13.7585 | 13.992 | 15.0344 | 14.5708 | 14.5704 | 14.6214 | 12.557 | 11.8956 | 12.138 | 11.3734 | 11.395 | 9.6021 | 11.6359 | 12.9356 | 11.9372 | 11.7948 | 11.4913 | 11.2424 | 11.6494 | 10.9499 | 11.233 | [7 empty] | 9.5604 | 9.4731 | 6.7632 | 7.5041
|  | RSD |  |  | 3.521% | 1.558% | 12.72% | 0.31% | 2.552% | 1.782% | 4.088% | 4.162% | 3.817% | 0.8657% | 1.042% | 3.068% | 1.401% | 4.096% | 2.311% | 3.634% | 0.7136% | 1.003% | 2.384% | 5.94% | 6.486% | 4.228% | 4.947% | 4.066% | 1.923% | 0.736% | 2.108% | 1.743% | 0.2951% | 3.179% | [7 empty] | 0.8994% | 0.4301% | 0.8668% | 0.7729%
|  | % Recovery |  |  | 240% | 213.1% | 192.4% | 196.9% | 210.3% | 191.8% | 184.1% | 147.3% | 142.5% | 137.6% | 139.9% | 150.3% | 145.7% | 145.7% | 146.2% | 125.6% | 119% | 121.4% | 113.7% | 114% | 96.02% | 116.4% | 129.4% | 119.4% | 117.9% | 114.9% | 112.4% | 116.5% | 109.5% | 112.3% | 82.56% | 68.5% | 75.92% | 67.74% | 71.86% | 66.85% | 80.37% | 95.6% | 94.73% | 67.63% | 75.04%
| 112 | CCV 10 ppb MAJORS | R | 06/03/2023 02:49:56 | 22860.7575 | 20634.5933 | 18.2955 | 19.3139 | 19326.5307 | 17904.7891 | 16952.1775 | 14.0346 | 13.5491 | 14.3133 | 14.6135 | 15.1383 | 13.9184 | 13.9297 | 13.828 | 11.9315 | 12.3968 | 12.5668 | 10.7862 | 10.649 | 10.0675 | 11.8767 | 11.6157 | 12.0555 | 12.0109 | 11.6614 | 10.6018 | 11.4443 | 11.4496 | 11.3598 | [7 empty] | 10.0635 | 9.9163 | 6.5032 | 7.4578
|  | RSD |  |  | 4.187% | 5.734% | 3.78% | 1.775% | 3.484% | 2.948% | 5.583% | 5.234% | 4.433% | 2.651% | 3.604% | 2.718% | 5.387% | 3.912% | 3.249% | 3.922% | 3.29% | 3.495% | 10.33% | 1.573% | 1.812% | 1.554% | 10.43% | 2.432% | 2.5% | 3.137% | 5.798% | 8% | 2.784% | 1.282% | [7 empty] | 4.619% | 2.197% | 10.49% | 0.8466%
|  | % Recovery |  |  | 228.6% | 206.3% | 183% | 193.1% | 193.3% | 179% | 169.5% | 140.3% | 135.5% | 143.1% | 146.1% | 151.4% | 139.2% | 139.3% | 138.3% | 119.3% | 124% | 125.7% | 107.9% | 106.5% | 100.7% | 118.8% | 116.2% | 120.6% | 120.1% | 116.6% | 106% | 114.4% | 114.5% | 113.6% | 86.8% | 68.61% | 80.76% | 66.92% | 73.18% | 67.76% | 78.26% | 100.6% | 99.16% | 65.03% | 74.58%
| 129 | CCV 10 ppb MAJORS | R | 06/03/2023 04:34:42 | 21093.3047 | 18897.1161 | 18.4399 | 20.276 | 19070.8214 | 16932.2362 | 15843.9077 | 13.5792 | 12.7624 | 14.4052 | 14.519 | 20.7953 | 13.133 | 13.0339 | 13.3477 | 11.3365 | 12.2772 | 12.5252 | 11.4074 | 10.3838 | 13.2541 | 16.3241 | 12.7961 | 16.5329 | 16.4322 | 15.7281 | 11.567 | 16.0732 | 11.3664 | 15.6033 | [7 empty] | 9.8015 | 9.7809 | 6.6986 | 7.4315
|  | RSD |  |  | 17.06% | 17.78% | 12.01% | 0.6695% | 19.36% | 15.46% | 17.94% | 16.05% | 16.74% | 3.114% | 1.364% | 14.14% | 18.74% | 17.84% | 17.48% | 15.59% | 2.199% | 2.163% | 3.628% | 14.9% | 16.24% | 13.97% | 2.727% | 10.53% | 12.14% | 11.74% | 4.725% | 11.87% | 1.189% | 13.28% | [7 empty] | 1.788% | 1.813% | 0.8238% | 0.2824%
|  | % Recovery |  |  | 210.9% | 189% | 184.4% | 202.8% | 190.7% | 169.3% | 158.4% | 135.8% | 127.6% | 144.1% | 145.2% | 208% | 131.3% | 130.3% | 133.5% | 113.4% | 122.8% | 125.3% | 114.1% | 103.8% | 132.5% | 163.2% | 128% | 165.3% | 164.3% | 157.3% | 115.7% | 160.7% | 113.7% | 156% | 80.85% | 51.42% | 83.32% | 65.69% | 70.42% | 65.61% | 61.51% | 98.01% | 97.81% | 66.99% | 74.31%
| 130 | CCV 10 ppb MAJORS | R | 06/03/2023 04:41:13 | 21761.6891 | 19380.299 | 18.1671 | 19.6343 | 19339.3999 | 17668.7693 | 16603.32 | 14.1108 | 13.3492 | 14.2487 | 14.9199 | 15.1752 | 13.8782 | 13.6736 | 13.8828 | 11.8883 | 12.3379 | 12.3305 | 11.4732 | 10.8783 | 9.7348 | 11.7961 | 12.6374 | 12.3362 | 11.7149 | 11.8037 | 11.4107 | 11.5964 | 11.3804 | 11.4001 | [7 empty] | 9.77 | 9.817 | 6.5307 | 7.9157
|  | RSD |  |  | 1.886% | 3.729% | 6.583% | 1.519% | 3.408% | 2.21% | 4.438% | 3.545% | 2.756% | 2.549% | 4.905% | 1.891% | 2.408% | 1.387% | 2.797% | 1.824% | 4.024% | 3.939% | 1.24% | 5.405% | 2.658% | 1.699% | 2.477% | 1.585% | 3.474% | 1.607% | 4.392% | 0.08698% | 2.917% | 1.429% | [7 empty] | 0.8087% | 2.53% | 0.7924% | 17.07%
|  | % Recovery |  |  | 217.6% | 193.8% | 181.7% | 196.3% | 193.4% | 176.7% | 166% | 141.1% | 133.5% | 142.5% | 149.2% | 151.8% | 138.8% | 136.7% | 138.8% | 118.9% | 123.4% | 123.3% | 114.7% | 108.8% | 97.35% | 118% | 126.4% | 123.4% | 117.1% | 118% | 114.1% | 116% | 113.8% | 114% | 82.84% | 66.19% | 79.11% | 66.44% | 68.24% | 64.95% | 77.56% | 97.7% | 98.17% | 65.31% | 79.16%
| 144 | CCV 10 ppb MAJORS | R | 06/03/2023 06:06:21 | 23421.8549 | 21366.2416 | 19.7732 | 20.4635 | 20739.4614 | 18765.838 | 17788.2699 | 14.8099 | 14.3279 | 14.0774 | 14.3105 | 15.2628 | 14.539 | 14.6014 | 14.6472 | 12.3025 | 11.9855 | 12.3486 | 11.8731 | 11.5123 | 9.8349 | 12.0184 | 12.9632 | 11.7748 | 12.1936 | 11.8886 | 11.966 | 11.3574 | 10.997 | 11.4155 | [7 empty] | 9.6105 | 9.5282 | 6.9007 | 7.4512
|  | RSD |  |  | 1.753% | 3.469% | 3.909% | 0.4774% | 2.865% | 1.867% | 2.547% | 0.6524% | 1.112% | 1.267% | 1.062% | 1.464% | 1.067% | 2.983% | 2.914% | 1.602% | 0.531% | 1.444% | 2.364% | 3.024% | 3.875% | 1.96% | 5.14% | 6.383% | 1.951% | 1.729% | 2.465% | 4.278% | 0.821% | 2.333% | [7 empty] | 1.577% | 0.6331% | 4.737% | 0.3781%
|  | % Recovery |  |  | 234.2% | 213.7% | 197.7% | 204.6% | 207.4% | 187.7% | 177.9% | 148.1% | 143.3% | 140.8% | 143.1% | 152.6% | 145.4% | 146% | 146.5% | 123% | 119.9% | 123.5% | 118.7% | 115.1% | 98.35% | 120.2% | 129.6% | 117.7% | 121.9% | 118.9% | 119.7% | 113.6% | 110% | 114.2% | 76.69% | 63.08% | 72.91% | 62.65% | 67.06% | 61.8% | 74.57% | 96.1% | 95.28% | 69.01% | 74.51%
| 145 | CCV 10 ppb MAJORS | R | 06/03/2023 06:12:53 | 22335.7108 | 20007.3012 | 19.9551 | 19.6122 | 20925.1471 | 17941.292 | 17143.2219 | 14.2424 | 13.8579 | 14.0292 | 14.0306 | 15.5376 | 14.3482 | 13.928 | 14.1099 | 12.1434 | 12.1387 | 12.3662 | 11.3409 | 11.0035 | 10.1228 | 11.8766 | 12.6111 | 12.3797 | 12.3669 | 11.7063 | 11.0055 | 11.4094 | 11.055 | 11.3472 | [7 empty] | 9.736 | 9.5107 | 6.6961 | 7.4202
|  | RSD |  |  | 3.199% | 2.121% | 6.239% | 4.136% | 4.862% | 0.5411% | 0.9494% | 2.337% | 1.839% | 2.444% | 1.926% | 2.813% | 1.297% | 2.308% | 2.161% | 1.688% | 0.8851% | 0.2873% | 2.187% | 1.007% | 1.572% | 1.955% | 2.708% | 7.177% | 1.1% | 1.8% | 0.3137% | 3.387% | 0.8533% | 1.453% | [7 empty] | 1.003% | 1.227% | 1.18% | 0.5548%
|  | % Recovery |  |  | 223.4% | 200.1% | 199.6% | 196.1% | 209.3% | 179.4% | 171.4% | 142.4% | 138.6% | 140.3% | 140.3% | 155.4% | 143.5% | 139.3% | 141.1% | 121.4% | 121.4% | 123.7% | 113.4% | 110% | 101.2% | 118.8% | 126.1% | 123.8% | 123.7% | 117.1% | 110.1% | 114.1% | 110.6% | 113.5% | 79.46% | 64.21% | 75.62% | 63.64% | 67.53% | 65.41% | 72.84% | 97.36% | 95.11% | 66.96% | 74.2%
| 161 | CCV 10 ppb MAJORS | R | 06/03/2023 07:50:31 | 22358.4488 | 20023.5587 | 18.7649 | 21.1372 | 20285.393 | 17789.0904 | 16721.3055 | 14.3841 | 13.8884 | 14.4481 | 14.6342 | 15.4213 | 14.215 | 14.0465 | 13.8967 | 12.0496 | 12.2425 | 12.5482 | 13.6506 | 10.808 | 9.6335 | 12.0002 | 15.1006 | 12.1458 | 12.0296 | 11.7464 | 13.5256 | 11.2244 | 11.2821 | 11.4594 | [7 empty] | 9.7471 | 9.6685 | 7.8487 | 8.1409
|  | RSD |  |  | 2.728% | 2.435% | 2.871% | 0.5512% | 1.203% | 0.8872% | 2.329% | 2.323% | 1.701% | 3.105% | 2.051% | 1.468% | 3.434% | 1.156% | 3.434% | 0.6154% | 3.211% | 1.705% | 14.42% | 2.932% | 4.248% | 1.094% | 18.05% | 4.241% | 1.327% | 0.6735% | 14.53% | 2.254% | 2.59% | 2.01% | [7 empty] | 2.311% | 0.1801% | 15.23% | 1.535%
|  | % Recovery |  |  | 223.6% | 200.2% | 187.6% | 211.4% | 202.9% | 177.9% | 167.2% | 143.8% | 138.9% | 144.5% | 146.3% | 154.2% | 142.2% | 140.5% | 139% | 120.5% | 122.4% | 125.5% | 136.5% | 108.1% | 96.33% | 120% | 151% | 121.5% | 120.3% | 117.5% | 135.3% | 112.2% | 112.8% | 114.6% | 67.95% | 63.43% | 73.75% | 62.01% | 61.54% | 61.52% | 76.51% | 97.47% | 96.68% | 78.49% | 81.41%
| 162 | CCV 10 ppb MAJORS | R | 06/03/2023 07:57:03 | 22680.5735 | 20310.9018 | 18.1759 | 20.0575 | 20410.5381 | 18361.8529 | 17484.8905 | 14.8212 | 13.8826 | 14.2275 | 14.4131 | 13.3342 | 14.4153 | 14.1999 | 14.2842 | 12.3249 | 12.066 | 12.2722 | 11.0053 | 11.796 | 8.5796 | 10.4726 | 11.8505 | 10.6998 | 10.5519 | 10.2262 | 10.8794 | 10.3456 | 11.1253 | 10.0612 | [7 empty] | 9.6075 | 9.4285 | 6.2012 | 8.1981
|  | RSD |  |  | 5.536% | 7.793% | 9.761% | 2.522% | 7.306% | 5.908% | 5.716% | 5.82% | 5.847% | 0.9588% | 0.8408% | 9.362% | 4.207% | 3.824% | 4.363% | 3.21% | 0.2883% | 0.4082% | 6.592% | 1.38% | 8.013% | 10.3% | 7.777% | 14.4% | 6.294% | 8.512% | 10.22% | 8.431% | 1.174% | 9.647% | [7 empty] | 1.159% | 1.336% | 7.418% | 18.6%
|  | % Recovery |  |  | 226.8% | 203.1% | 181.8% | 200.6% | 204.1% | 183.6% | 174.8% | 148.2% | 138.8% | 142.3% | 144.1% | 133.3% | 144.2% | 142% | 142.8% | 123.2% | 120.7% | 122.7% | 110.1% | 118% | 85.8% | 104.7% | 118.5% | 107% | 105.5% | 102.3% | 108.8% | 103.5% | 111.3% | 100.6% | 82.69% | 71.94% | 73.64% | 63.98% | 61.26% | 62.79% | 84.7% | 96.08% | 94.28% | 62.01% | 81.98%
QC STD 8
|  | Sample Id | R | Acquisition Time | Na KED 23/23 | Mg KED 24/24 | Al KED 27/27 | P Oshift 31/47 | S (34) -1 34/34 | K KED 39/39 | Ca KED 43/43 | V KED 51/51 | Cr (52) 52/52 | Cr (53) NH3 53/53 | Fe (54) NH3 54/54 | Mn Oshift 55/71 | Fe (56) KED 56/56 | Co KED 59/59 | Ni KED 60/60 | Cu (63) KED 63/63 | Cu (65) NH3 65/65 | Zn NH3 66/66 | As 75/75 | As KED 75/75 | As RH Oshift 75/91 | As Oshift 75/91 | Se (77) 77/77 | Se (77) Oshift 77/93 | Se (78) Oshift 78/94 | Se (80) Oshift 80/96 | Se (82) 82/82 | Se (82) Oshift 82/98 | Mo (95) NH3 95/95 | Mo (95) OShift 95/127 | Rh (IS1) 103/103 (IS) | Rh (IS2) 103/103 (IS) | Rh (IS3) 103/103 (IS) | Rh (IS4) 103/103 (IS) | Rh (IS5) 103/103 (IS) | Rh (IS6) 103/103 (IS) | Rh OShift (IS1) 103/119 (IS) | Cd (111) NH3 111/111 | Cd (114) NH3 114/114 | Pb 208/208 | Pb NH3 208/208
| 22 | LLOQ 0.1 | R | 06/02/2023 17:19:57 | -21.9855 | 50.9662 | -0.41297 | 0.28369 | 1215.184 | -46.7358 | 125.3788 | 0.10957 | 0.11069 | 0.098192 | -0.2392 | 0.11878 | -0.0053462 | 0.11058 | 0.11615 | 0.10939 | 0.10227 | -0.010861 | 0.078753 | 0.095414 | 0.081962 | 0.09817 | 0.02152 | 0.13285 | 0.071288 | 0.10055 | 0.084762 | 0.078256 | 0.090107 | 0.098494 | [7 empty] | 0.090996 | 0.0939 | 0.084129 | 0.087
|  | RSD |  |  | 81.77% | 15.1% | 20.71% | 17.95% | 14.47% | 42.91% | 21.02% | 15.81% | 8.459% | 14.22% | 15.23% | 4.229% | 1720% | 2.256% | 17.79% | 2.909% | 15.08% | 295.9% | 64.16% | 20.07% | 7.595% | 6.002% | 760.6% | 65.96% | 16.73% | 1.758% | 60.5% | 23.49% | 14.11% | 10.74% | [7 empty] | 13.62% | 9.962% | 9.904% | 2.003%
|  | % Recovery |  |  | -21.99% | 50.97% | -413% | 283.7% | 1215% | -46.74% | 125.4% | 109.6% | 110.7% | 98.19% | -239.2% | 118.8% | -5.346% | 110.6% | 116.1% | 109.4% | 102.3% | -10.86% | 78.75% | 95.41% | 81.96% | 98.17% | 21.52% | 132.9% | 71.29% | 100.5% | 84.76% | 78.26% | 90.11% | 98.49% | 120.2% | 103.2% | 117.8% | 110.5% | 108.8% | 100.8% | 118.5% | 91% | 93.9% | 84.13% | 87%
QC Spikes (% Recovery)
|  | Sample Id | R | Acquisition Time | Na KED 23/23 | Mg KED 24/24 | Al KED 27/27 | P Oshift 31/47 | S (34) -1 34/34 | K KED 39/39 | Ca KED 43/43 | V KED 51/51 | Cr (52) 52/52 | Cr (53) NH3 53/53 | Fe (54) NH3 54/54 | Mn Oshift 55/71 | Fe (56) KED 56/56 | Co KED 59/59 | Ni KED 60/60 | Cu (63) KED 63/63 | Cu (65) NH3 65/65 | Zn NH3 66/66 | As 75/75 | As KED 75/75 | As RH Oshift 75/91 | As Oshift 75/91 | Se (77) 77/77 | Se (77) Oshift 77/93 | Se (78) Oshift 78/94 | Se (80) Oshift 80/96 | Se (82) 82/82 | Se (82) Oshift 82/98 | Mo (95) NH3 95/95 | Mo (95) OShift 95/127 | Rh (IS1) 103/103 (IS) | Rh (IS2) 103/103 (IS) | Rh (IS3) 103/103 (IS) | Rh (IS4) 103/103 (IS) | Rh (IS5) 103/103 (IS) | Rh (IS6) 103/103 (IS) | Rh OShift (IS1) 103/119 (IS) | Cd (111) NH3 111/111 | Cd (114) NH3 114/114 | Pb 208/208 | Pb NH3 208/208
| 39 | MT 2 LSP, Spike 1 of Row Index 38 | R | 06/02/2023 19:19:58 | 1534% | 1409% | 810.2% | 7622% | 1352% | 1390% | 1383% | 126.9% | 123.6% | 119.4% | 252.9% | 134% | 278.9% | 126.8% | 127.9% | 111.3% | 108% | 129.3% | 113.2% | 110.2% | 99.16% | 114.5% | 122.5% | 120.5% | 120.1% | 115.5% | 115.8% | 123% | 112.4% | 111.9% | [7 empty] | 100.8% | 99.48% | 83.82% | 85.33%
| 40 | MT 2 LSP, Spike 1 of Row Index 38 | R | 06/02/2023 19:26:09 | 1417% | 1304% | 226.9% | 5446% | 1277% | 1271% | 1300% | 120.3% | 114.6% | 118.1% | 222.7% | 134.2% | 175.6% | 122.3% | 123.1% | 104.8% | 107.5% | 125.3% | 113.9% | 105.5% | 98.31% | 116.9% | 120.8% | 124.3% | 119.2% | 118.3% | 114.9% | 120.4% | 112.3% | 116% | [7 empty] | 99.84% | 98.69% | 84.37% | 84.82%
| 58 | MT 12 LSP, Spike 1 of Row Index 57 | R | 06/02/2023 21:16:23 | 1711% | 1494% | 1097% | 9089% | 1500% | 1540% | 1488% | 136.4% | 127.5% | 130.8% | 195.2% | 143.8% | 277.5% | 133.9% | 136.3% | 111.9% | 99.01% | 111.1% | 129.4% | 115.4% | 101.5% | 124.6% | 144.5% | 126.7% | 130.2% | 125.3% | 132.3% | 130.3% | 117.3% | 117.6% | [7 empty] | 100.8% | 99.77% | 82.79% | 83.37%
| 82 | MT 22 LSP, Spike 1 of Row Index 81 | R | 06/02/2023 23:44:15 | 2207% | 1926% | 203.8% | -177532% | 1664% | 1784% | 1675% | 144.8% | 145.7% | 137.9% | 847.5% | 156.4% | 428.1% | 138.6% | 139.1% | 118% | 111.7% | 142% | 124.4% | 112% | 100.7% | 124.6% | 130.6% | 127.7% | 132.9% | 125.2% | 129% | 123.2% | 115.1% | 118.8% | [7 empty] | 99.25% | 100.6% | 75.88% | 78.13%
| 83 | MT 22 LSP, Spike 1 of Row Index 81 | R | 06/02/2023 23:50:45 | 1970% | 1695% | 110.2% | 23772% | 1422% | 1594% | 1486% | 126.4% | 122.3% | 138.4% | 823.4% | 158.1% | -1053% | 123.8% | 121.2% | 79.49% | 110% | 138.6% | 124.7% | 97.11% | 101.9% | 123.9% | 131.1% | 132% | 131.6% | 125.2% | 127.2% | 122% | 115.6% | 118.9% | [7 empty] | 101.8% | 100.3% | 77.04% | 75.51%
| 103 | MT 32 LSP, Spike 1 of Row Index 102 | R | 06/03/2023 01:54:03 | 2430% | 2052% | 235.8% | 278.6% | 1763% | 1921% | 1767% | 151% | 141.9% | 147.6% | 188.6% | 166.4% | 177.8% | 142.7% | 207.1% | 121.5% | 123.6% | 499.5% | 123.7% | 113.7% | 103.5% | 123.7% | 132.6% | 120.2% | 128.5% | 126% | 129.2% | 122.6% | 115.2% | 118.2% | [7 empty] | 100% | 98.51% | 73.69% | 75.6%
| 104 | MT 32 LSP, Spike 1 of Row Index 102 | R | 06/03/2023 02:00:34 | 2369% | 2006% | 245% | 189.7% | 1722% | 1850% | 1707% | 148% | 141% | 146.3% | 186.7% | 159.8% | 177.8% | 142.9% | 207.1% | 123.8% | 124.6% | 504.6% | 119.6% | 118.2% | 106.8% | 122.7% | 125.8% | 125.7% | 127.6% | 121% | 119.4% | 118.7% | 116.3% | 115.3% | [7 empty] | 100.2% | 99.58% | 72.24% | 74.61%
| 122 | MT 32 LSP, Spike 1 of Row Index 121 | R | 06/03/2023 03:50:58 | 2402% | 2061% | 253.9% | 103.7% | 1810% | 1893% | 1813% | 150.5% | 144.8% | 146.8% | 187.7% | 167% | 178.6% | 142.5% | 210.3% | 123.1% | 123.7% | 500.6% | 122.5% | 110.7% | 102.2% | 125.2% | 131.5% | 128% | 129.8% | 123.7% | 128.5% | 122.3% | 115.3% | 118.3% | [7 empty] | 100.6% | 98.16% | 72.59% | 85.06%
| 123 | MT 32 LSP, Spike 1 of Row Index 121 | R | 06/03/2023 03:57:29 | 2313% | 2015% | 235.5% | 224.3% | 1785% | 1882% | 1710% | 147.4% | 143.4% | 133.2% | 173% | 162.2% | 177% | 141.7% | 206.1% | 122.5% | 114.7% | 467.2% | 110% | 113.4% | 102.1% | 121.7% | 115.7% | 122.8% | 122.7% | 120% | 108.7% | 117.6% | 106.7% | 116.3% | [7 empty] | 92.27% | 92.08% | 64.84% | 79.93%
| 147 | MT 33 B LSP, Spike 2 of Row Index 146 | R | 06/03/2023 06:24:56 | 240.7% | 209.1% | 198.5% | 206.9% | 183.4% | 192.7% | 186.7% | 152.2% | 143.3% | 131.2% | 139.9% | 155.9% | 147.5% | 145.5% | 143.6% | 121.5% | 106.9% | 108.7% | 112.5% | 109.3% | 96.07% | 115.5% | 118.9% | 115.1% | 114.4% | 111.2% | 109.1% | 109.6% | 103.7% | 118.6% | [7 empty] | 84.87% | 83.87% | 58.43% | 71.6%
| 165 | MT 33 C LSP, Spike 2 of Row Index 164 | R | 06/03/2023 08:15:09 | 236.8% | 204.5% | 204.1% | 216.1% | 174.2% | 185.1% | 182.3% | 152.1% | 142.7% | 112% | 111.1% | 152.8% | 144.5% | 142.2% | 142% | 118.7% | 89.85% | 104.1% | 128% | 110.2% | 92.66% | 111.9% | 134.8% | 114.6% | 113.1% | 107.7% | 121.2% | 107% | 0.1907% | 116% | [7 empty] | 104.7% | 105.3% | 48.79% | 125.8%
| 166 | MT 33 C LSP, Spike 2 of Row Index 164 | R | 06/03/2023 08:21:40 | 234.8% | 206.1% | 196.1% | 206.2% | 175.5% | 189% | 180.3% | 152.8% | 142.8% | 146.8% | 157.3% | 155.8% | 145.9% | 145.2% | 143.3% | 120.9% | 120.9% | 121.2% | 116.3% | 113.6% | 92.31% | 114.3% | 123% | 115% | 114.8% | 112.4% | 112.9% | 107.3% | 116% | 118.2% | [7 empty] | 93.73% | 93.4% | 60.04% | 71.74%
QC Duplicates (% Difference)
|  | Sample Id | R | Acquisition Time | Na KED 23/23 | Mg KED 24/24 | Al KED 27/27 | P Oshift 31/47 | S (34) -1 34/34 | K KED 39/39 | Ca KED 43/43 | V KED 51/51 | Cr (52) 52/52 | Cr (53) NH3 53/53 | Fe (54) NH3 54/54 | Mn Oshift 55/71 | Fe (56) KED 56/56 | Co KED 59/59 | Ni KED 60/60 | Cu (63) KED 63/63 | Cu (65) NH3 65/65 | Zn NH3 66/66 | As 75/75 | As KED 75/75 | As RH Oshift 75/91 | As Oshift 75/91 | Se (77) 77/77 | Se (77) Oshift 77/93 | Se (78) Oshift 78/94 | Se (80) Oshift 80/96 | Se (82) 82/82 | Se (82) Oshift 82/98 | Mo (95) NH3 95/95 | Mo (95) OShift 95/127 | Rh (IS1) 103/103 (IS) | Rh (IS2) 103/103 (IS) | Rh (IS3) 103/103 (IS) | Rh (IS4) 103/103 (IS) | Rh (IS5) 103/103 (IS) | Rh (IS6) 103/103 (IS) | Rh OShift (IS1) 103/119 (IS) | Cd (111) NH3 111/111 | Cd (114) NH3 114/114 | Pb 208/208 | Pb NH3 208/208
| 26 | MT 1 LD, Duplicate of Row Index 25 | R | 06/02/2023 17:44:04 | 6.758% | 4.189% | 6.456% | 4.245% | 5.418% | 8.109% | 6.277% | 0.2153% | 0.5603% | 8.3% | 3.757% | 3.863% | 2.3% | 7.41% | 1.104% | 0.3667% | 6.369% | 9.475% | 9.263% | 1.425% | 7.882% | 4.661% | 21.97% | 11.14% | 5.908% | 11.04% | 13.66% | 7.667% | 6.932% | 3.461% | [7 empty] | 13.87% | 7.506% | 6.378% | 4.349%
| 37 | MT 1 LD, Duplicate of Row Index 36 | R | 06/02/2023 19:07:53 | 7.602% | 8.805% | 8.525% | 4.607% | 3.79% | 9.834% | 8.59% | 0.2925% | 2.217% | 0.9642% | 1.032% | 8.145% | 2.594% | 2.784% | 15.52% | 6.327% | 4.626% | 5.777% | 5.763% | 4.06% | 11.24% | 7.752% | 2.886% | 4.373% | 4.691% | 5.275% | 9.519% | 8.254% | 6.209% | 17.92% | [7 empty] | 9.308% | 2.752% | 2.077% | 2.406%
| 55 | MT 11 LD, Duplicate of Row Index 54 | R | 06/02/2023 20:57:47 | 1.484% | 4.258% | 0.2192% | 4.691% | 1.496% | 2.166% | 1.923% | 2.912% | 2.512% | 0.4891% | 1.827% | 3.196% | 1.815% | 84% | 2.365% | 0.00896% | 0.3657% | 1.073% | 0.7859% | 2.987% | 1.015% | 2.958% | 3.508% | 5.295% | 1.267% | 1.675% | 3.235% | 2.652% | 5.787% | 3.544% | [7 empty] | 3.682% | 2.152% | 2.861% | 1.323%
| 56 | MT 11 LD, Duplicate of Row Index 54 | R | 06/02/2023 21:04:18 | 6.314% | 6.343% | 6.894% | 1.555% | 5.596% | 6.44% | 7.947% | 4.337% | 10.96% | 1.669% | 1.041% | 0.8046% | 6.093% | 88.87% | 5.934% | 7.419% | 1.54% | 2.019% | 0.4645% | 9.507% | 1.575% | 0.5325% | 8.517% | 7.49% | 0.3595% | 3.144% | 4.385% | 7.024% | 10.58% | 1.053% | [7 empty] | 5.556% | 4.559% | 0.668% | 4.673%
| 75 | MT 21 LD, Duplicate of Row Index 74 | R | 06/02/2023 23:00:59 | 205.4% | 214.8% | 216.1% | 200% | 264.9% | 206.2% | 203.4% | 189.4% | 198.8% | 199.9% | 200.2% | 200.2% | 200.1% | 199.1% | 201.6% | 199.6% | 200.2% | 201.4% | 164% | 199.6% | 200.6% | 200.5% | 475.5% | 198.4% | 202.4% | 200.5% | 207.5% | 200.5% | 199.7% | 202.3% | [7 empty] | 200% | 200% | 204.8% | 202.5%
| 100 | MT 31 LD, Duplicate of Row Index 99 | R | 06/03/2023 01:35:29 | 16.49% | 6.051% | 3.478% | 2.684% | 6.87% | 5.199% | 3.322% | 6.06% | 0.6064% | 5.072% | 3.001% | 2.183% | 1.842% | 0.3758% | 0.2243% | 2.001% | 1.78% | 4.665% | 2.538% | 3.996% | 2.066% | 2.748% | 6.844% | 11.35% | 17.29% | 2.503% | 11.99% | 33.05% | 4.471% | 3.016% | [7 empty] | 2.543% | 3.278% | 0.2845% | 0.2432%
| 101 | MT 31 LD, Duplicate of Row Index 99 | R | 06/03/2023 01:42:00 | 15.71% | 12.59% | 0.3923% | 2.429% | 2.288% | 4.008% | 2.34% | 0.3139% | 3.676% | 2.779% | 2.657% | 2.443% | 1.801% | 2.18% | 5.742% | 2.582% | 1.658% | 3.216% | 3.577% | 18.29% | 1.432% | 2.532% | 6.084% | 8.899% | 6.772% | 0.6779% | 9.679% | 0.8242% | 2.325% | 3.022% | [7 empty] | 2.213% | 2.173% | 1.738% | 6.168%
| 119 | MT 31 LD, Duplicate of Row Index 118 | R | 06/03/2023 03:32:23 | 13.27% | 18.29% | 5.536% | 3.429% | 15.88% | 5.25% | 4.542% | 11.2% | 6.391% | 10.29% | 2.724% | 4.945% | 4.058% | 4.669% | 1.831% | 5.3% | 2.766% | 3.003% | 6.467% | 1.605% | 1.338% | 3.54% | 2.916% | 10.07% | 1.57% | 1.529% | 23.19% | 39.51% | 8.86% | 6.153% | [7 empty] | 1.91% | 4.183% | 2.191% | 14.53%
| 120 | MT 31 LD, Duplicate of Row Index 118 | R | 06/03/2023 03:38:54 | 21.92% | 21.73% | 6.831% | 1.014% | 12.32% | 1.558% | 5.595% | 6.708% | 0.7617% | 4.435% | 10.73% | 0.6868% | 1.122% | 2.267% | 5.015% | 2.348% | 10.52% | 9.791% | 3.754% | 4.258% | 4.229% | 2.247% | 54.99% | 15.69% | 16.35% | 8.04% | 43.93% | 24.24% | 7.666% | 0.2179% | [7 empty] | 10.41% | 9.288% | 8.962% | 6.134%
Samples, Blanks & Calibration Standards
|  | Sample Id | R | Acquisition Time | Na KED 23/23 | Mg KED 24/24 | Al KED 27/27 | P Oshift 31/47 | S (34) -1 34/34 | K KED 39/39 | Ca KED 43/43 | V KED 51/51 | Cr (52) 52/52 | Cr (53) NH3 53/53 | Fe (54) NH3 54/54 | Mn Oshift 55/71 | Fe (56) KED 56/56 | Co KED 59/59 | Ni KED 60/60 | Cu (63) KED 63/63 | Cu (65) NH3 65/65 | Zn NH3 66/66 | As 75/75 | As KED 75/75 | As RH Oshift 75/91 | As Oshift 75/91 | Se (77) 77/77 | Se (77) Oshift 77/93 | Se (78) Oshift 78/94 | Se (80) Oshift 80/96 | Se (82) 82/82 | Se (82) Oshift 82/98 | Mo (95) NH3 95/95 | Mo (95) OShift 95/127 | Rh (IS1) 103/103 (IS) | Rh (IS2) 103/103 (IS) | Rh (IS3) 103/103 (IS) | Rh (IS4) 103/103 (IS) | Rh (IS5) 103/103 (IS) | Rh (IS6) 103/103 (IS) | Rh OShift (IS1) 103/119 (IS) | Cd (111) NH3 111/111 | Cd (114) NH3 114/114 | Pb 208/208 | Pb NH3 208/208
| 25 | MT 1 | R | 06/02/2023 17:38:02 | 3938.0267 | 228.4614 | 422.4351 | 2517.3629 | 3274.1776 | 3290.2344 | 673.5447 | 0.45867 | 0.55734 | 0.53116 | 91.6986 | 0.8713 | 98.3796 | 0.072481 | 0.37763 | 4.0977 | 4.1035 | 12.9328 | 1.5666 | 1.3648 | 1.3534 | 1.6312 | 0.7488 | 0.92616 | 0.84418 | 0.86674 | 0.97459 | 0.93372 | 0.067212 | 0.062977 | [7 empty] | 0.068459 | 0.069387 | 0.15119 | 0.14031
|  | IS % Recovery | [32 empty] | 106.4% | 98.46% | 111.5% | 99.27% | 102.9% | 87.68% | 115%
|  | Sample RSD |  |  | 0.9883% | 12.24% | 0.7114% | 2.795% | 2.432% | 1.772% | 1.579% | 6.933% | 0.7885% | 1.407% | 0.6656% | 2.394% | 0.8796% | 10.48% | 0.6753% | 1.332% | 1.13% | 0.6094% | 5.394% | 9.042% | 8.331% | 2.113% | 8.12% | 21.28% | 8.056% | 7.17% | 9.414% | 12.29% | 5.247% | 2.409% | [7 empty] | 1.453% | 5.151% | 1.375% | 1.424%
|  | Sample Concentration SD |  |  | 38.9201 | 27.9669 | 3.0052 | 70.3597 | 79.6425 | 58.2989 | 10.6331 | 0.0318 | 0.0043947 | 0.0074758 | 0.61034 | 0.020857 | 0.86533 | 0.0075979 | 0.0025501 | 0.054564 | 0.046365 | 0.078815 | 0.084495 | 0.12341 | 0.11275 | 0.03446 | 0.060803 | 0.19712 | 0.068004 | 0.062147 | 0.091749 | 0.11471 | 0.0035263 | 0.001517 | [7 empty] | 0.0009949 | 0.0035741 | 0.0020796 | 0.0019984
| 36 | MT 1 | R | 06/02/2023 19:01:50 | 4390.7358 | 251.0427 | 480.1772 | 2566.2268 | 3789.2436 | 3596.5431 | 806.9309 | 0.48063 | 0.58234 | 0.55929 | 97.0559 | 1.0589 | 103.677 | 0.082321 | 0.37773 | 4.1209 | 4.1489 | 13.3781 | 1.5977 | 1.3501 | 1.6209 | 1.9139 | 0.81317 | 1.0787 | 1.028 | 1.0591 | 0.97694 | 0.97873 | 0.070888 | 0.078951 | [7 empty] | 0.072153 | 0.072524 | 0.14317 | 0.1374
|  | IS % Recovery | [32 empty] | 100.1% | 83.62% | 107.2% | 92.09% | 95% | 90.82% | 95.62%
|  | Sample RSD |  |  | 1.673% | 2.569% | 0.8845% | 0.8532% | 6.07% | 1.041% | 13.74% | 3.842% | 8.806% | 3.574% | 1.413% | 12.7% | 0.2749% | 6.175% | 5.916% | 0.628% | 1.93% | 0.6891% | 6.162% | 8.469% | 13.82% | 12.32% | 12.09% | 4.791% | 13.53% | 17.62% | 8.741% | 16.66% | 2.977% | 8.974% | [7 empty] | 12.9% | 2.507% | 2.291% | 3.282%
|  | Sample Concentration SD |  |  | 73.4351 | 6.4494 | 4.2469 | 21.8949 | 230.025 | 37.4382 | 110.895 | 0.018464 | 0.051283 | 0.01999 | 1.3715 | 0.13445 | 0.285 | 0.0050836 | 0.022347 | 0.025879 | 0.080059 | 0.09219 | 0.098447 | 0.11434 | 0.22397 | 0.23571 | 0.098335 | 0.051679 | 0.13909 | 0.18656 | 0.085398 | 0.1631 | 0.0021104 | 0.0070852 | [7 empty] | 0.0093052 | 0.001818 | 0.0032799 | 0.0045098
| 38 | MT 2 | R | 06/02/2023 19:13:55 | 5158.7022 | 252.262 | 560.1251 | 2891.4774 | 4116.3693 | 4139.4865 | 851.201 | 0.50451 | 0.62417 | 0.61414 | 99.9173 | 0.95328 | 104.0469 | 0.084898 | 0.40195 | 4.5614 | 4.7273 | 14.9376 | 1.7748 | 1.6273 | 1.6088 | 1.9035 | 0.87084 | 0.95784 | 1.0829 | 1.0914 | 1.1047 | 1.0527 | 0.098176 | 0.098976 | [7 empty] | 0.076314 | 0.076029 | 0.14735 | 0.14372
|  | IS % Recovery | [32 empty] | 97.04% | 88.5% | 104.4% | 88.15% | 93.28% | 87.34% | 101.6%
|  | Sample RSD |  |  | 1.721% | 8.3% | 3.278% | 0.9013% | 2.322% | 1.762% | 13.24% | 1.374% | 3.427% | 1.495% | 1.947% | 7.358% | 3.119% | 7.157% | 10.91% | 2.963% | 2.019% | 1.397% | 5.045% | 2.099% | 12.48% | 9.316% | 10.83% | 18.63% | 11.66% | 11.94% | 5.997% | 8.844% | 6.488% | 12.47% | [7 empty] | 4.782% | 3.225% | 1.12% | 2.084%
|  | Sample Concentration SD |  |  | 88.7673 | 20.9388 | 18.3597 | 26.0601 | 95.5944 | 72.9267 | 112.7078 | 0.00693 | 0.021388 | 0.0091833 | 1.9449 | 0.070141 | 3.2449 | 0.0060759 | 0.043838 | 0.13514 | 0.095447 | 0.20867 | 0.089534 | 0.034153 | 0.20079 | 0.17733 | 0.094291 | 0.1784 | 0.12631 | 0.1303 | 0.066249 | 0.0931 | 0.0063692 | 0.01234 | [7 empty] | 0.0036496 | 0.002452 | 0.001651 | 0.0029956
| 41 | MT 3 | R | 06/02/2023 19:32:11 | 4004.6025 | 212.2501 | 435.5836 | 2351.5468 | 3706.9383 | 3294.5661 | 619.4182 | 0.42402 | 0.67014 | 0.66024 | 83.7203 | 0.79953 | 84.5065 | 0.073066 | 0.42675 | 3.759 | 3.9661 | 12.4638 | 1.581 | 1.4894 | 1.3547 | 1.6234 | 0.8665 | 1.0681 | 0.90249 | 0.90039 | 1.0432 | 0.93624 | 0.071804 | 0.065333 | [7 empty] | 0.065051 | 0.068259 | 0.14473 | 0.13871
|  | IS % Recovery | [32 empty] | 99.83% | 94.09% | 110% | 93.25% | 94.5% | 90.87% | 109%
|  | Sample RSD |  |  | 1.171% | 9.207% | 3.629% | 1.546% | 7.048% | 4.028% | 6.089% | 4.878% | 4.502% | 4.711% | 0.6689% | 1.527% | 2.228% | 8.187% | 1.753% | 2.224% | 0.4469% | 1.077% | 0.8155% | 7.865% | 2.005% | 1.886% | 17.82% | 4.074% | 3.796% | 2.833% | 11.35% | 3.512% | 4.175% | 1.27% | [7 empty] | 7.641% | 7.416% | 0.5022% | 3.663%
|  | Sample Concentration SD |  |  | 46.8805 | 19.5409 | 15.8058 | 36.3466 | 261.2588 | 132.7036 | 37.7191 | 0.020683 | 0.030167 | 0.031104 | 0.56 | 0.012206 | 1.8829 | 0.0059816 | 0.0074816 | 0.083589 | 0.017723 | 0.13421 | 0.012892 | 0.11713 | 0.027168 | 0.030611 | 0.15443 | 0.043514 | 0.034261 | 0.025512 | 0.11837 | 0.032877 | 0.0029976 | 0.00082977 | [7 empty] | 0.0049703 | 0.0050622 | 0.00072681 | 0.0050807
| 42 | MT 4 | R | 06/02/2023 19:38:14 | 5342.3849 | 259.1695 | 587.2254 | 3022.489 | 4880.4377 | 4463.1677 | 866.4666 | 0.55984 | 0.66918 | 0.63449 | 134.5167 | 1.1584 | 145.0237 | 0.097116 | 0.43049 | 4.8579 | 4.8848 | 15.7073 | 1.9539 | 1.7358 | 1.6878 | 2.0419 | 0.92327 | 1.1444 | 1.0885 | 1.1468 | 1.1902 | 1.0775 | 0.083875 | 0.080313 | [7 empty] | 0.086431 | 0.085481 | 0.15863 | 0.15228
|  | IS % Recovery | [32 empty] | 98.19% | 89.72% | 103.4% | 90.12% | 93.81% | 88.11% | 104.9%
|  | Sample RSD |  |  | 2.677% | 4.705% | 0.3803% | 1.231% | 4.453% | 2.372% | 10.21% | 5.79% | 2.768% | 2.208% | 1.039% | 1.507% | 1.357% | 7.295% | 10.69% | 1.422% | 1.714% | 0.8096% | 4.391% | 6.247% | 1.62% | 0.807% | 4.294% | 8.625% | 8.66% | 1.124% | 4.367% | 6.264% | 0.316% | 4.216% | [7 empty] | 3.455% | 4.264% | 4.673% | 1.119%
|  | Sample Concentration SD |  |  | 143.0355 | 12.1945 | 2.2332 | 37.2107 | 217.3073 | 105.8785 | 88.4935 | 0.032414 | 0.018524 | 0.014007 | 1.3977 | 0.017462 | 1.9683 | 0.0070844 | 0.046017 | 0.0691 | 0.083703 | 0.12716 | 0.085793 | 0.10844 | 0.027344 | 0.016477 | 0.039649 | 0.098709 | 0.094263 | 0.012891 | 0.051977 | 0.067493 | 0.00026501 | 0.003386 | [7 empty] | 0.002986 | 0.0036449 | 0.0074127 | 0.0017046
| 48 | MT 5 | R | 06/02/2023 20:15:30 | 5444.3989 | 303.142 | 572.47 | 3141.2391 | 4613.6915 | 4290.7112 | 924.0429 | 0.53388 | 0.64217 | 0.61705 | 110.827 | 1.055 | 119.3673 | 0.097861 | 0.43402 | 4.6243 | 4.6274 | 15.2062 | 1.8163 | 1.514 | 1.4764 | 1.8379 | 1.0096 | 1.0629 | 1.0186 | 1.0114 | 1.1309 | 0.93376 | 0.075855 | 0.06983 | [7 empty] | 0.08105 | 0.079407 | 0.14204 | 0.14285
|  | IS % Recovery | [32 empty] | 96.07% | 89.31% | 102.8% | 87.6% | 90.9% | 84.45% | 107.4%
|  | Sample RSD |  |  | 1.325% | 9.614% | 2.915% | 1.672% | 2.369% | 1.143% | 3.957% | 3.447% | 2.702% | 0.6876% | 2.859% | 1.4% | 2.553% | 12.58% | 6.705% | 1.405% | 1.865% | 2.291% | 4.013% | 3.204% | 1.687% | 2.212% | 3.294% | 14.98% | 7.055% | 8.706% | 10.14% | 11.87% | 3.641% | 3.449% | [7 empty] | 2.019% | 2.163% | 1.864% | 0.7342%
|  | Sample Concentration SD |  |  | 72.1515 | 29.1449 | 16.6871 | 52.5098 | 109.2754 | 49.0372 | 36.5648 | 0.018402 | 0.017351 | 0.0042427 | 3.1686 | 0.014773 | 3.0473 | 0.012306 | 0.029103 | 0.064973 | 0.086308 | 0.34834 | 0.072885 | 0.048505 | 0.024912 | 0.040645 | 0.033257 | 0.1592 | 0.071858 | 0.088049 | 0.11463 | 0.1108 | 0.002762 | 0.0024088 | [7 empty] | 0.0016361 | 0.0017173 | 0.0026477 | 0.0010488
| 49 | MT 6 | R | 06/02/2023 20:21:32 | 13254.0143 | 719.7689 | 1332.6762 | 7172.5836 | 9196.9916 | 10665.9048 | 2172.6291 | 1.2201 | 1.3237 | 1.2586 | 243.5425 | 2.3934 | 261.7472 | 0.20286 | 0.97031 | 10.8734 | 10.7793 | 35.549 | 4.2744 | 4.0632 | 3.7545 | 4.5278 | 2.0552 | 2.3986 | 2.5983 | 2.4247 | 2.5625 | 2.5676 | 0.18769 | 0.17022 | [7 empty] | 0.1924 | 0.18337 | 0.2689 | 0.27379
|  | IS % Recovery | [32 empty] | 93.37% | 84.1% | 98.96% | 84.57% | 86.13% | 84.71% | 98.05%
|  | Sample RSD |  |  | 1.792% | 1.468% | 3.426% | 1.208% | 3.699% | 1.669% | 2.289% | 4.549% | 1.637% | 1.1% | 1.151% | 1.779% | 1.792% | 2.732% | 1.064% | 2.077% | 1.647% | 1.111% | 1.982% | 3.564% | 1.55% | 0.1609% | 1.655% | 2.901% | 4.274% | 0.6331% | 3.41% | 1.215% | 5.42% | 3.901% | [7 empty] | 6.331% | 1.277% | 0.1671% | 6.235%
|  | Sample Concentration SD |  |  | 237.5007 | 10.5667 | 45.6622 | 86.6667 | 340.2054 | 178.0082 | 49.7362 | 0.055506 | 0.02167 | 0.013847 | 2.8028 | 0.042588 | 4.6903 | 0.0055425 | 0.010327 | 0.22588 | 0.17755 | 0.395 | 0.084708 | 0.14482 | 0.058186 | 0.0072864 | 0.034018 | 0.069589 | 0.11106 | 0.01535 | 0.087381 | 0.031187 | 0.010173 | 0.0066398 | [7 empty] | 0.012181 | 0.002342 | 0.00044944 | 0.01707
| 50 | MT 7 | R | 06/02/2023 20:27:35 | 13093.7643 | 758.1468 | 1388.7267 | 7189.1611 | 9166.6726 | 10793.4359 | 1935.6723 | 1.1856 | 1.3271 | 1.3165 | 238.1719 | 2.3275 | 255.2242 | 0.20901 | 1.0616 | 10.8057 | 10.8248 | 35.487 | 4.3018 | 3.8956 | 3.7332 | 4.5399 | 2.1609 | 2.3122 | 2.6021 | 2.4548 | 2.7629 | 2.3709 | 0.2028 | 0.17851 | [7 empty] | 0.17877 | 0.18277 | 0.31886 | 0.31049
|  | IS % Recovery | [32 empty] | 94.93% | 83.8% | 99.07% | 85.33% | 88.67% | 82.6% | 98.71%
|  | Sample RSD |  |  | 1.278% | 13.97% | 1.854% | 2.419% | 1.753% | 0.4474% | 4.398% | 0.975% | 0.9707% | 16.58% | 15.99% | 0.9864% | 1.308% | 2.209% | 2.617% | 1.088% | 16.92% | 15.47% | 2.367% | 1.596% | 5.993% | 1.432% | 14.33% | 3.708% | 5.082% | 1.007% | 4.687% | 3.234% | 18.86% | 1.672% | [7 empty] | 16.44% | 14.11% | 1.665% | 0.8361%
|  | Sample Concentration SD |  |  | 167.3163 | 105.9058 | 25.7415 | 173.9377 | 160.7292 | 48.2888 | 85.1375 | 0.01156 | 0.012882 | 0.21826 | 38.0916 | 0.022958 | 3.3378 | 0.0046174 | 0.027782 | 0.11761 | 1.832 | 5.4897 | 0.10183 | 0.062169 | 0.22374 | 0.065001 | 0.30966 | 0.085748 | 0.13224 | 0.024719 | 0.1295 | 0.076671 | 0.038253 | 0.0029843 | [7 empty] | 0.029383 | 0.025797 | 0.005309 | 0.002596
| 51 | MT 8 | R | 06/02/2023 20:33:38 | 12795.8848 | 718.3922 | 1342.167 | 6781.8018 | 9112.9461 | 10496.7611 | 2058.432 | 1.1633 | 1.3687 | 1.2594 | 217.3419 | 2.2428 | 241.7167 | 0.20255 | 1.0324 | 10.4651 | 10.0661 | 32.6747 | 3.9864 | 3.7723 | 3.6098 | 4.2297 | 1.992 | 2.2354 | 2.3749 | 2.3077 | 2.4231 | 2.3432 | 0.17724 | 0.17091 | [7 empty] | 0.17447 | 0.1666 | 0.26395 | 0.26116
|  | IS % Recovery | [32 empty] | 94.14% | 84.6% | 97.74% | 85.84% | 89.38% | 82.91% | 95.88%
|  | Sample RSD |  |  | 7.794% | 3.933% | 10.88% | 1.416% | 10.46% | 9.902% | 13.78% | 9.367% | 7.965% | 3.454% | 0.3211% | 0.5287% | 8.455% | 13.29% | 11.14% | 10.33% | 0.8466% | 0.3502% | 1.702% | 10.43% | 2.57% | 1.785% | 10.25% | 5.522% | 2.478% | 1.852% | 3.792% | 4.119% | 2.858% | 3.134% | [7 empty] | 8.653% | 2.748% | 2.388% | 2.755%
|  | Sample Concentration SD |  |  | 997.2934 | 28.2577 | 145.9768 | 96.0462 | 952.9979 | 1039.3631 | 283.6076 | 0.10897 | 0.10902 | 0.043493 | 0.69793 | 0.011858 | 20.4381 | 0.026919 | 0.11496 | 1.0811 | 0.08522 | 0.11443 | 0.067836 | 0.39339 | 0.092781 | 0.075482 | 0.20412 | 0.12345 | 0.058855 | 0.042738 | 0.091883 | 0.096526 | 0.005066 | 0.0053557 | [7 empty] | 0.015097 | 0.0045776 | 0.0063028 | 0.0071945
| 52 | MT 9 | R | 06/02/2023 20:39:40 | 12208.9519 | 653.6601 | 1285.0152 | 6982.8798 | 8663.2891 | 9719.0445 | 1911.2318 | 1.1438 | 1.2449 | 1.2333 | 224.1746 | 2.2126 | 233.4657 | 0.19984 | 0.93338 | 11.7302 | 11.9763 | 34.085 | 4.27 | 3.6746 | 3.5425 | 4.2531 | 2.2316 | 2.5109 | 2.3457 | 2.3332 | 2.7059 | 2.4277 | 0.18493 | 0.16324 | [7 empty] | 0.1752 | 0.1756 | 0.27053 | 0.25304
|  | IS % Recovery | [32 empty] | 88.84% | 83.4% | 99.46% | 82.93% | 87.81% | 80.55% | 96.8%
|  | Sample RSD |  |  | 0.01043% | 7.095% | 0.6402% | 1.861% | 2.69% | 1.774% | 10.03% | 5.304% | 2.845% | 0.9479% | 0.9172% | 2.504% | 0.8696% | 7.223% | 6.214% | 1.859% | 0.6775% | 0.745% | 3.664% | 8.357% | 2.86% | 0.8485% | 4.388% | 7.476% | 5.796% | 7.088% | 7.508% | 8.653% | 1.923% | 3.866% | [7 empty] | 3.792% | 3.769% | 4.957% | 1.822%
|  | Sample Concentration SD |  |  | 1.2739 | 46.3741 | 8.2269 | 129.9675 | 233.0812 | 172.464 | 191.7014 | 0.060672 | 0.035423 | 0.011691 | 2.0562 | 0.055405 | 2.0301 | 0.014434 | 0.058004 | 0.2181 | 0.081144 | 0.25392 | 0.15645 | 0.3071 | 0.10132 | 0.036088 | 0.097931 | 0.1877 | 0.13597 | 0.16538 | 0.20315 | 0.21007 | 0.0035563 | 0.0063101 | [7 empty] | 0.006643 | 0.0066186 | 0.01341 | 0.0046106
| 53 | MT 10 | R | 06/02/2023 20:45:42 | 12729.0397 | 760.0704 | 1375.8636 | 7377.1544 | 8982.8171 | 10488.9016 | 1993.7665 | 1.1939 | 1.5046 | 1.4337 | 246.053 | 2.4661 | 260.7206 | 0.20506 | 1.0248 | 11.9928 | 11.9431 | 36.2005 | 4.3419 | 3.758 | 3.8499 | 4.572 | 2.0949 | 2.3888 | 2.5134 | 2.543 | 2.593 | 2.4176 | 0.19157 | 0.18912 | [7 empty] | 0.18258 | 0.17426 | 0.26714 | 0.25831
|  | IS % Recovery | [32 empty] | 90.74% | 80.72% | 97.23% | 83.04% | 86.91% | 80.64% | 92.68%
|  | Sample RSD |  |  | 2.35% | 10.52% | 4.126% | 1.021% | 2.183% | 2.385% | 6.227% | 3.262% | 1.198% | 2.4% | 1.387% | 0.8303% | 1.623% | 2.518% | 2.305% | 2.335% | 1.904% | 1.348% | 2.222% | 8.493% | 1.668% | 0.3364% | 3.716% | 16.52% | 1.025% | 4.943% | 3.904% | 13.07% | 2.494% | 3.203% | [7 empty] | 4.831% | 2.237% | 1.986% | 1.592%
|  | Sample Concentration SD |  |  | 299.0979 | 79.9519 | 56.7705 | 75.3263 | 196.0817 | 250.1237 | 124.1485 | 0.038941 | 0.018026 | 0.034409 | 3.4137 | 0.020476 | 4.2323 | 0.0051632 | 0.02362 | 0.28001 | 0.22734 | 0.48812 | 0.096469 | 0.31917 | 0.064197 | 0.015379 | 0.077853 | 0.39465 | 0.025752 | 0.12571 | 0.10123 | 0.31592 | 0.0047769 | 0.0060574 | [7 empty] | 0.0088197 | 0.0038983 | 0.0053052 | 0.0041134
| 54 | MT 11 | R | 06/02/2023 20:51:44 | 25247.4894 | 1464.7179 | 2644.5768 | 14676.4934 | 16195.242 | 20678.5465 | 4332.207 | 2.4004 | 2.5326 | 2.5679 | 478.5436 | 4.9298 | 506.2743 | 0.40944 | 2.0135 | 20.152 | 20.4554 | 71.2293 | 8.4965 | 7.3843 | 7.5744 | 8.9291 | 4.1273 | 5.0453 | 4.9622 | 5.0111 | 5.5323 | 5.092 | 0.41223 | 0.35751 | [7 empty] | 0.36724 | 0.353 | 0.46897 | 0.47136
|  | IS % Recovery | [32 empty] | 89.06% | 79.53% | 95.37% | 80.34% | 83.98% | 79.08% | 90.64%
|  | Sample RSD |  |  | 1.668% | 2.111% | 1.275% | 0.6638% | 0.1169% | 3.331% | 5.79% | 3.46% | 0.5172% | 0.5408% | 0.05037% | 0.8465% | 1.359% | 0.9114% | 5.009% | 0.3242% | 0.4379% | 0.8144% | 1.858% | 4.613% | 2.153% | 1.502% | 4.704% | 6.519% | 3.285% | 3.012% | 0.8337% | 1.665% | 1.085% | 3.502% | [7 empty] | 0.4168% | 1.9% | 3.232% | 1.38%
|  | Sample Concentration SD |  |  | 421.0025 | 30.9197 | 33.7182 | 97.4238 | 18.9283 | 688.8292 | 250.8189 | 0.083046 | 0.013099 | 0.013888 | 0.24106 | 0.041731 | 6.8805 | 0.0037316 | 0.10086 | 0.065336 | 0.089567 | 0.58009 | 0.15791 | 0.34062 | 0.16307 | 0.1341 | 0.19417 | 0.3289 | 0.16299 | 0.15092 | 0.046121 | 0.084774 | 0.0044716 | 0.012521 | [7 empty] | 0.0015306 | 0.0067068 | 0.015159 | 0.0065027
| 57 | MT 12 | R | 06/02/2023 21:10:20 | 26342.5624 | 1560.1194 | 2680.7375 | 15291.7572 | 16365.7034 | 21281.8487 | 4253.8227 | 2.389 | 2.5949 | 2.7041 | 499.7453 | 4.7999 | 498.9965 | 0.43527 | 1.9143 | 20.7387 | 22.1678 | 75.2042 | 8.8558 | 7.7155 | 7.4343 | 9.0999 | 4.2069 | 5.3851 | 5.364 | 5.1638 | 5.537 | 5.2368 | 0.40404 | 0.35362 | [7 empty] | 0.39152 | 0.36931 | 0.48594 | 0.47577
|  | IS % Recovery | [32 empty] | 88.36% | 79.89% | 94.62% | 78.96% | 84.64% | 78.34% | 94.53%
|  | Sample RSD |  |  | 4.862% | 8.084% | 4.965% | 1.033% | 5.56% | 5.11% | 3.782% | 4.699% | 2.902% | 23.14% | 22.99% | 1.945% | 5.911% | 6.109% | 1.731% | 4.316% | 22.63% | 22.53% | 5.694% | 6.549% | 1.134% | 1.976% | 7.742% | 3.601% | 9.578% | 1.029% | 4.689% | 7.279% | 23.94% | 1.587% | [7 empty] | 21.04% | 20.79% | 3.753% | 2.09%
|  | Sample Concentration SD |  |  | 1280.882 | 126.1133 | 133.1042 | 158.0054 | 909.8681 | 1087.5181 | 160.8929 | 0.11225 | 0.075294 | 0.62559 | 114.9011 | 0.093354 | 29.4963 | 0.026591 | 0.033129 | 0.89513 | 5.0166 | 16.9457 | 0.50423 | 0.50532 | 0.084312 | 0.17977 | 0.3257 | 0.19394 | 0.51375 | 0.053116 | 0.25965 | 0.38118 | 0.096744 | 0.0056135 | [7 empty] | 0.082374 | 0.076783 | 0.018236 | 0.0099452
| 66 | MT 13 | R | 06/02/2023 22:06:09 | 30737.5572 | 1820.8405 | 3100.6238 | 17961.8263 | 19195.8508 | 24300.9309 | 4801.868 | 2.7657 | 2.9328 | 2.8446 | 531.2763 | 5.3806 | 557.2625 | 0.41973 | 2.1147 | 23.0437 | 22.8116 | 77.5063 | 10.0831 | 8.3222 | 8.201 | 9.8608 | 4.8127 | 5.5166 | 5.3644 | 5.3412 | 5.9779 | 5.0301 | 0.45202 | 0.36972 | [7 empty] | 0.39763 | 0.37004 | 0.51607 | 0.47666
|  | IS % Recovery | [32 empty] | 82.49% | 77.12% | 90.35% | 77.39% | 82.22% | 73.12% | 89.77%
|  | Sample RSD |  |  | 1.956% | 6.979% | 2.892% | 1.587% | 5.309% | 3.046% | 4.265% | 0.4151% | 1.618% | 1.913% | 2.022% | 2.177% | 1.238% | 2.618% | 3.629% | 0.812% | 1.249% | 1.982% | 13.54% | 1.427% | 4.58% | 0.3389% | 19.9% | 2.654% | 1.751% | 1.274% | 15.17% | 6.436% | 2.713% | 0.5155% | [7 empty] | 1.834% | 1.343% | 15.53% | 1.434%
|  | Sample Concentration SD |  |  | 601.1015 | 127.0697 | 89.6833 | 284.9736 | 1019.061 | 740.1771 | 204.8184 | 0.011482 | 0.047447 | 0.054414 | 10.7421 | 0.11712 | 6.8961 | 0.010988 | 0.076735 | 0.18712 | 0.28495 | 1.5362 | 1.3653 | 0.11874 | 0.37562 | 0.033414 | 0.95752 | 0.1464 | 0.093943 | 0.068022 | 0.90694 | 0.32374 | 0.012263 | 0.001906 | [7 empty] | 0.0072934 | 0.0049709 | 0.080163 | 0.0068344
| 67 | MT 14 | R | 06/02/2023 22:12:11 | 30703.9049 | 1770.0111 | 3009.2567 | 17140.4128 | 18615.777 | 24097.4289 | 4432.4352 | 2.6847 | 2.8559 | 3.0333 | 583.4065 | 5.853 | 549.5117 | 0.44979 | 2.19 | 22.6285 | 22.5197 | 75.6795 | 9.4708 | 8.3281 | 7.9865 | 9.6979 | 4.3 | 5.5278 | 5.431 | 5.4105 | 5.6735 | 5.4869 | 0.46434 | 0.37948 | [7 empty] | 0.38501 | 0.36444 | 0.49709 | 0.49228
|  | IS % Recovery | [32 empty] | 85.79% | 75.97% | 90.6% | 77.42% | 80.63% | 73.74% | 89.27%
|  | Sample RSD |  |  | 2.234% | 3.461% | 1.881% | 0.07655% | 2.423% | 0.889% | 7.072% | 1.412% | 3.078% | 11% | 7.306% | 0.6111% | 1.058% | 4.187% | 2.145% | 1.816% | 0.9116% | 0.9512% | 2.027% | 2.769% | 4.16% | 1.789% | 6.67% | 3.31% | 1.876% | 0.4904% | 1.614% | 9.852% | 3.438% | 5.8% | [7 empty] | 2.117% | 2.921% | 4.018% | 0.428%
|  | Sample Concentration SD |  |  | 686.0519 | 61.2577 | 56.6074 | 13.1218 | 451.07 | 214.2363 | 313.4665 | 0.037916 | 0.087906 | 0.33366 | 42.623 | 0.035764 | 5.8138 | 0.018832 | 0.046964 | 0.411 | 0.20528 | 0.71986 | 0.19196 | 0.23058 | 0.33226 | 0.1735 | 0.28681 | 0.18295 | 0.10189 | 0.026534 | 0.09155 | 0.54055 | 0.015962 | 0.022009 | [7 empty] | 0.0081514 | 0.010646 | 0.019971 | 0.0021069
| 68 | MT 15 | R | 06/02/2023 22:18:14 | 30329.0572 | 1775.8103 | 2986.0743 | 17146.9435 | 18819.1582 | 24220.4616 | 4441.4852 | 2.6714 | 2.9493 | 2.9503 | 517.054 | 9.7159 | 545.2228 | 0.43221 | 2.7244 | 22.5383 | 22.1493 | 76.1901 | 9.4913 | 7.9056 | 8.0273 | 9.7551 | 4.4638 | 5.4427 | 5.5721 | 5.4362 | 5.7642 | 5.4677 | 0.47372 | 0.39614 | [7 empty] | 0.39311 | 0.37653 | 0.49425 | 0.48125
|  | IS % Recovery | [32 empty] | 84.13% | 75.77% | 89.33% | 76.92% | 80.78% | 74.88% | 89.1%
|  | Sample RSD |  |  | 3.613% | 7.217% | 3.293% | 3.522% | 2.477% | 1.186% | 1.56% | 3.697% | 2.759% | 0.4389% | 0.8361% | 2.259% | 1.035% | 3.364% | 1.856% | 2.3% | 0.6576% | 1.646% | 0.5683% | 7.534% | 4.193% | 0.7789% | 0.8199% | 10.2% | 2.619% | 0.5279% | 2.475% | 11.18% | 1.429% | 3.255% | [7 empty] | 3.803% | 2.037% | 3.703% | 1.053%
|  | Sample Concentration SD |  |  | 1095.7762 | 128.1655 | 98.3391 | 603.8388 | 466.1854 | 287.3071 | 69.3007 | 0.098771 | 0.081371 | 0.012949 | 4.3229 | 0.2195 | 5.6453 | 0.01454 | 0.050559 | 0.5183 | 0.14566 | 1.254 | 0.053943 | 0.59558 | 0.33659 | 0.075982 | 0.0366 | 0.5553 | 0.14594 | 0.028699 | 0.14268 | 0.61146 | 0.0067697 | 0.012895 | [7 empty] | 0.01495 | 0.0076693 | 0.018301 | 0.0050652
| 69 | MT 16 | R | 06/02/2023 22:24:17 | 8952.299 | 1172.257 | 24.6276 | 13630.4333 | 13152.7928 | 9335.3707 | 502.9978 | 0.25554 | 2.7674 | 2.7594 | 1249.2524 | 8.199 | 1139.5852 | 0.2229 | 2.1867 | 22.6325 | 22.9954 | 66.5366 | 7.0708 | 5.1974 | 5.2779 | 6.4969 | 4.7078 | 4.9034 | 5.0627 | 5.0403 | 5.7142 | 4.8317 | 2.231 | 2.2188 | [7 empty] | 11.8326 | 11.2179 | 0.21164 | 0.17442
|  | IS % Recovery | [32 empty] | 78.27% | 78.76% | 94.92% | 79.14% | 84.23% | 79.1% | 93.82%
|  | Sample RSD |  |  | 2.344% | 1.658% | 6.436% | 0.3291% | 4.954% | 1.665% | 5.368% | 4.47% | 1.417% | 1.816% | 1.12% | 0.904% | 0.3393% | 1.829% | 4.706% | 0.4745% | 0.9314% | 1.23% | 16.45% | 5.98% | 4.705% | 1.313% | 26.4% | 5.59% | 1.69% | 3.815% | 22.51% | 4.501% | 2.519% | 1.387% | [7 empty] | 2.057% | 2.648% | 15.9% | 0.9175%
|  | Sample Concentration SD |  |  | 209.8075 | 19.4354 | 1.5851 | 44.8539 | 651.585 | 155.4171 | 27.0034 | 0.011422 | 0.039221 | 0.050111 | 13.9883 | 0.074121 | 3.867 | 0.0040765 | 0.1029 | 0.10738 | 0.21417 | 0.81811 | 1.1632 | 0.3108 | 0.2483 | 0.085302 | 1.243 | 0.2741 | 0.085581 | 0.19227 | 1.2865 | 0.21746 | 0.056199 | 0.030783 | [7 empty] | 0.24336 | 0.29703 | 0.033652 | 0.0016003
| 70 | MT 17 | R | 06/02/2023 22:30:20 | 21229.9516 | 2684.8274 | 52.6549 | 30446.6843 | 28066.5028 | 22237.4235 | 1132.122 | 0.57723 | 6.0914 | 5.8882 | 2749.1227 | 19.764 | 2738.849 | 0.51094 | 4.9551 | 50.6334 | 49.3095 | 144.4342 | 14.2194 | 12.3232 | 13.2062 | 15.7811 | 9.0058 | 12.269 | 12.6082 | 12.1684 | 11.3891 | 11.9798 | 4.7945 | 5.2846 | [7 empty] | 25.856 | 24.7805 | 0.37231 | 0.35988
|  | IS % Recovery | [32 empty] | 82.19% | 72.19% | 88.36% | 76.58% | 80.47% | 77.42% | 83.37%
|  | Sample RSD |  |  | 1.158% | 3.036% | 1.568% | 3.091% | 1.863% | 1.319% | 4.047% | 8.373% | 1.345% | 1.817% | 0.9843% | 10.45% | 0.7678% | 0.9783% | 3.884% | 1.048% | 1.616% | 1.018% | 1.474% | 4.437% | 9.783% | 10.26% | 4.657% | 5.031% | 8.216% | 10.53% | 1.298% | 6.49% | 1.406% | 9.888% | [7 empty] | 0.6166% | 0.8778% | 1.591% | 0.6524%
|  | Sample Concentration SD |  |  | 245.825 | 81.523 | 0.82538 | 941.031 | 522.7482 | 293.2938 | 45.8146 | 0.048329 | 0.081905 | 0.10696 | 27.0583 | 2.0652 | 21.0287 | 0.0049987 | 0.19248 | 0.53046 | 0.79699 | 1.4705 | 0.20959 | 0.54675 | 1.292 | 1.6194 | 0.41937 | 0.6172 | 1.0359 | 1.281 | 0.14778 | 0.77755 | 0.067392 | 0.52252 | [7 empty] | 0.15943 | 0.21752 | 0.0059232 | 0.0023477
| 71 | MT 18 | R | 06/02/2023 22:36:32 | 18623.5557 | 2281.0516 | 46.3191 | 27615.8101 | 24636.082 | 19477.6248 | 1004.3962 | 0.51498 | 4.7032 | 4.7108 | 2418.7587 | 15.5559 | 2185.5614 | 0.44386 | 4.4388 | 43.0445 | 43.2757 | 124.9763 | 12.5026 | 10.3976 | 10.5607 | 12.5271 | 7.7149 | 9.6203 | 9.9076 | 9.6502 | 9.9133 | 9.835 | 4.2291 | 4.2799 | [7 empty] | 22.396 | 21.2256 | 0.29733 | 0.28141
|  | IS % Recovery | [32 empty] | 81.35% | 75.75% | 87.51% | 75.93% | 79.75% | 73.85% | 86.87%
|  | Sample RSD |  |  | 2.927% | 6.235% | 3.94% | 1.868% | 4.343% | 3.974% | 14.97% | 1.436% | 1.523% | 1.305% | 1.414% | 2.313% | 1.495% | 4.212% | 3.125% | 2.324% | 1.031% | 1.321% | 1.544% | 5.128% | 0.368% | 0.9555% | 7.371% | 4.514% | 2.545% | 2.362% | 0.8598% | 4.995% | 0.3205% | 1.79% | [7 empty] | 0.8716% | 0.7466% | 2.133% | 2.252%
|  | Sample Concentration SD |  |  | 545.0682 | 142.2279 | 1.8251 | 515.8659 | 1069.8434 | 773.9583 | 150.3418 | 0.0073941 | 0.071615 | 0.061474 | 34.2131 | 0.35986 | 32.6744 | 0.018695 | 0.13869 | 1.0002 | 0.44638 | 1.6511 | 0.19302 | 0.53323 | 0.038858 | 0.1197 | 0.56863 | 0.43425 | 0.25215 | 0.22794 | 0.08524 | 0.49126 | 0.013553 | 0.076609 | [7 empty] | 0.19521 | 0.15847 | 0.0063421 | 0.0063388
| 72 | MT 19 | R | 06/02/2023 22:42:43 | 9214.5213 | 1187.8096 | 24.0498 | 13890.2763 | 13076.5655 | 9523.619 | 505.9479 | 0.25464 | 2.286 | 2.2286 | 1234.8446 | 7.8282 | 1084.3919 | 0.22943 | 3.4478 | 24.9884 | 25.5039 | 63.8711 | 5.9391 | 5.1523 | 5.0124 | 6.293 | 3.8344 | 4.5419 | 4.9302 | 4.797 | 4.6841 | 4.6831 | 2.0523 | 2.0806 | [7 empty] | 11.254 | 10.6989 | 0.18799 | 0.17659
|  | IS % Recovery | [32 empty] | 83.77% | 74.99% | 88.16% | 76.25% | 82.36% | 73.95% | 91.01%
|  | Sample RSD |  |  | 3.218% | 5.356% | 2.473% | 0.03002% | 3.965% | 3.621% | 16.34% | 8.23% | 17.38% | 21.83% | 20.06% | 2.729% | 3.941% | 8.51% | 1.886% | 3.94% | 18.3% | 18.8% | 1.054% | 7.567% | 1.895% | 1.74% | 8.559% | 9.774% | 2.513% | 5.508% | 1.792% | 0.9775% | 19.96% | 1.988% | [7 empty] | 18.13% | 19.46% | 0.8672% | 1.548%
|  | Sample Concentration SD |  |  | 296.541 | 63.6148 | 0.5948 | 4.1695 | 518.4888 | 344.8605 | 82.6867 | 0.020956 | 0.39741 | 0.48646 | 247.7038 | 0.21365 | 42.7351 | 0.019525 | 0.065015 | 0.98459 | 4.6665 | 12.0084 | 0.062624 | 0.38989 | 0.094968 | 0.10952 | 0.32817 | 0.44391 | 0.12391 | 0.26424 | 0.08395 | 0.045776 | 0.40957 | 0.041354 | [7 empty] | 2.0405 | 2.0822 | 0.0016303 | 0.0027341
| 73 | MT 20 | R | 06/02/2023 22:48:45 | 20144.4177 | 2380.2906 | 48.5462 | 28884.8142 | 25867.958 | 20396.1114 | 995.6703 | 0.50021 | 4.7815 | 4.8015 | 2499.7317 | 16.1044 | 2265.957 | 0.44006 | 4.0102 | 45.2108 | 45.3714 | 130.9351 | 12.2058 | 10.8696 | 10.7466 | 12.8257 | 7.7118 | 9.8549 | 10.3885 | 9.8474 | 9.8735 | 9.8703 | 4.1843 | 4.2133 | [7 empty] | 23.357 | 22.0995 | 0.30253 | 0.28984
|  | IS % Recovery | [32 empty] | 82.37% | 75.26% | 85.07% | 74.03% | 79.43% | 72.74% | 86.89%
|  | Sample RSD |  |  | 0.7387% | 5.287% | 6.148% | 2.169% | 2.172% | 2.764% | 5.283% | 4.053% | 5.742% | 2.826% | 0.9874% | 1.259% | 0.9009% | 8.308% | 1.717% | 0.8351% | 0.9958% | 0.877% | 0.8% | 4.6% | 3.251% | 1.371% | 5.313% | 4.487% | 1.507% | 4.232% | 3.583% | 0.9645% | 0.6743% | 0.7527% | [7 empty] | 0.2579% | 0.4491% | 2.264% | 1.34%
|  | Sample Concentration SD |  |  | 148.806 | 125.8354 | 2.9846 | 626.6289 | 561.7404 | 563.802 | 52.6038 | 0.020273 | 0.27457 | 0.13567 | 24.6822 | 0.20281 | 20.4141 | 0.036562 | 0.068858 | 0.37755 | 0.45181 | 1.1483 | 0.097649 | 0.50003 | 0.34938 | 0.17587 | 0.40974 | 0.44216 | 0.15654 | 0.41675 | 0.35381 | 0.095201 | 0.028214 | 0.031715 | [7 empty] | 0.060246 | 0.099243 | 0.0068492 | 0.0038838
| 74 | MT 21 | R | 06/02/2023 22:54:57 | 10368.6275 | 1389.894 | 26.1933 | 16791.9613 | 13708.9997 | 10586.5301 | 573.1915 | 0.26084 | 2.8347 | 2.8774 | 1356.5094 | 8.6074 | 1209.7909 | 0.23841 | 2.4675 | 23.6617 | 24.0391 | 72.0528 | 6.2961 | 5.5014 | 5.3945 | 6.5977 | 4.0134 | 5.1139 | 5.1603 | 5.0197 | 5.1573 | 4.7603 | 2.2103 | 2.2105 | [7 empty] | 12.2157 | 11.7715 | 0.21137 | 0.19643
|  | IS % Recovery | [32 empty] | 86.08% | 77.21% | 86.72% | 75.52% | 84.58% | 68.91% | 91.35%
|  | Sample RSD |  |  | 2.507% | 5.05% | 4.533% | 0.5244% | 2.28% | 2.197% | 1.873% | 4.271% | 1.865% | 7.809% | 1.307% | 3.111% | 0.8425% | 2.192% | 2.145% | 0.7357% | 1.315% | 1.409% | 3.468% | 3.006% | 5.297% | 3.952% | 7.12% | 2.228% | 1.102% | 6.659% | 1.441% | 4.495% | 1.918% | 2.771% | [7 empty] | 2.339% | 0.5922% | 6.683% | 8.456%
|  | Sample Concentration SD |  |  | 259.9156 | 70.1872 | 1.1874 | 88.0567 | 312.5811 | 232.5807 | 10.7362 | 0.01114 | 0.05286 | 0.22468 | 17.7266 | 0.26782 | 10.192 | 0.0052268 | 0.052935 | 0.17408 | 0.31603 | 1.0151 | 0.21837 | 0.16539 | 0.28573 | 0.26077 | 0.28574 | 0.11395 | 0.056849 | 0.33424 | 0.07432 | 0.21399 | 0.042394 | 0.061264 | [7 empty] | 0.28573 | 0.069706 | 0.014127 | 0.01661
| 81 | MT 22 | R | 06/02/2023 23:38:12 | 7350.3081 | 1011.1728 | 19.7143 | 10961.948 | 10871.874 | 7490.4282 | 374.7182 | 0.22444 | 1.6051 | 1.6466 | 941.3146 | 6.019 | 822.8003 | 0.16519 | 1.4798 | 16.2837 | 16.3145 | 46.6807 | 4.2972 | 3.6982 | 3.867 | 4.5236 | 2.8171 | 3.5258 | 3.6479 | 3.486 | 3.4347 | 3.3286 | 1.4823 | 1.4479 | [7 empty] | 8.1357 | 7.6787 | 0.17965 | 0.17139
|  | IS % Recovery | [32 empty] | 84.62% | 74.16% | 86.67% | 75.85% | 81.5% | 71.95% | 83.91%
|  | Sample RSD |  |  | 1.471% | 6.384% | 3.195% | 2.379% | 1.222% | 1.547% | 6.474% | 8.328% | 4.861% | 4.979% | 1.846% | 2.388% | 1.253% | 8.669% | 2.405% | 2.335% | 1.228% | 0.7782% | 1.362% | 3.302% | 3.277% | 1.826% | 13.12% | 6.75% | 1.864% | 0.502% | 3.353% | 6.862% | 2.254% | 1.443% | [7 empty] | 1.531% | 0.895% | 0.8977% | 3.239%
|  | Sample Concentration SD |  |  | 108.1141 | 64.5541 | 0.62996 | 260.7811 | 132.8089 | 115.8894 | 24.259 | 0.018692 | 0.078025 | 0.081978 | 17.3742 | 0.14374 | 10.3107 | 0.014321 | 0.035586 | 0.38026 | 0.2004 | 0.36327 | 0.058508 | 0.12212 | 0.12672 | 0.082611 | 0.36948 | 0.23799 | 0.068002 | 0.017498 | 0.11518 | 0.2284 | 0.033413 | 0.020899 | [7 empty] | 0.12458 | 0.068726 | 0.0016127 | 0.0055506
| 84 | MT 23 | R | 06/02/2023 23:56:47 | 4167.6084 | 531.5582 | 12.1815 | 6379.9953 | 7387.8538 | 4361.5988 | 237.885 | 0.11409 | 1.1276 | 1.0868 | 504.7472 | 3.5213 | 508.0426 | 0.10454 | 0.94733 | 9.5858 | 9.6734 | 28.5944 | 2.6441 | 2.2552 | 2.1525 | 2.647 | 2.0474 | 2.193 | 2.0502 | 1.9592 | 2.189 | 2.1011 | 0.96868 | 0.92851 | [7 empty] | 4.8808 | 4.6181 | 0.10234 | 0.083308
|  | IS % Recovery | [32 empty] | 81.67% | 76.92% | 86.97% | 75.88% | 90.86% | 73.65% | 91.63%
|  | Sample RSD |  |  | 2.139% | 8.176% | 7.676% | 2.205% | 2.583% | 1.703% | 2.773% | 15.09% | 3.953% | 3.137% | 0.7619% | 4.77% | 0.751% | 5.75% | 2.933% | 1.692% | 0.843% | 1.599% | 3.264% | 5.757% | 7.235% | 3.705% | 1.584% | 8.91% | 7.373% | 4.537% | 4.393% | 9.768% | 2.272% | 2.797% | [7 empty] | 2.117% | 1.348% | 1.472% | 21.39%
|  | Sample Concentration SD |  |  | 89.1371 | 43.4602 | 0.93502 | 140.6898 | 190.8542 | 74.2684 | 6.5964 | 0.017215 | 0.044571 | 0.034089 | 3.8458 | 0.16795 | 3.8155 | 0.0060108 | 0.02779 | 0.16217 | 0.08155 | 0.45723 | 0.086312 | 0.12983 | 0.15573 | 0.098061 | 0.032424 | 0.1954 | 0.15115 | 0.088889 | 0.096156 | 0.20523 | 0.022008 | 0.025974 | [7 empty] | 0.10332 | 0.062233 | 0.001506 | 0.017823
| 85 | MT 24 | R | 06/03/2023 00:02:49 | 3953.5362 | 461.5362 | 10.051 | 5575.9935 | 6759.4653 | 3941.4553 | 190.6957 | 0.11138 | 1.0913 | 1.1048 | 452.6408 | 3.3004 | 451.843 | 0.093842 | 0.91476 | 9.0554 | 9.0861 | 26.0721 | 2.4435 | 1.9862 | 2.0786 | 2.518 | 1.6085 | 2.2221 | 2.0217 | 1.8397 | 1.9844 | 1.9986 | 0.82636 | 0.84657 | [7 empty] | 4.4974 | 4.2525 | 0.074568 | 0.071418
|  | IS % Recovery | [32 empty] | 84.56% | 75% | 88.1% | 77.43% | 84.13% | 77.18% | 87.81%
|  | Sample RSD |  |  | 4.201% | 11.86% | 11.57% | 1.326% | 9.347% | 3.075% | 8.43% | 11.3% | 4.961% | 5.011% | 0.9568% | 1.859% | 2.521% | 2.769% | 9.63% | 3.016% | 0.5661% | 0.809% | 1.736% | 19.34% | 2.254% | 1.661% | 14.9% | 11.35% | 1.896% | 5.141% | 4.86% | 11.58% | 0.7362% | 0.7873% | [7 empty] | 1.484% | 0.5749% | 2.53% | 0.3167%
|  | Sample Concentration SD |  |  | 166.096 | 54.7474 | 1.1625 | 73.9154 | 631.8076 | 121.2158 | 16.0755 | 0.012581 | 0.054143 | 0.055358 | 4.3307 | 0.061355 | 11.3925 | 0.0025986 | 0.088096 | 0.27312 | 0.051438 | 0.21091 | 0.04243 | 0.38406 | 0.046854 | 0.041825 | 0.23965 | 0.25217 | 0.038328 | 0.094571 | 0.096443 | 0.23135 | 0.0060839 | 0.0066649 | [7 empty] | 0.066748 | 0.024448 | 0.0018864 | 0.0002262
| 86 | MT 25 | R | 06/03/2023 00:08:52 | 4647.6853 | 593.1143 | 13.2612 | 6963.9529 | 7198.9212 | 4765.6722 | 296.9423 | 0.13343 | 0.79615 | 0.85817 | 523.3037 | 3.8095 | 514.8923 | 0.10586 | 0.68377 | 10.5038 | 10.8708 | 30.6891 | 2.894 | 2.3309 | 2.4426 | 2.9203 | 2.3797 | 2.4727 | 2.1586 | 2.2642 | 2.383 | 2.0209 | 0.87431 | 0.8895 | [7 empty] | 5.1671 | 4.909 | 0.10707 | 0.095725
|  | IS % Recovery | [32 empty] | 78.45% | 74.42% | 87% | 76.11% | 82.48% | 72.87% | 86.03%
|  | Sample RSD |  |  | 2.215% | 4.165% | 2.756% | 2.359% | 1.975% | 3.877% | 10.77% | 10.29% | 4.287% | 3.606% | 0.8148% | 2.519% | 2.701% | 4.576% | 6.811% | 1.122% | 1.892% | 1.091% | 4.407% | 7.835% | 3.137% | 1.679% | 11.63% | 10.83% | 7.768% | 3.331% | 3.68% | 9.737% | 1.554% | 2.935% | [7 empty] | 0.59% | 0.745% | 6.13% | 1.564%
|  | Sample Concentration SD |  |  | 102.9453 | 24.705 | 0.3655 | 164.2911 | 142.1447 | 184.776 | 31.9723 | 0.013729 | 0.034128 | 0.030943 | 4.2638 | 0.095973 | 13.9089 | 0.0048445 | 0.04657 | 0.11788 | 0.20565 | 0.33479 | 0.12755 | 0.18262 | 0.076614 | 0.049041 | 0.27666 | 0.2677 | 0.16768 | 0.075426 | 0.087687 | 0.19679 | 0.013589 | 0.026106 | [7 empty] | 0.030484 | 0.036574 | 0.0065636 | 0.001497
| 87 | MT 26 | R | 06/03/2023 00:14:55 | 4284.898 | 525.8116 | 14.2656 | 6463.3928 | 7101.0936 | 4202.6383 | 206.2797 | 0.12336 | 1.2065 | 1.1962 | 486.3739 | 3.5335 | 473.6568 | 0.097103 | 0.9879 | 9.6872 | 9.9788 | 28.048 | 3.0733 | 2.2087 | 2.2832 | 2.6812 | 2.696 | 2.1316 | 2.017 | 2.0221 | 2.6589 | 1.9497 | 0.89117 | 0.91208 | [7 empty] | 4.7494 | 4.5586 | 0.14609 | 0.08439
|  | IS % Recovery | [32 empty] | 70.99% | 74.05% | 87.36% | 75.42% | 96.73% | 72.43% | 84.06%
|  | Sample RSD |  |  | 0.7031% | 10.79% | 4.823% | 0.8579% | 4.235% | 1.542% | 17.83% | 13.12% | 10.7% | 0.8781% | 1.015% | 2.133% | 0.6893% | 2.698% | 6.176% | 1.082% | 1.308% | 1.092% | 24.2% | 8.856% | 2.884% | 2.395% | 46.97% | 12.91% | 7.727% | 3.962% | 33.36% | 14.06% | 1.708% | 1.956% | [7 empty] | 2.539% | 1.031% | 28.61% | 8.269%
|  | Sample Concentration SD |  |  | 30.1252 | 56.7093 | 0.68796 | 55.4478 | 300.7392 | 64.7863 | 36.7868 | 0.016186 | 0.12907 | 0.010503 | 4.9377 | 0.075363 | 3.2649 | 0.0026196 | 0.061012 | 0.1048 | 0.13056 | 0.30617 | 0.74389 | 0.19561 | 0.065841 | 0.064218 | 1.2663 | 0.27529 | 0.15586 | 0.080117 | 0.88695 | 0.27419 | 0.015225 | 0.017842 | [7 empty] | 0.12059 | 0.047005 | 0.041797 | 0.0069778
| 88 | MT 27 | R | 06/03/2023 00:21:03 | 12290.216 | 1704.7488 | 29.9034 | 17410.0987 | 16171.658 | 12068.5986 | 733.8392 | 0.30382 | 2.7257 | 2.6016 | 1472.115 | 9.683 | 1274.2666 | 0.25533 | 2.361 | 24.8398 | 24.9139 | 72.3287 | 6.6712 | 5.7115 | 6.0217 | 7.1672 | 4.1722 | 5.7849 | 5.4283 | 5.4354 | 5.4076 | 5.2683 | 2.3074 | 2.3514 | [7 empty] | 12.4902 | 11.8986 | 0.1834 | 0.17704
|  | IS % Recovery | [32 empty] | 85.13% | 74.03% | 85.04% | 75.55% | 82.16% | 72.67% | 85.32%
|  | Sample RSD |  |  | 2.043% | 1.108% | 4.583% | 2.607% | 0.5448% | 1.286% | 5.504% | 3.12% | 1.677% | 1.913% | 0.7257% | 1.515% | 1.346% | 3.758% | 4.158% | 1.618% | 0.9869% | 1.014% | 3.993% | 4.545% | 4.862% | 0.448% | 9.357% | 4.399% | 5.274% | 2.699% | 8.413% | 6.543% | 0.6589% | 0.2412% | [7 empty] | 1.774% | 0.8392% | 3.835% | 0.2609%
|  | Sample Concentration SD |  |  | 251.0986 | 18.8815 | 1.3706 | 453.8726 | 88.0963 | 155.2295 | 40.3939 | 0.0094788 | 0.045714 | 0.049778 | 10.6834 | 0.14674 | 17.1476 | 0.0095955 | 0.098181 | 0.4018 | 0.24588 | 0.73359 | 0.26641 | 0.25961 | 0.29277 | 0.032112 | 0.39037 | 0.25447 | 0.2863 | 0.14668 | 0.45494 | 0.34472 | 0.015204 | 0.0056711 | [7 empty] | 0.22156 | 0.099851 | 0.0070329 | 0.00046195
| 89 | MT 28 | R | 06/03/2023 00:27:06 | 13748.68 | 1689.5292 | 32.522 | 18895.0879 | 17998.3442 | 13362.4792 | 689.1893 | 0.32406 | 2.5894 | 2.6384 | 1608.5639 | 10.0255 | 1429.693 | 0.28189 | 2.3779 | 28.2912 | 27.2659 | 80.1424 | 7.7783 | 6.8868 | 6.5432 | 7.6258 | 5.2128 | 6.0895 | 6.0755 | 5.9121 | 6.3238 | 5.9183 | 2.4737 | 2.4263 | [7 empty] | 13.7015 | 13.2527 | 0.19538 | 0.17245
|  | IS % Recovery | [32 empty] | 77.61% | 75.05% | 81.58% | 74.1% | 80.03% | 72.15% | 84.7%
|  | Sample RSD |  |  | 3.866% | 6.313% | 2.693% | 0.8543% | 6.312% | 4.65% | 10.96% | 9.451% | 5.521% | 1.111% | 1.083% | 15.39% | 5.883% | 5.313% | 6.422% | 5.057% | 1.695% | 1.146% | 5.546% | 11.3% | 15.81% | 15.09% | 2.687% | 23.65% | 13.55% | 16.35% | 6.987% | 13.71% | 2.118% | 15.25% | [7 empty] | 0.5278% | 0.8633% | 7.68% | 1.308%
|  | Sample Concentration SD |  |  | 531.4886 | 106.6594 | 0.8757 | 161.4205 | 1136.0461 | 621.4114 | 75.5378 | 0.030626 | 0.14297 | 0.029325 | 17.4134 | 1.5431 | 84.1078 | 0.014976 | 0.15272 | 1.4306 | 0.46216 | 0.91806 | 0.4314 | 0.77809 | 1.0342 | 1.1511 | 0.14005 | 1.4404 | 0.82296 | 0.96654 | 0.44182 | 0.81135 | 0.052391 | 0.36997 | [7 empty] | 0.072316 | 0.11441 | 0.015005 | 0.0022564
| 90 | MT 29 | R | 06/03/2023 00:33:08 | 25930.7988 | 3315.7679 | 59.2385 | 35007.4459 | 30986.87 | 25936.1238 | 1380.7219 | 0.60663 | 5.4587 | 5.6832 | 3144.7884 | 19.8541 | 2626.8061 | 0.52079 | 5.0038 | 51.8753 | 52.5969 | 154.718 | 14.9214 | 12.7158 | 13.0098 | 15.0609 | 9.4267 | 12.3204 | 11.871 | 11.6849 | 12.2219 | 11.6484 | 4.9995 | 5.0585 | [7 empty] | 26.9175 | 25.6433 | 0.33965 | 0.29001
|  | IS % Recovery | [32 empty] | 78.6% | 72.28% | 81.11% | 71.42% | 84.91% | 74.36% | 81.08%
|  | Sample RSD |  |  | 2.106% | 3.027% | 3.002% | 0.6309% | 1.438% | 1.397% | 13.94% | 8.096% | 3.78% | 4.424% | 3.251% | 2.08% | 3.354% | 1.728% | 0.9347% | 2.718% | 2.868% | 3.206% | 2.57% | 3.82% | 6.424% | 1.977% | 4.226% | 3.577% | 2.352% | 0.532% | 1.318% | 1.767% | 2.741% | 1.798% | [7 empty] | 3.484% | 3.874% | 1.861% | 20.24%
|  | Sample Concentration SD |  |  | 546.2137 | 100.3848 | 1.7783 | 220.8453 | 445.6703 | 362.2029 | 192.5151 | 0.049114 | 0.20636 | 0.2514 | 102.2307 | 0.41299 | 88.0906 | 0.0090003 | 0.046773 | 1.4097 | 1.5087 | 4.9608 | 0.38344 | 0.48578 | 0.8358 | 0.29774 | 0.39842 | 0.44072 | 0.27919 | 0.06216 | 0.16108 | 0.20583 | 0.13705 | 0.090942 | [7 empty] | 0.93779 | 0.99332 | 0.0063215 | 0.058696
| 98 | MT 30 | R | 06/03/2023 01:23:04 | -110.7365 | 221.1253 | 728.2531 | 2886.1545 | 4884.4676 | 23.8532 | 2048.7213 | 0.91961 | 0.6147 | 0.64576 | 768.3711 | 121.3975 | 585.3216 | 0.57087 | 1.006 | 70.9182 | 80.1046 | 154.8468 | 1.4167 | 1.4386 | 1.3133 | 1.5507 | 0.44731 | 0.42512 | 0.37839 | 0.44905 | 0.50002 | 0.44832 | 0.35519 | 0.30881 | [7 empty] | 1.8206 | 1.7278 | 2.2988 | 2.4146
|  | IS % Recovery | [32 empty] | 87.88% | 74.18% | 83.86% | 68.42% | 79.2% | 74.7% | 84.66%
|  | Sample RSD |  |  | 5.106% | 9.317% | 3.099% | 4.249% | 1.862% | 43.69% | 7.28% | 1.814% | 2.209% | 2.047% | 7.813% | 3.373% | 2.749% | 5.812% | 4.025% | 2.346% | 5.245% | 4.769% | 0.9263% | 8.015% | 3.952% | 4.201% | 16.94% | 15.67% | 3.221% | 12.63% | 8.191% | 23.22% | 6.608% | 5.629% | [7 empty] | 6.179% | 4.421% | 0.9682% | 0.5173%
|  | Sample Concentration SD |  |  | 5.6538 | 20.6025 | 22.569 | 122.6185 | 90.949 | 10.4216 | 149.1558 | 0.016682 | 0.013582 | 0.013215 | 60.0365 | 4.0943 | 16.0918 | 0.033181 | 0.040489 | 1.6635 | 4.2014 | 7.385 | 0.013123 | 0.1153 | 0.051898 | 0.065145 | 0.075753 | 0.066608 | 0.012186 | 0.056733 | 0.040956 | 0.1041 | 0.02347 | 0.017382 | [7 empty] | 0.1125 | 0.076378 | 0.022257 | 0.01249
| 99 | MT 31 | R | 06/03/2023 01:29:15 | 67.1371 | 422.5248 | 548.3727 | 2843.9193 | 3847.9022 | 742.1862 | 1148.704 | 0.66841 | 0.46758 | 0.45224 | 429.6306 | 87.6994 | 434.4449 | 0.56428 | 0.64272 | 52.3942 | 51.031 | 160.4809 | 1.273 | 1.2283 | 1.1539 | 1.3629 | 0.4243 | 0.25557 | 0.3279 | 0.29876 | 0.35966 | 0.25233 | 0.17954 | 0.19105 | [7 empty] | 1.7808 | 1.6966 | 1.6865 | 1.7046
|  | IS % Recovery | [32 empty] | 85.8% | 75.76% | 84.55% | 76.69% | 79.83% | 71.05% | 86.61%
|  | Sample RSD |  |  | 22.37% | 4.343% | 3.392% | 2.847% | 9.62% | 6.21% | 7.881% | 3.222% | 9.721% | 5.678% | 5.811% | 3.127% | 1.493% | 1.272% | 3.517% | 2.507% | 6.198% | 5.607% | 2.905% | 7.175% | 5.801% | 2.287% | 35.75% | 7.471% | 8.088% | 11.1% | 24.87% | 20.15% | 3.699% | 6.893% | [7 empty] | 5.362% | 4.596% | 1.479% | 1.366%
|  | Sample Concentration SD |  |  | 15.0179 | 18.3499 | 18.6017 | 80.9803 | 370.1788 | 46.089 | 90.531 | 0.021536 | 0.045452 | 0.025679 | 24.9637 | 2.7427 | 6.4875 | 0.0071799 | 0.022603 | 1.3133 | 3.1631 | 8.9976 | 0.036983 | 0.088128 | 0.066935 | 0.031173 | 0.1517 | 0.019092 | 0.02652 | 0.033153 | 0.089442 | 0.050843 | 0.0066403 | 0.013169 | [7 empty] | 0.095478 | 0.077979 | 0.024944 | 0.023287
| 118 | MT 31 | R | 06/03/2023 03:26:12 | 83.5412 | 466.9858 | 602.5057 | 3024.0423 | 4750.0796 | 754.125 | 1249.7838 | 0.73216 | 0.47644 | 0.43272 | 433.9076 | 90.3855 | 454.3924 | 0.58535 | 0.64747 | 54.4451 | 50.5312 | 162.0076 | 1.2794 | 1.1534 | 1.2055 | 1.4148 | 0.37912 | 0.3112 | 0.30343 | 0.3073 | 0.32009 | 0.26047 | 0.17255 | 0.18702 | [7 empty] | 1.7646 | 1.677 | 1.6107 | 1.6969
|  | IS % Recovery | [32 empty] | 86.63% | 70.82% | 80.75% | 72.78% | 74.91% | 68.26% | 80.3%
|  | Sample RSD |  |  | 17.83% | 13.18% | 2.139% | 1.99% | 4.064% | 9.164% | 3.943% | 4.525% | 4.884% | 17.74% | 15.91% | 1.564% | 1.023% | 4.959% | 3.793% | 1.854% | 15.16% | 14.15% | 1.32% | 1.576% | 1.051% | 2.538% | 38.74% | 26.83% | 13.14% | 17.31% | 17.44% | 23.44% | 20.31% | 5.563% | [7 empty] | 15.24% | 14.59% | 2.643% | 0.9184%
|  | Sample Concentration SD |  |  | 14.8962 | 61.5362 | 12.8902 | 60.1744 | 193.0571 | 69.1114 | 49.2736 | 0.033131 | 0.023271 | 0.076757 | 69.016 | 1.4132 | 4.6462 | 0.029026 | 0.024557 | 1.0092 | 7.6592 | 22.9227 | 0.016887 | 0.018175 | 0.012667 | 0.035911 | 0.14686 | 0.083484 | 0.039857 | 0.053209 | 0.055836 | 0.061045 | 0.035046 | 0.010403 | [7 empty] | 0.26886 | 0.24471 | 0.042567 | 0.015584

## Calibration Curves 12
Calibration Table - Na KED 23/23
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 0.33333 | 173.2051 | 0 |  | 0 | 0
Cal. Std.5 | 0.1139 | 2.1995 | 1000 | 990.5164 | -9.4836 | -0.94836
Cal. Std.6 | 0.53054 | 2.5241 | 5000 | 5157.6412 | 157.6412 | 3.1528
Cal. Std.7 | 1.04 | 0.77036 | 10000 | 10252.7747 | 252.7747 | 2.5277
Cal. Std.8 | 2.6621 | 6.5269 | 25000 | 26477.3189 | 1477.3189 | 5.9093
Cal. Std.9 | 4.6893 | 0.93998 | 50000 | 46753.0659 | -3246.9341 | -6.4939
Cal. Std.10 | 9.5983 | 2.2169 | 100000 | 95852.3792 | -4147.6208 | -4.1476
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99868 | 9.9981e-05 | 0.01487 | 2.5711e-06 | 0.0061399 | 0.002844 | 1.7028
Calibration Table - Mg KED 24/24
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 0 |  | 0 |  | 0 | 0
Cal. Std.5 | 0.044287 | 3.5676 | 1000 | 999.0291 | -0.97094 | -0.097094
Cal. Std.6 | 0.21514 | 3.0126 | 5000 | 5044.1152 | 44.1152 | 0.8823
Cal. Std.7 | 0.42109 | 1.7193 | 10000 | 9920.0692 | -79.9308 | -0.79931
Cal. Std.8 | 1.0582 | 5.2571 | 25000 | 25003.5244 | 3.5244 | 0.014098
Cal. Std.9 | 1.852 | 0.75839 | 50000 | 43797.3242 | -6202.6758 | -12.4054
Cal. Std.10 | 3.6634 | 1.7236 | 100000 | 86684.2364 | -13315.7636 | -13.3158
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99997 | 4.2237e-05 | 0.0020901 | 2.3563e-07 | 0.00045956 | 0 | 0
Calibration Table - Al KED 27/27
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 4 | 50 | 0 |  | 0 | 0
Cal. Std.5 | 0.0042525 | 28.762 | 1 | 1.0046 | 0.0046175 | 0.46175
Cal. Std.6 | 0.013117 | 3.1637 | 5 | 4.8911 | -0.10888 | -2.1776
Cal. Std.7 | 0.023857 | 1.7912 | 10 | 9.5999 | -0.40013 | -4.0013
Cal. Std.8 | 0.066153 | 5.6739 | 25 | 28.1438 | 3.1438 | 12.5751
Cal. Std.9 | 0.1091 | 4.3067 | 50 | 46.9741 | -3.0259 | -6.0517
Cal. Std.10 | 0.22821 | 3.1358 | 100 | 99.1938 | -0.80619 | -0.80619
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99744 | 0.0022809 | 0.0019611 | 8.1699e-05 | 0.0001951 | -0.00039195 | 0.25319
Calibration Table - P Oshift 31/47
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 2280.1821 | 3.2494 | 0 |  | 0 | 0
Cal. Std.4 | 0.023222 | 4.773 | 0.5 | 1.8636 | 1.3636 | 272.7113
Cal. Std.5 | 0.024893 | 3.7138 | 1 | 1.9976 | 0.99762 | 99.7621
Cal. Std.6 | 0.079664 | 8.0236 | 5 | 6.3929 | 1.3929 | 27.8584
Cal. Std.7 | 0.13764 | 1.0309 | 10 | 11.0451 | 1.0451 | 10.4514
Cal. Std.8 | 0.32982 | 1.0494 | 25 | 26.4679 | 1.4679 | 5.8715
Cal. Std.9 | 0.62341 | 0.37528 | 50 | 50.028 | 0.027971 | 0.055942
Cal. Std.10 | 1.2391 | 0.47343 | 100 | 99.4349 | -0.56509 | -0.56509
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99975 | 0.012461 | 0 | 0.00012345 | 0 | 0.00047374 | 0.13596
Calibration Table - S (34) -1 34/34
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 372.3382 | 5.0699 | 0 |  | 0 | 0
Cal. Std.6 | 0.071543 | 5.826 | 5000 | 4949.6683 | -50.3317 | -1.0066
Cal. Std.7 | 0.14873 | 6.0781 | 10000 | 10079.4311 | 79.4311 | 0.79431
Cal. Std.8 | 0.3936 | 5.6309 | 25000 | 26352.1901 | 1352.1901 | 5.4088
Cal. Std.9 | 0.71747 | 1.6866 | 50000 | 47875.4863 | -2124.5137 | -4.249
Cal. Std.10 | 1.4876 | 2.5201 | 100000 | 99052.5906 | -947.4094 | -0.94741
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99895 | 1.5047e-05 | -0.0029373 | 3.9824e-07 | 0.0039013 | 0.69958 | 421.9484
Calibration Table - K KED 39/39
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 9.3333 | 26.9637 | 0 |  | 0 | 0
Cal. Std.5 | 0.060755 | 2.1916 | 1000 | 993.6008 | -6.3992 | -0.63992
Cal. Std.6 | 0.27785 | 2.4651 | 5000 | 5129.6499 | 129.6499 | 2.593
Cal. Std.7 | 0.53719 | 0.58088 | 10000 | 10070.508 | 70.508 | 0.70508
Cal. Std.8 | 1.389 | 6.4828 | 25000 | 26299.7528 | 1299.7528 | 5.199
Cal. Std.9 | 2.4268 | 1.0003 | 50000 | 46071.4173 | -3928.5827 | -7.8572
Cal. Std.10 | 4.7968 | 2.2646 | 100000 | 91223.4826 | -8776.5174 | -8.7765
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99852 | 5.2489e-05 | 0.0086024 | 1.6525e-06 | 0.0036026 | -0.016507 | 13.756
Calibration Table - Ca KED 43/43
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 0.33333 | 173.2051 | 0 |  | 0 | 0
Cal. Std.3 | 0.0013913 | 23.4758 | 0.1 | 72.2499 | 72.1499 | 72149.9109
Cal. Std.5 | 0.01856 | 4.2605 | 1000 | 1002.1128 | 2.1128 | 0.21128
Cal. Std.6 | 0.091125 | 4.5122 | 5000 | 4932.2957 | -67.7043 | -1.3541
Cal. Std.7 | 0.18335 | 1.789 | 10000 | 9927.3956 | -72.6044 | -0.72604
Cal. Std.8 | 0.48351 | 7.0793 | 25000 | 26184.4979 | 1184.4979 | 4.738
Cal. Std.9 | 0.89674 | 1.0009 | 50000 | 48565.4265 | -1434.5735 | -2.8691
Cal. Std.10 | 1.9285 | 1.3727 | 100000 | 104446.7551 | 4446.7551 | 4.4468
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99949 | 1.8463e-05 | 5.7271e-05 | 3.4055e-07 | 0.00074243 | -0.0059833 | 9.1098
Calibration Table - V KED 51/51
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 0.66667 | 86.6025 | 0 |  | 0 | 0
Cal. Std.2 | 0.0033327 | 8.9306 | 0.05 | 0.050034 | 3.392e-05 | 0.06784
Cal. Std.3 | 0.0072289 | 15.8189 | 0.1 | 0.10029 | 0.00028621 | 0.28621
Cal. Std.4 | 0.037736 | 15.1666 | 0.5 | 0.49375 | -0.0062512 | -1.2502
Cal. Std.5 | 0.076041 | 1.2652 | 1 | 0.9878 | -0.0122 | -1.22
Cal. Std.6 | 0.37794 | 3.9214 | 5 | 4.8816 | -0.11844 | -2.3689
Cal. Std.7 | 0.74842 | 0.67045 | 10 | 9.6599 | -0.34011 | -3.4011
Cal. Std.8 | 2.0907 | 6.0027 | 25 | 26.9715 | 1.9715 | 7.8862
Cal. Std.9 | 3.8593 | 1.3468 | 50 | 49.7824 | -0.21763 | -0.43526
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99914 | 0.077534 | -0.00054666 | 0.0014377 | 0.00016926 | 1.4731e-06 | 0.0021807
Calibration Table - Cr (52) 52/52
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 4.3333 | 35.2506 | 0 |  | 0 | 0
Cal. Std.1 | 0.00088954 | 47.5375 | 0.01 | 0.0097594 | -0.00024063 | -2.4063
Cal. Std.2 | 0.0060549 | 11.9229 | 0.05 | 0.057447 | 0.0074468 | 14.8935
Cal. Std.3 | 0.010132 | 10.5678 | 0.1 | 0.095084 | -0.0049165 | -4.9165
Cal. Std.4 | 0.052442 | 9.4284 | 0.5 | 0.48569 | -0.014307 | -2.8613
Cal. Std.5 | 0.10649 | 1.6047 | 1 | 0.98469 | -0.015308 | -1.5308
Cal. Std.6 | 0.5215 | 3.6482 | 5 | 4.8161 | -0.18395 | -3.679
Cal. Std.7 | 1.0532 | 1.0418 | 10 | 9.7247 | -0.27529 | -2.7529
Cal. Std.8 | 2.8766 | 5.5276 | 25 | 26.5589 | 1.5589 | 6.2357
Cal. Std.9 | 5.2542 | 1.0221 | 50 | 48.5088 | -1.4912 | -2.9824
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99774 | 0.10832 | -0.00016758 | 0.0027563 | 8.0677e-05 | 6.4082e-06 | 0.0042182
Calibration Table - Cr (53) NH3 53/53
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 18.6667 | 16.3664 | 0 |  | 0 | 0
Cal. Std.1 | 0.00034964 | 16.0647 | 0.01 | 0.0097333 | -0.0002667 | -2.667
Cal. Std.2 | 0.001966 | 11.2792 | 0.05 | 0.056826 | 0.0068258 | 13.6516
Cal. Std.3 | 0.0034383 | 6.9134 | 0.1 | 0.099722 | -0.00027767 | -0.27767
Cal. Std.4 | 0.016785 | 0.3793 | 0.5 | 0.48858 | -0.011416 | -2.2832
Cal. Std.5 | 0.034859 | 1.6995 | 1 | 1.0152 | 0.015176 | 1.5176
Cal. Std.6 | 0.17092 | 2.9182 | 5 | 4.9794 | -0.020571 | -0.41142
Cal. Std.7 | 0.33602 | 1.0875 | 10 | 9.7897 | -0.2103 | -2.103
Cal. Std.8 | 0.8314 | 1.2295 | 25 | 24.2227 | -0.7773 | -3.1092
Cal. Std.9 | 1.642 | 1.5062 | 50 | 47.8411 | -2.1589 | -4.3177
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.9984 | 0.034323 | 1.5568e-05 | 0.00073408 | 2.1487e-05 | 3.0108e-07 | 0.0012143
Calibration Table - Fe (54) NH3 54/54
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 718.1719 | 2.584 | 0 |  | 0 | 0
Cal. Std.4 | 0.023585 | 3.4539 | 0.5 | 0.50147 | 0.0014715 | 0.2943
Cal. Std.5 | 0.047991 | 3.2964 | 1 | 1.6265 | 0.62652 | 62.6519
Cal. Std.6 | 0.1181 | 0.53866 | 5 | 4.8582 | -0.14183 | -2.8366
Cal. Std.7 | 0.22777 | 2.5407 | 10 | 9.9137 | -0.08633 | -0.8633
Cal. Std.8 | 0.56036 | 1.1381 | 25 | 25.2453 | 0.24532 | 0.98128
Cal. Std.9 | 1.0991 | 1.0676 | 50 | 50.0775 | 0.077527 | 0.15505
Cal. Std.10 | 2.2313 | 0.88273 | 100 | 102.2693 | 2.2693 | 2.2693
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99984 | 0.021693 | 0.012706 | 0.00019567 | 0.0002381 | -8.9636e-06 | 0.0097587
Calibration Table - Mn Oshift 55/71
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 8.6667 | 35.2506 | 0 |  | 0 | 0
Cal. Std.3 | 0.0062396 | 5.7482 | 0.1 | 0.10038 | 0.00037648 | 0.37648
Cal. Std.4 | 0.029293 | 11.7728 | 0.5 | 0.48956 | -0.01044 | -2.088
Cal. Std.5 | 0.059699 | 0.60458 | 1 | 1.0029 | 0.0028781 | 0.28781
Cal. Std.6 | 0.2962 | 0.64139 | 5 | 4.9955 | -0.0045219 | -0.090439
Cal. Std.7 | 0.60107 | 1.8338 | 10 | 10.1422 | 0.14218 | 1.4218
Cal. Std.8 | 1.4792 | 1.6128 | 25 | 24.9675 | -0.032513 | -0.13005
Cal. Std.9 | 2.9686 | 1.4394 | 50 | 50.1112 | 0.11118 | 0.22236
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99994 | 0.059235 | 0.0002938 | 0.00029738 | 7.6765e-05 | 4.4605e-07 | 0.00078013
Calibration Table - Fe (56) KED 56/56
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 952.3651 | 3.0961 | 0 |  | 0 | 0
Cal. Std.4 | 0.092662 | 14.6806 | 0.5 | 0.50256 | 0.0025623 | 0.51245
Cal. Std.5 | 0.19121 | 4.6396 | 1 | 1.6127 | 0.61271 | 61.2708
Cal. Std.6 | 0.47153 | 2.1495 | 5 | 4.7704 | -0.22961 | -4.5923
Cal. Std.7 | 0.90358 | 1.0008 | 10 | 9.6373 | -0.36271 | -3.6271
Cal. Std.8 | 2.4216 | 6.7702 | 25 | 26.7371 | 1.7371 | 6.9484
Cal. Std.9 | 4.3568 | 0.88842 | 50 | 48.5368 | -1.4632 | -2.9265
Cal. Std.10 | 9.2524 | 1.8322 | 100 | 103.685 | 3.685 | 3.685
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99886 | 0.088773 | 0.048049 | 0.0021167 | 0.0025758 | 0.00013378 | 0.11251
Calibration Table - Co KED 59/59
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 1.6667 | 34.641 | 0 |  | 0 | 0
Cal. Std.1 | 0.0019258 | 37.9709 | 0.01 | 0.0098975 | -0.00010248 | -1.0248
Cal. Std.2 | 0.010838 | 6.9316 | 0.05 | 0.052577 | 0.0025767 | 5.1533
Cal. Std.3 | 0.020914 | 4.6609 | 0.1 | 0.10083 | 0.0008267 | 0.8267
Cal. Std.4 | 0.10106 | 3.3667 | 0.5 | 0.48466 | -0.015342 | -3.0683
Cal. Std.5 | 0.20434 | 3.695 | 1 | 0.97921 | -0.020792 | -2.0792
Cal. Std.6 | 1.0145 | 2.0161 | 5 | 4.8589 | -0.14107 | -2.8213
Cal. Std.7 | 2.0213 | 1.0885 | 10 | 9.6803 | -0.31965 | -3.1965
Cal. Std.8 | 5.5445 | 6.8128 | 25 | 26.5526 | 1.5526 | 6.2102
Cal. Std.9 | 10.0095 | 1.1933 | 50 | 47.9347 | -2.0653 | -4.1307
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99922 | 0.20882 | -0.00014097 | 0.0033735 | 9.3096e-05 | 5.6486e-07 | 0.00081405
Calibration Table - Ni KED 60/60
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 5.6667 | 71.3198 | 0 |  | 0 | 0
Cal. Std.2 | 0.0036324 | 11.0432 | 0.05 | 0.053941 | 0.0039405 | 7.8811
Cal. Std.3 | 0.005328 | 13.8822 | 0.1 | 0.084412 | -0.015588 | -15.5876
Cal. Std.4 | 0.027755 | 4.2413 | 0.5 | 0.48745 | -0.01255 | -2.5099
Cal. Std.5 | 0.0579 | 1.3742 | 1 | 1.0292 | 0.029196 | 2.9196
Cal. Std.6 | 0.27767 | 2.715 | 5 | 4.9788 | -0.021236 | -0.42472
Cal. Std.7 | 0.56416 | 0.65975 | 10 | 10.1272 | 0.1272 | 1.272
Cal. Std.8 | 1.5196 | 6.7764 | 25 | 27.2977 | 2.2977 | 9.191
Cal. Std.9 | 2.7066 | 0.91792 | 50 | 48.6293 | -1.3707 | -2.7414
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99645 | 0.055645 | 0.00063094 | 0.0019191 | 0.00024154 | -3.0281e-05 | 0.021049
Calibration Table - Cu (63) KED 63/63
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 7.3333 | 69.9764 | 0 |  | 0 | 0
Cal. Std.2 | 0.0095589 | 26.9199 | 0.05 | 0.058436 | 0.0084363 | 16.8726
Cal. Std.3 | 0.015686 | 10.2068 | 0.1 | 0.097952 | -0.0020476 | -2.0476
Cal. Std.4 | 0.08462 | 4.5737 | 0.5 | 0.54254 | 0.042538 | 8.5077
Cal. Std.5 | 0.16219 | 5.1308 | 1 | 1.0428 | 0.042798 | 4.2798
Cal. Std.6 | 0.76205 | 3.5132 | 5 | 4.9116 | -0.08842 | -1.7684
Cal. Std.7 | 1.4847 | 0.23442 | 10 | 9.5721 | -0.42789 | -4.2789
Cal. Std.8 | 3.9855 | 5.293 | 25 | 25.7008 | 0.70079 | 2.8032
Cal. Std.9 | 7.172 | 0.99857 | 50 | 46.2521 | -3.7479 | -7.4957
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.9983 | 0.15505 | 0.0004982 | 0.0040448 | 0.0010441 | 1.6273e-05 | 0.0098107
Calibration Table - Cu (65) NH3 65/65
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 48.0001 | 8.3333 | 0 |  | 0 | 0
Cal. Std.2 | 0.0041343 | 6.8621 | 0.05 | 0.048671 | -0.0013288 | -2.6576
Cal. Std.3 | 0.0077799 | 8.7628 | 0.1 | 0.10351 | 0.0035086 | 3.5086
Cal. Std.4 | 0.036004 | 2.273 | 0.5 | 0.52806 | 0.028063 | 5.6126
Cal. Std.5 | 0.072257 | 1.7128 | 1 | 1.0734 | 0.073379 | 7.3379
Cal. Std.6 | 0.33452 | 1.8712 | 5 | 5.0183 | 0.018328 | 0.36656
Cal. Std.7 | 0.64926 | 1.2728 | 10 | 9.7527 | -0.24727 | -2.4727
Cal. Std.8 | 1.5883 | 1.0785 | 25 | 23.8781 | -1.1219 | -4.4876
Cal. Std.9 | 3.0853 | 1.0271 | 50 | 46.3961 | -3.6039 | -7.2078
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99842 | 0.06648 | 0.00089866 | 0.0015252 | 0.00019196 | 4.6786e-07 | 0.00084243
Calibration Table - Zn NH3 66/66
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 676.683 | 16.9854 | 0 |  | 0 | 0
Cal. Std.2 | 0.0028949 | 4.1149 | 0.05 | -0.11474 | -0.16474 | -329.4893
Cal. Std.3 | 0.0066006 | 3.1675 | 0.1 | 0.0052757 | -0.094724 | -94.7243
Cal. Std.4 | 0.020829 | 1.2068 | 0.5 | 0.46612 | -0.033879 | -6.7759
Cal. Std.5 | 0.04146 | 2.3354 | 1 | 1.1343 | 0.13433 | 13.4326
Cal. Std.6 | 0.16238 | 1.2796 | 5 | 5.0507 | 0.050655 | 1.0131
Cal. Std.7 | 0.31527 | 1.4974 | 10 | 10.0025 | 0.0025 | 0.025
Cal. Std.8 | 0.78279 | 0.22004 | 25 | 25.1448 | 0.14475 | 0.57901
Cal. Std.9 | 1.5094 | 0.57751 | 50 | 48.6774 | -1.3226 | -2.6453
Cal. Std.10 | 2.9202 | 0.35282 | 100 | 94.3714 | -5.6286 | -5.6286
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99721 | 0.030875 | 0.0064377 | 0.0010338 | 0.0012169 | -4.0484e-05 | 0.049553
Calibration Table - As 75/75
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | -2539.7451 | 7.525 | 0 |  | 0 | 0
Cal. Std.2 | 0.0017998 | 141.6305 | 0.05 | 0.049538 | -0.000462 | -0.924
Cal. Std.3 | 0.0036447 | 25.6395 | 0.1 | 0.102 | 0.0020047 | 2.0047
Cal. Std.4 | 0.017294 | 4.5082 | 0.5 | 0.49015 | -0.0098541 | -1.9708
Cal. Std.5 | 0.036012 | 3.978 | 1 | 1.0225 | 0.022463 | 2.2463
Cal. Std.6 | 0.1758 | 2.6451 | 5 | 4.9976 | -0.0024064 | -0.048128
Cal. Std.7 | 0.35708 | 2.1835 | 10 | 10.153 | 0.15299 | 1.5299
Cal. Std.8 | 0.8974 | 2.7673 | 25 | 25.5185 | 0.51852 | 2.0741
Cal. Std.9 | 1.6719 | 3.1971 | 50 | 47.544 | -2.456 | -4.912
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99962 | 0.035164 | 5.7803e-05 | 0.0003981 | 5.0104e-05 | -0.00058958 | 0.14332
Calibration Table - As KED 75/75
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 1 | 173.2051 | 0 |  | 0 | 0
Cal. Std.2 | 0.00036958 | 47.9672 | 0.05 | 0.052376 | 0.0023761 | 4.7522
Cal. Std.3 | 0.00067787 | 16.1143 | 0.1 | 0.089003 | -0.010997 | -10.9974
Cal. Std.4 | 0.0045714 | 17.1283 | 0.5 | 0.55156 | 0.051564 | 10.3129
Cal. Std.5 | 0.0079006 | 9.9212 | 1 | 0.94708 | -0.052918 | -5.2918
Cal. Std.6 | 0.040801 | 3.2154 | 5 | 4.8557 | -0.14425 | -2.885
Cal. Std.7 | 0.083161 | 2.5541 | 10 | 9.8883 | -0.11174 | -1.1174
Cal. Std.8 | 0.22938 | 4.5447 | 25 | 27.26 | 2.26 | 9.0398
Cal. Std.9 | 0.40474 | 1.5823 | 50 | 48.0934 | -1.9066 | -3.8132
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99671 | 0.0084173 | -7.1289e-05 | 0.00027959 | 3.5188e-05 | -6.0636e-05 | 0.059837
Calibration Table - As RH Oshift 75/91
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 24.6667 | 20.405 | 0 |  | 0 | 0
Cal. Std.2 | 0.12247 | 12.0043 | 0.05 | 0.050696 | 0.00069582 | 1.3916
Cal. Std.3 | 0.22393 | 2.3342 | 0.1 | 0.097469 | -0.0025311 | -2.5311
Cal. Std.4 | 1.0758 | 10.735 | 0.5 | 0.49019 | -0.0098109 | -1.9622
Cal. Std.5 | 2.204 | 3.054 | 1 | 1.0103 | 0.01026 | 1.026
Cal. Std.6 | 10.9293 | 2.3245 | 5 | 5.0325 | 0.032548 | 0.65096
Cal. Std.7 | 22.1249 | 2.1028 | 10 | 10.1937 | 0.19366 | 1.9366
Cal. Std.8 | 55.9463 | 3.1257 | 25 | 25.7851 | 0.78506 | 3.1403
Cal. Std.9 | 104.5128 | 2.1282 | 50 | 48.1739 | -1.8261 | -3.6522
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99965 | 2.1692 | 0.0125 | 0.023443 | 0.0029505 | 1.0809e-06 | 0.0022768
Calibration Table - As Oshift 75/91
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 24.6667 | 20.405 | 0 |  | 0 | 0
Cal. Std.2 | 0.0019004 | 10.3362 | 0.05 | 0.051819 | 0.0018193 | 3.6386
Cal. Std.3 | 0.0032415 | 5.4748 | 0.1 | 0.092992 | -0.007008 | -7.008
Cal. Std.4 | 0.016197 | 10.8298 | 0.5 | 0.49073 | -0.0092719 | -1.8544
Cal. Std.5 | 0.033104 | 2.3922 | 1 | 1.0098 | 0.0097935 | 0.97935
Cal. Std.6 | 0.16241 | 1.4573 | 5 | 4.9796 | -0.020414 | -0.40829
Cal. Std.7 | 0.32461 | 1.9249 | 10 | 9.9592 | -0.040772 | -0.40772
Cal. Std.8 | 0.83833 | 1.5294 | 25 | 25.7309 | 0.73092 | 2.9237
Cal. Std.9 | 1.6636 | 0.73088 | 50 | 51.0684 | 1.0684 | 2.1367
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99931 | 0.032573 | 0.00021256 | 0.00049462 | 6.2251e-05 | -9.6894e-07 | 0.002297
Calibration Table - Se (77) 77/77
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 980.0336 | 1.594 | 0 |  | 0 | 0
Cal. Std.3 | 0.00013418 | 332.1896 | 0.1 | 0.10129 | 0.0012925 | 1.2925
Cal. Std.4 | 0.0014701 | 29.2121 | 0.5 | 0.4747 | -0.025297 | -5.0593
Cal. Std.5 | 0.0031966 | 17.5913 | 1 | 0.95732 | -0.04268 | -4.268
Cal. Std.6 | 0.01866 | 1.7873 | 5 | 5.2796 | 0.2796 | 5.592
Cal. Std.7 | 0.036588 | 4.6375 | 10 | 10.2911 | 0.29108 | 2.9108
Cal. Std.8 | 0.091949 | 4.8056 | 25 | 25.7657 | 0.76572 | 3.0629
Cal. Std.9 | 0.17233 | 3.4604 | 50 | 48.2345 | -1.7655 | -3.5309
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99899 | 0.0035775 | -0.0002282 | 7.1906e-05 | 1.8561e-05 | 0.0020165 | 0.28112
Calibration Table - Se (77) Oshift 77/93
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 0 |  | 0 |  | 0 | 0
Cal. Std.3 | 6.0641e-05 | 16.0892 | 0.1 | 0.099969 | -3.1233e-05 | -0.031233
Cal. Std.4 | 0.00033487 | 7.3617 | 0.5 | 0.48465 | -0.015354 | -3.0707
Cal. Std.5 | 0.0007581 | 10.8763 | 1 | 1.0784 | 0.078356 | 7.8356
Cal. Std.6 | 0.0033396 | 6.2577 | 5 | 4.6996 | -0.30041 | -6.0082
Cal. Std.7 | 0.0069237 | 3.6164 | 10 | 9.7273 | -0.27271 | -2.7271
Cal. Std.8 | 0.017923 | 0.28371 | 25 | 25.1569 | 0.15689 | 0.62755
Cal. Std.9 | 0.036836 | 0.21345 | 50 | 51.687 | 1.687 | 3.3741
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99881 | 0.00071287 | -1.0624e-05 | 1.5552e-05 | 4.0144e-06 | 0 | 0
Calibration Table - Se (78) Oshift 78/94
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 2 | 100 | 0 |  | 0 | 0
Cal. Std.3 | 0.00030015 | 11.6964 | 0.1 | 0.10235 | 0.0023476 | 2.3476
Cal. Std.4 | 0.0010509 | 9.9475 | 0.5 | 0.44418 | -0.05582 | -11.1641
Cal. Std.5 | 0.0022346 | 6.1729 | 1 | 0.98316 | -0.016835 | -1.6835
Cal. Std.6 | 0.011232 | 1.8837 | 5 | 5.08 | 0.079999 | 1.6
Cal. Std.7 | 0.021884 | 3.4323 | 10 | 9.9298 | -0.070159 | -0.70159
Cal. Std.8 | 0.057569 | 2.7728 | 25 | 26.178 | 1.178 | 4.712
Cal. Std.9 | 0.11526 | 0.85977 | 50 | 52.4448 | 2.4448 | 4.8897
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99829 | 0.0021962 | 7.5368e-05 | 5.7574e-05 | 1.4862e-05 | 6.6875e-06 | 0.0137
Calibration Table - Se (80) Oshift 80/96
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 3.3333 | 69.282 | 0 |  | 0 | 0
Cal. Std.3 | 0.00051948 | 5.6567 | 0.1 | 0.10061 | 0.00060539 | 0.60539
Cal. Std.4 | 0.0023645 | 11.4642 | 0.5 | 0.48801 | -0.011992 | -2.3983
Cal. Std.5 | 0.0047427 | 3.3273 | 1 | 0.98738 | -0.012621 | -1.2621
Cal. Std.6 | 0.024074 | 3.8649 | 5 | 5.0466 | 0.046552 | 0.93104
Cal. Std.7 | 0.046042 | 1.756 | 10 | 9.6592 | -0.34082 | -3.4082
Cal. Std.8 | 0.12201 | 2.1667 | 25 | 25.6105 | 0.61049 | 2.442
Cal. Std.9 | 0.24552 | 1.5453 | 50 | 51.5451 | 1.5451 | 3.0903
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99966 | 0.0047625 | 4.0349e-05 | 5.5278e-05 | 1.4269e-05 | 4.8607e-06 | 0.0073139
Calibration Table - Se (82) 82/82
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 29.0906 | 65.5978 | 0 |  | 0 | 0
Cal. Std.3 | 0.00045436 | 79.7737 | 0.1 | 0.10058 | 0.00057514 | 0.57514
Cal. Std.4 | 0.0016346 | 20.7474 | 0.5 | 0.46669 | -0.033311 | -6.6621
Cal. Std.5 | 0.0035746 | 19.1663 | 1 | 1.0685 | 0.068529 | 6.8529
Cal. Std.6 | 0.016803 | 8.5397 | 5 | 5.1723 | 0.1723 | 3.446
Cal. Std.7 | 0.032695 | 1.7289 | 10 | 10.1021 | 0.10208 | 1.0208
Cal. Std.8 | 0.082066 | 1.5124 | 25 | 25.4178 | 0.41778 | 1.6711
Cal. Std.9 | 0.15018 | 4.0219 | 50 | 46.5481 | -3.4519 | -6.9038
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99854 | 0.0032236 | 0.00013015 | 7.7967e-05 | 2.0126e-05 | -0.00052711 | 0.12898
Calibration Table - Se (82) Oshift 82/98
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 0 |  | 0 |  | 0 | 0
Cal. Std.3 | 9.1072e-05 | 11.8156 | 0.1 | 0.1017 | 0.0017039 | 1.7039
Cal. Std.4 | 0.00037965 | 32.0113 | 0.5 | 0.44124 | -0.058762 | -11.7523
Cal. Std.5 | 0.00091102 | 10.1229 | 1 | 1.0665 | 0.066452 | 6.6452
Cal. Std.6 | 0.0042258 | 5.2813 | 5 | 4.9666 | -0.033372 | -0.66743
Cal. Std.7 | 0.0083206 | 5.913 | 10 | 9.7844 | -0.21556 | -2.1556
Cal. Std.8 | 0.021731 | 3.4691 | 25 | 25.563 | 0.56296 | 2.2519
Cal. Std.9 | 0.044189 | 0.82518 | 50 | 51.9872 | 1.9872 | 3.9744
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99801 | 0.00084991 | 4.6325e-06 | 2.3988e-05 | 6.1922e-06 | 0 | 0
Calibration Table - Mo (95) NH3 95/95
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 9.3333 | 44.6071 | 0 |  | 0 | 0
Cal. Std.1 | 0.0015026 | 1.8284 | 0.01 | 0.010433 | 0.00043268 | 4.3268
Cal. Std.2 | 0.0035566 | 2.3371 | 0.05 | 0.041516 | -0.0084839 | -16.9678
Cal. Std.3 | 0.0068009 | 2.2591 | 0.1 | 0.090612 | -0.009388 | -9.388
Cal. Std.4 | 0.033379 | 3.9039 | 0.5 | 0.49283 | -0.0071698 | -1.434
Cal. Std.5 | 0.068347 | 1.9977 | 1 | 1.022 | 0.022004 | 2.2004
Cal. Std.6 | 0.33828 | 1.783 | 5 | 5.1069 | 0.10693 | 2.1386
Cal. Std.7 | 0.68282 | 1.0186 | 10 | 10.321 | 0.32103 | 3.2103
Cal. Std.8 | 1.7833 | 0.8785 | 25 | 26.9742 | 1.9742 | 7.8968
Cal. Std.9 | 3.5697 | 0.74549 | 50 | 54.0085 | 4.0085 | 8.0169
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.9963 | 0.06608 | 0.00081326 | 0.0021534 | 6.3029e-05 | 7.8e-07 | 0.00086441
Calibration Table - Mo (95) OShift 95/127
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 4.6667 | 65.4654 | 0 |  | 0 | 0
Cal. Std.1 | 0.0014669 | 3.0082 | 0.01 | 0.010311 | 0.00031122 | 3.1122
Cal. Std.2 | 0.0035828 | 5.811 | 0.05 | 0.043823 | -0.0061771 | -12.3543
Cal. Std.3 | 0.0067405 | 1.9976 | 0.1 | 0.093836 | -0.0061643 | -6.1643
Cal. Std.4 | 0.031184 | 10.9683 | 0.5 | 0.48098 | -0.019022 | -3.8044
Cal. Std.5 | 0.06631 | 0.62608 | 1 | 1.0373 | 0.037329 | 3.7329
Cal. Std.6 | 0.32647 | 1.2359 | 5 | 5.1578 | 0.15781 | 3.1562
Cal. Std.7 | 0.65972 | 1.426 | 10 | 10.436 | 0.43604 | 4.3604
Cal. Std.8 | 1.7049 | 0.95471 | 25 | 26.9903 | 1.9903 | 7.9614
Cal. Std.9 | 3.5001 | 0.87516 | 50 | 55.4237 | 5.4237 | 10.8473
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99746 | 0.063137 | 0.00081593 | 0.001839 | 5.075e-05 | 5.0067e-07 | 0.00072963
Calibration Table - Cd (111) NH3 111/111
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 0 |  | 0 |  | 0 | 0
Cal. Std.1 | 0.00040055 | 14.2459 | 0.01 | 0.0099689 | -3.1147e-05 | -0.31147
Cal. Std.2 | 0.0018146 | 5.3542 | 0.05 | 0.049858 | -0.00014238 | -0.28477
Cal. Std.3 | 0.0037066 | 0.96956 | 0.1 | 0.10323 | 0.003229 | 3.229
Cal. Std.4 | 0.018162 | 2.2798 | 0.5 | 0.51101 | 0.011009 | 2.2018
Cal. Std.5 | 0.035696 | 1.2988 | 1 | 1.0056 | 0.0056097 | 0.56097
Cal. Std.6 | 0.17637 | 2.3035 | 5 | 4.9739 | -0.026053 | -0.52106
Cal. Std.7 | 0.34574 | 0.54908 | 10 | 9.7515 | -0.24847 | -2.4847
Cal. Std.8 | 0.88101 | 0.10981 | 25 | 24.851 | -0.14902 | -0.59608
Cal. Std.9 | 1.7407 | 1.1014 | 50 | 49.1032 | -0.89682 | -1.7936
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99982 | 0.03545 | 4.7157e-05 | 0.0002517 | 7.3672e-06 | 0 | 0
Calibration Table - Cd (114) NH3 114/114
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 5.3333 | 21.6506 | 0 |  | 0 | 0
Cal. Std.1 | 0.00087683 | 8.6166 | 0.01 | 0.0099199 | -8.012e-05 | -0.8012
Cal. Std.2 | 0.0043644 | 9.1915 | 0.05 | 0.050899 | 0.00089904 | 1.7981
Cal. Std.3 | 0.0088892 | 3.2754 | 0.1 | 0.10407 | 0.004065 | 4.065
Cal. Std.4 | 0.043608 | 0.49346 | 0.5 | 0.51201 | 0.012012 | 2.4025
Cal. Std.5 | 0.084585 | 1.4969 | 1 | 0.99349 | -0.0065126 | -0.65126
Cal. Std.6 | 0.41713 | 1.3954 | 5 | 4.9009 | -0.099148 | -1.983
Cal. Std.7 | 0.83659 | 1.0252 | 10 | 9.8295 | -0.17053 | -1.7053
Cal. Std.8 | 2.1105 | 0.84356 | 25 | 24.7974 | -0.20258 | -0.81034
Cal. Std.9 | 4.1569 | 1.0617 | 50 | 48.8427 | -1.1573 | -2.3145
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99973 | 0.085107 | 3.2587e-05 | 0.00074742 | 2.1877e-05 | 1.4234e-07 | 0.00018674
Calibration Table - Pb 208/208
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 26.6667 | 30.3109 | 0 |  | 0 | 0
Cal. Std.1 | 0.0023835 | 6.5843 | 0.01 | 0.0097229 | -0.00027708 | -2.7708
Cal. Std.2 | 0.011652 | 8.3219 | 0.05 | 0.054906 | 0.0049057 | 9.8113
Cal. Std.3 | 0.022281 | 0.96162 | 0.1 | 0.10672 | 0.0067227 | 6.7227
Cal. Std.4 | 0.10827 | 0.61733 | 0.5 | 0.52592 | 0.02592 | 5.1839
Cal. Std.5 | 0.21557 | 0.28412 | 1 | 1.049 | 0.049019 | 4.9019
Cal. Std.6 | 1.0061 | 2.2274 | 5 | 4.9029 | -0.097113 | -1.9423
Cal. Std.7 | 1.922 | 0.58706 | 10 | 9.368 | -0.632 | -6.32
Cal. Std.8 | 4.3293 | 1.0161 | 25 | 21.1033 | -3.8967 | -15.5867
Cal. Std.9 | 6.1368 | 6.2025 | 50 | 29.915 | -20.085 | -40.1701
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.99616 | 0.20513 | 0.00038902 | 0.0073561 | 0.000203 | -5.6069e-06 | 0.00078098
Calibration Table - Pb NH3 208/208
| Net Intensity (cps) | RSD (%) | STD Conc (ppb) | Apparent Conc. (ppb) | Residuals (ppb) | Residual Error (% RE)
Blank | 57.3334 | 2.014 | 0 |  | 0 | 0
Cal. Std.1 | 0.0045514 | 5.3534 | 0.01 | 0.0099954 | -4.5875e-06 | -0.045875
Cal. Std.2 | 0.021995 | 3.9951 | 0.05 | 0.051244 | 0.0012442 | 2.4885
Cal. Std.3 | 0.039589 | 1.5124 | 0.1 | 0.092847 | -0.0071532 | -7.1532
Cal. Std.4 | 0.21289 | 1.5252 | 0.5 | 0.50265 | 0.002648 | 0.52959
Cal. Std.5 | 0.53685 | 10.9052 | 1 | 1.2687 | 0.26871 | 26.8711
Cal. Std.6 | 2.04 | 0.87977 | 5 | 4.8232 | -0.17681 | -3.5361
Cal. Std.7 | 4.1045 | 1.8979 | 10 | 9.7051 | -0.29494 | -2.9494
Cal. Std.8 | 9.2284 | 0.99078 | 25 | 21.8212 | -3.1788 | -12.715
Cal. Std.9 | 20.4072 | 1.3891 | 50 | 48.2553 | -1.7447 | -3.4895
Cor.Coeff. : | Slope | Intercept | Slope SD | Intercept SD | BEC | DL
0.9934 | 0.42289 | 0.00032436 | 0.018461 | 0.00054036 | -4.3748e-08 | 3.5861e-05